## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 5 Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |
| 5 October 2007 | Recalibrated TSG data based on results of 2007-42 |

## PROCESSING NOTES

Cruise: 2007-09

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, BC

Location: North-East Pacific

Project: High Seas Salmon

Party Chief: Morris J.

Platform: W.E. Ricker

Date: February 27, 2007 – March 27, 2007

Processed by: Germaine Gatien

Date of Processing: May 28, 2007 – 9 July 2007

Number of original CTD casts: 120 	Number of CTD casts processed: 120

Number of TSG files: 1			Number of TSG files processed: 1

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with a Seapoint Fluorometer (#2229) on the primary pump with a 10X cable. The deck unit was a model 911 (#0471). The data logging computer was PACOSAPFS03. The salinometer used was a model 8400B Autosal.

# SUMMARY OF QUALITY AND CONCERNS

There was only one deep salinity calibration sample. There were many samples from ~10m.

The primary sensor calibration drifted during a previous cruise so the secondary sensors were selected for casts #1 through 142. However, after that cast the primary sensors were selected because of bad data between the surface and 75db of the downcasts. The only recalibration done to the CTD data was to bring the primary and secondary sensors into agreement. There were doubts about the comparison with bottles, so it was considered wise to put off further salinity recalibration until more cruises have used this instrument or until a post-cruise calibration of the sensors.

There was no mention in the appropriate section of the log book that the Thermosalinograph was used. But there was some loop sampling and a TSG file was found on the CD ROM from sea. However, that file covers only the time between cast #353 and arrival in Nanaimo. No recalibration was applied since there are no loop samples and only 2 CTD casts that overlap with this file. The temperature is likely to be high by up to 0.2C˚ because of ship heating and the salinity is likely a little low but there is no history on which to base an estimate for that. Hugh Maclean believed that TSG data was gathered throughout. He checked the Ricker computer and found no file for that period. Should an earlier TSG file be found, it would be possible to do a comparison with loop samples allowing recalibration of this data.

# PROCESSING SUMMARY

##### 1. Seasave

This step was completed at sea; the raw data files have extension DAT.

##### 2. Preliminary Steps

The Log Book was obtained and the only problems noted were that the NMEA time did not update when the time changed from PST to PDT and that a Niskin bottle broke and was replaced. There is mention of loop sampling, but no TSG data, so it is presumed this was done as a surface sampling substitute.

The bottle salinity, nutrient and chlorophyll data were obtained in spreadsheet format. There was only one deep salinity bottle sample.

The cruise summary sheet was completed. The histories of the conductivity and pressure sensors were obtained as well as that of the thermosalinograph conductivity and temperature.

Calibration constants were checked. The only errors were in the transmissometer constants; these were fixed and the con file was named 2007-09-ctd.con.

##### 3. Conversion of Raw Data

All data were converted using the con file given above.

A few casts were examined and all expected channels are present and look reasonable. It is assumed that sampling was done by a Niskin bottle mounted 5m above the CTD as usual. However, a test attempt to create rosette files produced files from the beginning of the downcasts. Examination of the CTD data makes it clear that the Niskin bottles were fired at about 10m, so these surface files are not useful for comparison with the bottles. Surface data may be useful for comparison with the TSG, but the data in the ROS files are from a time when the pumps were not operating and the differences between the two T and S channels are large. The ROS files will not be processed.

##### 4. WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### 5. CELLTM

Tests were run on 4 casts using a variety of settings for CELLTM to determine the best choice of parameters. The changes were slight for the primary sensors among no adjustment, (0.02, 9) and (0.02, 7). For the secondary the choice of (0.02, 7) and (0.0245,9.5) are very close. CELLTM was run using (0.02, 7) for the primary and (0.0245, 9.5) for the secondary conductivity for all casts.

##### 6. DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### 7. Test Plots and Channel Check

Three casts were plotted to check for agreement between the pairs of T and C sensors. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 40 | 250 | -0.0009 | -0.00044 | -0.0038 | Mod, very steady |
| 139 | 250 | -0.0007 | -0.00042 | -0.004 v.noisy | Mod, very steady |
| 160 | 250 | -0.0005 | -0.00042 | -0.0039 | High, steady |
| 160 | 500 | -0.0005 | -0.00042 | -0.0041 | High, steady |
| 239 | 250 | -0.0005 | -0.00038 | -0.0042 v noisy | Mod, very steady |
| 356 | 250 | -0.0006 | -0.00036 | -0.0034 | High, steady |

There may be a slight drift in conductivity differences.

##### 8. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### 9. Checking Headers

The header check was run and no errors found.

The header summary was run and station names were checked against the log as well as a random sampling of positions and times to ensure that the GPS was working well. No errors were found.

The track plots (using event #s and station names) were produced and added to the end of this report.

The average surface pressure is 3.2db which is reasonable for the Ricker. One cast was examined that had surface pressures close to zero during the upcast and the conductivity is very low as expected.

The time changed from PST to PDT during the cruise and the computer clock was adjusted. The header times agree with the log entries.

##### 10. SHIFT

Conductivity

Tests were run on the primary and secondary conductivity channels using a variety of settings. The best results overall were with a choice of -0.6s for the primary and +1.5 for the secondary which is the same as found during 2006-31. All casts were put through SHIFT using those settings.

Fluorescence

In the past it has been found appropriate to advance the Fluorescence by +24 records (1s) relative to the other channels. A few casts were examined to compare the offset between the upcast and downcast fluorescence with that of the temperature traces to ensure this setting is appropriate. The CTD was stopped at 15db during the upcast, so this is a little hard to judge, but +24s looks like a reasonable choice.

A shift of +24 records (1s) was applied since this is the usual for this equipment. (Output: SHFFL)

##### 11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              		Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: All the warnings pertain to upcast data, so are of no concern.

##### 12. COMPARISON WITH BOTTLES

Salinity

There was only one deep bottle (500db). A comparison with the CTD data from 5db above the bottom suggests that the CTD is reading high by 0.009 for the primary and by 0.005 for the secondary. The salinity gradient is low at that level and to find salinity around 33.3388 you have to look as high as 460db for the primary and 473db for the secondary. So either there is a problem with the Niskin sample, or the CTD error is much larger than expected.

Surface samples were examined. First, a random selection of casts were studied to determine the depth at which the CTD stopped. An estimate of 14.5db suits most of the casts with a range of 13.5 to 15db covering all that were checked. A quick check of a few casts showed that the downcast data from 8 to 10.5db were reasonably close to the bottles. Files were thinned to values at 8.5, 9, 9.5, 10 and 10.5db and the salinity and fluorescence values exported to a spreadsheet (2007-09-bot.csv).

That file was rearranged and then combined with the salinity spreadsheet to produce 2007-09_bottle_comp_all-data.xls with separate columns for the different pressures. One pad value (-99) was replaced with 0. Average differences and standard deviations of differences were found for the different CTD depths. The differences were lowest at 10db for the primary salinity and 10.5 for the secondary. The standard deviations were lowest at 9.5 and 10db respectively. Looking at plots of differences versus cast number shows there are many outliers and that the secondary salinity differences are much noisier than the primary.

A second version of the file was saved as 2007-09_bottle_comp_excl.xls. Outliers were identified from plots and those casts were excluded from the comparison. The differences were least at 9db for the primary and 10db for the secondary with standard deviations lowest at 9.5 and 10.5db. The primary appears to be very close to the bottles and the secondary low by about 0.04. However, if we look at 5 casts that are well-mixed to at least 11db, so that knowing the depth of sampling is not critical, the primary appears to be high by about 0.009 and the secondary high by about 0.0045. This is close to the results of the single deep bottle. However, the median differences are 0.0066 and 0.002.

The secondary is closer to the bottles, but the differences are much noisier.

Fluorescence

Extracted chlorophyll data were obtained in spreadsheet format; this was simplified and saved as file 2009-09chl.csv. That file was converted to individual CHL files which were then compared with fluorescence values from 10db. For the 5 casts identified as well mixed at 10m from the salinity comparison, it was found that the fluorescence and chlorophyll were close with a ratio of from 0.7 to 1.2. The highest ratio was associated with the highest fluorescence.

When all values were used in COMPARE there is reasonable correspondence between the fluorescence and chlorophyll. When all data is included CHL ~ 0.85 times FL, but when a few outliers are excluded the slope is close to 1. There are only a few values above 1.5 and there is considerably more scatter among those points.

##### 13. DETAILED EDITING

CTDEDIT was used to clean the salinity data.

The choice of which sensor pair to use for the archive was complicated. The secondary salinity was closer to the bottles, but the differences were noisier in the second half of the cruise. During 2006-31 the primary was found to have a significant time-dependence in its calibration, with errors varying from -0.001 to +0.006. A comparison of salinity traces showed that from cast #145 to the end there were marked differences between the two salinity traces. Investigation showed that the secondary salinity data were bad between the surface and 75db.

The secondary sensors were selected for casts #1 through 142 and the primary sensors were selected for casts #145 to the end of the cruise.

On-screen plots of descent rate and pump status were used to guide editing.

The following casts required no editing: 55, 121 and 311

All other casts required light editing.

Note was made of the editing details in the headers of the relevant files.

##### 14. Other Comparisons

Previous experience with these sensors –The pressure sensor was recalibrated in May 2006 and during 2006-31 the recommended offset worked well. The temperature and conductivity sensors were also used during 2006-31 when the primary salinity showed some time dependence with differences from bottles varying from -0.001 to +0.006. The secondary salinity was found to be low by about 0.002. Calibration sampling was all from 10m bottles, so the analysis was limited to 14 bottles from well-mixed surface waters.

Historic ranges – Profile plots were made of T and S with local climatology superimposed where available. There were some excursions, mostly salinity looking low near the surface. Most of the casts are very close to the shore and the climatology covers a fairly large area, so is probably not reflective of near-shore stations. (The secondary salinity is known to be low for the later casts in the top 75db, but even the primary salinity falls a little below the minimum value.)

##### 15. Initial Recalibration

The salinity comparisons suggest that the primary channel gives salinity that is high by about 0.009 and the secondary by about 0.0045, but there are some doubts about this. During 2006-31 the secondary was found to be high by 0.002 while the primary was low early in the cruise and high by 0.006 at the end. The drift does not appear to be high during this cruise since the difference between the two salinity channels varies only slightly. Given the limitations in the comparison the only recalibration that will be applied will be to subtract 0.0045 from the primary salinity to bring it into line with the secondary. This is a little higher than the differences found in section 8, but most of those casts are ones for which the noise level was high. Examination of a few more casts from quieter sections (identified from plots done in the bottle analysis) found differences closer to 0.0045. When these sensors are next recalibrated this issue can be revisited.

File 2007-09-recal1.ccf was used to subtract 0.0045 from the primary salinity for all EDT files.

##### 16. Fluorescence Processing and special files for Angelica Peña

The EDT files were put through a median filter, size 11, applied to the fluorescence channel only.

Five sets of files were prepared for the use of Angelica Peña:

TWO DOWNCAST DATA SETS

1. The EDT files were clipped to 100db, bin-averaged with 0.25db bins, unwanted channels removed and HEADEDIT used to fix headers, formats, channel names and to add comments. (Output: FCTD1)

2. The EDT files were clipped to 100db, median filtered (width 11 applied to fluorescence only), bin-averaged with 0.25db bins, unwanted channels removed and HEADEDIT used to fix headers, formats, channel names and to add comments. (Output: FCTD2)

THREE UPCAST DATA SETS - Because there were no rosette files special files were prepared with upcast data from around the depths at which chlorophyll was sampled. The SHFFL files were put through REVERSE. (Output: REV) In a few cases there were negative values near the surface and a text editor was used to either remove those records or if there were just a few, the “-“ sign was removed.

3. The reversed files were put through DELETE to create files with upcast data only. Those were clipped so they contain only the top 16db of data, then REMOVE and HEADEDIT were run to remove unwanted channels and fix formats and add a comment. (Output: UPFCTD1)

4. Another set of upcast files were created by applying a median filter of fixed width 11 to the fluorescence channel of the REV files followed by DELETE, CLIP, REMOVE and HEADEDIT. (Output: UPFCTD2)

5. Finally a set of upcast files were created by running DELETE, CLIP, REMOVE, BIN AVERAGE and HEADEDIT. (Output: UPFCTD3) The bins were size 0.25db.

##### 17.  BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. No further editing was considered necessary.

Transmissivity and Fluorescence profiles were plotted on-screen and no problems noted.

##### 18. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from casts 1-142: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

The following channels were removed from casts 145-356: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag.

HEADER EDIT was used to fix formats and channel names, to fix the platform name and geographic description and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### 19. Thermosalinograph Data

a.) Checking calibrations

Configuration file 2007-09-0002.con was found. The calibrations were checked and no errors were found.

There are no files for 2007-09-0001 though there are loop samples suggesting that the TSG was in use. Perhaps the file has been mislaid or the TSG failed to record.

b.) Converting to IOS Headers and adding position headers and time channels

The data were converted to a CNV file using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Latitude, Longitude, Salinity:T0:C0 and Time Julian. There was no channel for Flow Rate, no secondary temperature and no fluorometer.

The files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

The file starts at the time of CTD cast #353 and continues through the trip to Nanaimo.

c.) CTDEDIT

A time-series plot was produced. There are a few spikes in temperature and salinity but they are in the Johnstone Strait / Discovery Passage region where such variability is expected. No editing was applied.

d.) Checking Time Channel

Matching times of casts 343 and 346 with times in the TSG file, the latitude and longitude readings are very close, so the clock appears to have worked well.

e.) Comparison of T and S from TSG and CTD data

There are two CTD readings that coincide with the TSG record, but they come from a little deeper than the TSG intake (~3db).

| CTD cast | TSG Temp | CTD Temp | TSG Salinity | CTD Salinity |
| --- | --- | --- | --- | --- |
| 343 | 7.0429 | 6.9685 | 30.2372 | 30.2510 (4.4db) |
| 346 | 6.8395 | 6.7598 | 26.5540 | 26.6936 (3.4db) |

The readings are reasonably close.

There is no record of any loop samples from the time of the TSG record.

f.) Calibration history

The TSG has not been used since the last recalibration of sensors.

Conclusions

There is insufficient information to justify recalibration at this time. If another hex file is found from this cruise or another cruise is processed that uses this instrument, this decision could be reconsidered.

f.) Recalibration

None was applied.

h.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary and Flag.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and the following comment about the data.

The temperature and salinity data have not been recalibrated due to a

lack of information .

As a final check plots were made of the cruise track and data; no problems were noted.

REVISION: October 5, 2007  Based on the results of cruise 2007-42 when the same TSG equipment was used the TOB file was recalibrated to apply offsets of -0.07C˚ and +0.024 to the Temperature:Primary and Salinity:T0:C0. The header entry was edited to explain what was done.

##### 20. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the TSG and CTD sensors.

Particulars: (Notes from log book)

160. Deep salinity sample

163. Time changed from PST to PDT after this cast. Clock on computer was adjusted.

215. Niskin bottle replaced.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2007-09 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Dates: Start: 27 February 2007 End: 27 March 2007 |  |  |  |  |  |
| Location: NE Pacific |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: Morris J |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506         Cruise ID#:		2007-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19/03/05 | Factory |  |  |
| Conductivity | 2280 | 23/05/06 | “ |  |  |
| Secondary Temp. | 2710 | 07/04/05 | “ |  |  |
| Secondary Cond. | 1764 | 23/05/06 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Transmissometer | 498DR | 28/12/06 | IOS |  |  |
| Pressure Sensor | 69698 | 26/05/2006 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2007-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 06/01/05 | Factory | 29/12/06 | Factory |
| Conductivity | 2488 | 06/01/05 | “ | 29/12/06 | Factory |