## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2007-11

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Broughton Archipelago

Project: Aquaculture Oceanography

Party Chief: Stucchi D.

Platform: Vector

Date: March 26, 2007 – April 1, 2007

Processed by: Germaine Gatien

Date of Processing: 11 April 2007 – 30 April 2007

Number of CTD casts:  	51 (+3 SBE19)	Number of CTD casts processed: 51

Number of rosette casts: 23	Number of rosette casts processed: 22 (no samples from 1 bottle)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea/Seatech transmissometer (#1005DR) and an Altimeter OA-916D (#1024). A SeaBird dissolved oxygen sensor (#1117) and a Seapoint Fluorometer (#2845) with a 10X cable were mounted on the pumps. The deck unit is unknown. The salinometer was an 8400B model Autosal (S/N 68572).

A SeaBird Model SBE 19 CTD (serial number 1294) was mounted on the SBE 911+ for 3 casts.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book and rosette sheets were available.

The SBE dissolved oxygen data in the CTD files should be considered, roughly:

- ±0.3ml/l from the surface to 250m

- ±0.1ml/l below 200m

The Thermosalinograph data was gathered and recalibrated by R. Pawlowicz (UBC). Formats were changed and a few channels removed; then the data was converted to IOS HEADER format. The file was thinned and was still too large for our plot routines to work, so it was divided into 3 separate TOB files. Note that this data covers the period March 13 to April 29, 2007.

Tests were run comparing SBE19 (Serial Number 1294) with the SBE911+ to assist with the calibration of data from other cruises. The conclusions are at the end of this report. Of special note is the suggestion that the SBE19 should be allowed to soak for 2-3 minutes for more accurate salinity data. This was a very useful experiment and it is recommended that it be repeated occasionally, preferably in fairly well-mixed water, since in areas of high variability the time-response characteristics limit the calibration accuracy.

# PROCESSING SUMMARY – SBE 911+

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained and read; no significant problems were mentioned.

The first cast had an error in the name so this was corrected.

An SBE19 was strapped to the SBE911+ for 3 casts for the purpose of calibrating the SBE19. The processing notes for these 3 casts and the comparison with the SBE911+ are at the end of the report.

Nutrient and salinity data were each obtained in spreadsheet format.

Individual dissolved oxygen bottle files were obtained with flag channels and comments.

There was no chlorophyll sampling.

The 911+ CTD calibrations were checked and the only problem noted is that the pressure offset was set to 0. Recent cruises using this instrument have led to the conclusion that a pressure offset of 1.2db is appropriate. A copy of file 2007-11-0001 with that adjustment was saved as 2007-11-CTD.con.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format. CLEAN was used to add event numbers and the extensions were changed to BOT. Header Check was run on the BOT files and the only error found was the absence of a station name for cast #1. This was added to the BOT file.

All BOT files were plotted and no significant outliers were found.

The SBE conversion routine was then run to create CNV files. A few files were checked and there are some pressures <0db which appear to correspond to in-water values. Cast #22 has a pressure of -0.3 at the top of the upcast with pumps on and in-water values of salinity. An offset of +1.2db has been applied to pressure for all cruises since July 2006 and looks appropriate for this cruise as well. So the con file was changed to add that offset and the data were reconverted.

Plots were made to ensure the data look reasonable. The descent rate was generally quite high and steady. Bottle stops typically lasted for at least 30s and there is clear evidence of the wisdom of that, as evidence of shed wake corruption is clearly visible through the first 10 to 20s of the stop.

The temperature and conductivity channels are farther apart than usual during downcasts and significantly different during upcasts. There appears to be a lot of noise in all channels, but this could be real variation due to active mixing.

There are significant differences between upcast and downcast transmissivity and fluorescence as well.

The oxygen trace shows the usual hysteresis but the vertical offset between downcast and upcast is only slightly greater than for temperature with about a 3s delay.

The fluorescence dark values are ~0.12ug/l.

The altimetry was sometimes extremely noisy but a signal is usually clear near the bottom.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Three casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried. A choice of (0.03, 9) was best for the primary sensors but there was little difference among a number of choices for the secondary sensors, with no correction seeming marginally better than other choices.

CELLTM was run using (0.03, 9) for the primary sensors.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 10 | 200 | +0.0008 | +0.00012 | +0.0008 | Moderate, very steady |
| 30 | 140 175 | +0.0007 +0.0007XN | +0.00015 +0.00015 XN | +0.0012 +0.0011 XN | Moderate, very steady |
| 44 | 160 | +0.0008 | +0.0002 | +0.0012 | Moderate, very steady |

While there is a slight hint of a trend in salinity differences, the data are too noisy to conclude that this is real. The differences are small and there is no sign of pressure-dependence in the salinity differences though that is hard to judge from such shallow casts.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

A station name was added to the header of cast #1.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and no errors were found in the headers. There is evidence of pressure spikes.

The header summary was checked carefully against the log book entries and two errors in station name were found and corrected in both full files and bottle files.

Track plots were produced and added to the end of this report. Cast #1 was excluded from those plots so the details in Broughton Archipelago are clear.

The average surface pressure is 2.8db which is a little deeper than usual for the Vector, but as the pressure offset was not in the configuration used at sea, it would have looked like 1.6db on-screen. There are no negative surface values. The mixed layer calculation shows that surface salinity samples will not be very useful for calibration purposes as the gradients are likely to be large there.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. That file was edited to add sample numbers taken from the rosette sheets. This was saved as addsamp.csv. There were a few problems in doing this. Cast #20 is said to have 2 bottles but no rosette file was created. There is a note in the log about communication errors and no bottle confirmation. The BL file is empty, so a fake bl file was created by adding two lines for the two bottles; scan numbers were estimated by examining a plot of descent rate versus scan numbers and times were estimated from the log entries. For the deepest bottle it was decided to avoid a part of the record that had large spikes. This file was then used to create a bottle file and produced reasonable results. The descent rate is low and pressure varies little for each bottle. Nonetheless, this bottle should be treated with care since the data are probably not from exactly the time of firing. So far, no bottle data have been found for this cast, so probably the bottles didn’t close. However, the BOT and SAM files will be kept in case any bottle data do turn up.

There was no sample #167 assigned.

The ADDSAMP file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data were provided in EXCEL format with flag and comment channels. Flags and comments had been added. Event #s were entered and channel names were changed to standard format and the file was saved as 2007-11-sal.csv. The following problems were noted:

- Cast #33. The rosette log indicates there was sampling from the two deepest bottles with sample #s 153 and 154. The analysis has values for samples #166 and 167. The first corresponds to the surface bottle of cast #33 and the second sample number is skipped in both the CTD Daily Log and rosette file, so presumably was never assigned. It will be assumed that 166 and 167 are really 153 and 154 and the samples were flagged “c” and a note made in the headers about this reassignment. This can be reconsidered after COMPARE is run.

- Cast #20. There were two samples shown on the rosette log and in the CTD log but there were no bottles in the analysis. There were also no oxygen or nutrient samples so presumably there was no sampling from this cast. There is mention in the log of problems in communication and no bottle confirmations, so maybe the bottles did not close.

- Cast #31 – The rosette log indicates that there should be a salinity sample for #140. No such sample is in the analysis spreadsheet.

- Cast #24 – after a preliminary run of COMPARE it was suspected that the two samples were out of order, with #117 having lower salinity water than that of #118, so the values were reversed and flagged “c” and a note put in the header of the SAL file.

There were two duplicate samples and the differences were each 0.0003 indicating excellent repeatability.

There were loop samples; these were saved separately as 2007-11-loop-sal.csv.

File 2007-11-sal.csv was converted into individual SAL files.

The dissolved oxygen data were received in individual files with flag channel and comments. The DO data for cast #13 were split into two files. These were combined using a text editor, put through CLEAN to fix the headers and renamed 2007-11-0013.ADD. The header of cast #49 had a formatting error in the headers which was fixed. File 2007-11-0005.ADD was relabelled 2007-11-0009.ADD because there was no rosette sampling from cast #5. The rosette sheet was labelled incorrectly.

The nutrient data were received in spreadsheet format; this was simplified, edited and saved as 2007-11nuts.csv and converted to individual NUT files.

The SAL, ADD and NUT files were merged with the CST files in three steps. (Output: MRG1, MRG3, MRG4) The MRG4 files were cleaned to reduce the headers to File and Comments. (Output:MRGCLN1) and those files were merged with the SAMAVG files to produce MRG files.

The MRG files were put through CLEAN, this time to remove SeaBird headers and comments from the secondary files. (Output: MRGCLN2)

##### 11. COMPARE

Salinity

COMPARE was run and there was a lot of scatter in the differences. When samples with differences >0.02 were excluded a flat trend line was achieved for both salinity channels when plotted against pressure. Plots against time are noisy with no obvious evidence of drift. Excluding a few more outliers provided extremely flat fits and with the primary salinity within 0.0015 of the CTD and the secondary within 0.0004. The bottles that had sample numbers reassigned were examined closely:

- Cast #24 – Bottles were fired at 296db and 251db. The bottle originally labelled as 117 differs from the CTD value at 296db by about 0.06 and from the CTD at 251db by 0.02. The bottle is more likely from 251db though the difference is still quite large. The bottle originally labelled 118 differs from the CTD at 251db by 0.046 and from that at 296db by 0.006. The standard deviation in the CTD data are fairly high especially for the 296db level, so it is possible that the differences are due to the CTD, but there is too much doubt here, so the sample numbers were returned to the original. A note was made in the headers and the flags changed to “d” because the uncertainly is so great.

- Cast #33 – There two samples compare well with the CTD data, especially sample #143. The differences between CTD and bottles were +0.0020 and -0.0012. The flags were left at “c” and a note added to the headers to indicate that the reassignment of labels appears to be correct.

The only other sample that was flagged was sample #37 from cast #11 which was well out of line in COMPARE and the CTD data were not noisy at that level. It was flagged “c”.

Dissolved Oxygen

Plots of Titrated DO versus CTD Salinity from the bottle files were examined. There were 3 significant outliers: sample #28 from cast #9, sample #158 from cast #33 and sample #178 from cast #49.

COMPARE was run. Plotting the differences versus CTD DO and pressure and excluding a few outliers based on residuals led to reasonably tight fits against pressure and CTD_DO. The three outliers in the plots mentioned above also show up as the largest outliers in this comparison. The three were flagged “c” and notes put in the headers of the relevant files. A simple offset of +0.16ml/l looks reasonable, but there is a tighter fit against DO:

DOX-BOT = 1.0101 * DOX-CTD + 0.1109

During 2007-11 when this sensor was used the fit was:

DOX-BOT = 0.0968 * DOX-CTD - 0.0142

During that cruise there were issues of anoxic sampling and the DO range was quite different. For both cruises the average correction is smaller than we are used to seeing.

The fit versus time shows little variation when the same outliers are excluded as for the other fits.

The lowest DO values during this cruise were about 1.7ml/l so hypoxia is not an issue.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors were recalibrated since previous uses. SeaBird reported that there was no drift in the primary calibration and little in the secondary. Problems noted in the past in the salinity from the secondary conductivity were found to have been largely due to the temperature sensor used with it; a different secondary temperature sensor was used for this cruise.

- This DO sensor is a new one that has been used only twice before, for 2006-35 and 2007-01. The results of 2006-35 were considered unreliable because the stops were too short for the sensor to equilibrate.

For 2007-01 the fit was found to be

CTD-BOT = 0.9893 DOX-CTD -0.0142

with an offset of -0.047ml/l applied as a second calibration for the early casts and -0.156ml/l for later casts. Low oxygen water was sampled and an anoxic plug was used prior to use which probably accounted for most of the drift during the cruise.

Historic ranges – There is very little local climatology for this region. For the few casts with local climatology available, the temperatures are mostly within the ranges while the salinity often falls below the historic minima. This is presumed to be because the climatology data are from Queen Charlotte Strait whereas these casts are close to shore in areas close to shore and fresh water sources.

##### 13. SHIFT

Fluorescence

The usual method for finding what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. A shift of 1s has been used in most other cruises and looks appropriate for 2007-11 as well.

All data were shifted by +24 records and a few casts were checked afterwards and found to be shifted appropriately. (Output: SHFFL)

Dissolved Oxygen

Based on the results of other recent uses of this instrument, tests were run using +70 to +120 records on three casts. Judging by how the downcast vs upcast trace offset compares with that of temperature and how features matched, the best choice varied from one depth to another, but overall a choice of +110 records appears to be to be best. A choice of +90 records was found best for when this equipment was used during 2007-01 but that included much deeper sampling and less variability. SHIFT was used to advance the DO channel by +110 records (about 4.5s).

After this step a few casts were plotted and the results are reasonable though there is clearly hysteresis.

Conductivity

Tests were run on three casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be +0.5 for the primary sensor and -0.5 for the secondary. All casts were shifted by those amounts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There are warnings for cast #20 that correspond to pressure spikes in the data that were not caught by WILDEDIT because they occur in pairs rather than single wild records. The output of DELETE misses a lot of data. So a text editor was used to replace the bad pressure records with pad values; this had to be done on the CLN file before running SHIFT. SHIFT and DELETE were then rerun on that file. There were still lots of warnings, but the output file did not skip as much data. This was a problem cast with no BL file created and communications problems are noted in the log.

##### DETAILED EDITING

The secondary sensors were selected for archiving because the data are slightly less noisy in T-S space and the secondary salinity is closest to the bottles.

CTDEDIT was used to edit the data. On-screen plots of descent rate and pump status were used to help distinguish local variability from corruption by shed wakes.

The following casts required fairly heavy editing: 48.

The following casts required no editing: 14, 18.

All other casts required only lightly editing mostly at the surface and bottom of casts.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

Salinity will not be recalibrated.

File 2007-11-recal1.ccf was used to recalibrate dissolved oxygen using the results of section 11:

DOX_CTD_NEW = 1.0101 * DOX-CTD_ORIGINAL + 0.1109

That correction was applied to the SAM and MRGCLN2 files. COMPARE was rerun for dissolved oxygen and showed that the recalibration worked well. (See 2007-11-dox-comp2.xls.)

(This step was first done on temporary files because nutrients were not yet available. When the nutrient data became available, this step was repeated on the new MRG files with nutrient data in them to ensure no errors had crept in – no problems were found.)

The edited downcast files, EDT, were recalibrated using 2007-11-recal1.ccf. (Output: COR1)

The COR1 files were clipped to 100db and set aside to be processed later for Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time, and the comparison with titrated samples handles those due to drift in calibration. But there remains an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged (0.5db bins) & thinned to the usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX and outliers were excluded. A fairly flat fit was found by comparing differences against pressure values; when outliers were excluded the average difference showed the CTD values to be low by 0.1368ml/l. (See 2007-11-dox-comp3.xls.) The CTD DO data in the thinned files were recalibrated using file 2007-11-recal2.ccf to apply an offset +0.1368ml/l and COMPARE was rerun. The results were satisfactory. (See 2007-11-dox-comp4.xls.)

File 2007-11-recal2.ccf was then applied to the AVG files to raise the SBE DO values by 0.1368ml/l, but not to the bottle files since this error does not apply to data collected while stopped. (Output: COR2)

##### Special Fluorometer Processing

All COR1 files were put through CLIP to produce files with data to 100db only for the use of A. Peña.

Those files were then bin-averaged (1/4db bins), recalibrated suing 2007-11-recal2.ccf and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD1) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data were put through a median filter with fixed width 11, before bin-averaging.

The recalibrated CTD bottle files (SAMCOR1) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered, roughly:

- ±0.3ml/l from the surface to 250m

- ±0.1ml/l below 200m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~85% to 110% with the highest values in DO saturation occurring at sites of high fluorescence values.

##### Final Bottle Files

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE,  Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### 22. Producing final files

The following steps were used to look for processing errors:

a) Track plots were produced using the final CTD and CHE files.

b) A cross-reference listing and header check was produced for CTD and CHE files.

The header checks (2007-11 and 2007-11- were saved to provide a list of channels for the CTD and CHE files to be archived.

The sensor history was updated for the CTD pressure, conductivity and dissolved oxygen sensors.

23. Thermosalinograph Files

A Thermosalinograph was used during this cruise by Rich Pawlowich from UBC. The raw data was not available but a file from Dr. Pawlowich with recalibrated salinity and CHL-a channels was obtained (2007-11-0001.dat.). Note that this file contains data from March 13, 2007 to April 29, 2007. The file was edited to remove headers and the MATLAB DATE and CONDUCTIVITY channels and to change the format of the date and time from separate year, month, day, hour, minute and second channels to YYYY/MM/DD and HH:MM:SS channels. ADD IOS HEADER was then used to convert the file to one with IOS HEADERS. The file was extremely large with sampling approximately every 3s, so the file was thinned by choosing every 10th point. That file was then put through HEADEDIT to change formats and add a comment; the output was 2007-11-0001.HDR. That file was too large to enable track plot to work, so FRACTURE was used to divide it into 3 files, which were then named 2007-11-0001.tob, -0002.tob and -0003.tob. Track plots were then produced and added to the end of the report.

Particulars from logs

1.  Bottles fired for testing only – no sampling done.

14. May have touched bottom while drifting.

20. Problems with bottle firing – no samples.

PROCESSING SUMMARY – SBE19

An SBE19 was strapped to the SBE911+ for 3 casts.

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The cast names for the 3 SBE19 casts were named using 3 digit event numbers, presumably to distinguish them from the SBE911+ casts, but since a 4-digit number is required in IOS SHELL a 9 was used as the first digit, so 9049, 9050 and 9051 are the corresponding casts to 0049, 0050 and 0051.

There was no con file with the data for the SBE19 so one was prepared using the latest calibrations and named 1294-Dec05.con. Some files had been converted at sea and the same results were found in a test conversion using this con file.

After conversions the times were compared with those of the SBE911 and they differed significantly, but since the only purpose was to compare T and S, no attempt was made to fix the times.

- Conversion of Raw Data

- The raw data were converted using conversion file 1294-Dec05.con.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. ALIGNCTD

RUN 1: Tests were run on a few casts using advancements of from 0.5s to 1.5s to temperature. A setting of +1.3s gave had the best overall effect on reducing salinity spikes.

RUN 2: This step was skipped and alignment done later in IOSSHELL using SHIFT. This is the method that was used for 2005-19 and 2005-31 except that conductivity was shifted. It is easier to fine-tune later.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

Plots were made and the data look fine. No effort was put into adding times and positions since these data are only for the purpose of calibration of the instrument. There are SBE911+ casts at the same locations and times so it is not anticipated that anyone will want to use this data.

The pressure was checked by plotting conductivity near the surface. It appears that the conductivity becomes non-zero at about -0.1db. On the upcast the conductivity drops suddenly at about -0.2db. In the most recent previous use the average surface pressure was +0.2db and the negative values were associated with very low salinity. No adjustment was made. Pressure should be considered ±0.2db.

11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

12. CTDEDIT

CTDEDIT was used to remove records from the surface for all 3 files and from the bottom of 2 of them.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and no problems noted. No further editing will be done.

14. Intercomparisons

PLOTS

Plots were made of T and S versus pressure for both CTDs. The strongest impression is that there is a significant alignment difference. There are two different issues to be considered. One is the simple matter of pressure range and whether one or the other is wrong. Looking at the top of the casts, the SBE19 has pressures that are lower than the 911+ by 2, 2.6 and 2.7db. It is believed that it was no more than 1db higher. An initial study of the surface pressures of the SBE19 would suggest that pressure is low by at least 0.4db based on “in-water” values for negative pressures. So a recalibration of pressure by something adding between 0.5db and 1.5db is reasonable; the difference was split and recalibration of the SBE19 was applied by adding 1db.

The second issue is the differences in the traces due to transit time or other non-pressure errors. The SBE19 is unpumped and in the wake of the SBE911, so we expect some differences, but these differences look similar independent of the local gradient. Looking for the depth at which a feature occurs shows that the SBE19 is seeing things TOO SOON in some sense. The alignment step applied earlier advanced the temperature, so this appears to have been too large an adjustment. It was decided to return to the initial processing and skip the alignment step. When plots were then rerun the traces were closer with features about 1 to 2m apart whereas they were more than 2m before.

Tests were run to see if shifting the temperature could fix the alignment. A setting of -0.3 records did make the features line up well with those in the 911+ data. However, it made the salinity look worse if salinity was recalculated. If salinity was not recalculated it looks ok. However, the temperature and salinity vary little for these casts, so this might not be the best test. It may be that fine-tuning the conductivity alignment with recalculation of salinity is required as well for the best salinity. There was just one salinity bottle for these 3 casts and it is deeper than the SBE19. However, since the SBE19 salinity is higher than the bottle 2db below it, it is safe to presume that the SBE19 salinity is high.

COMPARE – The DELETE files from the SBE911+ and from the SBE19 were edited, bin-averaged (1db) and then thinned to every 3db. To enable the comparison of the files in COMPARE the names were changed so that instead of 2007-11-9049.thn and -0049.thn, they were -0049.thn and -0049.thn1. They were then imported to COMPARE. Plots were produced of the differences versus pressure, salinity (secondary for the 911+) and file pair numbers. When a few outliers were removed based on residuals, T was found to be low by 0.0064Cº. There were more outliers in the salinity differences but when enough were removed to achieve a flat trendline, the SBE19 salinity was found to be high by 0.0127. The error in the temperature would account for almost half of the salinity error. While there is some uncertainty in the alignment that might affect salinity values, this effect would be much smaller than the error in salinity so is not considered very significant.

When plotted against time there is some suggestion of drift, but there are only 3 casts and the second two look very similar. It is noted that most of the salinity outliers were from the beginning of cast #49. Separate plots were made of T and S differences for each cast. The average difference for casts #50 and 51 are very close to each other at -0.0058 Cº and -0.0056 Cº and 0.0127 and 0.0131 salinity units, so it was decided to use the average of these two for recalibration purposes. This closer examination of cast #49 suggests that the differences versus pressure do not settle down to the expected flat trendline until about the 85db level, which was at scan #426. For cast #50 the SeaCat was at 2.6db for scan #426 having just completed its soak. For cast #51 the soak time was about 50% longer than for cast #49. For cast #49 the soak time was about 2 minutes. Perhaps a 3 to 4 minute soak is needed for this instrument. . The manual recommends 1 to 2 minutes soaking except where there are large differences between the temperature of the place where the CTD was stored and the ocean temperatures. The air temperatures reported for Campbell River on March 31, 2007 ranged from 5.5 to 11.3 Cº and closer to the higher end during these casts, so not notably lower than the water temperatures. If the CTD was stored inside, the error due to it being warm would be of the opposite sign, so this does not look like a problem due to particular conditions at the time.

Another possibility is that the differences for the first cast could be due to its location in a strong mixing regime, but the differences reduce quite gradually over the top 85db and variability in the 911+CTD settles down below 30db to a level similar to the other casts.

16. Producing final files

CALIBRATE was used to increase the temperature by 0.0057 Cº` and lower the salinity by 0.0129 and plots made to ensure that the corrections were appropriate. The temperature differences are reduced for cast #49 but are still significant. The differences for casts #50 and 51 are very similar. The thinned files were recalibrated and COMPARE rerun and the results show these settings are effective.

No final files were prepared since they are not intended for archiving.

17. CONCLUSIONS RE SBE19

Leaving alignment of the temperature signal to later in the processing is a better idea. The tests are easier to run in IOSSHELL and a better judgment can be made about in the choice of parameters in SHIFT. For this data shifting temperature by approximately -0.4 records aligns it with the 911+ temperature features. If salinity is recalculated this has a bad effect on salinity, so either don’t recalculate, or realign conductivity to smooth salinity and then recalculate salinity.

The pressure appears to be low by about 1db.

The temperature appears to be low by about 0.056Cº.

The salinity appears to be high by about 0.127, but about half of this error is due to the temperature error.

The SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature.

# Institute of Ocean Sciences  CRUISE SUMMARY

| Cruise ID#: 2007-11 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 26 March 2007 End: 1 April 2007 |  |  |  |  |  |  |
| Location: Broughton Archipelago |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Stucchi D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |
| 2 | SEABIRD | 19 | 1294 | No | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550        Cruise ID#:		2007-11

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23/12/06 | Factory “ |  |  |
| Conductivity | 1766 | 23/12/06 | “ |  |  |
| Secondary Temp. | 2106 | 21/12/06 | “ |  |  |
| Secondary Cond. | 1729 | 23/12/06 | “ |  |  |
| Transmissometer | 498DR | 28/12/06 | Factory |  |  |
| SBE 43 DO sensor | 0766 | 17/10/06 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| Pressure Sensor | 75636 | 04/06/1999 | Factory |  |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 /1294        Cruise ID#:		2007-11

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 29/11/06 | Factory “ |  |  |
| Conductivity | 1294 | 29/11/06 | “ |  |  |
| Secondary Temp. | 1294 | 06/12/06 | “ |  |  |