## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 11-Jun-2013 | Added Iron profile files with cast numbers 8xxx from Keith Johnson’s spreadsheet file which can be found in the cruise .DOC directory. |

## PROCESSING NOTES

Cruise: 2007-12

Agency: OSD

Location: North-East Pacific

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: May 22, 2007 – May 29, 2007

Processed by: Germaine Gatien

Date of Processing: 25 September 2007 – 27 October 2007

Number of original CTD casts: 66	Number of CTD casts processed: 66

Number of bottle casts: 	30

Number of bottle casts processed: 29 (No sampling from cast #107)

TSG files: 4

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Chelsea/Seatech transmissometer (#953), an SBE 43 DO sensor (#1117, probably on the primary pump), a Seapoint Fluorometer (#2845) with a 10X cable (probably on the secondary pump) and an altimeter (#1252). The deck unit was an SBE 11+ model (#0424) and there was a mid-ship winch. The guesses about the pumps on which the external sensors were mounted are based on how they were done for cruise 2007-13 which followed.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order and there were some sampling notes.

There were serious problems in both primary and secondary temperature and conductivity data, and hence in the salinity and dissolved oxygen data, probably due to a pump cable leading to flow problems. The pump cable was replaced part-way through cruise 2007-13 that followed and the problem disappeared. A lot of data from below 1000db was removed. The data that remain below 1000db should be used with care since it may have been affected in a less obvious fashion. The upcast data are also bad in large sections.

There were duplications of some sample numbers. There were no notes of sample numbers in the Daily Log Book during the time period when the errors occurred, no doubt contributing to the errors. There were apparent errors in the labelling of a few salinity samples.

The altimetry was unusually noisy so that the algorithm that calculates altimeter header entries often worked poorly. Bad values were replaced in the headers where a reasonable estimate could be made, or removed where this was not possible. Header comments were adjusted where appropriate.

As of Oct. 27, 2007 ammonia samples had been analysed but the data were not in final form, so were not ready for addition to the chemistry files.

Temperatures between 80 and 400db just offshore of the shelf break were often below the historic minima for the region. The values look much like the stations to seaward where temperatures are within the historic range. This is assumed to be due to the movement of offshore water closer to shore and not due to instrumental error and probably reflects the fact that the ranges are inadequate.

The dissolved oxygen data in the CTD files should be considered

- ±0.5ml/l from 0 – 50db (except in very high gradient areas where ~±1ml/l)

- ±0.25ml/l from 50 – 250db

- ±0.1ml/l from 250 – 1200db

- data below 1200db are considered unreliable by the manufacturer

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets and sampling notes. The same equipment was used as during 2007-13 which immediately preceded this cruise, except that no PAR sensor was used for this one.

A number of problems were noted from the log. The sounder was not working well and the CTD hit bottom during cast #37. There were problems in the headers of casts #1, 16 and 37.

Titrated chlorophyll, nutrients and salinity data were obtained in spreadsheet format.

The oxygen files were provided in individual ADD files with flags and comments. There were initially a lot of problems with these files, but they were corrected.

A pressure test was done on deck and the reading was 0.26db.

The cruise summary sheet was completed. The histories of the conductivity, DO and pressure sensors were obtained. The calibration constants were checked for all instruments and a pressure offset was entered. The file was saved as 2007-12-CTD.con

##### Conversion of Raw Data

Data were converted using the configuration file listed above. Errors in the headers of casts #1, 16 and 37 were fixed after conversion.

A few casts were examined and all expected channels are present. During 2007-13 the upcast temperature traces were much noisier and further apart than in the downcasts, and that is also seen in this data. The fluorescence, oxygen and transmissivity data have the usual appearance and the conductivity channels are close during downcasts but are quite different during upcasts. The altimetry is noisy but looks usable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. Errors in the headers of casts #1, 16 and 37 were fixed after conversion. Only 65 rosette casts were produced but the sampling notes give 66 as the number. This is probably because cast #4 is shown as a rosette cast in the log book, but no bottles were fired and no rosette file produced.

The rosette files were then converted to IOS SHELL files and renamed with extensions BOT.

All BOT files were plotted and outliers were found for cast #52. CTDEDIT was used to remove some data from the secondary salinity channel for Niskin #2. The output file ED1 was copied to BOT.

The data from casts #70 and 73 are very noisy below 1000db but it is not a matter of a few outliers. Editing will not help.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on many casts using settings (α = 0.02, β=7), (0.03, 7), (0.02, 9), (0.03, 9) and (0.0245, 9.5) to see what settings look best for this cruise. Overall, (0.03, 9.0) looked best for the primary and (0.02, 7.0) for the secondary conductivity channel. During 2007-13 using the same equipment, the best choices were (0.03, 9.0) for the primary and (0.02, 9.0) for the secondary and for 2007-14 it was (0.03, 9.0) for the primary and (0.02, 7.0).

CELLTM was run using (0.03, 9.0) and (0.02, 7.0).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on some deep casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 34 | 600 890 1400 1900 | X noisy -0.0008 data bad data bad | +0.0001 +0.00008 data bad data bad | +0.0023 +0.0022 data bad data bad | High, Noisy |
| 69 | 500 890 1400 1900 | -0.002 noisy -0.0007 noisy data bad data bad | +0.0001 +0.0001 data bad data bad | +0.0014 +0.0017 data bad data bad | High, fairly steady |
| 86 | 500 890 | -0.0009 -0.0005 | +0.0001 +0.00007 | +0.0015 +0.0014 | High, noisy |

In examining the differences it was found that the temperature, salinity and conductivity data are frequently bad below 1000db during both downcast and upcast. Temperature and salinity gradually shift to higher values and then suddenly revert to expected values. Examination of T-S plots shows problems are in both channel pairs. The other pumped channels look ok, but since the oxygen value depends on temperature those values must also be unreliable. The problems look very much like the flow rate problems observed during the cruise that followed, 2007-13. For details on the problem see the processing report for that cruise, 2007-13_proc.doc. Note that the problems did not occur during 2007-14, so the replacement of a pump cable during 2007-13 was probably the solution.

There is no obvious pressure or time dependence, but given there is little reliable data below 1000m the conclusion is weak. Another problem noted during 2007-13 was with shifts in values occurring below 2000db, but none of these casts is deep enough for that to be seen.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run and the only problem noted is that the time in the log book is one hour different from that in the header for the first cast. This is assumed to be due to the fact that the computer clock was in error, since the System Upload and NMEA times differ by 1 hour for the first two casts. The NMEA time is believed to be correct. The speed between casts #107 and 109 is fairly high, but this was a long steam so the speed is reasonable.

The cruise track was plotted and no problems found. There was one odd jog in the track, but the positions are correct according to the log and the station name indicates that it was intended to be off the CS line.

The average surface pressure is 3.4db which is reasonable for the Tully. The lowest readings are associated with very low salinity as expected.

The altimeter readings from the header were exported to a spreadsheet and a few casts were checked. There is a lot of noise in the data for many casts, and the algorithm worked badly for some casts and looks a little low for others. This will be checked later in the processing and estimates will be made and the headers corrected.

##### BOTTLE FILE PREPARATION

Some sample numbers were repeated. For cast #88 the samples named 278 through 288 have been renamed as 9278 through 9288. This affects salinity, dissolved oxygen and nutrient data, but not titrated chlorophyll data. Sample #277 was also used for both cast #77 and a loop sample. It is noted that there were no sample numbers recorded for several casts in the Daily Log Book for the time during which the errors occurred. Cast #88 will need rearranging from time to time since some merges are based on sample # and some on bottle #.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Cast #1 had some bottles for which no sample numbers were recorded, so they were removed from the list. Cast #107 had no sample numbers, so all bottles were removed from the list. The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

The salinity spreadsheet (2007-12 sals.xls) included loop data and rosette data. The file was edited to change headers to standard names and remove unnecessary columns, and then saved in two files, one containing rosette data only (2007-12-sal.csv) and the other loop data (2007-12-sal-loop.csv). There were neither sample numbers nor event numbers for the salinity bottles; the station name and bottle number were recorded from the labels. This was used to determine sample numbers and event numbers, but it is clearer and more efficient if the sample numbers are put on the labels, especially in cases where the rosette log sheet is not completed. The salinity spreadsheet was then rearranged on sample number and then converted to individual SAL files.

Dissolved oxygen files (*.add) were provided with a flag channel and comments entered in the headers. There are many problems with the OXY files that led to problems in the creation of the ADD files so some files had to be sent back for correction.

The nutrient spreadsheet was simplified and saved as 2007-12nuts.csv.  A loop file was prepared separately. Extraneous columns were removed and header names were changed to standard format. Data were sorted on sample number. File 2007-12nuts.csv was then converted to NUT files.

Extracted chlorophyll data were obtained and saved as 2007-12CHL.csv. The file was edited to remove extraneous lines and columns, header names were changed to standard format, and two loop samples were removed and saved separately as 2007-12-chl-loop.csv. 2007-12CHL.csv was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

There is a file with Ammonia data. It is not ready for inclusion in the chemistry files as the duplicates have not been averaged and no flags have been assigned. It can be added later.

##### 12. COMPARE

Salinity

COMPARE was run with pressure as the reference channel. During the first run a few errors were found and corrected in the SAL files:

- For cast #70 COMPARE suggested that two samples had been reversed (labels: LBP 7-18 and LBP 7-17). When reversed the results were much better, so the salinity files were changed but “c” flags were inserted and a note to indicate these were not the original values from the analyst. No problems were noted in other samples. This was probably just a typo rather than sampling from the wrong bottles, since the samples were recorded out of order in the spreadsheet – i.e. values are probably in right order, just labels reversed.

- For cast #93 the rosette log indicates that sampling was from Niskin bottles #2, 6 and 9, but the labels indicate samples were from bottles #4, 6 and 9. A “c” flag was attached but later removed since COMPARE suggests #4 is correct. A note was left in the header explaining that there is some confusion.

- For cast #106 an error was found in the assignment of sample number and this was corrected.

COMPARE was then rerun. The only extreme outlier was the surface bottle from cast #106 with the CTD salinity lower than the bottle by 1.3. This bottle was flagged “d” as it appears there must have been a misfire or the sample came from a different bottle than shown on the records; the salinity looks like it is from between ~75 and 100db. The chlorophyll, dissolved oxygen and nutrients also look suspicious. When bottles from above 100db are excluded there are 5 significant outliers and they all have large standard deviations in the CTD salinity. They are from 1000db to 1500db (casts 31, 70 and 73) and include CTD data that are clearly bad. When those points are excluded, the primary salinity is low by an average of 0.003 and the secondary by 0.001.The trendlines are flat and there is no obvious temporal variation.

Dissolved Oxygen

COMPARE was run using pressure as the reference variable. The fit of differences versus dissolved oxygen provides a better fit than against pressure. The samples flagged “d” by the analyst all appear as outliers and the one from cast #58 was an extreme outlier. Because the latter sample was run without a stirrer, the value was replaced with a pad value as it is very clearly bad. None of the samples flagged “c” are severe outliers. Sample #321 from the surface of cast #106 was a severe outlier; it was flagged “d” as are all the samples from that bottle; it seems probable that there was a misfire.

The best fit is against CTD DO values though the fit against pressure is reasonably flat as well:

CTD-BOT = 1.0928 * DOX-CTD + 0.0291

During 2007-35 which preceded this cruise and for 2007-13 and 2007-14 which followed it, the fits were found to be, respectively:

CTD-BOT = 1.0522 * DOX-CTD - 0.05 (2007-35)

CTD-BOT = 1.0838 * DOX-CTD + 0.0601 (2007-13)

CTD-BOT = 1.0821 * DOX-CTD + 0.0293 (2007-14)

There was a lot of anoxic sampling during 2007-35 so it is expected to have quite a different fit.

The fit against file pair shows no evidence of significant temporal drift.

Plots were made of DO_CTD and DO_BOT versus Salinity as a final check. The only outliers had already been flagged except for samples #9285 and 9286 from cast #88. The profile is very complex for that cast so that it is not obvious which sample might be the outlier, nor is it clear that there is not a problem with the CTD data since the sensor has a slow response time. So no further flags were added.

Fluorescence

COMPARE was run using the CTD Fluorescence and the Titrated Chlorophyll from bottles. When all data were plotted there was a lot of scatter with FL roughly 112% of CHL. That ratio was higher offshore where there were lower chlorophyll values and lower in the near-shore where the CHL was higher.

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values from 0.6 to 1.6s were found with most at the lower end of that range. The shift of +24 records (1s) which has been used in other recent cruises is thus appropriate. Fluorescence was shifted by +24 records.

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because of hysteresis there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly; matching features from downcast to upcast is useful. There were not many such features in this data. For 2007-13, +130 was used, so values from +110 to +150 were tried for this data. For other cruises in the past year the choices were +90, +90, +110, +120 and +130 records with a gradual trend to higher values. A choice of +130 looks reasonable for this data. SHIFT was run using +130 records.

Conductivity

Tests were run on 2 casts with few or no stops for bottles using shifts between -1s and +1s and T-S plots were prepared to compare the results. A setting of +0.6s worked best for the primary conductivity and -0.6s looked best for the secondary. All casts were put through two runs of SHIFT using the settings of +0.6 and -0.6s.

##### DELETE

Before running DELETE the initial 5000 and 5500 records were removed from casts #72 and 82 respectively because there was an initial soak period during which the CTD was lowered a little and then returned to the top before the full lowering occurred. DELETE would probably select the data from the first lowering so it was removed.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #4 which had little data in it. Something had gone wrong during the final SHIFT run and after that was rerun, and the cast run through DELETE again, there were no further warnings.

##### DETAILED EDITING

The secondary salinity is closest to the bottles as it was when the same sensors were used for all other cruises since they were last recalibrated, so secondary sensors were selected for editing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

All casts required some editing and the editing was very heavy below 1000db for 31, 32, 33, 65, 66, 69, 70 and 73.

It is believed that these patches of bad data were due to a pump cable problem that was resolved during the cruise that followed this one.

Note was made of the editing details in the headers of the relevant files.

At this stage the altimetry readings in the header were edited in the SAMAVG and EDT files. To determine an appropriate entry, plots of altimetry near the bottom of the edited files were examined. In many cases the header entry was appropriate but especially in the first part of the cruise there were many cases where the entry is clearly wrong, usually too low. Such entries were removed and where possible replaced and the note explaining the header calculation was replaced with an appropriate note. If there was no useful data, the entry was removed. For the SAMAVG files the value found for the EDT files was used with an adjustment, where appropriate, to allow for the bottom bottle being a little higher than the lowest data from the full files.

##### Initial Recalibration

File 2007-12-recal1.ccf was to apply the following correction to the CTD Dissolved Oxygen channel:

CTD corrected = 1.0928 DOX-CTD + 0.0291

COMPARE was then rerun to check that the results were as expected and they were. (See 2007-12-dox-comp2.xls.) The same DO calibration was applied to the EDT files.

##### Special Fluorometer Processing

The COR1 files were clipped to 100db and processed separately for A. Peña. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other Comparisons

Previous experience with these sensors –

Salinity: The primary sensors were found to be low by 0.0015, <0.001 and 0.0018 during cruises before this one and 0.007 and 0.006 during cruises in May-June. The latter two of these were ones with large patches of bad data. The secondary sensors were found to be low by < or ~0.001 for the 3 cruises before this one and low by 0.004 during two cruises that followed, those two being the ones with pump problems.

Dissolved Oxygen: This sensor has been used for many other cruises since it was last recalibrated. The calibration for this cruise is fairly close to that of 2007-13 and 2007-14 which followed, but quite different from the cruise that immediately preceded it. That cruise sampled anoxic waters so significant differences are expected.

Pressure: This pressure sensor has been drifting slowly and during all cruises in 2006 and 2007 an offset of +1.2db was found appropriate.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursions for salinity were a few near-surface casts with salinity slightly below the minima. The temperature had many values below the historic minima mostly between 80db and 400db. These occurred primarily mid-way along the LG and LBP lines. The most offshore casts have very similar bulges of low temperature water, but the ranges used for those casts are different than the ones used for the casts closer to shore. It seems likely that offshore water has moved closer to shore than usual. The excursions from the ranges more likely reflect inadequacy in the ranges than in the CTD.

##### Final Calibration of DO

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. Data from below 1200m were excluded plus outliers identified by residuals. Neither the fit against pressure nor that against DO are flat. The offsets are smaller at the top and bottom of the DO range presumably because the gradients are low there. The average difference for all data in the fit is 0.18ml/l. During 2007-13 the average difference was ~0.09ml/l. For this cruise a polynomial fit against DO value looks satisfactory so was applied:

DOX_CORRECTED = 0.0143* DOX*DOX + 0.8844*DOX - 0.054

COMPARE was rerun after applying the above correction to the bottle files. The results were good. There is no significant time-dependence. (See 2007-12-dox-comp3.xls and 2007-12-dox-comp4.xls.)

The recalibration was also applied to the AVG and CLIP files. (Output: COR2 and CLIPCOR2)

The clipped files were recalibrated. One set was then bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were filtered before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added and the channels reordered to put the two SBE DO channels together.

At this point all data were plotted and it was noted that the transmissivity had a few zero values at depth in casts #88 and 109. Pad values were substituted for the zero values in those two casts (around 350 and 360db respectively.)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.5ml/l from 0 – 50db (except in very high gradient areas where ~±1ml/l)

- ±0.25ml/l from 50 – 250db

- ±0.1ml/l from 250 – 1200db

- data below 1200db are considered unreliable by the manufacturer

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values for most casts were between 105% and 125% with values as low as low as 80% near the mouth of Juan de Fuca Strait and up to 145% in Saanich Inlet. These values do not suggest any problems with the data.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats, to change the project name, geographic area and units and to add a comment about quality flags and analysis methods. (Output: CHE)

##### 22. Thermosalinograph Data

a.) Checking calibrations

There were 2 files containing TSG data. A report was printed for the con file, the fluorometer calibration was corrected and serial number were entered and the primary temperature calibration was corrected; the resulting file was saved as 2007-12-TSG.con.

The history of the sensor was obtained.

b.) Converting to IOS Headers, adding position headers and time channels, preliminary checks

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC.

Time-series plots were produced. The salinity data had just a few spikes that are not associated with spikes in any other channel including the flow rate. Editing will be required to interpolate or remove these spikes.

The flow rate is higher than usual (~1.4) which was true of 2007-13 as well.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within .3db of 4db and exported to a spreadsheet. The TSG files were opened in EXCEL, median and standard deviations were calculated for temperature, salinity and fluorescence, and the file was then reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2007-12-ctd-tsg-comp.xls). The positions were compared and were very close, with average differences for both latitude and longitude of <0.0001º and no difference greater than 0.0003º so the clock appears to have worked well.

This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T0 The average difference over the whole record shows the TSG lab temperature to be high by about 0.14.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There were 34 casts that could be used. Graphs were prepared comparing the TSG temperature, salinity and fluorescence with those of the CTD.

Because the TSG intake and the CTD data are not from exactly the same depth, it is best to compare casts for which the surface gradients are low; for this cruise the mixed layer was shallow for all casts, but based on an estimate of mixed-layer depth from the CTD files, 4 casts were identified as best for comparison with the TSG data. Using all casts the intake temperature was found to be high by 0.05C˚ but using only the 4 best-mixed ones it is high by only 0.004C˚. Using the 10 casts for which the TSG temperature had the lowest standard deviation over 2 minutes, the TSG is high by less <0.01 C˚.  Given that the CTD data and the TSG data may not be from exactly the same depth or time, this is a good agreement.

When all casts are included the TSG salinity is lower than the CTD salinity by an average of 0.041 units. When only the 4 best-mixed casts are used it is low by 0.031 units but when the 10 with the lowest standard deviation in the TSG data are used it appears low by 0.052 units.

The ratio of TSG fluorescence to CTD fluorescence ranges from 1.1 to 4.7, an average of about 2.0 and an average of 1.4 for the 4 best-mixed casts. When the 10 records with the lowest standard deviations in TSG fluorescence are used then the average ratio is 1.7. The ratio was not dependent on the TSG fluorescence value. (See 2007-12-ctd-tsg-comp.xls)

- Loop Bottle Comparisons There were 3 loop bottles but the time of the third has a question mark by it, and it does not seem to fit the data from the time given. Checks were made to see if the data made sense if the date was wrong or the time was local, but the data still look wrong.

The salinity values (using a median over a 2-minute window) were compared with the TSG files and for one sample the TSG salinity was high by 0.01 and for the other it was low by 0.04.

There were 2 loop chlorophyll samples and the TSG FL is found to be approximately 1.7 and 0.7 times the loop CHL. (See 2007-12-chl-loop-TSG-comp.xls.)

- Calibration History This was the first use of the TSG after recalibration in April 2007. It was used for cruise 2007-13 in May/June when the temperature error was found to be dependent on the intake temperature (Lab Temperature correction = 0.01*Intake Temperature – 0.23). For the temperatures seen during this cruise that would give lab temperatures high by 0.12 to 0.13C˚ and the salinity was found to be low by 0.07. During 2007-14 the salinity was found to be low by 0.05. The lab temperature was found to be high by 0.12 or 0.13C˚.  Both cruises encountered difficulties with TSG and CTD, so the data are not as reliable as usual.

Conclusions

The intake temperature is higher than that of the CTD by less than 0.01Cº and by about 0.004Cº when only well-mixed casts used. This is as close as we can expect given the sampling depths and time of sampling each is not exactly the same. The differences between the lab temperature and the intake temperature averaged about 0.14 Cº. The intake temperature will be archived.

The TSG fluorescence is higher than the CTD fluorescence by an average factor of 1.7 with the ratios for most casts between 1.4 and 2. The ratio of TSG fluorescence to loop chlorophyll is 1.7 for one sample and 0.7 for the other.

The TSG salinity is lower than the CTD salinity by ~0.05. It is lower than one loop sample by 0.04 and higher than the other by 0.01. The results of two cruises that followed this one indicated that the salinity was low by 0.07 and 0.05 but there were severe problems with both the TSG and CTD during those cruises. The addition of 0.05 units looks appropriate.

e.) Editing

The time-series plots were examined and there were a few single-point salinity spikes with no associated temperature spikes in the first and fourth files. A graphical editor was used to clean the salinity spikes.

f.) Recalibration CALIBRATE was used to apply an offset +0.05 to channel Salinity:T0:C0.

g.) Preparing Final Files

REMOVE was used to remove the following channels: Scan_Number, Temperature:Difference, Conductivity:Primary, Flag and UPloy0 (flow rate).

HEADEDIT was used to was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, add the depth of sampling to the header and change the names Temperature:Primary to Temperature:Lab and Temperature:Secondary to Temperature:Intake.

REORDER was used to put Temperature:Intake before Temperature:Lab so that programs will selectively pick that channel.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars from log and cruise notes:

1. Wrong cruise # in header. Test cast. All bottles fired.

2. CTD down to 25m, up to 18 for long stop, back to 20m to close 2 Niskin bottles, then the cast was run.

16. Scrambled SeaBird position/time headers.

37. Glitch with the cable connecting power to the CTD and then in getting pump started. Hit bottom. No station name. Should be LC11.

39. Oxygen chemicals are delivering bubbles to the sample.

66. Large difference between two CTD temperature channels.

72. Initial lowering to 30m with pumps off, returned to surface and then ran full cast.

82. Large pressure spike at beginning of file.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2007-12 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 22 May 2007 End: 29 May 2007 |  |  |  |  |  |  |
| Location: West Coast Vancouver Island |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Yelland D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443        Cruise ID#:		2007-12

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23Dec06 | Factory “ |  |  |
| Conductivity | 1766 | 23Dec06 | “ |  |  |
| Secondary Temp. | 2374 | 21Dec06 | “ |  |  |
| Secondary Cond. | 1729 | 23Dec06 | “ |  |  |
| Transmissometer | 953DR | 23April07 | IOS |  |  |
| SBE 43 DO sensor | 1117 | 17/Oct/2006 | Factory |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| Pressure Sensor | 75636 | 04/Jun/1999 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2007-12

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 10/04/07 | Factory |  |  |
| Conductivity | 2487 | 10/04/07 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Temperature 2 | 2416 | 23/Dec/06 |  |  |  |
| Flow Meter | ? | ? | ? |  |  |