## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 2 April 2025 | ### Updated channel names & formats in TOB and loop files. GG/SH |
| ### 3 Dec 2024 | ### Added Data Description G.G. |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle lost during addition of HPLC. S.H. |
| ### 26 Nov 2020 | ### Added HPLC data. S.H. |
| 19-Dec-2014 | Fixed some flags and removed few lines with no sampling. G.G. |
| 6-Feb-2014 | Fixed some flags and replaced some bad bottle values with pad values. Prepared loop file. |
| 11-Jun-2013 | Added Iron profile files with cast numbers 8xxx from Keith Johnson’s spreadsheet file which can be found in the cruise .DOC directory. |
| 18-April-2010 | Added Lisa Miller’s Dissolved Inorganic Carbon and Alkalinity data to the rosette files. J.L. |
| Dec. 2, 2008 | Missing CHL & nutrients from surface bottles added to CHE files. Comment fixed for nutrient analysis method. |
| July 31, 2008 | Corrected CHE file for cast #68 – fixed errors in nutrient and chl samples. |
| April 8, 2008 | DMS data added to CHE files |
| Oct. 27, 2007 | TSG Salinity recalibrated. See note at end of section 24(g). |

## PROCESSING NOTES

Cruise: 2007-13

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: 30 May 2007 – 19 June 2007

Processed by: Germaine Gatien

Date of Processing: 5 July 2007 – 13 August 2007

Number of original CTD casts: 72	Number of CTD casts processed: 68 (1 upcast only, 3 aborted)

Number of bottle casts: 	66		Number of bottle casts processed: 66

TSG files: 2				Number of TSG	 casts processed: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Chelsea/Seatech transmissometer (#953), an SBE 43 DO sensor (#1117, on the primary pump), a Seapoint Fluorometer (#2845) with a 10X cable (on the secondary pump) and an altimeter (#1252). The deck unit was an SBE 11+ model (#0424) and there was a mid-ship winch.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order and notes from the chief scientist greatly simplified the processing job.

Chemistry files were prepared for casts #32, 41 and 67 but there was no sampling for IOS; sampling was done by researchers from other institutions, so these files were provided for them.

There were problems with salinity analysis at sea. Two different Portasals were used; the use of one was abandoned because of instability and the second produced more outliers than expected. The majority of samples were analyzed at IOS using an Autosal. When outliers were removed, the Portasal values did not look out of line.

During this cruise 23 bottles were fired rapidly at 3000db during 1 cast. Two bottles (from Niskin #20 and 23) are outliers; the others fall within ±0.001 salinity units. No other samples from the same Niskin bottles showed up as outliers. The processor is unable to assess whether bad bottles come from a single session of analysis. The Portasal files show the time of analysis, but the Autosal files do not. If it is possible to have this included in the Autosal records, it would be useful.

There were a number of problems with the CTD data:

- Through this cruise and the one that followed there were unusual shifts in the differences between sensor pairs at about 2500db. The shifts appear to be associated with a change in primary values. The shifts are small, but systematic and could be significant in deep waters. No explanation was found for the shifts. See APPENDIX I for a fuller discussion of this problem.

- A problem seen during the first half of the cruise (up to cast #50) appears to be caused by interruption in the flow to the sensors. It affected both downcasts and upcasts and both temperature and conductivity sensors. At depth bad data were seen clearly in editing and were removed, but in the upper 800m they may remain undetected because of higher local gradients. During cast #50 the data at depth became much worse. After that cast some repairs were made and the problem disappeared. See APPENDIX II for a fuller discussion of this problem.

- Throughout the cruise there were many small spikes in salinity which look like they are caused by small variations in alignment; most of the spikes have been smoothed in editing.

Because some of these problems produce systematic errors, the salinity must be considered to be of lower quality than usual even after metre-averaging, especially for casts 1 to 50.

The salinity analysis was made more difficult by the fact that the salinity sampling late in the cruise was only done at the surface and very deep. A few casts spread through the cruise with samples from a variety of depths would be useful especially when it is clear that there are problems with the salinity data.

There were many loop samples and CTD casts with which to recalibrate the thermosalinograph data. Tests show no obvious difference between comparisons while stopped and those while the ship was underway. For the first time it was found that the temperature recalibration was dependent on intake temperature; it has been noted often that the temperature correction varies with season, but we don’t usually have such a large range of temperatures sampled, nor enough calibration sampling to make a correction like this one. It will be interesting to note in future whether this recalibration equation (or a fine-tuning of it) may be suitable for all data. The comparison of thermosalinograph salinity with loop samples and CTD data showed a lot of scatter, so TSG salinity should be considered ±0.02.

The dissolved oxygen files had the wrong format for the flag channel – only 2 columns should be used.

The dissolved oxygen data in the CTD files should be considered

- ±0.5ml/l from 0 –75db

- ±0.8ml/l from 75 to 300db; data tend to be somewhat high in this high gradient region.

- ±0.25ml/l below 300db

- data below 800db tend to be low.

- data below 1200db are considered unreliable by the manufacturer.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets, a cruise report and notes from the chief scientist on particular issues that would affect the processing job. There are a number of problems noted including transmissivity spikes and occasional large differences between sensor pairs which may have been due to pump problems.

Titrated chlorophyll, nutrients and salinity data were obtained in spreadsheet format.

The oxygen files were provided in individual ADD files.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and a few errors were found and corrected. The pressure and PAR offsets were missing and the PAR date of calibration was wrong. Two files were created:

2007-13-CTD1.con – Configuration without PAR (for cast #1 only)

2007-13-CTD2.con – Configuration with PAR. (for all other casts, though many didn’t actually have the  PAR sensor mounted.) PAR will be converted for all but cast #1 even though there will be no signal for many casts; it is easier to remove PAR later.

##### Conversion of Raw Data

Data were converted using the configuration files as listed above. PAR was converted for all the CTD casts except for #1. After conversion corrections were made to station names and file names and files for casts 96 and 97 were deleted based on the chief scientist’s notes.

A few casts were examined and all expected channels are present and look reasonable when plotted. The upcast temperature and conductivity traces are much noisier and further apart than in the downcasts. In some spots the primary traces look noisier, in others the secondary. As mentioned in the log the altimetry is very noisy. Fluorescence, PAR and DO data look fine and while there are a lot of spikes in transmissivity there is little hysteresis and downcast and upcast data are very close.

Rosette files were converted using a start time of -2s and duration of 5s.  After conversion corrections were made to station names and file names and casts 96 and 97 were deleted based on the chief scientist’s notes.

The rosette files were then converted to IOS SHELL files. Those files were copied to *.BOT. All BOT files were plotted. CTDEDIT was used to clean spikes in salinity in casts #30 and 68. The output edited files were copied to *.BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on many casts using settings (α = 0.02, β=7), (0.03, 7), (0.02, 9), (0.03, 9) and (0.0245, 9.5) to see what settings look best for this cruise. Several choices looked very similar, but overall (0.03, 9.0) looked best for the primary and (0.02, 9.0) for the secondary conductivity channel. CELLTM was run using those values.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on some deep casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 18 | 500 1950 | -0.0008 -0.0010 | +0.00024 +0.0002 | +0.0034 +0.0025 | High High |
| 48 | 500 1950 | -0.0008 -0.0002 | +0.00019 +0.00023 | +0.0032 +0.0030 | High High, Noisy |
| 67 | 500 1950 3900 | -0.001 XN -0.0003 -0.0001 | +0.00025 +0.00023 +0.00027 | +0.0035 +0.0032 +0.0035 | High, Noisy High, Noisy High |
| 91 | 500 1950 3900 | -0.0008 -0.0002 -0.0002 | +0.00021 +0.00023 +0.00026 | +0.0032 +0.0029 +0.0034 | High High, Noisy High, Noisy |
| 93 | 500 1950 3900 | -0.00020 -0.00023 -0.00025 | +0.00020 +0.00031 +0.00025 | +0.0026 +0.0031 +0.0034 | High, Noisy High, Noisy High |

The differences are not large and there is no indication of temporal variation. Overall there is no significant pressure dependence, but there is a strange feature in temperature and salinity differences for the deepest casts. Below 2000db the differences shift by about 0.0006C˚ and 0.0004 salinity units. Sometimes there are several shifts between 2100 and 2600db though it is not absolutely clear that some of these are not due to local variations in descent speed. These are not large shifts, but they are unusual. For a report on further investigation see APPENDIX I.

There were notes in the log book about the primary temperature being quite different between downcast and upcast. Examining a few of these casts, sometimes the downcast data looks unreliable and for others the upcast looks bad. The descent rate was very noisy for some of these casts and will account for some differences. However, it looks as though at least one pump was not operating well. And for cast #32 both temperature sensors are full of spikes below 900db of the downcast, though the secondary spikes are larger. The primary pump was changed at that point and no further mention is made of large differences until cast #50. During that cast both sensor pairs have spiky data below 850db during both upcast and downcast. Further repairs were done after cast #50.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. Problems arose because there were 2 files that were too large to convert. STRIP was used on the CNV files to remove channels that are useful but not critical: Oxygen:Voltage:SBE,  Status:Pump and Descent_Rate.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run. An error was found in the geographic area in the headers – that will be fixed later when HEADEDIT is run on the final files. The cruise track was plotted and no problems found.

The average surface pressure is 3.6db, which is reasonable for the Tully. The mixed-layer depth calculation indicated only a few well-mixed casts; after examining a few plots the reason for this appears to be very noisy data at 4db which was the reference pressure chosen for the calculation. When 5db was used many more casts showed as being well-mixed. Nonetheless, for comparison with the TSG the best choices are casts 3, 18, 60, 68, 72, 75, 77, 80, 85, 90, 94 and 100 since the intake is at about 4.5db.

The altimeter readings from the header were exported to a spreadsheet and checking a few casts made it obvious there were problems. The algorithm is supposed to only enter a reading in the headers if the CTD gets within 15m of the bottom. There were entries for many shallow casts in deep water and this is clearly caused by many spikes in the altimetry. Every cast with a header entry was checked to ensure the signal looks believable based on the signal and the log book entries for bottom depth. The following casts had entries that didn’t look reasonable so they were removed from the header of the CLN and BOT files: 8, 11, 21, 30, 33, 39, 51, 53, 58, 63, 65, 66, 72, 74, 78, 80, 81, 83, 93. The header was removed from other BOT files that had only surface sampling but altimeter readings that suggested they were very close to bottom.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The sample numbers for cast 32 were changed from 175-182 to 9175-9182 because the former set of numbers were used for cast #33 as well. There was no IOS sampling from cast #32, but using pad values for the sample numbers causes problems in some routines and we do want to create a bottle file for UBC use. For casts #47 and 99 the log book says there was only 1 bottle fired, but there were in fact 3 for each.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

The salinity analysis from this cruise is complex because three salinometers were used:

- Portasal #16776960 – Used for 4 loop samples and 3 Niskin samples (samples 27-29) from casts 3 and 5. Proved to be unstable so no further analysis was done with that instrument. All samples will be flagged “c” or “e” but should be examined again after COMPARE is run.

- Portasal #58879 – used for Niskin samples (samples 30-90) from P2 to P8

- Autosal #68572	- Casts #17-#100 (samples 71-671) plus loops

The raw data is in two files: 2007-13 Salinity.xls contains the Autosal results; 2007-13sal.xls contains the results from the two Portasals. These files were simplified and saved as 2007-13 SALINITY-autosal.csv and 2007-13 SALINITY-portasal.csv.  The rosette sampling results from those two files were then combined into file 2007-13 SALINITY-CTD.csv and the loop samples were combined as 2007-13-salinity-loop.csv. Quality flag columns were added and Portasal results were averaged. One error was corrected: sample #204 was said to be from station P18, but is really from P19. The results from Portasal 16776960 were flagged “c” or “e” with a pad value entered for the “e” point. The CTD file was converted into individual SAL files. The LOOP file was set aside for later work on the TSG data.

Because of stability problems with the first Portasal, the analyst ran a test taking 9 samples from a single Niskin bottle. The analysis shows a standard deviation of 0.0008 with the difference between the minimum and maximum values being 0.0021. However, there is a trend from higher values to lower ones which is a little worrisome, possibly due to temperature variations during analysis. (32.5857, 32.5861, 32.863, 32.5850, 32.5848, 32.5842, 32.5842, 32.5845, 32.5845.)

Dissolved oxygen files (*.add) were provided with a flag channel and comments entered in the headers. The format is different from the usual, with 3 columns for the flag channel. Since that format is given in the header this was not changed, but it later proved to be a problem. There was no flag for a sample from cast #100 which has a note in the rosette log saying “Problem with colorimeter don’t trust value”. A “c” flag was entered along with a comment.

After a first run of COMPARE some outliers were identified in the ADD files. For further evidence profile plots were examined of salinity versus DO for all bottle files and DO versus pressure. The following actions were taken:

#48  flagged “c” - outlier in COMPARE and in DO vs Sal plot.

#144 flagged "d" - outlier in comparison to CTD, and in plots of DO vs Pressure and DO vs salinity

#147 flagged "d" - same as 144

#169 - originally named #158 in ADD file but rosette log shows no DO sampling for #158 and this value is entered on log sheet for 169. Changed sample number and put note in header, no flag assigned

File for cast #68 contains DO data from casts 68 and 70. I separated them.

#390 - One of the cast #70 bottles was entered in ADD file as 391 but it is clear from rosette log that it is 390. Entered note in header but no flag.

#649 - analyst flagged "c" - I changed it to "d" based on comparison to CTD and profiles of DO vs pressure and DO vs salinity

The nutrient spreadsheet was simplified and saved as 2007-13nuts.csv. Loop files were prepared separately. Extraneous columns were removed and header names were changed to standard format. Data were sorted on sample number. File 2007-13nut.csv was then converted to NUT files.

Extracted chlorophyll data were obtained and saved as 2007-13CHL.csv. The file was edited to remove extraneous lines and columns, header names were changed to standard format, data were sorted on sample number. 2007-13CHL.csv was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

##### 11. COMPARE

Salinity

Because the other bottle data were not ready yet and a comparison of salinity with bottles was required in order to know which sensor pair to choose for editing salinity, temporary merged bottle files were prepared with salinity only. (MRGTEMP) Later, after the final MRG files were ready this step was rerun; by then the only salinity value from cast #47 had been flagged so that cast was dropped from the comparison.

COMPARE was run with pressure as the reference channel. A few outliers were excluded from the fit. Most correspond to CTD data with a high standard deviation, so it is not clear that there is anything wrong with the bottles. The salinity sample #247 from cast #23 was flagged “c” since there is a large difference from the CTD salinity and the bottle salinity is very close to that of sample #246. Most of the samples from cast #5 were flagged. The standard deviations in the CTD salinity were large, but these differences seem way out of line with the other casts. These samples were analyzed on Portasal #58879.

With outliers excluded, the trendlines are fairly flat but the differences seem larger from 0 – 400m and below 3000m than at mid-depths. The primary salinity is low by an average of 0.0068 and the secondary by 0.0041. (The difference between these two results is reasonably close to the observations of section 7.)

A fit versus file pair number hints at some time dependence, but since there was little deep sampling early and only deep sampling late in the cruise, this is not clear. To try to determine if there is time or pressure dependence many plots were produced:

- Choosing only the casts after pump repairs shows the same pattern with a change between 2000db and 3000db.

- Cast #70 was the last cast with any sampling between 200 and 2500db and there is little deep sampling early in the cruise. So the apparent depth dependence could be confused with time dependence. To test this, plots were made of bottles from a particular level against time. The differences from 200db – 400db show lots of variability but no notable trend but there is no sampling at that level after cast #70. At 1500db there are few bottles and none after cast #70; the differences are lowest for the 2 before the pump was replaced and quite steady thereafter. At 2500db the differences get larger with time with the largest shift between #70 and #86. At 3000db there is a lot of variability with the biggest differences after cast #72. At 3500db there is again a lot of variability but if 2 outliers are excluded it looks quite flat with time. We do not usually rely on surface bottles since the gradients are larger, but the plot for bottles between 0 and 15db is the one with the most information since there were shallow bottles throughout the cruise. That plot really shows a lot of noise in the early casts, but does suggest a change between casts #76 and 77.  However, when the bottles are reduced to only those with a well-mixed surface layer the results look quite steady with time (but there are no samples between casts #18 and 68.)

- Single casts with sampling from many depths were examined to see if there was a significant change in the pressure dependence at any point. Casts 25, 33, 43, 70 and 100 were chosen. All casts showed a lot of variability especially in the top 1500db. Cast #70 has the flattest fit and the lowest average difference. There is no evidence of time dependence.

- Another cause of confusion is that there are a lot of bottles at 3000db from one cast. They all have unusually low values so they will skew the averages. When one sample was chosen to represent the 23 from that cast (with a value close to the average for all the 3000db samples) the average difference below 200db is -0.0063 for the primary and -0.0037 for the secondary.

There is no evidence of pressure-dependence or time-dependence, but there is a lot of noise in the comparison. That may be due partly to pump problems and partly to the pattern of sampling.

COMPARE was run two other times, with sample number and Niskin bottle number as the reference channels. The plots against sample number confirm that there is no notable problem with the first Portasal. There are a number of outliers from the group analyzed on the 2nd Portasal, but there are other samples that look ok. From the bottle # plots, no Niskin bottle shows up as bad; there are more outliers for the higher numbered bottles, but that is expected since they are usually from near the surface.

There were 23 bottles fired at 3000db during cast #84. The CTD showed very low variability through the firing time with standard deviations as seen in COMPARE <0.00033. When 2 outliers are excluded (samples 521 and 524 from Niskins #20 and 23) the secondary salinity is low by 0.0059 with a range of -0.0051 to -0.0068, suggesting salinity errors of <±0.001 for the salinity analysis when outliers are excluded. The two outliers come from samples #521 and 524. The first 7 bottles sampled (and presumably the first 7 analyzed) are remarkably close. The same pattern is seen in the differences between primary salinity and bottles. From the COMPARE run described in the previous paragraph there were no other outliers from those particular Niskin bottles, so the problem would not appear to be with the bottles themselves. It would be useful in processing data to be able to see the order of analysis and whether a group like this were all done at once or multiple sessions. It might shed light on the source of poor correspondence between bottles and CTD. With the Portasal this information was available, but not from the Autosal.

The shifts in T and S differences noted in section 7 do not show clearly in the salinity comparison, but these were small shifts (~0.0005) and the noise in the comparison makes it unlikely that we could resolve such a change. For both primary and secondary, the differences are larger below 2900db than above 2000db, but whereas the secondary CTD salinity goes down by an average of about 0.0016 the primary is down by about 0.0022. This suggests that the primary has moved lower than the secondary by a further 0.0006. The shifts noted in section 7 were on the order of 0.0004, so this result is roughly consistent. These COMPARE results are discussed further in APPENDIX I.

Outliers were assigned as follows:

- Cast 33, samples #27 and 28: These had been flagged “c” because the Portasal used proved unstable soon afterwards. However, they look ok in COMPARE, so the flags were removed.

- Cast #43, sample #247: Salinity value flagged “d” since it differs from CTD by 0.88 and because it has a value very close to the bottle below. Probably came from wrong bottle.

- Cast #84, sample #524: Out of line with other bottles fired at the same time. Flagged “d”. Problem assumed to be in bottle or analysis.

- Cast #86, samples #527-529: CTD higher than bottles by 0.02, 0.01 and 0.02. No sign of problem with CTD data. One bottle did not fire, so investigated possibility that samples came from wrong Niskin bottles. No problems noted by analysts looking at other variables. If we presume that the sample said to be from Niskin #1 is really from #2 and so on, we find the primary to be low by an average of 0.0074 and the secondary by 0.0046. The standard deviations are reduced from 0.009 for the original differences to <0.0005 for the differences after reassignment. Marie Robert confirms that the sampling would have been from Niskin Bottles #2, 3 and 4. The sample numbers were changed to 528, 529 and 530 and a note put in the headers, but no flags were assigned.

The only other major outliers are associated with very noisy CTD data so no flags are justified.

Dissolved Oxygen

COMPARE was run using pressure as the reference variable. The best fit is against CTD DO values. When a few outliers and bottles from below 1525db are excluded the fit was:

CTD-BOT = 1.0838 * DOX-CTD + 0.0601

There was no evidence of temporal variations.

This compares with the following recent results using this sensor:

CTD-BOT = 1.0274 * DOX-CTD + 0.023 (2007-26 shallow casts only – short bottle stops)

CTD-BOT = 1.0101 * DOX-CTD + 0.1109 (2007-11 mostly shallow casts)

While this seems like a fairly large drift between April and June, 2007-26 had very short bottle stops so that the DO in the bottles may effectively come from deeper water and hence be closer to the CTD than if the stops were longer. Also there was no deep sampling for 2007-26 or 2007-11 so that the fits would not be as reliable. There may have been other cruises using this equipment that have not been processed yet; the more it has been used the more drift is expected. Another factor to be kept in mind is that there were pump problems during this cruise. This sensor was mounted on the primary pump. However, the problems appear to have been mostly at depths at which the CTD dissolved oxygen is not very reliable and not included in the comparison. Another unknown is whether an anoxic plug was used between casts on this cruise, though the offset has been negative on other instruments when a plug was used. During 2007-26, since there was little time between casts it is most unlikely that a plug was used. Whatever the cause of the drift, the bottle data is considered reliable and using it to calibrate the CTD data is justified.

The only outliers, beyond those identified in step 10, were either very deep or very close to the surface and do not require flags as the CTD is the more likely cause of the mismatches.  (See 2007-13-dox-comp1.xls.)

Fluorescence

COMPARE was run using the CTD Fluorescence and the Titrated Chlorophyll from bottles. When all data were plotted there was a lot of scatter, with the data from the first few casts standing out from the rest. Those were the only casts with high values of CHL. When the casts with lower CHL are plotted together the slope of FL vs CHL is ~0.6 when the intercept is set to 0. Near-shore the slope is ~1.2. There are a number of differences <0 suggesting that the dark values of the fluorometer might be higher than usual.

##### 12. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values of from 1 to 1.6s were found, with most close to 1. For most cruises a shift of +24 records (1s) has been found to be appropriate. SHIFT was run using +24records. (Output: SHFFL)

Conductivity

Tests were run on 4 casts with few or no stops for bottles using shifts between -1s and +1s and T-S plots were prepared to compare the results. A setting of +0.6s worked best for the primary conductivity and -0.7s looked best for the secondary. When these sensors were used during 2007-26 the settings used were +0.5 and -0.7s. All casts were put through two runs of SHIFT using the settings of +0.6 and -0.7s. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Values from +80 to +160 were tried and the best overall match of features was with a choice of +130 records. For other cruises in the past year the choices were +90, +90, +110 and +120 records with a gradual trend to higher values. SHIFT was run using +130 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings about 4 casts.

30. Three records had bad pressure values – these were replaced by interpolated values and    DELETE was rerun.

32. The warnings all pertain to upcast data, so are no problem.

84. There were many questionable near-surface pressure values. DELETE was rerun after replacing some by interpolation; there was still one warning, but the DEL files looks ok.

86. Upcast only, so the warning is of no significance since this file will not be processed further.

In editing it was found that cast #30 contained bad data from the soak period. A text editor was used to remove the first 2482 records and DELETE was rerun.

##### DETAILED EDITING

The secondary salinity is closest to the bottles and was chosen for other recent cruises using these sensors. The primary salinity showed some odd shifts in values. Examination of T-S plots for a few casts shows that the secondary sensors provide smoother data. Since cruise 2007-14 used the same sensors it was decided to process that cruise up to this point in order to have more evidence as to which sensors are better. While the primary sensors produced salinity that was a little closer to the bottles than during this cruise, the secondary was closer and very similar to this cruise. Secondary sensors were selected for archiving.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

There were problems with the salinity near the surface of many casts and at depth for some, with unstable excursions to higher values that are obvious in the salinity channels. There appear to be some excursions to lower temperature values in some places, higher in others, and often the temperature is suspiciously constant; however, any features in temperature are much less obvious and somewhat masked by the local gradients. To edit the T-S excursions it was decided to remove the records from the suspect sections since all pumped channels may be affected. It is possible that there were similar excursions in the upper ocean that were not recognized because of larger gradients, so the data from casts 1 to 50 should be considered of lower quality than usual. See the APPENDIX for a fuller discussion of these features.

Sections of bad data as described above were found in: 8, 12, 17-20, 22-29,32,50.

Many casts had a lot of small spikes (1 to 3 points) in salinity (~.02 to .2 units); these were cleaned by interpolation.

All casts required some editing. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

The secondary salinity is lower than the bottles and a value of 0.004 looks like a reasonable estimate based on this cruise and 2007-14.

File 2007-13-recal1.ccf was prepared to add 0.004 to the secondary salinity and to apply the following equation to the dissolved oxygen sample:

DOX-CTD corrected = 1.0838 * DOX-CTD + 0.0601

These corrections were first applied to the rosette and chemistry files and then COMPARE was rerun to check that the results were as expected and they were. The downcast files were recalibrated. (See 2007-13-sal-comp2.xls and 2007-13-dox-comp2.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 100db and processed separately for A. Peña. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

Primary sensors were used on 2007-11 and 2007-26 when the primary salinity was found to be low by 0.0015 and by <0.001 and the secondary low by 0.0004 and <0.001; these were not great comparisons due to short stops and shallow sampling. The sensors were also used during 2007-14 which followed this cast. At that time the primary was found to be low by 0.0053 and the secondary by 0.0039.

2. Dissolved Oxygen: This sensor has been used 4 times. During cruise 2006-35 there was some hypoxic effects on the calibration and during 2007-11 and 2007-26 the sampling was all fairly shallow and many stops were very short.

3. Pressure: The pressure sensor has shown slow drift over the past 5 years with an offset of +1.2db applied over the past year.

Historic ranges –  Profile plots were made with historic ranges of T and S superimposed. All salinity data was within the ranges. The deep temperatures from stations P16 to P26 and S1-S5 & PIE93 are very close to and sometimes below the historic minima. (This may only reflect a paucity of data in the ranges.) The gradients of salinity and dissolved oxygen are impressive at about 100db with salinity values moving from near-minimum to near-maximum in about 20db.

Repeat Casts – There were many repeat casts at P26 including a few deep ones. There was a lot of variability near the surface but at ~800m the differences along σt-lines were about 0.002Cº and 0.0004 units of salinity with similar differences at 3000db.

Other intercomparisons – Plots were made of very deep casts in the line from PIE93 to north of P26 and the differences along σt-lines at about 3500db were very small, almost undetectable between casts #75 and 76. Casts #75, 76, 84 and 85 (over 2º of latitude) have minimum temperatures that are within 0.005Cº of each other. Cast #77 stands out as warmer than those by about 0.005Cº.  The minimum temperature at PIE93 when sampled in 1999 (1999-10-0031) was 1.4850ºC and during this cruise (2007-13-0075) was 1.4901ºC. This implies a warming of about 0.005Cº which, according to Howard Freeland, is about what is expected based on the warming noted up to 1999. PIE93 had the lowest minimum of all those sampled. A Japanese expedition is expected to visit this site in the near future which may allow a further intercalibration.

##### Final Calibration of DO

SHIFT corrects for the error in DO alignment with temperature due to transit time. The first recalibration corrects for the calibration drift in the sensor. A further correction is needed for the downcast data only, to allow for the “in-motion” error due to poor sensor time response. To check for this downcast data is compared with upcast bottles.

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When data from below 1500m are excluded plus outliers identified by residuals, it was found that the differences were quite flat when plotted against CTD DO value, with the CTD data being high by an average of 0.16ml/l. However, the differences look lower for high and low DO values. This is presumably because the DO gradient was low down to 75 or 100db, then high from 75 to 200db (very high at P26) and fairly low below 300db. In recalibrating this data we must make a compromise between fitting the high gradient zone and the rest of the profile. There is every reason to expect a poor fit between 75 and 200 since a minor vertical movement in the interface will produce a major mismatch in the upcast and downcast data. While most of the data comes from the lower gradient areas, using a global average does not really give us a good balance, since the differences in the high gradient region tend to be large thus have an overly large influence on the result. So points were gradually removed based on residuals until a fit that is quite flat with DO value was achieved; that fit went through the middle of the high and low values and included some data from the high gradient region but not many. The result was that the downcast DO data was found to be high by 0.09ml/l.

Plots were made to see if the DO was affected by the pump problems noted in the salinity data. There is no evidence that the DO data was affected, though small differences would not be obvious. The slow time response may have been helpful in this regard. Also, the area with the sharpest DO gradients was sampled after the pump problems cleared up. (See 2007-13-dox-comp3.xls.)

The thinned downcast files were recalibrated using file 2007-13-recal2.ccf to reduce the dissolved oxygen channel by 0.09ml/l. COMPARE was rerun and the results were satisfactory. The differences are near zero at the surface, a little high from 75 to 500db and a little low from 800db down. (See 2007-13-comp4.xls.)

A second recalibration was applied to the downcast files only to subtract 0.09ml/l. (Output: COR2 )

The clipped files were recalibrated. One set was then bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were filtered before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from all casts except:  #11, 21, 30, 39, 53-66, 68, 71-74, 78-83.

A second SBE DO channel was added and the channels reordered to put the two SBE DO channels together.

Profile plots were made. The only problems noted were a few large spikes in transmissivity at depth in casts 7, 8, 67, 70 and 71. The transmissivity data were not edited.

HEADER EDIT was used to fix formats and channel names, to change the geographic area and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.5ml/l from 0 –75db.

- ±0.8ml/l from 75 to 300db; data tend to be somewhat high in this high gradient region.

- ±0.25ml/l below 300.

- data below 800db tend to be  low.

- data below 1200db are considered unreliable by the manufacturer.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The near-surface values were between 108% and 120% for casts out as far as P13 and further offshore the values were from 103 to 107%. This is within expected values. Cast #1 had 150% saturation and this was true of the upcast as well as downcast data, so not likely due to the sensor not having been used for a while. However, it is possible that the recalibration applied was not appropriate for this cast if an anoxic plug was used during storage but not between other casts. There was no DO sampling for this cast. Nonetheless, very high values are often seen at this site.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

The PAR channel was removed from all casts except ##11, 21, 30, 39, 53-66, 68, 71-74, 78-83 because the instrument was not mounted on the CTD for those casts. Plots show no PAR signal for a few of these casts but that is because there was no sampling close to the surface where there was a signal.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units, to change the geographic area and to add a comment about quality flags and analysis methods. The standard comments for hydro files were adapted to reflect the fact that the chlorophyll was analyzed on board. The following comments were entered:

The dissolved oxygen data is considered to be:

- ±0.5ml/l from 0 –75db

- ±0.8ml/l from 75 to 300db; data tend to be somewhat high in this high gradient region.

- ±0.25ml/l below 300

- data below 800db tend to be on low.

- data below 1200db are considered unreliable by the manufacturer.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### 22. Thermosalinograph Data

a.) Checking calibrations

There were 2 files containing TSG data. A report was printed for the con file, the fluorometer calibration was corrected and serial number were entered and the primary temperature calibration was corrected; the resulting file was saved as 2007-13-TSG.con.

The primary and secondary temperature sensors had been recalibrated since the last data processed to date, so there is no history available.

b.) Converting to IOS Headers, adding position headers and time channels, preliminary checks

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC.

Time-series plots were produced. The salinity data have many spikes and jumps that are not associated with spikes in any other channel including the flow rate. Editing will be required to interpolate or remove these spikes. The flow rate starts at ~1.03 and after a short time goes up suddenly to ~1.15 and stays at that value until the end of the record.

Preliminary track plots show that there are spikes in the positions.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within .3db of 4db and exported to a spreadsheet.

The TSG file was opened in EXCEL, median and standard deviations were calculated for temperature, salinity and fluorescence, and the file was then reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2007-13-ctd-tsg-comp.xls). The positions were compared and were very close, with average differences for both latitude and longitude of <0.0001º and no difference greater than 0.0003º so the clock appears to have worked well.

This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are variations in alignment, but they are not systematic. This step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 The average difference over the whole record shows the TSG temperature to be high by 0.150 C˚, but there is considerable variation. For the early near-shore casts the differences are ~0.147, go up to ~0.152 offshore, average ~0.155 for the North-South line and are ~0.127 as the ship returned to near-shore. For the whole east-west line at the end of the cruise, the average is 0.131. When plotted against temperature it appears that the reason for the variation is the in situ temperature; in the warmer water near-shore there is less internal heating than in cooler offshore waters. There was a larger range of temperature during this cruise than we usually see, with temperatures varying from about 7˚C to 16˚C. In spring 2005 the range was ~9˚C to 16˚C. Plots were made after outliers were removed based on the standard deviation in the TSG temperature data and differences between TSG intake and lab temperature. These show that the warming by the ship is inversely related to the CTD temperature and to the intake temperature.

Correction to lab temp to account for ship heating = ~0.01*Intake temperature – 0.23.

The error varies from 0.07C˚ to 0.17C˚ for the temperature range encountered in this cruise; had the intake temperature ever reached 23˚C there would have been no correction, which makes sense given that is probably close to the temperature of the ship. (See 2007-13-ctd-tsg-comp-temp-reduced.xls.)

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There were 50 casts that could be used. Graphs were prepared comparing the TSG temperature, salinity and fluorescence with those of the CTD.

After a few outliers were removed the average difference between the TSG lab temperature and the CTD temperature was 0.172 C˚. Records were gradually removed based on standard deviations in the TSG temperature and the TSG temperature was found to be high by about 0.154C˚ with 17 records included. Looking at casts with a well-mixed surface layer (11 casts) the average difference was ~0.016, and ~0.0164 for the 9 best mixed surfaces. Overall, 0.016 looks like a reasonable estimate. (See 2007-13-ctd-tsg-comp-temp-reduced.xls)

For salinity there was only 1 significant outlier; the average of all other casts showed the salinity to be low by 0.083. Removing more points based on standard deviations in the TSG salinity brings the difference down to about 0.068 when only 17 records are included. Looking at casts with a well-mixed surface layer (11 casts) the average difference was 0.076, and 0.069 based on the 9 casts with the best mixed surface waters. (See 2007-13-ctd-tsg-comp-sal-reduced.xls)

The ratio of TSG fluorescence to CTD fluorescence ranges from 2 to 7. Even after restricting the choice to the 9 best mixed casts the range was 2.5 to 5.4 with an average of 3.3.

(See 2007-13-ctd-tsg-comp-sal-reduced.xls)

- Loop Bottle Comparisons There were 22 loop salinity samples, but one was not analyzed and 2 run on Portasal #68627 were unstable. The range of temperatures associated with the loop samples is much more limited than from the whole record (7.3 to 11.3 C˚.)

The salinity values (using a median over a 2-minute window) were compared with the TSG files and there is a lot of scatter, but in all cases the TSG salinity is lower than the loop samples. The TSG salinity is low by an average of 0.068 (when the 2 unstable samples and one other outlier are excluded) with a range of from 0.054 to 0.089.

Plots of differences against standard deviation in the TSG salinity do not suggest that noisy TSG data is the explanation for the scatter. A plot of differences against intake temperature does show a trend with the differences larger for higher temperatures. It is unknown what would cause a temperature-dependent error in the salinity. When plots were made of the differences against salinity no pattern emerges, but plots against standard deviation in TSG salinity and intake temperature both show slight dependence. A further plot of intake temperature against standard deviation in salinity suggests that at higher temperatures the data is noisier (the near-shore data is both noisy and warm), so the temperature dependence is probably just coincidental. Using an offset of 0.068 produces differences within ±0.022 units of salinity.

The fluorescence/chlorophyll ratio is noisy with an average of 3.2 and a range of 0.9 to 7 when one point is excluded. The first 4 samples from Juan de Fuca Strait have an average of 1.4 and offshore the ratio is higher. For the samples furthest offshore the average is ~5.8. The ratio increases as FL or CHL increase, but there is a lot of scatter, so no clear trendline.

To study differences in the TSG system depending on whether the ship was moving or not, pairs of samples were taken at 4 sites, 1 at the end of a CTD cast and another after 10 or 15 minutes when the ship was underway and moving at 10knots. The differences between the TSG and loop salinities were larger when stopped than when underway in 3 of the 4 cases, but the differences were not large. In the chlorophyll data there is a similar variation with the CHL/FL ratio higher when underway for 3 of the 4 cases, but again the differences were not large.

- Calibration History The TSG was recalibrated shortly before this cruise so there is no history for the salinity. However, we can judge from the history that the TSG temperature is usually high by about 0.18 to 0.2Cº at this time of year.

Conclusions

The temperature is higher than that of the intake temperature by an average of 0.15Cº and the CTD temperature by about 0.160Cº averaged over the whole record, but there is evidence of temperature dependence in the differences. This is somewhat lower than the results of recent cruises but the sensors had been recalibrated recently. Using the following correction

T correction = 0.01* Intake Temperature – 0.23

produces good results. The fact that there would be no correction for an intake temperature of 23˚C makes sense since that would be close to the ship temperature.

The salinity is lower than the loop by about 0.068 based on the loop comparison and by about 0.07 compared to the CTD salinity. Using a correction of 0.068 produces salinity that is very close to the loop samples on average, and all except one outlier are within ±0.02.

The average TSG FLUOR / Loop CHL ratio is very close to that for the TSG FLUOR / CTD FLUOR. The values range from ~1 to 7 for the former and ~2 to 7 for the latter; the former includes samples in Juan de Fuca Strait where the ratio was lowest, whereas there were no CTD casts in the Strait. For the loop samples from farthest offshore the average ratio is ~5.8.

f.) Editing

The time-series plots were examined and there was a lot of salinity spiking (all to lower values) with no associated temperature spikes. CTDEDIT was used twice, the first run was to remove one spike in latitude and longitude in each file; the second run was to remove large, single-point spikes. A small group of bad salinity values were also removed from the first file. After this step plots were made and no further editing was needed.

g.) Preparing Final Files

It was requested by the Chief Scientist that instead of recalibrating the primary temperature we rename it as Temperature:Lab and retain the secondary temperature renaming it as Temperature:Intake. That way those taking samples from the lab will know the temperature at the time of sampling and the in situ temperature will be well represented by the intake temperature.

HEADEDIT was used to was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, add the depth of sampling to the header and change the name Temperature:Primary to Temperature:Lab and Temperature:Secondary to Temperature:Intake.

REORDER was used to ensure that Temperature:Intake channel was before Temperature:Lab.

REMOVE was used to remove the following channels: Scan_Number, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Flag and UPloy0 (flow rate).

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

Oct. 27, 2007: It was discovered that the salinity had not been recalibrated. Channel Salinity:T0:C0 was recalibrated by adding 0.07 using file 2007-13-tsg-recal.ccf.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

LOOP File preparation – February 6, 2014

An initial surface/loop file 2007-13-che-loop.csv was prepared by the chief scientist and contains 5m rosette samples plus loop samples.

A 6-line header was added which involved removing 1 column and rearranging some others.

The original header line was removed leaving the 6-line header.

The file break column was filled with value 1 so all data will be in a single file when converted later.

The file was then saved as 2007-13-che-loop-6linehdr.csv.

CONVERT was run to get an IOS Header file, followed by CLEAN to get start and stop times and positions.

Comments were added based on the CHE file headers plus some extra information that was in the csv file. Comments were added concerning flags on samples from Niskin bottles..

The final file was named 2007-13-surface.loop.

A track plot looks reasonable and a plot of salinity versus date looks right.

APPENDIX I – Notes on shift in differences between sensors

In section 7 it was mentioned that there were shifts in the differences between pairs of sensors that occurred below 2000db. On further investigation it appears that these shifts were generally around 2500db. Profile plots of the sensors pairs and T-S plots were prepared for a few deep casts.

- For cast #93 the shift is associated with features in the secondary temperature, so this may indicate a problem with either the pump or sensor. However, it could also mean that the secondary is measuring something that the primary is not detecting. Looking on a T-S surface and plotting just 2500 scans around the shift showed that the primary sensors have a distinct (and slightly unstable) shift but the secondary do not. The T-S curve is a straight line for the secondary and for the primary it is two parallel straight lines of the same slope as for the primary but the break between the two lines is at the point of the shift. On the upcast there is too much noise to see such a shift, but the differences do appear to become smaller when the CTD reaches about 3000db.

- There is a similar but smaller feature for cast #67. The T-S curve is very straight through this section with the same slope before and after the shift, but the line is shifted towards lower salinity.

- For cast #85 there is a clear kink in the primary salinity trace at about 2520db, just where the salinity difference becomes higher. The primary density shows a slight instability between 2510 and 2520db for cast #85.

- Cast #86/87 was investigated in depth to try to understand the shift. Looking at salinity differences during the downcast, they are ~0.003 down to 2510db, and then shift to ~0.0035 for a short time and then settle to about 0.033. Once the CTD stopped at the bottom the differences went down very slightly to 0.0032. In the upcast (#87) there is too much noise in the differences to pick out a shift, but it does look as though the differences are larger below 3000db (~0.0028 above and ~0.0032 below 3000db).

- COMPARE shows that the differences between CTD and bottles increased at depth. Between 200 and 2000m the differences are extremely flat (especially for the secondary) and show the primary low by 0.0056 and the secondary low by 0.0033. Between 2900 and 4000 the differences are fairly flat and show the primary low by 0.0078 and the secondary by 0.0049. The difference between the two CTD salinity channels thus is 0.0023 above 2000m and 0.0029 between 2900 and 4000m. Those differences are smaller than seen in section 7, but those are from a CTD in motion and these are not, so there are different errors possible. The fact that the primary error increases by 0.0006 units more than the secondary is probably due to the shifts being discussed here. But why do both sensors have larger errors at depth? One possibility is that cast #84 with 23 samples from one bottle is overrepresented in this calculation and that 1 bottle did provide larger differences than most. If we only use 1 of those deep samples in the comparison the error is smaller.

Available evidence suggests that the primary sensors account for the shift in differences and the secondary temperature and salinity are probably the more reliable.

APPENDIX II – Notes on flow rate problems

There were patches of unstable T/S data as described in section 13. Some investigation was done to try to understand the source of this problem. This does not seem to have anything to do with the issues raised in section 7, namely the shift in differences between sensors seen below 2000db. Those shifts occur throughout the cruise, whereas the problems detailed in section 13 are not seen after cast #50.

A few general observations about these unstable features:

- They are generally seen below 800db. This could just mean they are not obvious due to larger gradients above that, though that seems unlikely. Alternately, high pressure may be part of the problem. However, they do not occur during some deep casts; no such problems were noted in the following deep casts: 13-14, 43-48.

- The features were noted during casts: 8, 12, 17, 18, 19, 20, 22, 25, 26, 27, 29, 32, 50.

- The features are seen in both downcast and upcast traces and in both sensor pairs.

- The features in downcasts are typically marked by a gradual move to higher salinity values and a sudden return to expected values. The conductivity moves to higher values and has an even sharper return to normal. The temperature behaviour is not as easy to see but it appears, judging from a comparison of downcast and upcast traces, that it sometimes moves to lower values, at other times to higher values and often looks suspiciously constant. During the upcasts the same patterns are seen but with the opposite signs. The features can be very large, ~5 to 10m and sometimes there are a series of them over 100s of metres. The temperature and salinity in the features are low/high by up to 0.15 units each, though the errors are usually closer to 0.02 units. (There are occasionally smaller features that look more like two-sided noise such as caused by flow-rate irregularity, but those are most often found close to major excursions.)

- The pump status = 1 throughout these features, showing that the pumps were not turned off.

- The features may be associated with high acceleration of the CTD, but this is not clear since the acceleration was so often high that it may not be significant.

- The features do not seem dependent on the descent rate itself, but during this cruise it was generally high at the pressures where the features are seen. The features are noted in casts with noisy descent rates and others with quiet rates. Cast #26 had a high and fairly steady descent rate, but is one of the most corrupted by these features. An investigation to see if very high rates were implicated found no evidence of that.

- There are other features seen in this data that are commonly seen in other cruises. They appear to be a combination of flow variation followed immediately by shed wake corruption. When the CTD is slowing down very quickly the high acceleration would lead to a slight rise in salinity due to reduced flow, followed by a sudden reduction as a shed wake overwhelms the CTD. Such features are smaller, usually ~1m. There are probably also lots of small rises in salinity in high acceleration areas that are not associated with shed wakes, but they are rarely large enough to be noticeable.

- During the cruise there were a variety of repairs made, including replacing the primary pump after cast #32 and replacing a pump cable after cast #50.

- Looking at the COMPARE results for salinity, it was found that for the deep bottles from casts with serious flow problems, the CTD secondary salinity was lower than the bottles by about 0.0033, whereas it was lower by 0.0046 for the deep bottles from casts with no such problems. This result is based on very few bottles, so at most we can say that the problems MAY affect the CTD even when it is stopped.

The conclusion is that there were flow problems due to pump malfunctions of some sort. It is known that flow problems can arise from either low descent rate or high acceleration, and mounting the CTD horizontally makes the latter effect worse. During this cruise the effects are much larger than those noted in the past, yet these descent rates and accelerations are not unusual. So clearly this was not a normal response, but rather, a significant pump malfunction. After the replacement of the primary pump following cast #32 there was a distinct improvement to both channels. Then the problem re-occurred during cast #50. A pump cable was replaced at that point and the problem was not seen again. Why a repair to one pump should affect both is a puzzle; perhaps in doing the repair something else was tightened or adjusted in some way.

Particulars:

3 – station name wrong; should be P1. Fixed.

7 – Transmissometer spikes

8 – Transmissometer spikes

12 – Large temperature difference 700-1300db down and up

17 – station name wrong; should be P8

20 – rosette out, then back to 5 for Niskin

22 – Fuzzy salinity signal

30, 32, 84, 85 – major pressure and temperature spikes

32 – Swapped PR1 pump at about this time.

44 – station name wrong; should be P21

48 – dissolved oxygen sample recorded in log, but no file found for it.

49 – aborted

50 – Large sensor difference. Secondary pump or blockage? Pump cable changed after cast.

59 – error in file name, should be named 2007-13-0058

67 – Transmissometer spikes

70 – Transmissometer spikes, altimeter moved to other side of rosette

71 – Transmissometer spikes

72 – Changed transmissometer and oxy Y cable

86 – upcast of downcast 85; station name missing – should be S2.

96, 97 – bad – station was redone and is named 0098

Casts with PAR: 11, 21, 30, 39, 53-66, 68, 71-74, 78-83

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2007-13 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 19 June 2007 End: 3 July 2007 |  |  |  |  |  |  |
| Location: N.E. Pacific |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Robert M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443        Cruise ID#:		2007-13

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23Dec06 | Factory “ |  |  |
| Conductivity | 1766 | 23Dec06 | “ |  |  |
| Secondary Temp. | 2374 | 21Dec06 | “ |  |  |
| Secondary Cond. | 1729 | 23Dec06 | “ |  |  |
| Transmissometer | 953DR | 23April07 | IOS |  |  |
| SBE 43 DO sensor | 1117 | 17/Oct/2006 | Factory |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| PAR | 4656 | 11/Feb/2003 |  |  |  |
| Pressure Sensor | 75636 | 04/Jun/1999 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2007-13

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 10/04/07 | Factory |  |  |
| Conductivity | 2487 | 10/04/07 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Temperature 2 | 2416 | 23/Dec/06 |  |  |  |
| Flow Meter | ? | ? | ? |  |  |