## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2007-14

Agency: OSD

Location: North-East Pacific

Project: Line Z

Party Chief: Robert M.

Platform: John P. Tully

Date: June 19, 2007 – July 3, 2007

Processed by: Germaine Gatien

Date of Processing: 10 July 2007 – 24 August 2007

Number of original CTD casts: 34	Number of CTD casts processed: 34

Number of bottle casts: 	33

Number of bottle casts processed: 33 (but 1 has no sampling so only 32 to be archived)

TSG files: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Chelsea/Seatech transmissometer (#953), an SBE 43 DO sensor (#1117, on the primary pump), a Seapoint Fluorometer (#2845) with a 10X cable (on the secondary pump) and an altimeter (#1252). The deck unit was an SBE 11+ model (#0424) and there was a mid-ship winch.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order and notes from the chief scientist greatly simplified the processing job.

Syringes were not attached to the sensors between casts following recent recommendations from SeaBird. The salinity comparison for this cruise compares well with that of 2007-13 so there is no evidence that this was a problem.

The dissolved oxygen data in the CTD files should be considered

- ±0.4ml/l from 0 –250db (except in very high gradient areas where ~±1ml/l)

- ±0.2ml/l from 250 – 500db

- ±0.1ml/l from 500 – 1500db

- data below 1200db are considered unreliable by the manufacturer,  but errors do not appear large compared to bottles until below 1500db.

About a quarter of the salinity data in the Thermosalinograph record were bad and were removed.

The TSG flow rate is higher than usual and was increased part way through the first file; the change of flow rate does not appear to have changed the noise level in the salinity, but did reduce the temperature change between the intake and lab.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets and a cruise report. There was no mention of serious problems with the CTD. The same equipment was used as during 2007-13 which immediately preceded this cruise, except that no PAR sensor was used for this one.

Titrated chlorophyll, nutrients and salinity data were obtained in spreadsheet format. The oxygen files were provided in individual ADD files. Quality flags had been entered but the comments were entered in the wrong way. A text editor was used to fix those.

A pressure test was done on deck and the reading was 0.44db.

The cruise summary sheet was completed. The histories of the conductivity, DO and pressure sensors were obtained. The calibration constants were checked for all instruments and a pressure offset was entered. The file was saved as 2007-14-CTD.con

##### Conversion of Raw Data

Data were converted using the configuration file listed above. File names for 4 casts had “dat” added to the standard name – those names were corrected after conversion.

Two of the files are misnamed: 2007-14-0002.cnv was renamed as 2007-14-0004.cnv and 2007-14-0003.cnv as 2007-14-0005.cnv.

A few casts were examined and all expected channels are present and look reasonable when plotted. During 2007-13 the upcast temperature traces were much noisier and further apart than in the downcasts, but there is no evidence of that in this data.

Rosette files were converted using a start time of -2s and duration of 5s. Two files were misnamed as noted for the CNV files and those names were corrected.

The rosette files were then converted to IOS SHELL files and renamed with extensions BOT. All BOT files were plotted and no outliers found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on many casts using settings (α = 0.02, β=7), (0.03, 7), (0.02, 9), (0.03, 9) and (0.0245, 9.5) to see what settings look best for this cruise. Overall, (0.03, 9.0) looked best for the primary and (0.02, 7.0) for the secondary conductivity channel. During 2007-13 using the same equipment, the best choices were (0.03, 9.0) for the primary and (0.02, 9.0) for the secondary.

CELLTM was run using (0.03, 9.0) and (0.02, 7.).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on some deep casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 26 | 1000 2000 3000 3700 | -0.0006 -0.0002 -0.0007 -0.0003 | +0.00008 +0.00011 +0.00011 +0.00013 | +0.0016 +0.0016 +0.0021 +0.0019 | High, noisy |
| 28 | 1000 2000 3000 3700 | -0.0006 -0.0003 -0.0006 -0.0004 | +0.00007 +0.00011 +0.00012 +0.00013 | +0.0014 +0.0017 +0.0020 +0.0020 | High, noisy |
| 37 | 1000 2000 3000 | -0.0009 -0.0005 -0.0006 | +0.00007 +0.00008 +0.0001 | +0.0015 +0.0014 +0.0019 | High, noisy |

The salinity differences go up slightly with pressure and temperatures differences go up between 2000 and 3000db. When differences are plotted there is a clear shift of about 0.005 at about 2500db. Similar shifts were seen in the 2007-13 data. On a T-S surface the primary data looks a little odd at the point of these shifts but it is not as clear as in the 2007-13 data.

The conductivity differences are significantly lower than those observed during 2007-13. The temperature differences are similar to that cruise so the resulting salinity differences are lower.

The descent rate was a little steadier for the 2007-14 casts used above, than those used for 2007-13; it could be that one sensor handled rapid changes better than the other, resulting in smaller differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. Problems arose because files 2007-14-0020.cnv and -0032.cnv were too large to convert. STRIP was used on the CNV files to remove channels that are useful but not critical: Oxygen:Voltage:SBE,  Status:Pump and Descent_Rate, then conversion was successful.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and header summary were run and the only problem noted is in the Project name and geographic headers – this will be fixed later when HEADER EDIT is run on the final files. The header summary was run and no errors were found. The cruise track was plotted and no problems found. The average surface pressure is 1.8db which is a little low for the Tully. However, many of the lowest readings are associated with very low salinity so they probably are from very close to the surface.

The altimeter readings from the header were exported to a spreadsheet and a few casts were checked. There is a lot of noise in the data for many casts, but the algorithm worked well near the bottom.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

The salinity spreadsheet (2007-14 sals.xls) included loop data and rosette data. The file was edited to change headers to standard names and to add a flag channel and then saved in two files, one containing rosette data only, 2007-14-sal-ctd.csv, and the other loop data, 2007-14-sal-loop.csv. The CTD salinity spreadsheet was then converted to individual SAL files.

Dissolved oxygen files (*.add) were provided with a flag channel and comments entered in the headers. The flags had the wrong format with one extra space. This will be fixed later in HEADEDIT.

The nutrient spreadsheet was simplified and saved as 2007-14nuts.csv. Loop files were prepared separately. Extraneous columns were removed, header names were changed to standard format, pad values were entered for blank values. Data were sorted on sample number. File 2007-14nuts.csv was then converted to NUT files.

Extracted chlorophyll data were obtained and saved as 2007-14 CHL.csv. The file was edited to remove extraneous lines and columns, header names were changed to standard format, data were sorted on sample number. 2007-14CHL.csv was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

##### 12. COMPARE

Salinity

COMPARE was run with pressure as the reference channel. There were only 4 major outliers, one at 75db of cast #9, one at 3499db of cast #27 and two surface samples with very large standard deviations in the CTD data. The outlier from cast #9 also appears as an outlier in the salinity profile plot, so was flagged “d”. The outlier from cast #27 is not as far out, but at this depth it is significant. Since both sensors agree and the standard deviations are low, it is assumed to be a bad bottle, so the value was flagged “c”. When records with pressure < 300db were removed plus one deep outlier, the trendline was very flat for the differences between the secondary salinity and the bottles and fairly flat for the primary. The primary salinity was found to be low by an average of 0.0054 and the secondary low by an average of 0.0039.  During 2007-13 which immediately preceded this cruise and used the same equipment, the primary was found to be low by 0.0068 and the secondary low by 0.0041. The results are similar for the secondary sensors. The primary values are closer to the bottles than during 2007-13; during the earlier cruise there were occasional large differences between sensor pairs and there is some evidence of particular problems with the primary sensors or pump. The CTD system may have been performing better for this cruise. A fit versus file pair number shows little time dependence, with the secondary especially flat.

Comments were put in the headers about the outliers mentioned above and for sample #235, cast #34. The pumps were turned off before that sample was taken so the bottle is probably ok, but CTD salinity is not reliable.

In section 7 it was reported that when the temperature and salinity differences between sensor pairs were plotted against pressure there was a shift around 2500db. This was also noted during 2007-13. The primary differences in COMPARE have some pressure dependence; it looks more gradual rather than a distinct shift, but such a small shift is unlikely to be detectable here.

The primary sensors are closer to the bottles than in 2007-13 and the conductivity differences are smaller as noted in section 7. This could be due to repairs to the CTD (problems had been noted with the primary pump during the earlier cruise and a number of steps were taken to fix it) or perhaps not attaching syringes had a good effect on the system.

After the comparison was run it was discovered that two casts had not been included; each had only a surface sample and these were reasonably close to the CTD values, so COMPARE was not rerun.

Dissolved Oxygen

COMPARE was run using pressure as the reference variable. The fit of differences versus dissolved oxygen provides a better fit than against pressure. The only severe outlier is from a depth at which the sensor is not expected to perform well. When a few minor outliers were excluded and all samples from below 1500db, the following fit was found:

The best fit is against CTD DO values though the fit against pressure is reasonably flat as well:

CTD-BOT = 1.0821 * DOX-CTD + 0.0293

During 2007-13 the fit was found to be:

CTD-BOT = 1.0838 * DOX-CTD + 0.0601

When the data from this cruise were broken down into groups there was little difference in the resulting fits, except that for the 3 casts near Vancouver Island the offset was negative; there was no deep sampling and no DO values <1 for those casts, so the fit does not cover the complete DO range of the other casts. When the Line Z group were compared with the other casts which were closer to shore, there was only a slight difference (slope/offset = 1.0832/0.0292 for Line Z and 1.0815/0.0297 for the near-shore group.) So the fit using all casts looks appropriate for the whole cruise. (See 2007-14-dox-comp1.xls.)

The analyst had flagged sample #46 from cast #13 “c” with the comment that there was a large plankton in the sample; that flag was changed from “c” to “d” because it was also a significant outlier in COMPARE. It is noted that the silicate value was also considered out of line.

No further flags were assigned to the DO data as all other outliers were either from >1525db where the sensor does not work well, or from samples for which the standard deviation in the CTD DO data was relatively high, or they were only slightly out of line.

Plots were made of DO_CTD and DO_BOT versus Salinity as a final check. The only outlier was sample #46 which has already been flagged “d”.

Fluorescence

COMPARE was run using the CTD Fluorescence and the Titrated Chlorophyll from bottles. When all data were plotted there was a lot of scatter with FL roughly 107% of CHL. When the data were roughly divided into Offshore and Nearshore groups there is a tighter fit (especially for offshore) with the FL roughly 80% for the offshore and 130% for the nearshore. These ratios are slightly higher than during 2007-13. The offset suggests that fluorescence is 0.21 for zero chlorophyll, but checking a few casts shows that the dark values are ~0.04, so perhaps the problem is in measuring CHL at low values.

##### 14. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values from 1.1 to 1.6s were found with most at the lower end of that range. The shift of +24 records (1s) which has been used in other recent cruises is thus appropriate. Fluorescence was shifted by +24 records.

Conductivity

Tests were run on 3 casts with few or no stops for bottles using shifts between -1s and +1s and T-S plots were prepared to compare the results. A setting of +0.6s worked best for the primary conductivity and -0.7s looked best for the secondary. (When these sensors were used during 2007-13 the settings used were +0.6 and -0.6s). All casts were put through two runs of SHIFT using the settings of +0.6 and -0.6s.

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because of hysteresis there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly; matching features from downcast to upcast is useful. There were not many such features in this data. For 2007-13, +130 was used, so values from +110 to +150 were tried for this data. For other cruises in the past year the choices were +90, +90, +110, +120 and +130 records with a gradual trend to higher values. A choice of +130 looks reasonable for this data. SHIFT was run using +130 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: No warnings.

##### DETAILED EDITING

The secondary salinity is closest to the bottles as it was when the same sensors were used for 2007-13. The differences between bottles and secondary are also flatter against pressure than those of the primary salinity. There are also some suggestions of a problem in the primary temperature as described in section 7, and as noted during 2007-13. So the secondary sensors will be selected for further processing and archiving. However, for cast #13 there was corruption of data from 235db to 255db in the secondary channels only, so the primary were selected for that cast only.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

All casts required some editing. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

File 2007-14-recal1.ccf was applied to add 0.0054 to the primary salinity and 0.0039 to the secondary salinity and to apply the following correction to the CTD Dissolved Oxygen channel:

CTD corrected = 1.0821 DOX-CTD + 0.0293

COMPARE was then rerun to check that the results were as expected and they were. (See 2007-14-sal-comp2.xls and 2007-14-dox-comp2.xls.)

The same DO and Salinity calibration was applied to the EDT files.

##### Special Fluorometer Processing

The COR1 files were clipped to 150db and processed separately for A. Peña. This is deeper than usual, but as seen during some casts from 2007-14 the signal is larger at depth than usual. For example, for cast #5 fluorescence does not reach “dark values” until about 155db. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other Comparisons

Previous experience with these sensors –

Salinity: The primary sensors were found to be low by 0.0015, <0.001 and 0.0072 during cruises in March, April and May-June. The second of these cruises had a very noisy comparison with short bottle stops. The last of these was the spring Line P cruise, during which there were large patches of bad data and the primary pump was replaced part way through the cruise.

Dissolved Oxygen: This sensor has been used for 5 other cruises since it was last recalibrated. The calibration for this cruise is close to that of 2007-13 which immediately preceded it, but quite different from earlier cruises.

Pressure: This pressure sensor has been drifting slowly and during all cruises in 2006 and 2007 an offset of +1.2db was found appropriate.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. The only excursion of salinity from those ranges was at cast #47 with salinity being slightly low from 50 to 90db. The temperature was above the historic maxima from 300 to 600db of cast #9 and near the maximum from 500 to 800db of cast #13; these were both casts in the Sitka Eddy area. For casts 45 and 47 (both close to shore) the temperature was a little low near the bottom (around 200db). These excursions do not look indicative of calibration problems as they are not systematic. It is more likely that they reflect real conditions in eddies and near-shore that are not well represented in the climatology.

Repeat Casts – There were no repeat casts.

##### Final Calibration of DO

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When data from below 1500m are excluded plus outliers identified by residuals, it was found that the differences were quite flat when plotted against CTD DO value, with the CTD data being high by an average of ~0.12ml/l.

The fit is better for high and low DO values than intermediate ones as was also seen during 2007-13, presumably because there is same very sharp DO gradient between 100 and 200db. During 2007-13 the average difference was ~0.09ml/l. A trial recalibration using a simple offset of -0.12 did not produce good results (most values below 800db were low and near-surface values mostly high), so a second attempt was made doing a fit against pressure. This produced the following fit:

DOX_CORRECTED = DOX_CTD -0.01626 +0.0001*PRESSURE

Doing a recalibration using that fit produced much better results.

There is no significant time-dependence. (See 2007-14-dox-comp3.xls and 2007-14-dox-comp4.xls.)

That recalibration was also applied to the downcast AVG and CLIP files (Output: COR2 and CLIPCOR.)

The clipped files were recalibrated. One set was then bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were filtered before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts except #13: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The following channels were removed from cast #13 only: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added and the channels reordered to put the two SBE DO channels together.

At this point all data was plotted and it was noted that the transmissivity had unbelievable spikes at depth in casts #1 and 21. Pad values were substituted for the bad values in those two casts (around 1800 and 940db respectively.)

HEADER EDIT was used to fix formats, to change the geographic area and the project name and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.4ml/l from 0 –250db (except in very high gradient areas where ~±1ml/l)

- ±0.2ml/l from 250 – 500db

- ±0.1ml/l from 500 – 1500db

- data below 1200db are considered unreliable by the manufacturer,  but errors do not appear large compared to bottles until below 1500db.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The near-surface values for most casts were between 105% and 115%; for casts 8, 16, 17, 18 and 22 it was between 115% and 130%. The casts with higher values had more variability in the DO near the surface than those with lower values. These values do not suggest any problems with the data.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts except #13 to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

REMOVE was run on cast #13 only to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats, to change the project name, geographic area and units and to add a comment about quality flags and analysis methods.

Because the dissolved oxygen flag had the wrong format, that information was lost in the HEADEDIT stage. Only two bottles had been flagged, so that information was reinserted using a text editor.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

Note that cast #23 was processed to this point, but as no sample numbers were assigned this will not be archived. (Output: CHE)

##### 22. Thermosalinograph Data

During the cruise problems were noted in the TSG salinity. The flow was increased at 2237 on June 24. On June 26 after sampling (Loop6) the Thermosalinograph was stopped, drained and then restarted.

a.) Checking calibrations

There were 2 files containing TSG data. A report was printed for the con file, the fluorometer calibration was corrected and serial number were entered and the primary temperature calibration was corrected; the resulting file was saved as 2007-14-TSG.con.

The history of the sensor was obtained.

b.) Converting to IOS Headers, adding position headers and time channels, preliminary checks

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Temperature:Secondary, Temperature:Difference, Conductivity:Primary, Fluorescence:URU:Wetlabs, UPloy0, Latitude, Longitude, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC.

Time-series plots were produced. The salinity data have many spikes and jumps that are not associated with spikes in any other channel including the flow rate. Editing will be required to interpolate or remove these spikes.

The flow rate is higher than usual (~1.4 for the first 34 hours and then 2.2 later).

A preliminary track plot showed there was one bad position record. The latitude and longitude values for record #11535 from the first file were replaced with interpolated values using a text editor. Following that the track plot looks as expected. CLEAN was rerun after this step to fix the headers.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within .3db of 4db and exported to a spreadsheet. The TSG file was opened in EXCEL, median and standard deviations were calculated for temperature, salinity and fluorescence, and the file was then reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2007-14-ctd-tsg-comp.xls). The positions were compared and were very close, with average differences for both latitude and longitude of <0.0001º and no difference greater than 0.0003º so the clock appears to have worked well.

This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are variations in alignment, but they are not systematic. This step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 The average difference over the whole record shows the TSG lab temperature to be high by 0.137 in the first file and 0.110 in the second file. Both files are extremely noisy through part of their records. When the comparison is made in the quieter parts the differences are 0.0132 and 0.115, so not much closer. The flow rate was changed partway through the first file so a plot was made of the temperature difference versus flow rate and that correlates well, the water does not warm up as much with the higher flow rate which makes sense. The salinity has many spikes in the latter half of the 1st file and the early part of the 2nd file with no noticeable change associated with the flow rate change. Only 17 hours after that did the salinity signal settle down.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There were 33 casts that could be used. Graphs were prepared comparing the TSG temperature, salinity and fluorescence with those of the CTD.

Based on an estimate of mixed layer depth from the CTD files, 11 casts were identified as best for comparison with the TSG data. Using all casts the intake temperature was found to be low by 0.16C˚ but using only the 11 well-mixed ones it is low by 0.0025C˚. When a few more casts were eliminated due to high standard deviations in the TSG temperature, the average difference was 0.002C˚. Given that the CTD data and the TSG data may not be from exactly the same depth this is a good agreement; also the CTD data are from the beginning of the cast and the loop samples from the end.

For salinity there is severe noise in the TSG data so any approach leaves a lot of noise in the comparison. The standard deviation over 2 minutes is of limited value since sometimes a larger section is offset. Looking at all casts the TSG salinity is low by 0.39 units. When only the 11 well-mixed casts are used it is low by 0.21 units. When casts with a lower standard deviation in salinity are used it appears to be low by 1 unit. When including only points with a lower standard deviation in TSG salinity and well-mixed near the surface the TSG salinity appears to be low by 0.05.

The ratio of TSG fluorescence to CTD fluorescence ranges from 0.5 to 6.8 and an average of about 3.1. When the points with high standard deviation in TSG fluorescence are removed then the average ratio is 3.5. (See 2007-14-ctd-tsg-comp.xls)

- Loop Bottle Comparisons There were 17 loop bottles. Note that none of the loop samples correspond to the part of the TSG record that was heavily corrupted with salinity spikes.

The salinity values (using a median over a 2-minute window) were compared with the TSG files and for all the TSG salinity is low, by from 0.02 to 0.26. When all points are used the TSG appears low by 0.08 but rejecting 4 with high standard deviation in the salinity suggests it is low by 0.05.

For 2 stations loop samples were taken at the end of a cast while the ship was stopped and a few minutes later when it was up to cruising speed. The differences were identical for the first pair at station Haida 6 and there was a small difference (0.002) for the second at Haida 7. The second pair both had higher standard deviations than the first. As expected the standard deviations were higher underway then while stopped.

There is insufficient data to reach a firm conclusion but there is some suggestion that the difference between bottles and TSG may have gone down when the flow rate went up. It is hard to think of explanation for this that would account for the TSG being closer to both the CTD and the loop samples. (See 2007-14-sal-tsg-loop-comp.xls.)

When loop chlorophyll is plotted against the TSG fluorescence (median over 2 minutes) the TSG FL is found to be approximately 4.5 times the loop CHL, or 4.8 times if one outlier is excluded. The range of the ratio is from 0.4 to 8.4. That outlier is one of only 2 or 3 points from well offshore, so the value may well be valid, reflecting different conditions. There is no dependence seen in a plot of TSG FL/Loop CHL versus standard deviation in the TSG data. (See 2007-14-chl-tsg-loop-comp.xls.)

- Calibration History The TSG was recalibrated in April 2007. Since then it has only been used for 2007-13 when the salinity was found to be low by 0.07 and the lab temperature was found to be high by an average of 0.15Cº but the range of errors was high and correlated to in situ temperatures.

Conclusions

The intake temperature is lower than that of the CTD by about 0.002Cº when only well-mixed casts with low standard deviation in the TSG temperature are used. This is as close as we can expect given the sampling depths are not exactly the same and time of sampling each is not exactly the same.

The differences between the lab temperature and the intake temperature changed during the cruise when the flow rate of the loop was increased.

The TSG fluorescence is higher than both the CTD fluorescence and the loop chlorophyll, with a wide range of ratios in each case and averages of about 3.5 and 4.8, respectively.

The salinity is lower than both the CTD and the loop by about 0.05 when only low standard deviation samples in well-mixed waters were used. The results of cruise 2006-13 showed salinity to be low by 0.07. Both cruises were affected by bad salinity data and it is possible that one or the other is a better comparison. 2007-13 had better TSG salinity but poorer CTD data. The flow rate differences might possibly be a factor as well. For the time being each cruise will be recalibrated based on the sampling from that cruise. Time may make it clearer that one or the other was better and the data could be recalibrated later. For now the salinity will be increased by 0.05.

f.) Editing

The time-series plots were examined and there was a lot of salinity spiking with no associated temperature spikes. Where there are spikes with only a few points in them interpolation was done, but where there was a large section of bad data the points were replaced with pad values. A lot of salinity points were removed from the latter part of the 1st file and the first quarter of the 2nd file.

g.) Recalibration CALIBRATE was used to apply an offset +0.05 to channel Salinity:T0:C0.

h.) Preparing Final Files

HEADEDIT was used to was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH, add the depth of sampling to the header and change the names Temperature:Primary to Temperature:Lab and Temperature:Secondary to Temperature:Intake.

REORDER was used to put Temperature:Intake before Temperature:Lab so that programs will selectively pick that channel.

REMOVE was used to remove the following channels: Scan_Number, Temperature:Difference, Conductivity:Primary, Flag and UPloy0 (flow rate).

REORDER was used to ensure that Temperature:Intake channel was before Temperature:Lab.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars from log and cruise notes:

1. Station name should be Eddy 1. CTD went past 150 on upcast, then went back down to sample.

4. Station name should be Eddy 2

4. File originally named 2007-14-0002, but is event #4 according to log and Chief Scientist’s notes.

5. File originally named 2007-14-0003, but is event #5 according to log and Chief Scientist’s notes.

5. Winch spooling problems so some “ups and downs”.

9. Bottles 13 and 15 did not close properly – top valve slightly open.

16. Depth should be 170 not 200.

20. Depth should be 4038.

21. Station name should be Z17. Transmissometer signal lost on way down.

23. 24 bottles tripped at 2000m but no sampling

26, 27, 28, 29 – extraneous “dat” in file name removed

34. Pump turned off at 5m before sample was taken

37. Temperature spike at beginning of cast.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2007-14 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 19 June 2007 End: 3 July 2007 |  |  |  |  |  |  |
| Location: N.E. Pacific |  |  |  |  |  |  |
| Vessel: John P. Tully Party Chief: Robert M. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443        Cruise ID#:		2007-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23Dec06 | Factory “ |  |  |
| Conductivity | 1766 | 23Dec06 | “ |  |  |
| Secondary Temp. | 2374 | 21Dec06 | “ |  |  |
| Secondary Cond. | 1729 | 23Dec06 | “ |  |  |
| Transmissometer | 953DR | 23April07 | IOS |  |  |
| SBE 43 DO sensor | 1117 | 17/Oct/2006 | Factory |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| Pressure Sensor | 75636 | 04/Jun/1999 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2007-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 10/04/07 | Factory |  |  |
| Conductivity | 2487 | 10/04/07 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 18/01/01 | “ |  |  |
| Temperature 2 | 2416 | 23/Dec/06 |  |  |  |
| Flow Meter | ? | ? | ? |  |  |