## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 25 Feb 2025 | Corrected channel name for channel Nitrate_plus_nitrite:ISUS in CHE files. G.G. |

## PROCESSING NOTES

Cruise: 2007-19

Agency: OSD

Location: Halifax to Kugluktuk

Project: C3O

Party Chief: Sarah Zimmermann

Platform: Louis S. St. Laurent

Date: 4 July 2007 – 26 July 2007

Processed by: Germaine Gatien

Date of Processing: 6 March 2017 – 27 March 2017

Number of original DAT files:	67	Number of CTD files: 61

Number of Ros files: 	62		Number of bottle casts processed: 58

# INSTRUMENT SUMMARY

The primary CTD system used on board was a Seabird SBE9+ CTD (serial number 0724) configured with a 24- position SBE-32 pylon and ice-strengthened rosette frame with 10L Niskin bottles fitted with internal stainless steel springs.  The data were collected real-time using a SBE 11+ deck unit and a computer running Seasave Win32 V 5.37d acquisition software.  The CTD was set up with two temperature sensors, two conductivity sensors, two oxygen sensors, fluorometer, transmissometer, altimeter, PAR sensor and surface reference PAR.  On casts shallower than 1000m, there were also nitrate and PAR sensors. The NMEA system was not turned on until cast #38.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 69086.

Typical Deployment Method:

The transmissometer(s) windows and nitrate sensor window were wiped with deionized water soaked Kim wipe or Q-tip prior to each deployment.  The PAR sensor was wiped periodically.

The package was lowered to 5m, taking a PAR measurement just below the surface if required.  The pumps were turned on the system soaked for 3 minutes.  The package was then brought up to just below the surface to begin a clean cast, and lowered to the requested depth at 60m/m.  The casts were typically to the bottom or 1000m whichever was shallower.  Full depth casts (over 1000m) were performed near the Ocean Station Bravo site and along the Baffin Bay section.

Niskin bottles were closed during the upcast after waiting 30seconds at each trip depth. There were 4 exceptions when there were no stops: casts 46, 47, 49 and 54.

The SPAR sensor was not put in its final location up above the hanger near the

incubation tanks until a few days into the program. Prior to this, the SPAR was placed

near the CTD deployment area where it could have been in shaded by the lab container. For 3 early casts there was no signal.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log and rosette log sheets were not in the usual covers so there were no lists of equipment or personnel and the usual identifying information about the cruise. While that information was available in other documents it is helpful to have it all in one place.

The use of consecutive numbers to name CTD files is very confusing since a number of different activities were identified by their own consecutive numbers, so that the cast number column is very confusing. The normal naming protocol is to use a 4-digit number corresponding to the event number which leads to a unique log entry. The file names were not changed because it would lead to great confusion because there are many associated documents that use the consecutive cast numbers.

The NMEA system was not in use for casts #1 to #37, but position information was available in spreadsheet format which enabled easy merging into the IOS Headers. For consistency those positions were used for all files, even those with NMEA entries. The NMEA data are downloaded when the file is opened while the log entries are usually made closer to the start of the cast.

It is recommended that a rosette sheet be prepared for every cast for which any bottles are sampled even if there is only 1 bottle sampled or the water was gathered for general use and no sample number attached. For cast #43 all bottles were fired but only 1 had a sample number assigned; there was no rosette sheet for this cast. While a note was found in the chemistry spreadsheet to explain this, it is much more efficient to have an annotated rosette sheet for every cast.

The rosette sheet for cast #55 was extremely confusing. The column said to be Niskin # contains numbers that are neither Niskin bottle #s nor the order of firing; they are meaningless. The pressure confirmation column does contain useful numbers but there is no indication on the sheet that these are actual Niskin bottle #s and also the firing order. A note has been entered on the rosette sheet to indicate which numbers should be used.

Before 2008 no hysteresis correction was available for the SBE43 sensors. There are few bottles below 2000m and a lot of scatter in the fits of CTD DO versus bottles, so evidence of hysteresis in the CTD dissolved oxygen data is weak. However, it is likely that values below 2000m are a little low. The deepest bottle came from 3500db and suggests the error could be as large as 0.2mL/L.

The comparison between salinity bottles and CTD salinity showed the primary to be low by ~0.0019 and the secondary low by ~0.0013. Analysis was quick so little evaporation or adsorption of samples is expected. While the secondary salinity calibration looked good, there were many bad data in the secondary channels that are believed to be due to the secondary pump system. The primary temperature,  salinity and dissolved oxygen data were selected for archiving.

There were many problems noted in the analysis documents for extracted chlorophyll and there is a lot of scatter in the comparison of the SeaPoint fluorometer with extracted chlorophyll. The overall pattern of the comparison is typical of this type of fluorometer with fluorescence values higher than CHL values when CHL values are low and lower when CHL is high.

PAR:Reference data were removed from casts 2, 4 and 5 because there was no signal.

Dissolved oxygen samples from 3 casts were replaced with pad values because they were obviously bad both in profile and compared to CTD data. Sampling errors could not explain the errors so they are assumed to be due to unknown analysis problems.

The Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7 mL/L from 0 to 50db

±0.3 mL/L from 50 to 200db

±0.15 mL/L from 200db to 300db

±0.05 mL/L below 300db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

The file names are not based on event numbers. Rather CTD casts were numbered consecutively.

Complicating things further there are some other entries in the log that use their own consecutive cast numbers.  So, for example, GRB 1 is after CTD 5.  Normally, these file names would be adjusted to match the log entries, but there are many associated documents, especially the chemistry files, that use the cast numbers, so the file names will not be changed.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as analysis sheets for salinity and nutrients.

Nutrients, extracted chlorophyll, NH4, DIC, Alkalinity and FW Alkalinity were obtained in a combined spreadsheet and dissolved oxygen and salinity data were obtained in separate spreadsheets.

The cruise summary sheet was completed.

The CON file was changed several times during the cruise:

- Cast #1 was a test cast; it was restarted as 1a after the altimeter scale factor was changed from 1 to 15.

- Cast #2 has the entries for PAR and SPAR corrected.

- Cast #10 has a change to the serial number for the fluorometer and the serial number and calibration date for one oxygen sensor and to the serial number and calibration parameters for the other dissolved oxygen sensor. No instruments were switched; these are just corrections.

- Cast #38 has the NMEA option chosen for the first time.

The calibration constants were checked for all instruments and a number of problems were found:

- The fluorometer serial number was incomplete.

- The altimeter serial number was incomplete.

- The transmissometer has entries for a March 20, 2003 calibration, but the appropriate choice is 28 August 2006. The parameters entered do not correspond to any entry that could be found in the calibration spreadsheet. Some values had been corrected after this cruise due to an error found in the formula used to derive them, but the values in this con file don’t look like either the original or corrected values so it is unknown where they came from.

- The PAR parameters were corrected after this cruise due to an error in the formula used.

- The SPAR serial number was incomplete.

One copy of the configuration file was saved as 2007-19-ctd2.con and corrections were made to that file.

The same file was saved as 2007-19-ctd1.con but with the NMEA option NOT chosen. While having NMEA chosen does not interfere with the conversion it might be confusing to someone later trying to understand why there are no NMEA data in the first half of the cruise.

##### Conversion of Full Files from Raw Data

Casts #1-37 were converted using file 2007-19-ctd1.xmlcon and for casts #38-62 file 2007-19-ctd2.xmlcon was used.

The hysteresis option is not available for the DO sensor calibrations run before 2008.

Voltage6 (ISUS) and PAR were converted for all files but will need to be removed from the deeper casts later as they were not actually mounted.

After conversion cast #1 was deleted and cast #1a was renamed as cast #1 since the first file was aborted. However, acquisition did not start until the CTD was at 40m so that cast will not be processed further.

Casts 24 and 24b were deleted and cast #24c was renamed as 24 because the log records say that the first two casts were aborted.

During cast #20 3 Niskin bottles did not close. Cast 2007-19-20-dianna.cnv is from a second cast to get samples from the depths that were missed. While bottles were fired during this cast the pumps were not on, so the CTD data from the first cast should be used.

A few casts were examined and all expected channels are present.

The two T/C pairs are in reasonable agreement but, as usual, upcasts are much noisier than downcasts.

The secondary dissolved oxygen looked bad for one cast – full scale throughout the cast. For a few others it looks fine.

Fluorescence and transmissivity look reasonable.

For one cast the SPAR looks lower than PAR but that was early in the cruise and there is a note that the SPAR sensor may have been shaded. Later the two are reasonably close.

Voltage6 is the ISUS channel and the profile has the right shape with some hysteresis.

The descent rates are mostly quite steady and close to 1m/s.

##### BOTTLE FILE PREPARATION

The ROS files were created using file 2007-19-ctd1.xmlcon and 2007-19-ctd2.xmlcon.

For most files data were selected for the 10s window around firing time, but for the casts fired on the fly the window was from 5s before firing until firing. There was no salinity or dissolved oxygen sampling for those casts to help judge what window to use. (Casts 46, 47, 49, 54)

The ROS files were converted to IOS format.

They were put through CLEAN to create BOT files. Temperature and salinity were plotted for all BOT files to check for outliers.

CTDEDIT was used to clean a few points in channel Salinity:T0:C0 for cast #3 and Salinity:T1:C1 for casts #22, 26 and 49. The edited files were copied to *.BOT.

There were many places where salinity was noisy but editing was only applied when there were spikes containing only 1 or 2 points.

It was found that the secondary salinity data in cast #32 were bad throughout. For some other casts the secondary salinity had quite different values from the primary; they look suspicious but not clearly wrong. This will be investigated later.

There were ROS files for casts #1, 1a and 19 but no sampling was done so those files will not be processed further. There was also a ROS file called 2007-19-0020-dianna which contains data collected during a second shallow cast at the site of cast #20 to collect samples missed in the main cast. As explained in the previous section the pumps were not on and the data from 2007-19-0020-dianna should not be used, so this file was not processed further.

Cast #43 has 24 bottles fired but only 1 sample number assigned. There is no rosette sheet for this cast and no sample data were found. A note was found that water from all the bottles were combined for a virus sample and that the CTD data from just the first bottle would be used. So all other lines were removed from the SAMAVG file and the final MERGE was repeated.

At this stage the position and water depth data were added to the headers using program Merge:CSV file to Headers. First the data were found in the chemistry spreadsheet and entered in file 2007-19-merge-hdr.csv. The header names and data formats had to be adjusted. The positions in the spreadsheet differ slightly from those that are present in the files that did have NMEA download. The differences are small and match those in the log book. This is probably due to the fact that the NMEA data are recorded when the file starts and the log entries when the cast starts, generally just a few minutes later. Since the water depths are missing from all casts, and the positions in the log are appropriate, it was decided to use the log data for all casts.

To check that the merge worked well a track plot was made and it looks fine.

A preliminary header check and a cross-reference check were run. There are no negative values in fluorescence. The second dissolved oxygen sensor had bad values for many casts.

The MRH files were bin-averaged on bottle number and the output was used to create file ADDSAMP.csv. Sample numbers were added to the file based on the rosette log records.

There were many adjustments needed to the file and/or sample numbers:

- Sample numbers 77, 78 and 79 were used for the top 3 bottles of cast #6 and the bottom 3 of cast #7. One or the other set needs to be renamed to enable merging with CTD data and choosing cast #6 means that the sample numbers will still be increasing in the sampling order which makes sorting simpler and less error-prone. So the samples for cast #6 were named 9077, 9078, 9079. All samples will have to be renamed.

- There were no sample #s 190, 249, 703, 725-727, 757-759, 804-805, 822, 869.

- During cast 26 there were 24 bottles fired but no sampling was done from Niskin #24 and no sample number assigned. So this line was removed from the addsamp file.

- Niskin 22 failed to close during cast #27 so sample #452 was taken from Niskin #24. But since the pumps were off for Niskin 24, the CTD data from the intended bottle #22 will be used. This will complicate sorting so care will be needed.

- Cast #29 had sample numbers that had been used on a previous cast; this had already been noted and most analysis for these samples are reported with a leading nine for sample numbers, i.e. 9454-9463. However, the dissolved oxygen samples needed adjusting. There will be no problem sorting.

- Cast #30 had sample numbers that had been used on a previous cast (464-477). Most of them were changed on the rosette sheet so those are correct everywhere except the dissolved oxygen final data file. However samples 475-477 were not adjusted and will need to be changed to 9475-9477 in all files since there is duplication. This cast will take care in sorting.

- There is no rosette sheet for event 43 – there was just 1 bottle.

- Cast #51 had no rosette sheet. 24 bottles were fired but only 1 was assigned a sample number.

- Sample #839 was used for both casts #52 and 53. The sample number for event 52 was changed to #9839 which will not complicate processing and the only sampling noted for this was “phage”.

- There was no rosette file for cast #54 but there should be one. The BL file was empty so conversion was repeated using “Scans marked with bottle confirm bit”. This worked. This was a cast with bottles fired on the fly, so the first attempt was made using data from the 5s immediately before the bottle was fired.

- Cast #55 is confusing since bottles were fired out of order. Once the list is edited on sample number it is ok.

- Cast #57 had no rosette sheet. All bottles were fired but sample #s were only assigned for 1.

The addsamp.csv file was converted to CST files, which will form the framework for the bottle files.

The CST files for casts #28 and 30 were reordered on sample number using a text editor.

Those files were then bin-averaged and called SAMAVG.

Next, each of the analysis spreadsheets were examined to see what comments the analysts wanted included in the header file. These were used to create file 2007-19-bot-hdr.txt which will be updated as needed during processing.

Extracted Chlorophyll/Phaeo, Ammonium, DIC/Alkalinity, Nutrients

These bottle data were obtained in the general chemistry file QF2007-19_LSSL_Chem_RAW*.xls which included comments, flags and a precision study. The spreadsheet was simplified by removing some columns and the file was saved as 2007-19chem.csv, which was then converted to individual *.comb files. There were no comments about individual samples.

DISSOLVED OXGYEN

Dissolved oxygen data were provided in spreadsheet “DO_Values_2007-19*.xls” which includes flags, comments and a precision study. Draw temperatures were not available. The spreadsheet page with the final data was simplified and the file was then saved as 2007-19oxy.csv. There were some corrections based on the problems listed above. There is a comment about sample 158 but the flag is by sample 159. The flag was moved. Similarly the flag by 282 was moved to 287 that had a comment that would suggest need of a flag. After corrections the spreadsheet was converted to individual *.OXY files.

SALINITY

Salinity analysis was obtained in spreadsheet Salts 2007-19 for merging*.xls. Analysis was done within 9-22 days of collection. The file was simplified and saved as 2007-19sal.csv. There were no event numbers; those were added. There were no flags; since 3 sets of duplicates were rejected as Chauvenet outliers, those samples were flagged 3 with a comment. Changes were made to sample numbers in casts 7 and 8 as outlined above. Pad values, flag 9 and comments were added for missing data. The file was then ordered on cast number and sample number. Unnecessary columns were then removed. The file was then converted to individual SAL files.

The SAL, OXY and COMB files were merged with CST files in 3 steps.

After the 3rd step the files were put through CLEAN to reduce the headers to File and Comment sections only.

The merged files are ordered on sample number, but the SAMAVG files are ordered on bottle number, so one or the other set needs to be reordered in order to merge them. The MRGCLN1 files were reordered on Bottle_Number. The output files were named MRGCLN1s. Those files were then merged with SAMAVG files choosing the Bottle_Number from the SAMAVG files.

The output of the MRG files were exported to a spreadsheet and compared to the rosette log sheets to look for omissions and to see if the CTD data looked ok. A few problems were found:

- The secondary conductivity and salinity were bad for some of the cast #5 bottles. The pad value for the conductivity was non-standard (99.0000 rather than -99.0000). The SAMAVG files were edited using a text editor, but a check of the full profile shows that the secondary conductivity is bad in large parts of the cast, so it is most unlikely that the secondary salinity will be archived for this cast.

- As noted earlier the secondary dissolved oxygen sensor malfunctioned from cast 15 to the end.

- Cast #2, Niskin #1 was not used but is in the bottle file with no sample number, so that line was removed from the cast.

- Cast #27 has pumps off for bottle #24. Samples were taken from that bottle but should be linked to Niskin 22 because the CTD data for that bottle stop are better, having pumps on. An error had been made that associated the samples from Niskin 24. This was fixed.

- Casts #30 had 3 samples that repeated sample numbers – some were noted on the rosette sheet, but 475-477 had not been. They needed to be changed to 9475-9477.

- There are large differences between the CTD and Bottle DO at casts 17 and 22. The bottle data agree with the data in the spreadsheets. Some of the samples were flagged 3 but not most.

- The sampling for cast #1 does not match the log sheet, but the latter had many erasures so does not look reliable. Comparison with the rosette log sheets was abandoned because many of the samples ticked on the sheets are not found in the chemistry spreadsheet, so likely they were either not taken or were lost.

- The biggest problems were with cast #55; these were only discovered in the COMPARISON stage. The log sheet has a column for Niskin # but the numbers entered are neither Niskin Bottle 3s nor the order of firing. The bottles were fired out of order, so the ADDSAMP file had to be adjusted.

The ADDSAMP file was corrected and reconverted to CST files and the merge process was rerun after the corrections.

CLEAN was run to reset header values, reduce the headers to just history and comments and enter 0 in empty flag channels.

##### Compare

Salinity

Compare was run with pressure as reference channel.

First, outliers with differences >0.2psu were eliminated from the comparison.

For the primary salinity these were mostly isolated bottles and mostly near the surface where the standard deviations in the CTD data were high. One exception was cast #55 which had a group of bottles that are out of line. Examination of the files showed that the entries in the Niskin bottle column on the rosette log sheet are not really Niskin bottle #s nor are they firing order numbers. A return was made to the previous step, fixing an error in the ADDSAMP file and rerunning the MERGE steps for that file. That was fixed to correspond to the ROS file and the merges were rerun.

After the correction all the outliers in the primary salinity comparison with differences >0.2psu were found to be fairly shallow and to have high standard deviation in the CTD salinity.  When more data were removed based on the standard deviation in the CTD salinity being >0.0008psu, the CTD salinity was found to be lower than the bottle salinity by an average of 0.009psu, but there is significant pressure dependence in the fit. One of the outliers removed was more out of line than most despite the standard deviation being only slightly out of line so it was investigated. There was very high variability during the bottle stop, just not during the 10s included in the fit.

There are many points in the fit that have differences >0.01psu but they are all near the surface. This is likely due to incomplete flushing of Niskin bottles in the presence of higher salinity gradients. When only data from 500db down are included the primary salinity was low by an average of 0.0019psu with very little pressure dependence.

It is difficult to assess calibration drift with time but when bottles at 750db were compared, excluding cases with CTD primary salinity standard deviation >0.002, there is a hint of slight drift with the fit showing the CTD low by about 0.0014psu at the beginning and by about 0.0019psu at the time of the last bottle at that depth. Extended to the end of the cruise that would imply salinity was then low by about 0.0021psu. That fits quite well with the general fit of being low by 0.0019psu. Because Niskin flushing was likely better in the Atlantic than in the Arctic, the trend may not be due to calibration drift.

The secondary salinity comparison looks very different with many extreme outliers that are not associated with high standard deviation in the CTD salinity; rather, they have standard deviation values of 0. They are not restricted to the top 500db. As was noted during the cruise and in the initial assessment of the data, these appear to be due to instrument malfunction. Most of the problems were in the latter half of the cruise, but there was also a problem during cast #5. When the outliers are excluded the secondary salinity is lower than bottles by an average of 0.0013psu showing that sensor calibration drift is not significant. Problems encountered with the secondary dissolved oxygen sensor suggest a more general problem with the secondary system.

There was salinity and oxygen sampling for only 1 of the casts that were run with sampling on the fly. That cast does suggest the bottles contained water from somewhat deeper in the water column, but not far below. There are not enough data to justify adjusting the window used in conversion.

For full details for the COMPARE run see file 2007-19-sal-comp1.xls.

Dissolved Oxygen

COMPARE was run with pressure as the reference channel.

For the primary DO sensor the results look good overall, but for casts #17 (BB3), 18(BB4) and 22 (BB6) the results are bad with the CTD DO reading higher than the titrated DO. The bottom bottles look ok for those casts and a few mid-depth bottles are ok for #17. No error was found in the ADDSAMP file. The nutrients look normal. There is a mid-depth salinity maximum but the T-S plots show that both salinity bottles and the primary CTD salinity variations are stable. So there is no evidence of sample numbers being confused in salinity or nutrients. It is possible that there is a mix-up in the DO samples, but there is no evidence that happened. It is more likely that the sampling/analysis problems during those 3 casts were more severe than realized.

For the secondary CTD DO sensor results are bad for all bottles for 24 casts: 15, 17, 18, 21, 22, 23, 26, 27, 31, 32, 33 and 38 to 62.

Most data are good for casts #3, 5, 7, 8, 10, 11, 13, 16, 25, 29, 30, 34, 35 and 36.

The cases with bad dissolved oxygen have oxygen voltage values that are constant and at the maximum possible value.

The fit found when outliers are excluded based on CTD DO being higher than bottles and then based on residuals was:

CTD DO1 Corrected = CTD DO1 * 1.0604 + 0.0960

It is not anticipated that the secondary DO will be archived. When outliers were removed from those data in the same way there were only 56 values left (compared to 204 for the primary) and the fit was:

CTD DO2 Corrected = CTD DO2 * 1.0788 + 0.0539

Plots of Titrated DO and CTD DO against CTD salinity were examined. The only major outliers were from the 3 casts that were already discussed. Minor outliers all came from near the surface and for most the DO profile was complex with a sub-surface peak. For a few it looks like the Niskin bottle may not have flushed efficiently. There is no evidence of problems with analysis or sampling.

No further flagging is suggested.

A check for hysteresis in the deep casts showed some lower CTD values below 1500m, but there are not many bottles available to quantify the problem. It does appear that at 3500db the primary DO is reading low by about 0.2mL/L. For sensors calibrated before 2008 there is no hysteresis correction available.

Fluorescence

COMPARE was run with extracted chlorophyll and CTD Fluorescence using pressure as the reference variable. The CTD fluorometer was a SeaPoint sensor.

.

The ratio Fluor/CHL versus CHL was examined but cases of very low CHL were excluded since the ratios are very large due to the fluorometer never getting lower than 0.046ug/L. As is usually found for these sensors the fluorometer often reads high at low CHL, is close to the CHL between 1ug/L and 4ug/L and lower above 4ug/L.

For full details of the comparison see file 2007-19-fl-chl-comp1.xlsx.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only in the full cast files (*.CNV).

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

There are only 5 casts without stops for bottles and 3 of those have few distinguishing features to help judge alignment. Those were used to test which setting brings the SBE:DO trace into best alignment with temperature. For the primary DO a setting of +3.5s looks best overall. There was only 1 cast with enough good data to allow tests for the secondary DO sensor. For that one a setting of +5s looks like the best choice. ALIGNCTD was run on all casts using +3.5s for DO1 and +5s for DO2.

##### CELLTM

The noise in the upcast makes the tests for the best parameters for this routine very difficult to interpret. The default setting of (α = 0.0245, β=9.5) was used. One cast was checked for this cruise and the default setting does improve the data.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE and Channel Comparisons

Program DERIVE was run on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

DERIVE was run a second time on a few of the deeper casts to examine differences between sensor pairs. For casts #7 and 8 the differences varied greatly and results a short distance above or below are quite different, so the results are not significant. Only cast #7 is shown below as an example.  For cast #30 the conductivity data were bad around 500db.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2007-19-0007 | 500 | ~-0.0008 VN | 0.000 VN | +0.0004 VN | Noisy, High |
| “ | 1000 | ~-0.0005 VN | +0.004 VN | +0.0005 VN | “ |
| “ | 2000 | ~0.0000 VN | +0.004 VN | +0.0188 VN | “ |
|  | 2800 | ~0.0005 VN | +0.003 VN | +0.0416 VN | “ |
| 2007-19-0030 | 500 | +0.0015 | C1 off-scale | Sal1 off-scale | Steady, High |
| “ | 1000 | -0.0001 | +0.0001 | +0.001 | “ |
| “ | 2000 | +0.00005 | +0.0001 | +0.001 | “ |
| 2007-19-0031 | 500 | +0.0002 | +0.0001 | +0.001 | Steady, High |
| “ | 1000 | +0.00005 | +0.0001 | +0.001 | “ |
| “ | 2000 | +0.0001 | +0.0001 | +0.001 | “ |

The differences in temperature are small except for the 500db data from cast #30 – all channels there are bad. Conductivity differences were large for most depths of cast #7 and 2 levels of cast #30 but small at 2000db for #30 and all of cast #31. Similarly, salinity differences were very high for some of casts 7 and 30 but small for most of cast #30 and all of cast #31. The differences in salinity for casts #30 and 31 are reasonably close to the results of the salinity bottle comparison which suggested that the primary salinity was lower than the secondary by 0.0006psu.

##### Conversion to IOS Header Format

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number. Fluorescence profiles showed sections where the sensor had clearly gone off-scale with values staying constant for about 1m at 4.905ug/L, so CLEAN was used to replace values>4.090ug/L with pad values since fluorescence likely went above 5.

Since the NMEA positions were not acquired until cast #38 it is necessary to add them. The same approach was used as was applied to bottle files and described in section 4.  Program “Merge CSV files to Headers” using file 2007-19-merge-hdr.csv was run on all casts, even those that already had NMEA data. The entries are consistent with the log book entries.

##### Checking Headers

A cross-reference list was checked against the log book. There were station name errors in casts 2, 3, 4, 5, 6, 7, 9 and 49. Those were corrected in the header spreadsheet.

The cruise track was plotted twice, once with cast numbers and once with station names, and both were added to the end of this report. Cast #1 was a test cast and did not start until the CTD was at 37db, so it will not be processed further and was not included in the plots.

Surface check was run and shows an average surface pressure for the cruise was 0.755db which is lower than expected. There were a number of surface values <0.5db; it is likely that the CTD was deeper than that when acquisition was started. The shallowest data are recorded when the pumps are off so it is not clear when the CTD is in or out of water and pressure sensor may not be fully equilibrated at the beginning of files. At the end of casts the pumps are off and it is possible that the CTD looks like it is in water when it is really in air. There is no clear evidence of any error in pressure.

A header check was run. The header ranges show many signs of spikes, as expected. The lowest pressure is -0.0136db during cast #11. It occurred near the end of a cast with pumps off. The conductivity is very low and transmissivity close to 0, so it does appear that the CTD was either right at the surface or above the surface.

There were no negative fluorescence values and it looks like the CLEAN step removed off-scale data, though that needs to be checked on plots.

The altimeter and water depth readings from the headers of the MRH and SAMAVG files were exported to a spreadsheet. Plots were made of altimetry near the bottom for about half of the casts. In many cases there were many spikes in altimetry at the bottom, so an estimate was made by extrapolating the results in the 10 to 30m above the bottom. All the header entries look appropriate within the expectations for these data. At least one cast showed signs of shoaling while the CTD was at the bottom.

The entries in the bottle files were also checked and were close to those of the MRH files except for one cast which was checked and the deepest bottle had been ~3 metres above the bottom, so the entry looks correct.

A few water depth entries were checked and no errors were found. They were taken from the chemistry spreadsheet except for the few casts with no bottles and those were taken from the log book directly.

##### Shift

Before running the alignment steps, the initial soak records were removed. First all MRH files were copied to *.CLIP since a few casts had no subsurface soak.

The number of records to be removed varied from cast to cast and plots of pressure versus scan number were used to make estimates. CLIP was used to remove those data.

Fluorescence

The Fluorometer was pumped. Plots were made to see if alignment is required, but with most casts having many stops for bottles there is little reason to expect meaningful results by comparing upcasts with downcasts. Trying to match features in downcast profiles with those in temperature suggests that the a small advance might improve the data. Generally these sensors look better aligned with a 1s advance, so that setting was selected.

SHIFT was run on the SeaPoint fluorescence channel, advancing it by 24 records.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier. A few casts were checked to see if the alignment looked ok, and it did. No further alignment is needed for the DO concentration channel.

Conductivity

Tests were run on 4 casts using a variety of shifts. For the primary conductivity the best results were found with -0.2 or -0.3 records. For the secondary a setting or -0.2 looked best.

SHIFT was run on all casts using -0.25 records for the primary conductivity and -0.2 records for the secondary conductivity.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for 3 casts that were caused by pressure increasing very fast. These look like skipped records.

- Cast #8 involves downcast data. There are jumps in pressure of 5db, 10db and 2db at depths of about 1500m. In one case the conductivity are bad around the time of the jump.

- Cast #10 involves upcast data with a 2.5db jump during upcast around 415m.

- Cast #11 involves upcast data with a 2db jump around 925db.

The pressure increases are associated with impossibly high descent rates so they look like cases of missing records. The problems are seen in the raw data before any processing steps such as alignment. The upcast problems will not affect the profiles but for cast #8 there will be some gaps.

##### Other Comparisons

Other experience with these sensors –

The temperature, conductivity and dissolved oxygen sensors had factory calibrations before and after the cruise. During the cruise that followed the conductivity was adjusted in a way that would have raised the salinity by about 0.002.

Repeat Casts – There are some repeat casts but in most cases at least one is too shallow to expect good repeatability. A few pairs that reached to at least 1000db were examined. In each case the two casts are about 2.5 hours apart. Comparisons were made by looking at small, deep sections of data on T-S plots:

- Casts #7 and 8: Both got to at least 3000db. At about 2800db the temperatures vary by about 0.006C° and salinity by 0.001 along lines of constant sigma-T, though in some areas the two traces are indistinguishable.

- Casts #20 and 21: Both got to ~1000db. At 900db there were differences of ~0.001C° and 0.0001psu.

- Casts 27 and 28: Both got to ~1000db. At 900db there were differences of ~0.01C° and 0.001psu.

These comparisons show good repeatability.

Post-Cruise Calibration – There were post-cruise calibrations available for the temperature, conductivity and pressure. There was no significant drift in the temperature sensors. The primary conductivity had drifted down by 0.0003psu per month, so by the time of the cruise would be low by 0.0018psu. The secondary showed no drift. Pressure was recalibrated in March 2008 with a slope of 1.00001 and offset of -1.2740. The slope change is insignificant, but the offset shows pressure to be too high by ~1.3db. The previous calibration was in late 2002 so we might expect most of the drift would have occurred before this cruise. However, the next calibration in May 2009 shows pressure to be too high by ~0.7db, so perhaps the variability in the offset makes it unreliable to apply to a particular cruise. There were no post-cruise drift estimates for the dissolved oxygen since both had repairs done before calibration.

##### DETAILED EDITING

Since the secondary salinity channel is often bad, the primary temperature and salinity channels were selected for archiving, and hence, editing.

CTDEDIT was used to remove large spikes, remove or clean smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes. All files required some editing.

##### Initial Recalibration

The surface pressure study suggests that the pressure could be reading slightly low while the post-cruise calibration suggests it is too high by 1.3db. If the post-cruise offset were applied there would be many cases of the CTD acquiring data that is clearly in water but having negative pressures. Pressure will not be recalibrated.

There is no need to recalibrate temperature as drift was found to be negligible.

The two methods for estimating how salinity should be recalibrated are in remarkably good agreement; the comparison with bottles show the CTD primary salinity to be low by ~0.0019 while the post-cruise calibration drift suggests it would be low by ~0.0018 at the beginning of July and ~0.0021psu by the end of July. The post-drift estimate for the secondary salinity is ~0 and the comparison with bottles showed it to be low by an averaged of ~0.0013psu, so no recalibration will be applied to that channel.

Dissolved oxygen recalibration is based only on the comparison with bottles:

CTD DO1 Corrected = CTD DO1 * 1.0604 + 0.0960

It is not anticipated that the secondary DO will be archived, but the channel was recalibrated in case some of the data are found useful at a later date. The correction applied is:

CTD DO2 Corrected = CTD DO2 * 1.0788 + 0.0539

The MRGCCLN2 and SAM files were recalibrated using file 2007-19-recal1.ccf  to add 0.0019 to the primary salinity and to apply corrections to the two dissolved oxygen channels as noted above.

COMPARE was rerun on the salinity and dissolved oxygen data to check that the recalibrations were applied appropriately and they were.

For full details for the COMPARE runs see files 2007-19-sal-comp2.xlsx and 2007-19-dox-comp2.xlsx.

The EDT files were recalibrated.

##### Final Calibration of DO

The initial recalibration of dissolved oxygen corrects for sensor calibration drift. Alignctd corrects for transit time errors. Those 2 steps may partly correct for response time errors, but to see if a further correction is needed, a comparison is made of downcast CTD data to bottle data from the same pressure. Small differences are expected due to ship drift, temporal changes, incomplete flushing of Niskin bottles and noise in CTD data.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned and compared to the bottle values in the MRG files. COMPARE was run to study the differences between the downcast CTD DO data and the titrated samples from upcast bottles. When outliers were removed based on residuals, the CTD DO was higher than the bottles by an average of ~0.041mL/L and standard deviation of 0.051mL/L; the results vary with depth with the CTD DO being low by an average of.-0.0009mL/L below 400m and high by 0.07mL/L in the top 200db. Reading high in the upper ocean is likely to be due to slight inefficiency in flushing of bottle so that bottle samples come from a little deeper in the water column where values are lower. In deeper waters the vertical DO gradient is generally smaller so that any such error is slight. The small differences at depth show that the calibration is satisfactory.

Examination of the differences versus pressure was used to make a rough estimate of the errors in the downcast DO data. (See section 21 for error estimates.)

See 2007-19-dox-comp3.xlsx for the comparison details.

##### Fluorescence Processing

A median filter, size 11, was applied to the fluorescence channel in the COR1 files. Plots of a few casts showed that the filter was effective. (Output:*.FIL)

##### DERIVE and DO saturation

A second SBE DO channel (with umol/kg units) was added.

DERIVE was run twice – first to calculate dissolved oxygen saturation and second to calculate depth.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

On-screen plots were examined. The T-S plots have some unstable features near the surface of a few casts. These are mostly in narrow straits where there may be active mixing so such features may be real.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts: #7, 8, 11, 28, 29, 30, 31, 32, 33 & 60. (The cap was left on during cast #60.)

Nitrate:ISUS was removed from casts: #7, 8, 10, 28, 29, 30, 31, 32 & 33.

The PAR:Reference channel was removed from casts: #2, 4, 5 because there was no signal.

REORDER was run to get the two DO channels together.

At this point some further editing was applied to cast #38. The Transmissivity data below 500m were bad, but they were fine on the upcast. The bad data were replaced with pad values in the REO file.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence, PAR, PAR:Reference (surface PAR) and

Nitrate_plus_nitrite:ISUS:Voltage data are nominal and unedited except

that some records were removed in editing temperature and salinity.

For details on how the transmissivity calibration parameters were calculated

see the document in folder "\cruise_data\documents\transmissivity".

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples. It is

recommended that users check extracted chlorophyll values where available.

Dissolved oxygen was calibrated using the method described in SeaBird

Application Note #64-2, June 2012 revision, except that a small

offset in the fit was allowed.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be

especially true when vertical DO gradients are large. To get an estimate

of the accuracy of the SBE DO data during downcasts (after recalibration)

a rough comparison was made between downcast SBE DO and upcast titrated

samples. Some of the difference will be due to problems with flushing

of Niskin bottles and/or analysis errors, so the following statement

likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.7 mL/L from 0 to 50db

±0.3 mL/L from 50 to 200db

±0.15 mL/L from 200db to 300db

±0.05 mL/L below 300db

For details on the processing see document: 2007-19_Processing_Report.doc.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

A cross-reference list was produced; as noted in the particulars and log, cast #9 occurred before cast #7.

The track plot looks fine.

##### Dissolved Oxygen Study

As a final check of dissolved oxygen data, % saturation was calculated and plotted. Surface values varied from 95% in well-mixed surface waters and between 100 and 130% elsewhere with the highest values in the Canadian Arctic Archipelago where there were subsurface peaks. It is likely that if the CTD data started closes to the surface the values would be closer to 100%. These values suggest that the dissolved oxygen recalibration was reasonably good.

##### Final Bottle Files

The MRGCOR2 files were put through SORT to order on increasing pressure.

REMOVE was run on all casts to remove the following channels:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts: #7-8, 11 and 28-33.

The PAR:Reference channel was removed from casts: #2, 4 and 5 because there was no signal.

Editing was applied to a few files:

Cast #27 – Niskin #2 did not close, so that line was removed.

Cast #51 – All lines except #12 were removed from the MRGREM file – only bottle #12 was sampled but there are no samples in file. CTD data may be needed.

Cast #57 – All lines except #1 and 13 were removed from the MRGREM file – water from all bottles were combined but there are no samples in file. CTD data may be needed and data from 2 bottles were requested to be included in the file.

A second SBE DO channel was added for the CTD DO with mass units. Draw temperatures were not available, so mass units could not derived for the titrated DO.

DERIVE was run to add depth.

REORDER was run to get the 2 SBE DO channels together and put Depth after Pressure and to rename Voltage6 as Nitrate_plus_nitrite:ISUS:Voltage.

CLEAN was run to remove empty channels and to update channel ranges.

HEADER EDIT was run to ensure formats and units are correct, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and a few notes about the CTD data processing.

A separate header comment was used for casts 46, 47, 49 & 54 because there was no stop for those bottles.

One line was removed from cast #2 because there was no sample number and no sampling.

Standards check did not indicate any problems.

The header check showed that the ISUS data were missing from cast #54 for some unknown reason. That cast was reprocessed WITH that channel.

The track plot looks ok.

Plots of each file were examined to ensure no problems had crept in and none were found.

A cross-reference listing was produced for the CHE files.

Data were exported from the CHE files to file 2007-19-bottles-final.xlsx.

# Particulars

Note that the file names are based on cast numbers which were assigned consecutively to CTD casts. They are not based on the event numbers in the log book.

1 & 1a. Test cast – second has altimeter but acquisition started late. Neither to be archived.

2. Discard bottle #1 – Niskin slipped

2.  PAR and SPAR coefficients entered

7 & 9. Cast #9 was run before cast #7

10. Change of con file – same sensors but new coefficients for fluorometer and DO

24. Three attempts to get cast – use 24c.

27. Niskins #2 and 22 did not trip. Water was taken from Niskin #24 to replace #22 but the pumps were off when #24 was fired so the CTD data are from the firing for Niskin #22. The CTD data for Niskins #2 and #24 were removed from the bottle file.

29. Sample numbers were repeats of the previous CTD cast so a leading 9 was added to sample numbers.

38. NMEA turned on. Transmissivity bad during part of downcast but ok coming up.

46, 47, 49, 54 - No stop for bottle firing.

60. PAR cover left on and Nitrite =0 around 30m on way up.

# CRUISE SUMMARY

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 724 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4322 | 5Jan2007 | Factory | 9Jan08 | Factory |
| Conductivity | 2809 | 9Jan2007 | Factory | 28Dec07 | Factory |
| Secondary Temp. | 4239 | 30Dec06 | Factory | 3Jan08 | Factory |
| Secondary Cond. | 2810 | 9Jan07 | Factory | 4Jan08 | Factory |
| Fluorometer | SCF2569 | n/a |  |  |  |
| Transmissometer | CST-662DR | 28aug06 | Factory | 26May08 | Factory |
| SBE 43 DO sensor | 435 | 25Jan07 | Factory | * |  |
| SBE 43 DO sensor | 1115 | 17Oct06 | Factory | * |  |
| QSP2300 PAR | 70123 | 13Mar07 | Factory | 4Apr16 | Factory |
| QSR2200 SPAR | 20279 | 13Mar07 | Factory | 4Apr16 | Factory |
| PSA916D Altimeter | 1161 | n/a |  |  |  |
| Pressure Sensor | 90559 | 29Oct2002 | Factory | 11Mar08 | Factory |
| Primary Pump | 3610 |  |  |  |  |
| Secondary Pump | 3615 |  |  |  |  |

Note: * - repaired before next calibration after cruise