## Notes
2007-19 Oxygen data
2 Mar 2017 SZ
New terms for potassium iodate volume and flask volume. Terms were applied, generating new oxygen values. These values were reformated and are ready for merging.
Oxygen formula for excel provided by Kenny Scozzafava
Original data: The data have been added from files generated by IOS Shell. Data from N:\SHARE\DATA\2007\2007-LSSL-all\2007-19_C3O\Data\Chemistry\oxygen\02Nov2007\processed oxy data
New flask volumes have been measured Jan 2009 by Tara Lamothe and Bernard Minkley. These new volumes were entered in this spreadsheet and new oxygen values calculated. Flasks 76, 1194 and 3161 had stopper changes since the Jan 2008 so old values are best. New flask volumes for 3003,3004,3048,3082,3092,3123,3183 and 3186 were missing so old values kept.
Flask volume info based on information from: N:\SHARE\CHEMISTRY- Methods and Calibrations\Oxygen\Flask Calibrations\2009 Arctic flasks.xls Sheet Name: "Flasks 2009"
New Sp calculated: Sp - March 2017
"Final" data are in: Data - March 2017 version
Sheets: Timeline, Standards and Chemicals added to make a complete workbook. Sheets are copied from N:\SHARE\DATA\2007\2007-LSSL-all\2007-19_C3O\Data\Chemistry\oxygen\02Nov2007\Oxygen_Data_File_2007-19_final.xls
Analysis notes updated with these changes: AmendedC3O_OxygenAnalysis_Summary v2017-03-02.doc

## Data - March 2017 version
2 March 2017: Recalculated oxygen reformated for merging into chemistry sheet.
Notes comments are copied from earlier spreadsheet. Sometimes the comment refers to multiple samples in a cast but sometimes just one. Its a bit confusing.
Cast | Sample_ Number | Pressure | Oxy (mL/L) (2 Mar 2017) | Flag | Secondary | Secondary Flag | Comment, Nov 2007
3 | 16 | 75 | 7.8931 | 4 | 7.9063 | 4 | Colorimeter probe problems with drift for entire day's run
3 | 18 | 50 | 8.4986 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
3 | 20 | 25 | 8.7134 | 4 | 8.3627 | 4 | Colorimeter probe problems with drift for entire day's run
3 | 22 | 5 | 7.3353 | 4 | 7.8255 | 4 | Colorimeter probe problems with drift for entire day's run
5 | 45 | 214 | 7.5716 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
5 | 48 | 200 | 7.6139 | 4 | 7.8188 | 4 | Colorimeter probe problems with drift for entire day's run
5 | 49 | 150 | 7.6711 | 4 | 7.9293 | 4 | Colorimeter probe problems with drift for entire day's run
5 | 50 | 125 | 7.549 | 4 | 8.0135 | 4 | Colorimeter probe problems with drift for entire day's run
5 | 51 | 100 | 8.2914 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
5 | 54 | 75 | 8.1226 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
5 | 55 | 50 | 7.9362 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
5 | 56 | 35 | 7.7574 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
5 | 57 | 25 | 8.4056 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
5 | 61 | 5 | 10.0441 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
7 | 78 | 2000 | 6.7089 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
7 | 79 | 1500 | 7.0498 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 101 | 3459 | 7.0208 | 4 | 6.9979 | 4 | Colorimeter probe problems with drift for entire day's run
8 | 102 | 3000 | 5.8909 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 103 | 2500 | 6.3467 | 4 | 6.6389 | 4 | Colorimeter probe problems with drift for entire day's run
8 | 105 | 1000 | 6.8913 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 107 | 750 | 7.0676 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 109 | 500 | 7.0934 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 110 | 400 | 6.9364 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 111 | 300 | 6.9819 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 112 | 250 | 6.9825 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 114 | 200 | 6.9575 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 115 | 150 | 7.1029 | 4 | 7.0874 | 4 | Colorimeter probe problems with drift for entire day's run
8 | 116 | 125 | 7.3031 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 117 | 100 | 7.39 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 119 | 75 | 7.3522 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 120 | 50 | 7.3634 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 121 | 35 | 6.8814 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 122 | 25 | 7.6298 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
8 | 123 | 5 | 7.5377 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 126 | 1000 | 7.0099 | 4 | 6.8602 | 4 | Colorimeter probe problems with drift for entire day's run
10 | 128 | 750 | 6.9344 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 130 | 500 | 6.8956 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 132 | 350 | 7.059 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 133 | 300 | 7.0262 | 4 | 6.8631 | 4 | Colorimeter probe problems with drift for entire day's run
10 | 134 | 250 | 6.8442 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 136 | 200 | 6.8452 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 137 | 175 | 6.8681 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 138 | 150 | 6.9458 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 139 | 125 | 6.8948 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 140 | 100 | 7.0306 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 141 | 75 | 7.149 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 143 | 50 | 7.4393 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 144 | 35 | 7.5823 | 4 | 7.4344 | 4 | Colorimeter probe problems with drift for entire day's run
10 | 146 | 15 | 8.4753 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
10 | 147 | 5 | 8.3686 | 4 |  |  | Colorimeter probe problems with drift for entire day's run
11 | 149 | 1000 | 6.615 |  | 6.5855
11 | 150 | 750 | 6.8039
11 | 151 | 500 | 6.8892
11 | 152 | 350 | 7.1917
11 | 153 | 250 | 7.3968
11 | 154 | 225 | 7.466 | 3 |  |  | samples # 154,155,158 bubbles on probe at end of titration.
11 | 155 | 200 | 7.5348 | 3 | 7.4879 |  | samples # 154,155,158 bubbles on probe at end of titration.
11 | 157 | 150 | 7.5468
11 | 158 | 125 | 7.6004 |  |  |  | samples # 154,155,158 bubbles on probe at end of titration.
11 | 159 | 100 | 7.6699 | 3
11 | 160 | 75 | 7.6502
11 | 161 | 50 | 7.6942
11 | 162 | 35 | 7.8728
11 | 163 | 25 | 8.2014
11 | 164 | 15 | 8.1603
11 | 165 | 5 | 8.0147
13 | 192 | 585 | 5.2193 |  | 5.5606
13 | 196 | 500 | 5.4787
13 | 197 | 350 | 6.1152
13 | 199 | 250 | 6.4241
13 | 200 | 225 | 6.6948
13 | 201 | 200 | 6.1567
13 | 202 | 175 | 6.1217
13 | 203 | 150 | 6.5695
13 | 204 | 125 | 7.3553
13 | 205 | 100 | 7.5872 |  | 7.5611
13 | 207 | 75 | 8.0327
13 | 208 | 50 | 8.2463
13 | 209 | 35 | 8.4853
13 | 210 | 25 | 8.9089
13 | 211 | 15 | 9.1381
13 | 213 | 5 | 9.189
15 | 234 | 247 | 7.1053 |  | 7.1101
15 | 235 | 225 | 7.0826
15 | 236 | 200 | 7.0772
15 | 237 | 175 | 7.2455
15 | 238 | 150 | 7.3344 |  | 7.3721
15 | 239 | 125 | 7.3968
15 | 240 | 100 | 7.4648
15 | 241 | 75 | 7.7085
15 | 243 | 50 | 7.9102 | 3 |  |  | Sample #243 had bubble on probe, sample #246,247 had bugs in sample.
15 | 244 | 35 | 8.3959 |  | 8.3596
15 | 245 | 25 | 8.7255
15 | 246 | 15 | 8.603 | 3 |  |  | Sample #243 had bubble on probe, sample #246,247 had bugs in sample.
15 | 247 | 5 | 8.5794 | 3 |  |  | Sample #243 had bubble on probe, sample #246,247 had bugs in sample.
16 | 250 | 250 | 6.3593 |  | 6.3539
16 | 251 | 225 | 6.4727
16 | 252 | 200 | 6.7323 |  | 6.4882
16 | 253 | 175 | 6.8189
16 | 254 | 150 | 6.8732
16 | 255 | 125 | 7.0662
16 | 256 | 100 | 7.4943
16 | 257 | 75 | 7.1839 | 3 |  |  | Sample #257 had leaky burette.
16 | 258 | 50 | 7.7271 |  | 7.7261
16 | 259 | 35 | 8.2979
16 | 260 | 25 | 8.8304
16 | 261 | 15 | 8.547
16 | 262 | 5 | 8.4515
17 | 263 | 1000 | 4.3929 |  | 4.3926
17 | 264 | 750 | 4.2236
17 | 265 | 500 | 5.8182 |  | 5.7546
17 | 266 | 350 | 6.1751
17 | 267 | 250 | 6.0143
17 | 268 | 225 | 3.8896
17 | 269 | 200 | 4.1959
17 | 270 | 175 | 4.3651
17 | 271 | 150 | 4.6396
17 | 272 | 125 | 6.669
17 | 273 | 100 | 3.7918
17 | 274 | 75 | 6.9073
17 | 275 | 50 | 5.9832
17 | 276 | 35 | 4.4567
17 | 277 | 25 | 8.0818
17 | 278 | 15 | 4.5906
17 | 279 | 5 | 8.9241 | 3 | 5.8849 | 3 | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination?
18 | 280 | 1000 | 4.4671 |  | 2.567 | 4 | Sample #280 duplicate is poor. Triplicate agrees with primary.
18 | 281 | 750 | 3.8842
18 | 282 | 500 | 3.4017 | 3
18 | 283 | 350 | 4.4488
18 | 284 | 250 | 3.207
18 | 285 | 225 | 3.1694
18 | 286 | 200 | 3.2794
18 | 287 | 175 | 3.6268 |  | 4.2517 | 3 | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks before analysis. Duplicate sample had less precipitate in flask than others.
18 | 288 | 150 | 3.8595
18 | 289 | 125 | 4.4755
18 | 290 | 100 | 4.6779
18 | 291 | 75 | 4.2659 | 3 |  |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks before analysis. Duplicate sample had less precipitate in flask than others.
18 | 292 | 50 | 3.9136 | 3 |  |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks before analysis. Duplicate sample had less precipitate in flask than others.
18 | 293 | 35 | 5.4431 | 3 |  |  | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
18 | 294 | 25 | 7.0911 | 3 |  |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks before analysis. Duplicate sample had less precipitate in flask than others.
18 | 295 | 15 | 4.3938 | 3 | 4.9675 | 3 | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks before analysis. Duplicate sample had less precipitate in flask than others.
18 | 296 | 5 | 6.9675
21 | 322 | 1000 | 4.5747 |  | 4.5042
21 | 324 | 750 | 5.0423
21 | 326 | 500 | 5.4988 |  | 5.4555
21 | 327 | 350 | 5.5278
21 | 329 | 250 | 5.7271
21 | 330 | 225 | 5.8858
21 | 331 | 200 | 6.1502
21 | 332 | 175 | 6.5115
21 | 333 | 150 | 7.4268
21 | 334 | 125 | 7.5691
21 | 335 | 100 | 7.676
21 | 337 | 75 | 7.5585
21 | 338 | 50 | 7.7105
21 | 339 | 35 | 8.1046 | 3 |  |  | Sample #339 had small bubble on probe at end of titration.
21 | 340 | 25 | 8.3269 |  | 8.5609
21 | 341 | 15 | 8.6545
21 | 343 | 5 | 8.8026
22 | 344 | 1000 | 4.561 |  | 3.7262 | 4 | Sample #344 duplicate poor based on primary and triplicate values.
22 | 345 | 750 | 3.5993
22 | 346 | 500 | 3.7768
22 | 347 | 350 | 2.9788
22 | 348 | 250 | 3.6549 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 349 | 225 | 3.2382 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 350 | 200 | 3.5842 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 351 | 175 | 3.8766 |  | 3.7527
22 | 352 | 150 | 3.3595
22 | 353 | 125 | 4.0074 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 354 | 100 | 3.247 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 355 | 75 | 3.7549
22 | 356 | 50 | 3.3877 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 357 | 35 | 3.5134 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 358 | 25 | 4.1873
22 | 359 | 15 | 4.0574 | 3 |  |  | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
22 | 360 | 5 | 4.2298 |  | 4.6245
23 | 361 | 1000 | 4.5834 |  | 4.635
23 | 362 | 750 | 5.0793
23 | 363 | 500 | 5.4208
23 | 365 | 350 | 5.4648 |  | 5.5158
23 | 367 | 250 | 5.9533
23 | 368 | 225 | 6.4066
23 | 369 | 200 | 6.9664
23 | 370 | 175 | 7.4389
23 | 371 | 150 | 6.8765
23 | 372 | 125 | 7.3478 |  | 7.2413
23 | 373 | 100 | 7.3087
23 | 374 | 75 | 7.4137
23 | 375 | 50 | 7.6683
23 | 376 | 35 | 8.3616
23 | 377 | 25 | 8.4025 | 3 |  |  | Sample #377 had leaky stopper... maybe DMQ or air in sample.
23 | 378 | 15 | 8.7143
23 | 379 | 5 | 9.2958 |  | 9.3086
25 | 399 | 150 | 7.0588
25 | 400 | 125 | 7.1152 |  | 5.208 | 4 | Sample #400 primary swapped with duplicate as its likely bad based on triplicate value.
25 | 401 | 100 | 7.3572
25 | 402 | 75 | 7.7639
25 | 403 | 50 | 8.7745 |  | 7.914
25 | 404 | 35 | 8.5229
25 | 405 | 25 | 9.8174
25 | 406 | 15 | 9.7866
25 | 407 | 5 | 9.5148
26 | 408 | 1000 | 4.4915 |  | 4.4777
26 | 410 | 750 | 4.8946
26 | 412 | 500 | 5.4745
26 | 414 | 350 | 5.6014
26 | 415 | 250 | 5.749
26 | 417 | 225 | 5.827
26 | 418 | 200 | 6.0176 |  | 6.0282
26 | 419 | 175 | 6.3666
26 | 420 | 150 | 7.2438
26 | 421 | 125 | 7.6017 |  | 7.5318
26 | 422 | 100 | 7.7058
26 | 423 | 75 | 7.3241
26 | 425 | 50 | 7.6515
26 | 426 | 35 | 7.8734
26 | 427 | 25 | 8.144
26 | 429 | 5 | 9.2249
27 | 431 | 1000 | 4.4824 |  | 4.5133
27 | 433 | 750 | 4.9744
27 | 435 | 500 | 5.3891
27 | 437 | 350 | 5.5277
27 | 438 | 250 | 5.6202
27 | 440 | 225 | 5.7382 |  | 5.7616
27 | 441 | 200 | 6.0594 | 3 |  |  | Sample #441 had bugs in sample.
27 | 442 | 175 | 6.7129
27 | 443 | 150 | 6.8389
27 | 444 | 125 | 7.1927
27 | 445 | 100 | 7.3768
27 | 446 | 75 | 7.597
27 | 448 | 50 | 7.8707
27 | 449 | 35 | 8.1195
27 | 450 | 25 | 8.4352
27 | 451 | 15 | 8.8039
27 | 452 | 5 | 9.1751
29 | 454 | 150 | 6.13
29 | 456 | 125 | 6.8363
29 | 457 | 100 | 7.0703
29 | 458 | 75 | 7.3616
29 | 459 | 50 | 7.6789
29 | 460 | 30 | 8.0942
29 | 461 | 20 | 9.0054
29 | 462 | 10 | 9.8346 |  | 9.8723
29 | 463 | 5 | 9.7958
30 | 464 | 2350 | 2.7723
30 | 466 | 2250 | 2.7752 |  | 2.8181
30 | 467 | 2000 | 2.8209
30 | 468 | 1750 | 2.9579
30 | 469 | 1500 | 3.2675
30 | 470 | 1300 | 3.6439
30 | 471 | 1200 | 3.9478
30 | 472 | 1100 | 4.2221
30 | 473 | 1000 | 4.4675 |  | 4.5223
30 | 474 | 900 | 4.6487
30 | 475 | 800 | 4.8351
30 | 476 | 700 | 5.0188
30 | 477 | 600 | 5.1707
30 | 478 | 500 | 5.335
30 | 479 | 400 | 5.3764
30 | 480 | 350 | 5.3884
30 | 481 | 300 | 5.4667
30 | 482 | 275 | 5.4322
30 | 483 | 250 | 5.444
30 | 484 | 225 | 5.5027
30 | 486 | 200 | 5.5024
31 | 488 | 2045 | 2.742
31 | 489 | 2000 | 2.7847 |  | 2.7065
31 | 490 | 1750 | 2.8083
31 | 491 | 1500 | 3.1473
31 | 492 | 1250 | 3.7122
31 | 493 | 1000 | 4.4166
31 | 494 | 750 | 4.8983
31 | 495 | 500 | 5.4284
31 | 496 | 400 | 5.4711
31 | 497 | 350 | 5.5657
31 | 498 | 300 | 5.5543 |  | 5.5893
31 | 499 | 250 | 5.6873
31 | 500 | 225 | 5.8351
31 | 501 | 200 | 6.2986
31 | 502 | 175 | 7.0413
31 | 503 | 150 | 6.9724
31 | 504 | 125 | 7.135
31 | 505 | 100 | 7.1014 |  | 7.0586
31 | 506 | 75 | 7.479
31 | 507 | 50 | 7.8382
31 | 508 | 35 | 8.157
31 | 509 | 20 | 9.943
31 | 510 | 5 | 9.4548
32 | 511 | 1473 | 2.8876
32 | 512 | 1250 | 3.5723 |  | 3.5195
32 | 513 | 1000 | 4.438
32 | 514 | 750 | 4.7879
32 | 515 | 500 | 5.4522
32 | 516 | 400 | 5.554 |  | 5.6414
32 | 517 | 350 | 5.6425
32 | 518 | 300 | 5.602
32 | 519 | 250 | 5.7373
32 | 520 | 225 | 5.8566
32 | 521 | 200 | 6.2239 |  | 6.3107
32 | 522 | 175 | 6.8949
32 | 523 | 150 | 7.0707
32 | 524 | 125 | 7.2594
32 | 525 | 100 | 7.3295
32 | 526 | 75 | 8.0917
32 | 527 | 50 | 8.2245
32 | 528 | 35 | 9.5916
32 | 529 | 20 | 10.4628
32 | 530 | 5 | 9.3781
33 | 531 | 1042 | 3.3899
33 | 532 | 1000 | 3.5299 |  | 3.4445
33 | 535 | 400 | 5.7195 |  | 5.7832
34 | 558 | 200 | 7.3097 |  | 7.204
34 | 562 | 100 | 7.6632 |  | 7.6829
35 | 570 | 500 | 5.5969 |  | 5.7509 | 4 | Sample #570 replicate likely bad based on triplicate value.
36 | 571 | 150 | 7.4397 | 3 |  |  | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination?
36 | 572 | 150 | 7.4775 |  | 7.5198
36 | 573 | 125 | 7.4896
36 | 574 | 100 | 7.6337
36 | 575 | 95 | 7.6435
36 | 576 | 75 | 7.81
36 | 578 | 50 | 8.2811 |  | 8.3102
36 | 579 | 30 | 8.9475
36 | 580 | 21 | 9.3382
36 | 581 | 15 | 9.5835
36 | 582 | 10 | 9.7679 | 3 | 7.4771 | 3 | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination? Primary and duplicate swapped based on expected profile.
38 | 612 | 250 | 5.7585 |  | 5.7314
38 | 619 | 75 | 7.112 |  | 7.1
39 | 625 | 763 | 4.6537 |  | 4.637
39 | 628 | 500 | 5.4779
39 | 631 | 250 | 5.6895
39 | 633 | 200 | 6.1256
39 | 634 | 150 | 6.7897 |  | 6.7431
39 | 635 | 100 | 7.4385
39 | 638 | 75 | 7.7339
39 | 642 | 50 | 9.1359 |  | 9.0787
39 | 645 | 25 | 9.2035
39 | 648 | 5 | 8.7479
40 | 652 | 150 | 7.2933 |  | 7.251
40 | 655 | 50 | 7.7583 |  | 7.7504
41 | 674 | 300 | 6.0745 |  | 6.1213
41 | 677 | 200 | 6.1531
41 | 678 | 175 | 6.554 |  | 6.5854
41 | 679 | 150 | 6.9148
41 | 680 | 125 | 7.2159
41 | 681 | 100 | 7.1257 |  | 7.2284
41 | 682 | 75 | 7.0069
41 | 683 | 50 | 7.2849
41 | 684 | 35 | 9.6663
41 | 685 | 25 | 10.6197
41 | 686 | 15 | 10.342
42 | 693 | 175 | 6.7899
42 | 694 | 150 | 6.8125 |  | 6.8008
42 | 695 | 125 | 6.8921
42 | 696 | 100 | 6.4839 |  | 6.4655
42 | 697 | 75 | 6.6386
42 | 698 | 50 | 7.0143
42 | 699 | 35 | 8.9902
42 | 700 | 25 | 10.7758
42 | 701 | 15 | 10.2856
45 | 728 | 435 | 5.396
45 | 729 | 350 | 4.9805 |  | 4.9653
50 | 807 | 350 | 5.8846 |  | 5.761
50 | 809 | 250 | 5.8333
50 | 810 | 225 | 5.7657
50 | 811 | 200 | 5.8316
50 | 812 | 175 | 6.0133 |  | 6.0288
50 | 813 | 150 | 6.0898
50 | 814 | 125 | 6.445
50 | 815 | 100 | 6.3405
50 | 816 | 75 | 6.4735
50 | 818 | 35 | 7.4377 |  | 7.4696
50 | 819 | 25 | 7.7939
50 | 820 | 15 | 8.3529
50 | 821 | 5 | 9.7361 | 3 |  |  | Sample #821 had bugs in sample.
52 | 823 | 192 | 4.9041
52 | 824 | 175 | 4.859 |  | 4.8722
52 | 825 | 150 | 5.108
52 | 827 | 125 | 5.4361
52 | 829 | 100 | 5.9156
52 | 830 | 75 | 6.3743 |  | 6.378
52 | 834 | 35 | 7.8548 | 3 |  |  | Sample #834 had bugs in sample.
52 | 835 | 25 | 9.1035
52 | 837 | 15 | 9.5404
52 | 838 | 5 | 9.5521 |  | 9.4411
54 | 861 | 137 | 5.4633
54 | 862 | 100 | 5.7769 |  | 5.8176
54 | 863 | 75 | 6.3242
54 | 864 | 50 | 7.0514
54 | 865 | 35 | 8.3066 | 3 |  |  | Samples #865, 866, 867 had bugs in sample.
54 | 866 | 25 | 9.4846 | 3 |  |  | Samples #865, 866, 867 had bugs in sample.
54 | 867 | 15 | 9.4734 | 3 |  |  | Samples #865, 866, 867 had bugs in sample.
54 | 868 | 5 | 9.3708
55 | 870 | 102 | 8.0418
55 | 871 | 75 | 8.3966
55 | 873 | 50 | 8.315 | 3 | 8.3022 |  | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration.
55 | 875 | 35 | 9.051
55 | 876 | 25 | 9.2635 | 3 | 9.2389 | 3 | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration.
55 | 878 | 15 | 9.4377 |  | 9.476
55 | 879 | 5 | 9.628 | 3 |  |  | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration.
58 | 904 | 104 | 8.491
58 | 905 | 75 | 8.4416 | 3 | 8.5336 |  | Sample #905 had small bubble on probe at end of titration.
58 | 906 | 50 | 8.6576 |  | 8.66
58 | 907 | 35 | 9.0971
58 | 908 | 25 | 9.064
58 | 909 | 15 | 9.2763
58 | 910 | 5 | 8.9968
59 | 912 | 100 | 8.2391 |  | 8.237
59 | 915 | 35 | 8.8836 |  | 8.8917
62 | 941 | 210 | 6.9028 | 3 |  |  | Colorimeter drifting exceeding 100% for entire file.
62 | 942 | 200 | 7.131 | 3 |  |  | Colorimeter drifting exceeding 100% for entire file.
62 | 943 | 175 | 7.0786 | 3 | 7.0776 | 3 | Colorimeter drifting exceeding 100% for entire file.
62 | 949 | 35 | 7.6924 | 3 | 7.6879 | 3 | Colorimeter drifting exceeding 100% for entire file.

## Sp - March 2017
2 March 2017: Calculate Standard Pool from recalculated oxygen
Standard Pool is 0.055ml/l using 69 replicate pairs. Outliers not included in count.
21 outliers were removed in total:
20 outliers were removed due to problems identified with analysis or clearly poor (>1ml/l difference). | [11 empty] | For Chauvenet:
2 outliers were removed using Chauvenet criteria | [11 empty] | n | 71 | Ccritical | 2.6803
Triplicate was not used in standard pool as it was typically run the following day as a stability test. | [11 empty] | mean | 0.0073566
[12 empty] | std | 0.13642
Cast | Sample_ Number | Pressure | Oxy (mL/L) BOTH KIO3 STD vol and flask vol changes | Flag, Nov 2007 | Comment, Nov 2007 | Secondary |  |  | Triplicate |  |  | Diff (All) | Diff: Remove clearly bad values identified by analyst or using triplicates | Chauvenet (Z-score) | Accept /Reject | Diff^2 | Diff for STD comparison
3 | 16 | 75 | 7.8931 | d | Colorimeter probe problems with drift for entire day's run | 7.9063 | d | Colorimeter probe problems with drift for entire day's run | 7.8786 | d | Colorimeter probe problems with drift for entire day's run | -0.013204
3 | 20 | 25 | 8.7134 | d | Colorimeter probe problems with drift for entire day's run | 8.3627 | d | Colorimeter probe problems with drift for entire day's run | 8.498 | d | Colorimeter probe problems with drift for entire day's run | 0.35069
3 | 22 | 5 | 7.3353 | d | Colorimeter probe problems with drift for entire day's run | 7.8255 | d | Colorimeter probe problems with drift for entire day's run | 7.032 | d | Colorimeter probe problems with drift for entire day's run | -0.4902
5 | 48 | 200 | 7.6139 | d | Colorimeter probe problems with drift for entire day's run | 7.8188 | d | Colorimeter probe problems with drift for entire day's run | 7.3791 | d | Colorimeter probe problems with drift for entire day's run | -0.20492
5 | 49 | 150 | 7.6711 | d | Colorimeter probe problems with drift for entire day's run | 7.9293 | d | Colorimeter probe problems with drift for entire day's run |  |  |  | -0.25818
5 | 50 | 125 | 7.549 | d | Colorimeter probe problems with drift for entire day's run | 8.0135 | d | Colorimeter probe problems with drift for entire day's run |  |  |  | -0.46447
8 | 101 | 3459 | 7.0208 | d | Colorimeter probe problems with drift for entire day's run | 6.9979 | d | Colorimeter probe problems with drift for entire day's run | 7.0285 | d | Colorimeter probe problems with drift for entire day's run | 0.02296
8 | 103 | 2500 | 6.3467 | d | Colorimeter probe problems with drift for entire day's run | 6.6389 | d | Colorimeter probe problems with drift for entire day's run |  |  |  | -0.2922
8 | 115 | 150 | 7.1029 | d | Colorimeter probe problems with drift for entire day's run | 7.0874 | d | Colorimeter probe problems with drift for entire day's run |  |  |  | 0.0155
10 | 126 | 1000 | 7.0099 | d | Colorimeter probe problems with drift for entire day's run | 6.8602 | d | Colorimeter probe problems with drift for entire day's run | 6.89 | d | Colorimeter probe problems with drift for entire day's run | 0.14973
10 | 133 | 300 | 7.0262 | d | Colorimeter probe problems with drift for entire day's run | 6.8631 | d | Colorimeter probe problems with drift for entire day's run |  |  |  | 0.16312
10 | 144 | 35 | 7.5823 | d | Colorimeter probe problems with drift for entire day's run | 7.4344 | d | Colorimeter probe problems with drift for entire day's run |  |  |  | 0.14785
11 | 149 | 1000 | 6.615 |  |  | 6.5855 |  |  | 7.9325 |  |  | 0.029514 | 0.029514 | 0.16242 | ACCEPT | 0.0008711 | 0.029514
11 | 155 | 200 | 7.5348 | c | samples # 154,155,158 bubbles on probe at end of titration. | 7.4879 | [5 empty] | 0.046904 | 0.046904 | 0.28989 | ACCEPT | 0.0021999 | 0.046904
13 | 192 | 585 | 5.2193 |  |  | 5.5606 |  |  | 6.308 | c | Triplicate sample # 9192 had bubble on probe. | -0.3413 | -0.3413 | 2.5558 | ACCEPT | 0.11649 | -0.3413
13 | 205 | 100 | 7.5872 |  |  | 7.5611 | [5 empty] | 0.02612 | 0.02612 | 0.13754 | ACCEPT | 0.00068228 | 0.02612
15 | 234 | 247 | 7.1053 |  |  | 7.1101 |  |  | 7.1443 | c |  | -0.0047567 | -0.0047567 | 0.088795 | ACCEPT | 2.2626e-05 | -0.0047567
15 | 238 | 150 | 7.3344 |  |  | 7.3721 | [5 empty] | -0.037713 | -0.037713 | 0.33038 | ACCEPT | 0.0014223 | -0.037713
15 | 244 | 35 | 8.3959 | c |  | 8.3596 | [5 empty] | 0.036281 | 0.036281 | 0.21202 | ACCEPT | 0.0013163 | 0.036281
16 | 250 | 250 | 6.3593 |  |  | 6.3539 |  |  | 6.3444 |  |  | 0.0054305 | 0.0054305 | 0.01412 | ACCEPT | 2.949e-05 | 0.0054305
16 | 252 | 200 | 6.7323 |  |  | 6.4882 | [5 empty] | 0.24415 | 0.24415 | 1.7357 | ACCEPT | 0.059607 | 0.24415
16 | 258 | 50 | 7.7271 |  |  | 7.7261 | [5 empty] | 0.0010087 | 0.0010087 | 0.046532 | ACCEPT | 1.0175e-06 | 0.0010087
17 | 263 | 1000 | 4.3929 |  |  | 4.3926 |  |  | 4.2411 |  |  | 0.00028754 | 0.00028754 | 0.051819 | ACCEPT | 8.2678e-08 | 0.00028754
17 | 265 | 500 | 5.8182 |  |  | 5.7546 | [5 empty] | 0.06362 | 0.06362 | 0.41243 | ACCEPT | 0.0040475 | 0.06362
17 | 279 | 5 | 8.9241 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination? | 5.8849 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination? |  |  |  | 3.0392
18 | 280 | 1000 | 4.4671 |  |  | 2.567 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 4.4607 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks | 1.9002
18 | 287 | 175 | 3.6268 |  |  | 4.2517 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. |  |  |  | -0.62489
18 | 295 | 15 | 4.3938 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 4.9675 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. |  |  |  | -0.57369
21 | 322 | 1000 | 4.5747 |  |  | 4.5042 |  |  | 4.6675 |  |  | 0.070425 | 0.070425 | 0.46231 | ACCEPT | 0.0049597 | 0.070425
21 | 326 | 500 | 5.4988 |  |  | 5.4555 | [5 empty] | 0.043365 | 0.043365 | 0.26396 | ACCEPT | 0.0018806 | 0.043365
21 | 340 | 25 | 8.3269 |  |  | 8.5609 | [5 empty] | -0.234 | -0.234 | 1.7692 | ACCEPT | 0.054755 | -0.234
22 | 344 | 1000 | 4.561 |  |  | 3.7262 |  |  | 4.5114 |  |  | 0.83473
22 | 351 | 175 | 3.8766 |  |  | 3.7527 | [5 empty] | 0.12389 | 0.12389 | 0.85425 | ACCEPT | 0.015349 | 0.12389
22 | 360 | 5 | 4.2298 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 4.6245 | [5 empty] | -0.39472 | -0.39472 | 2.9474 | REJECT
23 | 361 | 1000 | 4.5834 |  |  | 4.635 | [5 empty] | -0.051586 | -0.051586 | 0.43207 | ACCEPT | 0.0026611 | -0.051586
23 | 365 | 350 | 5.4648 |  |  | 5.5158 | [5 empty] | -0.050964 | -0.050964 | 0.42751 | ACCEPT | 0.0025973 | -0.050964
23 | 372 | 125 | 7.3478 |  |  | 7.2413 | [5 empty] | 0.10651 | 0.10651 | 0.72682 | ACCEPT | 0.011344 | 0.10651
23 | 379 | 5 | 9.2958 |  |  | 9.3086 | [5 empty] | -0.012747 | -0.012747 | 0.14737 | ACCEPT | 0.0001625 | -0.012747
25 | 400 | 125 | 5.208 |  |  | 7.1152 |  |  | 7.1492 |  |  | -1.9072
25 | 403 | 50 | 8.7745 |  |  | 7.914 | [5 empty] | 0.86041 | 0.86041 | 6.2532 | REJECT
26 | 408 | 1000 | 4.4915 |  |  | 4.4777 |  |  | 4.6351 |  |  | 0.013813 | 0.013813 | 0.04733 | ACCEPT | 0.00019081 | 0.013813
26 | 418 | 200 | 6.0176 |  |  | 6.0282 | [5 empty] | -0.010613 | -0.010613 | 0.13172 | ACCEPT | 0.00011263 | -0.010613
26 | 421 | 125 | 7.6017 |  |  | 7.5318 | [5 empty] | 0.069949 | 0.069949 | 0.45883 | ACCEPT | 0.0048929 | 0.069949
27 | 431 | 1000 | 4.4824 |  |  | 4.5133 |  |  | 4.4901 |  |  | -0.030908 | -0.030908 | 0.28049 | ACCEPT | 0.00095531 | -0.030908
27 | 440 | 225 | 5.7382 |  |  | 5.7616 | [5 empty] | -0.023426 | -0.023426 | 0.22565 | ACCEPT | 0.00054878 | -0.023426
29 | 462 | 10 | 9.8346 |  |  | 9.8723 | [5 empty] | -0.037711 | -0.037711 | 0.33036 | ACCEPT | 0.0014221 | -0.037711
30 | 466 | 2250 | 2.7752 |  |  | 2.8181 |  |  | 2.815 |  |  | -0.04297 | -0.04297 | 0.36891 | ACCEPT | 0.0018464 | -0.04297
30 | 473 | 1000 | 4.4675 |  |  | 4.5223 | [5 empty] | -0.054803 | -0.054803 | 0.45565 | ACCEPT | 0.0030034 | -0.054803
31 | 489 | 2000 | 2.7847 |  |  | 2.7065 |  |  | 2.7541 |  |  | 0.078207 | 0.078207 | 0.51936 | ACCEPT | 0.0061164 | 0.078207
31 | 498 | 300 | 5.5543 |  |  | 5.5893 | [5 empty] | -0.034977 | -0.034977 | 0.31032 | ACCEPT | 0.0012234 | -0.034977
31 | 505 | 100 | 7.1014 |  |  | 7.0586 | [5 empty] | 0.042856 | 0.042856 | 0.26022 | ACCEPT | 0.0018366 | 0.042856
32 | 512 | 1250 | 3.5723 |  |  | 3.5195 |  |  | 3.589 |  |  | 0.052806 | 0.052806 | 0.33316 | ACCEPT | 0.0027884 | 0.052806
32 | 516 | 400 | 5.554 |  |  | 5.6414 | [5 empty] | -0.087382 | -0.087382 | 0.69446 | ACCEPT | 0.0076355 | -0.087382
32 | 521 | 200 | 6.2239 |  |  | 6.3107 | [5 empty] | -0.086837 | -0.086837 | 0.69047 | ACCEPT | 0.0075406 | -0.086837
33 | 532 | 1000 | 3.5299 |  |  | 3.4445 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others. | 3.4473 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others. | 0.085322 | 0.085322 | 0.57151 | ACCEPT | 0.0072798 | 0.085322
33 | 535 | 400 | 5.7195 |  |  | 5.7832 |  |  | 5.7361 |  |  | -0.063667 | -0.063667 | 0.52062 | ACCEPT | 0.0040534 | -0.063667
34 | 558 | 200 | 7.3097 |  |  | 7.204 |  |  | 7.2579 |  |  | 0.10575 | 0.10575 | 0.72126 | ACCEPT | 0.011183 | 0.10575
34 | 562 | 100 | 7.6632 |  |  | 7.6829 |  |  | 7.6268 |  |  | -0.019623 | -0.019623 | 0.19777 | ACCEPT | 0.00038507 | -0.019623
35 | 570 | 500 | 5.5969 |  |  | 5.7509 |  |  | 5.5988 |  |  | -0.15397
36 | 572 | 150 | 7.4775 |  |  | 7.5198 | [5 empty] | -0.042327 | -0.042327 | 0.3642 | ACCEPT | 0.0017916 | -0.042327
36 | 578 | 50 | 8.2811 |  |  | 8.3102 | [5 empty] | -0.029134 | -0.029134 | 0.26749 | ACCEPT | 0.00084879 | -0.029134
36 | 582 | 10 | 7.4771 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination? | 9.7679 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination? |  |  |  | -2.2908
38 | 612 | 250 | 5.7585 |  |  | 5.7314 |  |  | 5.773 |  |  | 0.02712 | 0.02712 | 0.14487 | ACCEPT | 0.00073551 | 0.02712
38 | 619 | 75 | 7.112 |  |  | 7.1 |  |  | 7.0961 |  |  | 0.012087 | 0.012087 | 0.034674 | ACCEPT | 0.00014609 | 0.012087
39 | 625 | 763 | 4.6537 |  |  | 4.637 |  |  | 4.7111 |  |  | 0.016651 | 0.016651 | 0.068131 | ACCEPT | 0.00027725 | 0.016651
39 | 634 | 150 | 6.7897 |  |  | 6.7431 | [5 empty] | 0.046605 | 0.046605 | 0.2877 | ACCEPT | 0.002172 | 0.046605
39 | 642 | 50 | 9.1359 |  |  | 9.0787 | [5 empty] | 0.057251 | 0.057251 | 0.36574 | ACCEPT | 0.0032776 | 0.057251
40 | 652 | 150 | 7.2933 |  |  | 7.251 |  |  | 7.3109 |  |  | 0.042286 | 0.042286 | 0.25604 | ACCEPT | 0.0017881 | 0.042286
40 | 655 | 50 | 7.7583 |  |  | 7.7504 |  |  | 7.7643 |  |  | 0.0079593 | 0.0079593 | 0.004418 | ACCEPT | 6.3351e-05 | 0.0079593
41 | 674 | 300 | 6.0745 |  |  | 6.1213 |  |  | 6.1021 |  |  | -0.046771 | -0.046771 | 0.39677 | ACCEPT | 0.0021875 | -0.046771
41 | 678 | 175 | 6.554 |  |  | 6.5854 | [5 empty] | -0.031422 | -0.031422 | 0.28426 | ACCEPT | 0.00098736 | -0.031422
41 | 681 | 100 | 7.1257 |  |  | 7.2284 | [5 empty] | -0.10271 | -0.10271 | 0.80684 | ACCEPT | 0.01055 | -0.10271
42 | 694 | 150 | 6.8125 |  |  | 6.8008 |  |  | 6.8366 |  |  | 0.011635 | 0.011635 | 0.031359 | ACCEPT | 0.00013536 | 0.011635
42 | 696 | 100 | 6.4839 |  |  | 6.4655 | [5 empty] | 0.01834 | 0.01834 | 0.080511 | ACCEPT | 0.00033635 | 0.01834
45 | 729 | 350 | 4.9805 |  |  | 4.9653 | [5 empty] | 0.015195 | 0.015195 | 0.057461 | ACCEPT | 0.0002309 | 0.015195
50 | 807 | 350 | 5.8846 |  |  | 5.761 |  |  | 5.7477 |  |  | 0.12364 | 0.12364 | 0.85243 | ACCEPT | 0.015288 | 0.12364
50 | 812 | 175 | 6.0133 |  |  | 6.0288 | [5 empty] | -0.015482 | -0.015482 | 0.16742 | ACCEPT | 0.00023971 | -0.015482
50 | 818 | 35 | 7.4377 |  |  | 7.4696 | [5 empty] | -0.03192 | -0.03192 | 0.28791 | ACCEPT | 0.0010189 | -0.03192
52 | 824 | 175 | 4.859 |  |  | 4.8722 |  |  | 5.0908 |  |  | -0.013195 | -0.013195 | 0.15065 | ACCEPT | 0.00017411 | -0.013195
52 | 830 | 75 | 6.3743 |  |  | 6.378 | [5 empty] | -0.0037628 | -0.0037628 | 0.081509 | ACCEPT | 1.4158e-05 | -0.0037628
52 | 838 | 5 | 9.5521 |  |  | 9.4411 | [5 empty] | 0.11097 | 0.11097 | 0.75956 | ACCEPT | 0.012315 | 0.11097
54 | 862 | 100 | 5.7769 |  |  | 5.8176 |  |  | 5.8409 |  |  | -0.040778 | -0.040778 | 0.35284 | ACCEPT | 0.0016628 | -0.040778
55 | 873 | 50 | 8.315 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration. | 8.3022 |  |  | 8.5094 |  |  | 0.012754 | 0.012754 | 0.039564 | ACCEPT | 0.00016266 | 0.012754
55 | 876 | 25 | 9.2635 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration. | 9.2389 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration. |  |  |  | 0.024592 | 0.024592 | 0.12634 | ACCEPT | 0.00060475 | 0.024592
55 | 878 | 15 | 9.4377 |  |  | 9.476 | [5 empty] | -0.038304 | -0.038304 | 0.33471 | ACCEPT | 0.0014672 | -0.038304
58 | 905 | 75 | 8.4416 | c | Sample #905 had small bubble on probe at end of titration. | 8.5336 |  |  | 8.4517 |  |  | -0.091969 | -0.091969 | 0.72809 | ACCEPT | 0.0084584 | -0.091969
58 | 906 | 50 | 8.6576 |  |  | 8.66 | [5 empty] | -0.0024023 | -0.0024023 | 0.071536 | ACCEPT | 5.7708e-06 | -0.0024023
59 | 912 | 100 | 8.2391 |  |  | 8.237 | [5 empty] | 0.0020846 | 0.0020846 | 0.038646 | ACCEPT | 4.3454e-06 | 0.0020846
59 | 915 | 35 | 8.8836 |  |  | 8.8917 |  |  | 8.8977 |  |  | -0.0081324 | -0.0081324 | 0.11354 | ACCEPT | 6.6136e-05 | -0.0081324
62 | 943 | 175 | 7.0786 | c | Colorimeter drifting exceeding 100% for entire file. | 7.0776 | c | Colorimeter drifting exceeding 100% for entire file. | 7.0938 | c | Colorimeter drifting exceeding 100% for entire file. | 0.0010557 | 0.0010557 | 0.046188 | ACCEPT | 1.1146e-06 | 0.0010557
62 | 949 | 35 | 7.6924 | c | Colorimeter drifting exceeding 100% for entire file. | 7.6879 | c | Colorimeter drifting exceeding 100% for entire file. |  |  |  | 0.0044816 | 0.0044816 | 0.021075 | ACCEPT | 2.0085e-05 | 0.0044816
Standard Pool:
[14 empty] | count w/o outliers |  | 69
[14 empty] | # outliers total |  | 22
[14 empty] | SS |  | 0.41044
[14 empty] | sp |  | 0.054536
[14 empty] | STD for compare |  |  | 0.077687
Non-Suitable Duplicate Samples
Casts #3,5,8,10 the endpoints were impossible to determine due to broken colorimiter probe.
For the remaining samples in this list were >0.06 ml/L DO apart and are suspect for contamination of at least one sample.
|  |  | Cast # | Sample # | DO (ml/L) | DO (ml/L) | Difference | Comments from notes
|  |  | 3 | 16 | 7.911 | 7.923 | -0.012 | bad colorimiter probe
|  |  |  | 18 | 8.516 | 8.374 | 0.142 | bad colorimiter probe
|  |  |  | 20 | 8.732 | 8.363 | 0.369 | bad colorimiter probe
|  |  |  | 22 | 7.348 | 7.784 | -0.436 | bad colorimiter probe
|  |  | 5 | 48 | 7.628 | 7.835 | -0.207 | bad colorimiter probe
|  |  |  | 49 | 7.687 | 7.943 | -0.256 | bad colorimiter probe
|  |  |  | 50 | 7.56 | 8.032 | -0.472 | bad colorimiter probe
|  |  | 8 | 101 | 7.037 | 7.013 | 0.024 | bad colorimiter probe
|  |  |  | 103 | 6.347 | 6.653 | -0.306 | bad colorimiter probe
|  |  |  | 115 | 7.114 | 7.107 | 0.007 | bad colorimiter probe
|  |  | 10 | 126 | 6.981 | 6.874 | 0.107 | bad colorimiter probe
|  |  |  | 133 | 7.041 | 6.878 | 0.163 | bad colorimiter probe
|  |  |  | 144 | 7.541 | 7.45 | 0.091 | bad colorimiter probe
|  |  | 17 | 279 | 8.946 | 5.898 | 3.048 | likely contamination
|  |  | 18 | 280 | 4.472 | 2.572 | 1.9 | bubble problems with entire cast
|  |  |  | 287 | 3.634 | 4.235 | -0.601 | sun
|  |  |  | 295 | 4.374 | 4.953 | -0.579
|  |  | 22 | 344 | 4.57 | 3.732 | 0.838 | bubble problems with entire cast
|  |  | 25 | 400 | 5.219 | 7.13 | -1.911 | triplicate shows problem with replicate
|  |  | 35 | 570 | 5.604 | 5.726 | -0.122 | duplicate sample precipitate was darker than sample. Triplicate show problem with replicate.
|  |  | 36 | 582 | 7.493 | 9.781 | -2.288 | likely contamination

## Chemicals
Arctic Cruise 2007-19 --Leg
Oxygen Chemicals for Leg 1
Analyst: Melissa Hennekes
Bottle # | Chemical | Volume | Bottle Dated | Other
1 | Potassium Iodate | 1L | 2007-05-16 | N = 0.01013306
2 | Manganious Chloride | 1L | 2007-03-09
5 | Manganious Chloride | 1L | 2005-08-24
3 | Alkaline Iodide | 1L | 2007-05-16
7 | Alkaline Iodide | 1L | 2005-08-24
4 | Sulfuric Acid 35% | 1L | 2007-03-12
10 | Sulfuric Acid 35% | 1L | 2007-03-12
19 | Sodium Thiosulfate | 1L | 2007-05-16
8 | Sodium Thiosulfate | 1L | 2007-03-08
9 | Total number of bottles
Arctic Cruise 2007-20 --Leg 2
Oxygen Chemicals for Leg 2
Analyst: Mary Steel
Bottle # | Chemical | Volume | Bottle Dated | Other
12 | Potassium Iodate | 1L | 2007-05-16 | N = 0.01013424
20 | Potassium Iodate | 1L | 2007-05-16 | N = 0.01013126
6 | Manganious Chloride | 1L | 2007-05-16
11 | Manganious Chloride | 1L | 2007-05-16
17 | Alkaline Iodide | 1L | 2005-04-06
18 | Alkaline Iodide | 1L | 2005-08-12
15 | Sulfuric Acid 35% | 1L | 2007-03-12
16 | Sulfuric Acid 35% | 1L | 2007-05-16
9 | Sodium Thiosulfate | 1L | 2007-05-16
13 | Sodium Thiosulfate | 1L | 2007-05-16
14 | Sodium Thiosulfate | 1L | 2007-03-08
11 | Total number of bottles

## timeline
Timeline
|  |  |  | DO values for Triplicates |  |  | TemperaturesOC at task time | Sodium Thiosulfate (ml)
Date | Cast Number | tripl. source cast | Station Name | Analysed first | analysed last | analysed 24 later | Lab Temperature control system | Initial Standard Volume | Standard Average Volume | #1 | #2 | #3 | KIO3 Standard N = Value | Date Standard was made at IOS | notes and log sheets
2007-07-03 | [6 empty] | 20 | 0.714 |  |  |  |  | 0.010133 | 2007-05-16
2007-07-06 | 3 |  | 1 |  |  |  | 21
| 3 | 3 | triplicate #16 | 7.911 | 7.923 | 7.885 | 21
| 3 | 3 | triplicate #18 | 8.516 | 8.374 | 8.303 | 21
| 3 | 3 | triplicate #20 | 8.732 | 8.363 | 8.517 | 21
|  | 3 | triplicate #22 | 7.348 | 7.784 | 7.059 | 21
2007-07-07 | [6 empty] | 21 |  | 0.72 | 0.721 | 0.721 | 0.719
| 5 |  | 2 |  |  |  | 21
2007-07-09 |  | 5 | triplicate #48 | 7.628 | 7.835 | 7.388 | 19.5
| 7 |  | ls4 |  |  |  | 19.5
| 8 |  | 4 |  |  |  | 20
| 10 | 8 | triplicate #101 | 7.037 | 7.013 | 7.044 | 19.5
| 10 |  | ls5 |  |  |  | 19.5
2007-07-10 | 10 |  | ls5 |  |  |  | 19.5
|  | 10 | triplicate #126 | 6.981 | 6.874 | 6.897 | 20 | [7 empty] | colorimiter probe replaced with new one [Standard check run afterwards. 2 Nov 2007 SZ]
[7 empty] | 20 |  | 0.724 | 0.725 | 0.722 | 0.726
| 11 |  | ls6 |  |  |  | 20.5 | [7 empty] | first station of good data!!!!!!
2007-07-11 | 13 | 11 | triplicate #149 | 6.638 | 6.598 | 7.95 | 20
| 13 |  | ls7 |  |  |  | 20
| 15 | 13 | triplicate #192 | 5.23 | 5.572 | 6.32 | 20.5 | [7 empty] | water installed in lab, maybe residual flux fumes [plumbing being installed. Doors opened for a few hours after plumbing work to vent the lab. Statoins run a couple hours after this. 2Nov 2007 SZ]
| 15 |  | bb1 |  |  |  | 20.5 | [7 empty] | water installed in lab, maybe residual flux fumes
2007-07-12 | 16 | 15 | triplicate #234 | 7.075 | 7.13 | 7.161 | 20.5
| 16 |  | bb2 |  |  |  | 20.5
| 17 |  | bb3 |  |  |  | 21
| 17 | 16 | triplicate#250 | 6.368 | 6.327 | 6.358 | 21
[7 empty] | 21 |  | 0.728 | 0.73 | 0.725 | 0.729
| 18 | 17 | triplicate #263 | 4.403 | 4.364 | 4.221 | 21
| 18 |  | bb4 |  |  |  | 21 | [7 empty] | sunny in lab, some bubble problems
2007-07-13 | 21 | 18 | triplicate #280 | 4.472 | 2.572 | 4.47 | 20
| 21 |  | bb5 |  |  |  | 20
| 22 | 21 | triplicate #322 | 4.55 | 4.513 | 4.675 | 19.5
| 22 |  | bb6 |  |  |  | 19.5 | [7 empty] | sunny in lab, some bubble problems
| 23 | 22 | triplicate #344 | 4.57 | 3.732 | 4.52 | 20.5
| 23 |  | bb7 |  |  |  | 20.5
| 25 | 23 | triplicate #361 | 4.562 | 4.61 | 4.557 | 20.5
| 25 |  | bb8 |  |  |  | 20.5
2007-07-14 | 26 | 25 | triplicate #400 | 5.219 | 7.13 | 7.166 | 21
| 26 |  | bb9 |  |  |  | 21
2007-07-15 |  | 26 | triplicate #408 | 4.496 | 4.487 | 4.645 | 22
| introduction of new chemicals: | [11 empty] | This could be Mary's first bottle used on Leg 2. However n= 0.01013424 is used in processing files which is different from Mary's normality. SZ | 2007-05-16 | new potassium iodate, pickling chemicals, sulfuric acid. SZ 2Nov 2007
| 27 |  | bb10 |  |  |  | 22
[7 empty] | 22 | 0.715 | 0.714 | 0.716 | 0.711 | 0.717
| 29 | 27 | triplicate #431 | 4.492 | 4.494 | 4.498 | 21
| 29 |  | bo1 |  |  |  | 21
| 30 |  | bo1 |  |  |  | 21.5
2007-07-16 | 31 | 30 | triplicate #466 | 2.781 | 2.824 | 2.825 | 21
| 31 |  | bew03 |  |  |  | 21
| 32 | 31 | triplicate #489 | 2.79 | 2.712 | 2.743 | 22
| 32 |  | bew05 |  |  |  | 22
| 33 | 32 | triplicate #512 | 3.556 | 3.526 | 3.568 | 22
| 33 |  | bew07 |  |  |  | 22
| 33 | 33 | triplicate #532 | 3.534 | 3.451 | 3.455 | 22
| 33 | 33 | triplicate #535 | 5.73 | 5.759 | 5.748 | 22
2007-07-17 | 34 |  | bew09 |  |  |  | 22 | [7 empty] | construction in lab all day
| 35 | 34 | triplicate #558 | 7.326 | 7.242 | 7.273 | 22 | [7 empty] | drafty hole open, heater running constantly
| 35 | 34 | triplicate #562 | 7.674 | 7.641 | 7.644 | 22
| 35 |  | bew09 |  |  |  | 22
| 36 | 35 | triplicate #570 | 5.604 | 5.726 | 5.611 | 22
| 36 |  | bew11 |  |  |  | 22
| 38 |  | bb11 |  |  |  | 21
[7 empty] | 21 |  |  | 0.724 | 0.714 | 0.733
| 39 | 38 | triplicate #612 | 5.771 | 5.745 | 5.785 | 22
| 39 | 38 | triplicate #619 | 7.127 | 7.114 | 7.111 | 22
| 39 |  | caa1 |  |  |  | 22 | [7 empty] | oxy was not sampled first this cast
2007-07-18 | 40 | 39 | triplicate #625 | 4.632 | 4.647 | 4.721 | 22
| 40 |  | caa2 |  |  |  | 22
| 41 | 40 | triplicate #652 | 7.251 | 7.259 | 7.281 | 21.5
| 41 | 40 | triplicate #655 | 7.773 | 7.765 | 7.782 | 21.5
| 41 |  | caa3 |  |  |  | 21.5
| 42 | 41 | triplicate #674 | 6.087 | 6.131 | 6.116 | 22
| 42 |  | caa4 |  |  |  | 22
| 42 | 42 | triplicate #694 | 6.828 | 6.816 | 6.799 | 22
2007-07-21 | 45 |  | caa5 |  |  |  | 20
[7 empty] | 20 |  | 0.715 | 0.718 | 0.713 | 0.713
2007-07-23 | 50 |  | caa9 |  |  |  | 22
| 52 | 50 | triplicate #807 | 5.86 | 5.774 | 5.758 | 22
| 52 |  | caa10 |  |  |  | 22
2007-07-24 | 54 | 52 | triplicate #824 | 4.87 | 4.882 | 5.098 | 21.5
| 54 |  | caa11 |  |  |  | 21.5 | [7 empty] | bottles tripped on the fly, samples dirty [bits of stuff in the sample, believe this was not from the ship, but rather algae and zooplankton from the water SZ 2Nov2007]
| 55 | 54 | triplicate #862 | 5.789 | 5.826 | 5.851 | 21.5
| 55 |  | caa12 |  |  |  | 21.5
2007-07-25 | 58 | 55 | triplicate #873 | 8.334 | 8.322 | 8.527 | 22
| 58 |  | caa13 |  |  |  | 22
| 59 | 58 | triplicate #905 | 8.458 | 8.492 | 8.469 | 21.5
| 59 |  | caa14 |  |  |  | 21.5
| 62 | 59 | triplicate #915 | 8.902 | 8.904 | 8.916 | 22
| 62 |  | caa16 |  |  |  | 22
| 62 | 62 | triplicate #943 | 7.095 | 7.091 | 7.109 | 22

## Standards
Standard Average on July 4 =0.714
Potassium Iodate Standard
| #1 | #2 | #3 | Average | Standard Deviation
Endpoint | 0.713 | 0.714 | 0.714 | 0.71367 | 0.00057735
standard checks
Date | Sodium Thiosulfate to Endpoint
| #1 | #2 | #3 | Average
2005-07-07 | 0.721 | 0.721 | 0.719 | 0.72
2005-07-10 | 0.725 | 0.722 | 0.726 | 0.724
2005-07-12 | 0.73 | 0.725 | 0.729 | 0.728
Standard Averageon July 15 =0.715 |  |  |  |  | [updated from 0.725 to 0.715 2Nov2007 SZ]
Potassium Iodate Standard
| #1 | #2 | #3 | Average | Standard Deviation
Endpoint | 0.716 | 0.716 | 0.714 | 0.71533 | 0.0011547
Date | Sodium Thiosulfate to Endpoint
| #1 | #2 | #3 | Average
2005-07-15 | 0.716 | 0.711 | 0.717 | 0.714
2005-07-17 | 0.724 | 0.714 | 0.733 | 0.724
2005-07-21 | 0.718 | 0.713 | 0.713 | 0.715

## FLASKS 2009
Current list of all flasks as of June 2009
Flasks #76, and #3161 had stoppers changed prior to Dec 2008 calibration, old volumes are valid prior to this date
Flask #1194 had stopper changed May 2009, old volume is valid prior to this date
Volume at 20C is based on Jan 2009 calibration for all flasks other than #800-883 performed May 2009
| June 2008 | Jan 2009 cal. BM&TL | VolJune -VolJan
FLASK ID# | Vol. @20C | Vol @ 20C |  | OWNER | comments
76 | 131.1978 | 132.3552 | 1.1573 | WP | new stopper
94 | 132.6816 | 133.1278 | 0.44611 | WP
196 | 136.6071 | 136.5979 | 0.0091753 | WP
329 | 133.5438 | 133.5225 | 0.021325 | A
371 | 130.4592 | 130.4126 | 0.046589 | WP
396 | 132.8267 | 132.7591 | 0.067648 | WP
425 | 132.8343 | 132.7958 | 0.03851 | A
439 | 134.0152 | 133.9785 | 0.036722 | A
456 | 130.8446 | 130.8177 | 0.026821 | A
458 | 132.6743 | 132.6468 | 0.027486 | A
462 | 134.7492 | 134.7236 | 0.025572 | A
468 | 134.8782 | 134.8388 | 0.039415 | WP
471 | 134.5299 | 134.5131 | 0.01682 | WP
505 | 130.7328 | 130.7264 | 0.0063647 | WP
510 | 134.0919 | 133.9136 | 0.17827 | WP
528 | 129.9467 | 129.947 | 0.00035926 | WP
547 | 132.4679 | 132.4457 | 0.022189 | A
552 | 131.8817 | 131.8262 | 0.05552 | WP
557 | 135.475 | 135.451 | 0.023944 | WP
559 | 135.7952 | 135.7462 | 0.048923 | WP
569 | 136.712 | 136.6912 | 0.020774 | WP
570 | 140.7332 | 140.6929 | 0.040273 | WP
575 | 133.1009 | 133.0695 | 0.031356 | A
586 | 133.1872 | 133.1518 | 0.035432 | A
589 | 128.8684 | 128.8344 | 0.03403 | WP
605 | 136.6466 | 136.6274 | 0.019141 | WP
611 | 131.7224 | 131.7049 | 0.017462 | WP
612 | 135.9699 | 135.9279 | 0.042019 | WP
620 | 133.7124 | 133.6889 | 0.023572 | WP
621 | 130.3053 | 130.3054 | 0.00015281 | WP
622 | 132.1297 | 132.1041 | 0.025577 | WP
631 | 136.4164 | 136.3771 | 0.039281 | WP
632 | 136.1879 | 136.1678 | 0.020133 | WP
638 | 132.6652 | 132.6424 | 0.02274 | WP
639 | 133.6203 | 133.5912 | 0.029029 | WP
647 | 129.1717 | 129.1242 | 0.047438 | WP
650 | 132.2339 | 132.1769 | 0.056994 | WP
657 | 135.4919 | 135.4826 | 0.0093198 | WP
658 | 131.7797 | 131.7153 | 0.064422 | A
668 | 130.1453 | 130.1276 | 0.017673 | WP
690 | 132.1392 | 132.1103 | 0.028965 | WP
700 | 127.5135 | 127.477 | 0.036526 | WP
705 | 134.9108 | 134.9098 | 0.00097038 | A
708 | missing | 139.9085 | #VALUE! | WP
712 | 130.6425 | 130.5769 | 0.065591 | WP
715 | 138.8641 | 138.8643 | 0.00024427 | WP
745 | 130.1754 | 130.0861 | 0.089221 | WP
800 |  | 129.7739 |  | A
801 |  | 132.0862 |  | A
802 |  | 132.1503 |  | A
803 |  | 130.5553 |  | A
804 |  | 132.4884 |  | A
805 |  | 131.5269 |  | A
806 |  | 131.1076 |  | A
807 |  | 128.3544 |  | A
808 |  | 129.2648 |  | A
809 |  | 131.2229 |  | A
810 |  | 131.9788 |  | A
811 |  | 131.2636 |  | A
812 |  | 130.2549 |  | A
813 |  | 127.2143 |  | A
814 |  | 129.9469 |  | A
815 |  | 133.0025 |  | A
816 |  | 132.0395 |  | A
817 |  | 132.6007 |  | A
818 |  | 129.666 |  | A
819 |  | 129.7704 |  | A
820 |  | 131.9075 |  | A
821 |  | 130.2262 |  | A
822 |  | 137.1124 |  | A
823 |  | 132.7893 |  | A
824 |  | 130.9125 |  | A
825 |  | 131.4536 |  | A
826 |  | 131.2345 |  | A
827 |  | 129.4608 |  | A
828 |  | 128.6674 |  | A
829 |  | 134.781 |  | A
830 |  | 131.0118 |  | A
831 |  | 132.5032 |  | A
832 |  | 131.7731 |  | A
833 |  | 130.1222 |  | A
834 |  | 130.2897 |  | A
835 |  | 130.8303 |  | A
836 |  | 130.0982 |  | A
837 |  | 129.9518 |  | A
838 |  | 130.4201 |  | A
839 |  | 133.6396 |  | A
840 |  | 130.1955 |  | A
841 |  | 131.5605 |  | A
842 |  | 131.057 |  | A
843 |  | 131.2576 |  | A
844 |  | 130.9697 |  | A
845 |  | 130.73 |  | A
846 |  | 129.8775 |  | A
847 |  | 130.7611 |  | A
848 |  | 131.7779 |  | A
849 |  | 130.7248 |  | A
850 |  | 132.524 |  | A
851 |  | 131.5111 |  | A
852 |  | 133.5621 |  | A
853 |  | 131.2523 |  | A
854 |  | 130.9434 |  | A
855 |  | 130.6325 |  | A
856 |  | 131.0588 |  | A
857 |  | 130.0408 |  | A
858 |  | 131.9404 |  | A
859 |  | 127.1393 |  | A
860 |  | 129.8013 |  | A
861 |  | 131.83 |  | A
862 |  | 129.717 |  | A
863 |  | 133.2759 |  | A
864 |  | 129.5716 |  | A
865 |  | 130.5686 |  | A
866 |  | 129.7536 |  | A
867 |  | 127.784 |  | A
868 |  | 125.9274 |  | A
869 |  | 142.8308 |  | A
870 |  | 130.7451 |  | A
871 |  | 130.0656 |  | A
872 |  | 130.2943 |  | A
873 |  | 127.3226 |  | A
874 |  | 132.5492 |  | A
875 |  | 131.0295 |  | A
876 |  | 128.0849 |  | A
877 |  | 129.2846 |  | A
878 |  | 127.7695 |  | A
879 |  | 137.0656 |  | A
880 |  | 130.1267 |  | A
881 |  | 129.9307 |  | A
882 |  | 130.871 |  | A
883 |  | 127.9922 |  | A
1023 | missing | 134.8307 | #VALUE! | A
1085 | missing | 132.1686 | #VALUE! | WP
1105 | 127.6261 | 127.6147 | 0.011355 | WP
1106 | 132.7336 | 132.6826 | 0.050966 | WP
1110 | 127.2882 | 127.2797 | 0.008534 | A
1119 | 132.8533 | 132.7789 | 0.07445 | A
1123 | 135.9616 | 135.9303 | 0.031279 | WP
1126 | 127.9897 | 127.9802 | 0.0095162 | WP
1128 | 127.2205 | 131.4381 | 4.2176 | WP
1130 | missing | 129.8804 | #VALUE! | A
1134 | 129.7388 | 137.2491 | 7.5103 | WP
1136 | missing | 129.6667 | #VALUE! | A
1141 | 128.5321 | 133.4436 | 4.9116 | A
1144 | 136.4967 | 136.4511 | 0.045547 | WP
1146 | 127.4429 | 127.4431 | 0.00014296 | WP
1150 | 134.5306 | 134.4947 | 0.035855 | A
1158 | 130.508 | 130.5202 | 0.012209 | WP
1174 | 133.3657 | 133.3278 | 0.037895 | WP
1175 | 143.1643 | 143.127 | 0.037288 | WP
1176 | 132.0526 | 131.975 | 0.077589 | WP
1180 | 134.4797 | 134.4337 | 0.046001 | WP
1188 | 143.6799 | 138.3565 | 5.3234 | A
1192 | 143.6164 | 133.6563 | 9.96 | A
1194 | 132.2582 | 131.74 | 0.51818 | A | new stopper May 2009
1201 | 136.4977 | 136.4448 | 0.052918 | WP
1202 | 146.3347 | 146.335 | 0.00025382 | A | added May 2009 found in calibration file
1204 | 133.6863 | 133.6583 | 0.028002 | A
1206 | 143.407 | 131.4627 | 11.9443 | WP
1209 | 145.2552 | 137.022 | 8.2332 | WP
1212 | 144.2071 | 144.1959 | 0.01124 | A
1213 | 136.1749 | 136.1141 | 0.060759 | WP
1216 | 142.509 | 142.4592 | 0.049783 | A
1218 | 146.5954 | 144.0216 | 2.5739 | WP
1219 | 147.9799 | 135.712 | 12.268 | A
1223 | 143.2435 | 143.2262 | 0.017238 | A
1224 | 144.6157 | 133.3049 | 11.3107 | A
1225 | 143.3014 | 136.8451 | 6.4564 | A
1228 | 143.0819 | 143.0396 | 0.042229 | A
1230 | 129.3001 | 129.2102 | 0.089965 | A
1231 | missing | 131.3241 | #VALUE! | A
1233 | missing | 145.1363 | #VALUE! | A
1238 | 144.8217 | 144.7496 | 0.072072 | A
1241 | 145.315 | 145.2695 | 0.045415 | A
1244 | 133.8685 | 133.8428 | 0.025666 | A
1247 | 144.1583 | 144.1256 | 0.032683 | WP
1248 | missing | 145.4153 | #VALUE! | WP
1252 | 136.9889 | 136.9675 | 0.021339 | A
1261 | 143.9569 | 131.881 | 12.0759 | WP
1267 | 143.77 | 143.7452 | 0.024872 | A
1268 | 142.5188 | 142.4983 | 0.020534 | A
1270 | missing | 129.7059 | #VALUE! | A
1271 | 137.6582 | 137.5493 | 0.10892 | A
1274 | 133.9498 | 133.9375 | 0.012334 | A
1276 | 130.8472 | 130.7703 | 0.076903 | A
1277 | 132.1328 | 134.7547 | 2.6218 | A
1278 | 131.6078 | 131.5842 | 0.023639 | A
1279 | 133.5465 | 133.5225 | 0.023992 | A
1280 | 136.8282 | 136.5566 | 0.27165 | A
1281 | 130.9424 | 130.9195 | 0.022952 | A
1282 | 133.7767 | 133.7368 | 0.039828 | A
1284 | missing | 144.0419 | #VALUE! | A
1286 | 142.4926 | 132.7798 | 9.7129 | A
1287 | 142.7166 | 128.0797 | 14.637 | A | BROKEN Sep 09
1288 | 136.6879 | 136.6566 | 0.031334 | A
1290 | 141.7185 | 147.2951 | 5.5767 | A
1291 | 140.3075 | 130.6108 | 9.6967 | A
1293 | 143.0929 | 143.0764 | 0.016463 | A
1295 | 142.2662 | 142.2474 | 0.018833 | A
1297 | 135.8107 | 135.494 | 0.31674 | A
1298 | 136.5303 | 142.9072 | 6.3769 | WP
1299 | 145.4258 | 143.9055 | 1.5203 | A
1300 | 137.3275 | 137.2077 | 0.11987 | WP
1301 | 128.2017 | 128.1754 | 0.026264 | WP
1302 | 132.7901 | 132.7491 | 0.041041 | WP
1306 | 134.0567 | 134.0211 | 0.035546 | WP
1307 | 133.1885 | 133.0741 | 0.11448 | WP
1309 | 134.6651 | 134.6569 | 0.0082012 | A
1311 | 129.7794 | 130.2229 | 0.44345 | WP
1312 | 132.0995 | 132.0667 | 0.032858 | WP
1313 | 136.454 | 136.3839 | 0.070139 | WP
1314 | 136.4258 | 136.3462 | 0.079594 | WP
1315 | 139.1722 | 139.0972 | 0.075017 | WP
1316 | 135.2038 | 133.0965 | 2.1073 | WP
1317 | 138.927 | 138.8662 | 0.060823 | WP
1318 | 133.6748 | 133.6157 | 0.059081 | WP
1319 | 133.309 | 133.2536 | 0.055412 | WP
1320 | 135.7885 | 135.7418 | 0.046643 | WP
1321 | 136.4274 | 136.3354 | 0.091991 | WP
1322 | 136.0494 | 136.0204 | 0.02898 | WP
1326 | 136.1468 | 136.175 | 0.028256 | WP
1327 | 134.4297 | 134.407 | 0.022688 | WP
1328 | 133.7864 | 133.7575 | 0.028861 | A
1330 | 134.4622 | 134.3999 | 0.062223 | WP
1331 | 133.9399 | 133.8906 | 0.04933 | A
1333 | 133.5609 | 133.5184 | 0.042531 | WP
1334 | missing | 132.6171 | #VALUE! | A | broken stopper Aug 09
1335 | 135.1042 | 135.1441 | 0.039915 | WP
1336 | missing | 133.1814 | #VALUE! | A
1338 | 136.2722 | 136.2589 | 0.013325 | WP
1340 | 135.8669 | 135.8127 | 0.054257 | A
1341 | 134.8089 | 134.8007 | 0.0082225 | A
1342 | 143.1831 | 143.1673 | 0.015837 | WP
1343 | 136.2947 | 147.4459 | 11.1513 | A
1344 | 134.3031 | 134.2848 | 0.018311 | A
1347 | 133.8205 | 133.7881 | 0.032443 | A
1351 | 134.5706 | 134.5543 | 0.016248 | A
1352 | missing | 132.8332 | #VALUE! | A
1354 | 134.3519 | 134.3363 | 0.015594 | A
1356 | 135.0479 | 135.0287 | 0.019255 | A
1358 | 128.6185 | 128.5982 | 0.020304 | A
1360 | 132.7285 | 132.7129 | 0.015637 | A
1361 | 136.1891 | 136.168 | 0.021049 | A
1362 | 140.1317 | 140.1201 | 0.011541 | A
1363 | 136.4104 | 136.3907 | 0.019698 | A
1364 | 132.1424 | 132.1556 | 0.013187 | A
1365 | 135.5105 | 135.4688 | 0.041689 | WP
1366 | 133.0947 | 133.0823 | 0.012381 | A
1367 | 136.1242 | 136.1031 | 0.021088 | WP
1368 | 135.4613 | 135.4334 | 0.027856 | WP
1373 | 135.2239 | 135.19 | 0.033882 | WP
1374 | 134.5371 | 134.513 | 0.024094 | WP
1375 | 137.1721 | 137.1432 | 0.028919 | A
1377 | 132.3597 | 132.3393 | 0.020436 | WP
1378 | 137.2454 | 132.5367 | 4.7087 | WP
1380 | 136.0885 | 136.0861 | 0.0024567 | WP
1381 | 136.4565 | 136.4374 | 0.019155 | A
1382 | 134.8658 | 133.9865 | 0.87926 | A
1384 | 131.9049 | 131.8608 | 0.044121 | WP
1386 | 134.4341 | 134.4218 | 0.012235 | WP
1387 | 136.188 | 136.1195 | 0.068496 | WP
1389 | 136.8342 | 136.8072 | 0.027075 | WP
1392 | 132.1929 | 132.1029 | 0.090013 | WP
1393 | 141.7089 | 141.6395 | 0.069443 | WP
1395 | 135.5867 | 135.5031 | 0.083654 | WP
1397 | 134.8778 | 134.8437 | 0.034127 | WP
1399 | 129.0287 | 128.9923 | 0.036369 | WP
1400 | 140.3856 | 140.3029 | 0.082678 | WP
1401 | 130.2671 | 130.2085 | 0.058676 | WP
1402 | 134.5065 | 134.4524 | 0.054042 | WP
1404 | 146.0066 | 145.9921 | 0.014493 | WP
1406 | 137.9307 | 137.9115 | 0.019178 | WP
1408 | 130.6699 | 130.6431 | 0.026779 | WP
1409 | 134.0012 | 133.9857 | 0.015461 | WP
1411 | 139.7352 | 139.6875 | 0.047765 | WP
1413 | 132.7787 | 132.6991 | 0.079588 | A
1414 | 132.6089 | 132.6186 | 0.0096793 | WP
1416 | 134.3188 | 134.2688 | 0.050011 | A
1417 | missing | 143.3551 | #VALUE! | A
1418 | 132.6121 | 132.5892 | 0.022973 | A
1421 | 139.4037 | 139.3636 | 0.040114 | WP
1422 | 134.2133 | 134.1581 | 0.05527 | A
1424 | 144.2029 | 144.1892 | 0.013732 | A
1426 | 132.3917 | 132.3658 | 0.025827 | A
1427 | 137.0189 | 136.975 | 0.043955 | A
1428 | 139.1401 | 139.0788 | 0.061325 | A
1429 | 132.4852 | 132.415 | 0.070128 | A
1430 | 132.0203 | 131.976 | 0.044287 | A
1432 | 132.5467 | 132.5323 | 0.014402 | A
1433 | 132.4378 | 132.4003 | 0.037499 | A
1434 | 131.4319 | 131.51 | 0.078149 | A
1440 | 131.0108 | 130.9774 | 0.033381 | A
1443 | 133.0467 | 133.0201 | 0.026556 | A
1445 | 133.3702 | 133.3771 | 0.0069168 | A
1446 | 135.3298 | 135.2985 | 0.031317 | A
1447 | 131.8778 | 131.8962 | 0.018385 | A
1448 | 133.185 | 133.1702 | 0.01485 | A
1451 | 131.4467 | 131.4161 | 0.030626 | A
1452 | 134.2047 | 134.1895 | 0.015232 | A
1453 | 130.367 | 130.3266 | 0.040407 | A
1454 | 132.1608 | 132.1275 | 0.033295 | A
1456 | 133.6171 | 133.5835 | 0.033638 | A
1457 | 129.7225 | 130.7424 | 1.0199 | A
1459 | 132.9432 | 132.9068 | 0.036324 | WP
1464 | 131.234 | 131.1891 | 0.044946 | WP
1465 | 132.1756 | 132.1265 | 0.049093 | WP
1468 | 130.9789 | 130.9946 | 0.015719 | A
1470 | 133.8021 | 133.6461 | 0.15591 | WP
1471 | 131.7981 | 131.7917 | 0.0063784 | A | broken Oct 09
1473 | 131.7349 | 131.7116 | 0.023296 | WP
1475 | 130.2767 | 130.2526 | 0.024046 | WP
1476 | 132.7598 | 132.7313 | 0.028476 | WP
1477 | 134.3976 | 134.3725 | 0.025063 | WP
1478 | 130.6114 | 130.6077 | 0.0037232 | WP
1481 | 131.2981 | 131.2763 | 0.021777 | WP
1482 | 131.2717 | 131.2497 | 0.021963 | A
1483 | 132.3839 | 132.4207 | 0.036843 | WP
1484 | 133.7492 | 133.7352 | 0.013974 | A
1485 | 133.1015 | 133.0854 | 0.016141 | WP
1486 | 143.245 | 143.2496 | 0.0046619 | A
1487 | 131.1597 | 131.1356 | 0.024127 | WP
1488 | 138.2414 | 138.2002 | 0.041162 | WP
1489 | 134.5347 | 133.5762 | 0.95854 | WP
1490 | 135.9464 | 135.869 | 0.077452 | WP
1491 | 140.7203 | 140.6427 | 0.077571 | WP
1492 | 140.1048 | 139.9858 | 0.11891 | WP
1493 | 139.5795 | 139.4414 | 0.13813 | A
1494 | 137.9871 | 137.8962 | 0.09097 | A
1495 | 139.0844 | 139.0385 | 0.045923 | A
1496 | 139.7368 | 139.6813 | 0.055462 | WP
1497 | 139.613 | 139.5977 | 0.015288 | WP
1498 | 137.6827 | 137.6666 | 0.016069 | WP
1499 | 136.9768 | 136.8962 | 0.080624 | A
2000 | 143.7448 | 131.6909 | 12.0539 | WP
2007 | 144.39 | 144.1373 | 0.25266 | WP
2014 | 143.8862 | 143.9159 | 0.029712 | WP
2015 | 142.4986 | 142.6118 | 0.11314 | WP
2017 | 144.6165 | 144.8189 | 0.20239 | WP
2019 | 134.4859 | 134.4022 | 0.083681 | WP
2020 | 145.9524 | 145.9497 | 0.0027367 | WP
2025 | 145.867 | 145.8939 | 0.026907 | WP
2027 | 142.11 | 141.8814 | 0.22864 | WP
2032 | 143.44 | 143.193 | 0.24696 | WP
2033 | 143.9244 | 143.5593 | 0.36515 | WP
2035 | 143.9923 | 134.2406 | 9.7517 | WP
2040 | 137.838 | 137.8055 | 0.03257 | WP
2044 | 140.7528 | 140.9981 | 0.24529 | WP
2047 | 143.6216 | 142.7708 | 0.85085 | WP
2049 | 144.0889 | 144.1547 | 0.065784 | WP
2051 | 143.0792 | 143.1094 | 0.030256 | WP
2055 | 142.8059 | 131.3389 | 11.467 | WP
2056 | 142.1405 | 142.0985 | 0.041987 | WP
2059 | 142.4462 | 142.4248 | 0.021352 | WP
2060 | 144.15 | 143.9012 | 0.24877 | WP
2061 | 143.7871 | 143.1712 | 0.61587 | WP
2063 | 142.8474 | 142.864 | 0.016616 | WP
2064 | 144.2318 | 144.2453 | 0.013472 | WP
2065 | 144.8806 | 131.9658 | 12.9147 | WP
2068 | 141.968 | 141.988 | 0.020017 | WP
2069 | 143.1785 | 143.1585 | 0.019936 | WP
2073 | 142.3191 | 141.6233 | 0.69576 | WP
2074 | 141.3861 | 141.3715 | 0.01463 | WP
2075 | 135.3243 | 135.2729 | 0.051446 | WP
2082 | missing | 131.8216 | #VALUE! | WP
2084 | 144.4425 | 144.5316 | 0.089084 | WP
2085 | 143.5931 | 143.5397 | 0.053399 | WP
2086 | 143.1551 | 143.1706 | 0.015506 | WP
2087 | 143.79 | 143.5166 | 0.27339 | WP
2088 | 144.2139 | 144.0937 | 0.12019 | WP
2089 | 145.4221 | 145.2288 | 0.19329 | WP
2090 | 134.7049 | 134.6712 | 0.0337 | WP
2091 | 143.7415 | 143.57 | 0.17149 | WP
2092 | 144.2848 | 144.1993 | 0.085415 | WP
2099 | 148.988 | 148.9069 | 0.081081 | WP
3006 | 143.6625 | 143.7048 | 0.042345 | WP
3007 | 142.36 | 142.1231 | 0.23691 | WP
3009 | 140.96 | 140.664 | 0.29596 | WP
3012 | 143.8798 | 143.9013 | 0.021536 | WP
3016 | 143.92 | 143.662 | 0.25803 | WP
3018 | 143.7226 | 143.7228 | 0.00020328 | WP
3019 | 137.4594 | 137.4714 | 0.012008 | WP
3024 | 145.1539 | 145.1762 | 0.022375 | WP
3028 | 141.7631 | 141.7907 | 0.027578 | WP
3033 | 139.4324 | 139.4588 | 0.026391 | WP
3034 | 138.6098 | 138.6146 | 0.0048678 | WP
3038 | 141.7531 | 141.7644 | 0.011331 | WP
3040 | 132.5271 | 132.5551 | 0.02796 | WP
3043 | 135.8543 | 135.8721 | 0.017767 | WP
3047 | 135.4564 | 135.5022 | 0.045789 | WP
3051 | 140.6161 | 140.6477 | 0.031543 | WP
3052 | 137.5464 | 137.5858 | 0.039377 | WP
3055 | 140.9701 | 140.9856 | 0.015438 | WP
3057 | 138.4989 | 138.4973 | 0.0015995 | WP
3058 | 138.2849 | 138.3074 | 0.022555 | WP
3061 | 143.1404 | 143.1344 | 0.0060007 | WP
3063 | 144.5081 | 144.5139 | 0.0057834 | WP
3067 | 141.5185 | 141.5236 | 0.0050595 | WP
3069 | 138.9437 | 138.9873 | 0.043635 | WP
3075 | 135.8805 | 135.884 | 0.0034846 | WP
3084 | 144.6184 | 144.6756 | 0.057205 | WP
3090 | 134.0191 | 133.9866 | 0.032432 | WP
3094 | 139.9703 | 139.9812 | 0.010919 | WP
3096 | 142.5284 | 142.5586 | 0.030164 | WP
3097 | 143.2039 | 143.1879 | 0.015999 | WP
3098 | 144.5883 | 144.5645 | 0.023816 | WP
3100 | 143.56 | 143.2848 | 0.27519 | WP
3104 | 143.1504 | 143.1468 | 0.0036643 | WP
3107 | 141.5687 | 141.5746 | 0.0059263 | WP
3108 | 143.2775 | 143.2932 | 0.015719 | WP
3109 | 143.51 | 143.2569 | 0.25312 | WP
3113 | 144.0566 | 144.0711 | 0.014485 | WP
3116 | 141.3 | 141.0862 | 0.21381 | WP
3118 | 141.04 | 140.7994 | 0.24059 | WP
3119 | 140.1107 | 140.1242 | 0.01345 | WP
3121 | 144.2138 | 144.2162 | 0.002394 | WP
3124 | 144.6044 | 144.6411 | 0.036679 | WP
3126 | 145.407 | 145.4256 | 0.018623 | WP
3130 | 143.78 | 143.5616 | 0.2184 | WP
3133 | 134.4067 | 134.3446 | 0.062116 | WP
3135 | 140.3472 | 140.392 | 0.044777 | WP
3137 | 141.91 | 141.6896 | 0.22038 | WP
3142 | 144.23 | 143.9688 | 0.26125 | WP
3151 | 143.1162 | 143.1505 | 0.034275 | WP
3157 | 133.07 | 132.8731 | 0.19686 | WP
3160 | 144.8987 | 144.9088 | 0.010066 | WP | volume mix-up with flask has been corrected Feb/09
3161 | 145.0258 | 144.3075 | 0.7183 | WP | new stopper
3164 | 141.9123 | 141.9355 | 0.023249 | WP
3166 | 145.7214 | 145.7412 | 0.019825 | WP
3171 | 143.0594 | 143.0623 | 0.0029829 | WP
3174 | 141.2903 | 141.3182 | 0.027889 | WP
3177 | 143.1296 | 143.1669 | 0.037276 | WP
3178 | 143.0995 | 143.1174 | 0.017939 | WP
3180 | 145.9485 | 145.9067 | 0.041854 | WP
3181 | 145.11 | 144.858 | 0.25198 | WP
3182 | 135.83 | 135.614 | 0.216 | WP | volume mix-up with flask has been corrected Feb/09
3184 | 135.9476 | 135.8762 | 0.071407 | WP
3187 | 131.5778 | 131.4873 | 0.090494 | WP
3188 | 131.82 | 131.6134 | 0.20662 | WP
3189 | 133.6824 | 133.587 | 0.095447 | WP
3190 | 131.8455 | 131.8213 | 0.024246 | WP
3191 | 135.0107 | 134.9518 | 0.058979 | WP
3193 | 131.705 | 131.6281 | 0.076906 | WP
3194 | 134.5925 | 134.541 | 0.051479 | WP
3196 | 132.4277 | 132.3992 | 0.028513 | WP
3198 | 136.5934 | 136.5025 | 0.090842 | WP
3199 | 135.7903 | 134.8688 | 0.92156 | WP
3200 | 134.5925 | 134.5035 | 0.088967 | WP
3201 | 137.2659 | 137.208 | 0.057844 | WP
3202 | 134.9974 | 134.9339 | 0.063422 | WP
3203 | 137.1019 | 137.025 | 0.076889 | WP
3204 | 132.8292 | 131.8575 | 0.97173 | WP
3206 | 133.8497 | 133.8167 | 0.032994 | WP
3207 | 138.76 | 138.4785 | 0.28149 | WP
3209 | 135.42 | 135.1873 | 0.23268 | WP

## Oxygen Recalc
2 March 2017: Oxygen recalculated using a) new KIO3 STD volume of 9.8 instead of 10.0ml b)new flask volume calibrated in 2008 and 2009.
Oxygen differences calculated using "Oxygen - Nov 2007 ("final")"
Oxygen - Nov 2007 | KIO3 volume, Nov 2007 | Oxygen - Nov 2007 ("final") | New Flag, Nov 2007 | Comment, Nov 2007 | Cast | Sample_ Number | Pressure | Flask_ Number1 | Flask_ Volume1 | Flask_ Volume 2009 Calibration | Flask_ Volume Best | Flask_ Volume Difference | Thio_ Titer1 | Thio_ Titer2 | Thio_ Titer3 | Blank Titer | Standard Titer | KIO3 Std Volume | KIO3 Normality | THIO Normality | Oxy (mL/L) with changed KIO3 STD | Oxy (mL/L) with changed flask volume | Oxy (mL/L) BOTH KIO3 STD vol and flask vol changes | Flag | Comment |  | Diff in Oxy: Changed KIO3 | Diff in Oxy: Changed Flask | Diff in Oxy: Change to both KIO3 and Flask | Comment:
7.909 | 10 | 7.911 | d | Colorimeter probe problems with drift for entire day's run | 3 | 16 | 75 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 1.403 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.8947 | 7.9089 | 7.8931 |  |  |  | -0.016301 | -0.002104 | -0.017946 | Oxy in column A seems to be edited? There are "0" values and +/-0.002
7.921 | 10 | 7.923 | d | Colorimeter probe problems with drift for entire day's run | 3 | 16 | 75 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 1.392 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.9071 | 7.9221 | 7.9063 |  |  |  | -0.015927 | -0.00087322 | -0.016742 | Calc oxy (AA) in spreadsheet is same as in column C using orinal terms.
0 | 10 | 7.885 | d | Colorimeter probe problems with drift for entire day's run | 3 | 116 | 75 | 2091 | 143.74 | 143.57 | 143.57 | -0.16996 | 1.405 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.8692 | 7.8944 | 7.8786 |  |  |  | -0.015811 | 0.0094491 | -0.0063639
7.85 | 10 | 7.852 | d | Colorimeter probe problems with drift for entire day's run | 3 | 216 | 75 | 2084 | 144.44 | 144.5316 | 144.5316 | 0.091584 | 1.406 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.8361 | 7.8468 | 7.8311 |  |  |  | -0.015885 | -0.0052027 | -0.02092
8.513 | 10 | 8.516 | d | Colorimeter probe problems with drift for entire day's run | 3 | 18 | 50 | 3057 | 138.5 | 138.4973 | 138.4973 | -0.0027015 | 1.461 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4984 | 8.5156 | 8.4986 |  |  |  | -0.017611 | -0.00038644 | -0.017443
8.372 | 10 | 8.374 | d | Colorimeter probe problems with drift for entire day's run | 3 | 18 | 50 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 1.411 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3576 | 8.4273 | 8.4104 |  |  |  | -0.016446 | 0.053313 | 0.036432
8.301 | 10 | 8.303 | d | Colorimeter probe problems with drift for entire day's run | 3 | 118 | 50 | 2033 | 143.92 | 143.5593 | 143.5593 | -0.36073 | 1.481 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.2862 | 8.324 | 8.3073 |  |  |  | -0.016803 | 0.020984 | 0.0043121
0 | 10 | 8.019 | d | Colorimeter probe problems with drift for entire day's run | 3 | 218 | 50 | 2068 | 141.97 | 141.988 | 141.988 | 0.017984 | 1.411 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.0029 | 8.0179 | 8.0019 |  |  |  | -0.016121 | -0.0010893 | -0.017149
0 | 10 | 8.732 | d | Colorimeter probe problems with drift for entire day's run | 3 | 20 | 25 | 3094 | 139.97 | 139.9812 | 139.9812 | 0.011182 | 1.514 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.7141 | 8.7309 | 8.7134 |  |  |  | -0.017864 | -0.0010836 | -0.01857
0 | 10 | 8.363 | d | Colorimeter probe problems with drift for entire day's run | 3 | 20 | 25 | 2032 | 143.48 | 143.193 | 143.193 | -0.28696 | 1.487 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3458 | 8.3795 | 8.3627 |  |  |  | -0.017219 | 0.016526 | -0.00025725
0 | 10 | 8.517 | d | Colorimeter probe problems with drift for entire day's run | 3 | 120 | 25 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.496 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4998 | 8.515 | 8.498 |  |  |  | -0.017163 | -0.0019741 | -0.019028
8.326 | 10 | 8.328 | d | Colorimeter probe problems with drift for entire day's run | 3 | 220 | 25 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 1.419 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3109 | 8.3254 | 8.3087 |  |  |  | -0.017106 | -0.0026221 | -0.019298
0 | 10 | 7.348 | d | Colorimeter probe problems with drift for entire day's run | 3 | 22 | 5 | 2056 | 142.14 | 142.0985 | 142.0985 | -0.041454 | 1.295 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3332 | 7.3501 | 7.3353 |  |  |  | -0.014826 | 0.0020685 | -0.012656
0 | 10 | 7.784 | d | Colorimeter probe problems with drift for entire day's run | 3 | 22 | 5 | 2027 | 142.91 | 141.8814 | 141.8814 | -1.0286 | 1.379 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.7684 | 7.8413 | 7.8255 |  |  |  | -0.015581 | 0.057252 | 0.041545
7.057 | 10 | 7.059 | d | Colorimeter probe problems with drift for entire day's run | 3 | 122 | 5 | 2044 | 140.75 | 140.9981 | 140.9981 | 0.2481 | 1.232 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0446 | 7.0461 | 7.032 |  |  |  | -0.014399 | -0.012856 | -0.026973
0 | 10 | 7.013 | d | Colorimeter probe problems with drift for entire day's run | 3 | 222 | 5 | 2085 | 143.59 | 143.5397 | 143.5397 | -0.050266 | 1.249 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.999 | 7.0156 | 7.0015 |  |  |  | -0.013966 | 0.0025747 | -0.01148
7.547 | 10 | 7.549 | d | Colorimeter probe problems with drift for entire day's run | 5 | 45 | 214 | 2073 | 142.32 | 141.6233 | 141.6233 | -0.69668 | 1.332 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.534 | 7.5868 | 7.5716 |  |  |  | -0.014987 | 0.037804 | 0.022606
0 | 10 | 7.628 | d | Colorimeter probe problems with drift for entire day's run | 5 | 48 | 200 | 2074 | 141.39 | 141.3715 | 141.3715 | -0.018512 | 1.337 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6129 | 7.6292 | 7.6139 |  |  |  | -0.015119 | 0.0011749 | -0.014108
7.686 | 10 | 7.687 | d | Colorimeter probe problems with drift for entire day's run | 5 | 49 | 150 | 3119 | 140.11 | 140.1242 | 140.1242 | 0.014162 | 1.335 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6719 | 7.6865 | 7.6711 |  |  |  | -0.01511 | -0.0004985 | -0.015896
0 | 10 | 7.56 | d | Colorimeter probe problems with drift for entire day's run | 5 | 50 | 125 | 2099 | 148.99 | 148.9069 | 148.9069 | -0.08305 | 1.397 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5447 | 7.5641 | 7.549 |  |  |  | -0.015268 | 0.0041484 | -0.011003
8.306 | 10 | 8.308 | d | Colorimeter probe problems with drift for entire day's run | 5 | 51 | 100 | 3121 | 144.21 | 144.2162 | 144.2162 | 0.0062065 | 1.485 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.2917 | 8.308 | 8.2914 |  |  |  | -0.016274 | 4.261e-06 | -0.016636
8.077 | 10 | 8.078 | d | Colorimeter probe problems with drift for entire day's run | 5 | 54 | 75 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.329 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.0622 | 8.1389 | 8.1226 |  |  |  | -0.015776 | 0.060907 | 0.044603
7.95 | 10 | 7.952 | d | Colorimeter probe problems with drift for entire day's run | 5 | 55 | 50 | 3104 | 143.15 | 143.1468 | 143.1468 | -0.0032473 | 1.411 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.936 | 7.9521 | 7.9362 |  |  |  | -0.016025 | 8.6268e-05 | -0.015842
7.796 | 10 | 7.798 | d | Colorimeter probe problems with drift for entire day's run | 5 | 56 | 35 | 2007 | 143.68 | 144.1373 | 144.1373 | 0.45734 | 1.389 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.7825 | 7.773 | 7.7574 |  |  |  | -0.015523 | -0.024994 | -0.040563
8.415 | 10 | 8.417 | d | Colorimeter probe problems with drift for entire day's run | 5 | 57 | 25 | 2087 | 143.61 | 143.5166 | 143.5166 | -0.093392 | 1.498 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4001 | 8.4225 | 8.4056 |  |  |  | -0.016945 | 0.005468 | -0.011401
10.062 | 10 | 10.065 | d | Colorimeter probe problems with drift for entire day's run | 5 | 61 | 5 | 3108 | 143.28 | 143.2932 | 143.2932 | 0.013209 | 1.786 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 10.0451 | 10.0643 | 10.0441 |  |  |  | -0.019934 | -0.00072054 | -0.020873
7.834 | 10 | 7.835 | d | Colorimeter probe problems with drift for entire day's run | 5 | 48 | 200 | 2064 | 144.23 | 144.2453 | 144.2453 | 0.015303 | 1.401 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.8197 | 7.8345 | 7.8188 |  |  |  | -0.015345 | -0.00049362 | -0.016187
0 | 10 | 7.943 | d | Colorimeter probe problems with drift for entire day's run | 5 | 49 | 150 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 1.317 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.9273 | 7.9452 | 7.9293 |  |  |  | -0.015714 | 0.0022035 | -0.013713
8.03 | 10 | 8.032 | d | Colorimeter probe problems with drift for entire day's run | 5 | 50 | 125 | 3135 | 140.35 | 140.392 | 140.392 | 0.041968 | 1.397 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.0159 | 8.0296 | 8.0135 |  |  |  | -0.016097 | -0.0024446 | -0.018528
7.387 | 10 | 7.388 | d | Colorimeter probe problems with drift for entire day's run | 5 | 48 | 200 | 3142 | 144.08 | 143.9688 | 143.9688 | -0.11125 | 1.32 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3733 | 7.3939 | 7.3791 |  |  |  | -0.014657 | 0.0059331 | -0.0088787
0 | 10 | 6.723 | d | Colorimeter probe problems with drift for entire day's run | 7 | 78 | 2000 | 2068 | 141.97 | 141.988 | 141.988 | 0.017984 | 1.184 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.7098 | 6.7224 | 6.7089 |  |  |  | -0.013235 | -0.00062763 | -0.014097
0 | 10 | 7.066 | d | Colorimeter probe problems with drift for entire day's run | 7 | 79 | 1500 | 3051 | 140.62 | 140.6477 | 140.6477 | 0.027665 | 1.232 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0512 | 7.064 | 7.0498 |  |  |  | -0.014793 | -0.0020472 | -0.0162
7.035 | 10 | 7.037 | d | Colorimeter probe problems with drift for entire day's run | 8 | 101 | 3459 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 1.2 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0227 | 7.0349 | 7.0208 |  |  |  | -0.014327 | -0.0020841 | -0.016179
5.887 | 10 | 5.888 | d | Colorimeter probe problems with drift for entire day's run | 8 | 102 | 3000 | 2033 | 143.92 | 143.5593 | 143.5593 | -0.36073 | 1.052 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.876 | 5.9028 | 5.8909 |  |  |  | -0.012039 | 0.014765 | 0.0029351
6.345 | 10 | 6.347 | d | Colorimeter probe problems with drift for entire day's run | 8 | 103 | 2500 | 2032 | 143.48 | 143.193 | 143.193 | -0.28696 | 1.13 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.3338 | 6.3594 | 6.3467 |  |  |  | -0.013177 | 0.012439 | -0.00030413
6.908 | 10 | 6.909 | d | Colorimeter probe problems with drift for entire day's run | 8 | 105 | 1000 | 2084 | 144.44 | 144.5316 | 144.5316 | 0.091584 | 1.238 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8957 | 6.9051 | 6.8913 |  |  |  | -0.013308 | -0.0039044 | -0.017738
7.072 | 10 | 7.073 | d | Colorimeter probe problems with drift for entire day's run | 8 | 107 | 750 | 2091 | 143.74 | 143.57 | 143.57 | -0.16996 | 1.261 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0591 | 7.0818 | 7.0676 |  |  |  | -0.013868 | 0.0087939 | -0.0053937
7.062 | 10 | 7.063 | d | Colorimeter probe problems with drift for entire day's run | 8 | 109 | 500 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 1.191 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0488 | 7.1076 | 7.0934 |  |  |  | -0.014225 | 0.044613 | 0.030373
6.95 | 10 | 6.951 | d | Colorimeter probe problems with drift for entire day's run | 8 | 110 | 400 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 1.222 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9371 | 6.9503 | 6.9364 |  |  |  | -0.013863 | -0.00065227 | -0.014577
6.966 | 10 | 6.945 | d | Colorimeter probe problems with drift for entire day's run | 8 | 111 | 300 | 2027 | 142.91 | 141.8814 | 141.8814 | -1.0286 | 1.231 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.931 | 6.9959 | 6.9819 |  |  |  | -0.014044 | 0.05094 | 0.036924
6.993 | 10 | 6.994 | d | Colorimeter probe problems with drift for entire day's run | 8 | 112 | 250 | 2056 | 142.14 | 142.0985 | 142.0985 | -0.041454 | 1.233 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9804 | 6.9965 | 6.9825 |  |  |  | -0.013582 | 0.0025002 | -0.011517
6.983 | 10 | 6.984 | d | Colorimeter probe problems with drift for entire day's run | 8 | 114 | 200 | 2044 | 140.75 | 140.9981 | 140.9981 | 0.2481 | 1.219 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9699 | 6.9714 | 6.9575 |  |  |  | -0.014105 | -0.012579 | -0.026546
7.181 | 10 | 7.114 | d | Colorimeter probe problems with drift for entire day's run | 8 | 115 | 150 | 2085 | 143.59 | 143.5397 | 143.5397 | -0.050266 | 1.267 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1004 | 7.1172 | 7.1029 |  |  |  | -0.013602 | 0.0031783 | -0.01108
7.306 | 10 | 7.308 | d | Colorimeter probe problems with drift for entire day's run | 8 | 116 | 125 | 2089 | 145.42 | 145.2288 | 145.2288 | -0.19118 | 1.318 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.2933 | 7.3177 | 7.3031 |  |  |  | -0.014666 | 0.0097282 | -0.0049309
7.404 | 10 | 7.405 | d | Colorimeter probe problems with drift for entire day's run | 8 | 117 | 100 | 3075 | 135.88 | 135.884 | 135.884 | 0.004022 | 1.247 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3902 | 7.4048 | 7.39 |  |  |  | -0.014811 | -0.00019785 | -0.015033
7.366 | 10 | 7.368 | d | Colorimeter probe problems with drift for entire day's run | 8 | 119 | 75 | 3058 | 138.29 | 138.3074 | 138.3074 | 0.017436 | 1.263 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3531 | 7.3669 | 7.3522 |  |  |  | -0.014885 | -0.0010669 | -0.015826
7.377 | 10 | 7.379 | d | Colorimeter probe problems with drift for entire day's run | 8 | 120 | 50 | 2063 | 142.85 | 142.864 | 142.864 | 0.014046 | 1.307 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3641 | 7.3782 | 7.3634 |  |  |  | -0.01486 | -0.00081348 | -0.015594
7.206 | 10 | 6.897 | d | Colorimeter probe problems with drift for entire day's run | 8 | 121 | 35 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 1.248 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8826 | 6.8952 | 6.8814 |  |  |  | -0.014426 | -0.0017551 | -0.015569
7.61 | 10 | 7.612 | d | Colorimeter probe problems with drift for entire day's run | 8 | 122 | 25 | 2061 | 143.79 | 143.1712 | 143.1712 | -0.61879 | 1.357 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5965 | 7.6451 | 7.6298 |  |  |  | -0.01551 | 0.033102 | 0.017788
7.555 | 10 | 7.556 | d | Colorimeter probe problems with drift for entire day's run | 8 | 123 | 5 | 2049 | 144.09 | 144.1547 | 144.1547 | 0.064699 | 1.35 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5412 | 7.5529 | 7.5377 |  |  |  | -0.014829 | -0.0031318 | -0.018261
7.012 | 10 | 7.013 | d | Colorimeter probe problems with drift for entire day's run | 8 | 101 | 3459 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.233 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9994 | 7.0119 | 6.9979 |  |  |  | -0.013603 | -0.0010909 | -0.015139
6.652 | 10 | 6.653 | d | Colorimeter probe problems with drift for entire day's run | 8 | 103 | 2500 | 3094 | 139.97 | 139.9812 | 139.9812 | 0.011182 | 1.155 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.6394 | 6.6522 | 6.6389 |  |  |  | -0.013565 | -0.0007737 | -0.014103
7.106 | 10 | 7.107 | d | Colorimeter probe problems with drift for entire day's run | 8 | 115 | 150 | 2015 | 142.5 | 142.6118 | 142.6118 | 0.11175 | 1.256 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0931 | 7.1016 | 7.0874 |  |  |  | -0.013942 | -0.0053521 | -0.01958
7.043 | 10 | 7.044 | d | Colorimeter probe problems with drift for entire day's run | 8 | 9101 | 3459 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 1.25 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.03 | 7.0426 | 7.0285 |  |  |  | -0.014013 | -0.0013688 | -0.015478
6.979 | 10 | 6.981 | d | Colorimeter probe problems with drift for entire day's run | 10 | 126 | 1000 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 1.248 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9668 | 7.024 | 7.0099 |  |  |  | -0.014168 | 0.042967 | 0.028895
6.906 | 10 | 6.907 | d | Colorimeter probe problems with drift for entire day's run | 10 | 128 | 750 | 3137 | 142.52 | 141.6896 | 141.6896 | -0.83038 | 1.221 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8934 | 6.9483 | 6.9344 |  |  |  | -0.01355 | 0.041349 | 0.027428
6.909 | 10 | 6.91 | d | Colorimeter probe problems with drift for entire day's run | 10 | 130 | 500 | 2086 | 143.16 | 143.1706 | 143.1706 | 0.01058 | 1.227 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8961 | 6.9094 | 6.8956 |  |  |  | -0.013913 | -0.00058689 | -0.01443
7.067 | 10 | 7.069 | d | Colorimeter probe problems with drift for entire day's run | 10 | 132 | 350 | 2087 | 143.61 | 143.5166 | 143.5166 | -0.093392 | 1.259 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0544 | 7.0732 | 7.059 |  |  |  | -0.014649 | 0.0041769 | -0.0099935
7.04 | 10 | 7.041 | d | Colorimeter probe problems with drift for entire day's run | 10 | 133 | 300 | 3126 | 145.41 | 145.4256 | 145.4256 | 0.01563 | 1.27 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.027 | 7.0403 | 7.0262 |  |  |  | -0.014032 | -0.00069377 | -0.014798
6.857 | 10 | 6.858 | d | Colorimeter probe problems with drift for entire day's run | 10 | 134 | 250 | 3123 | 142.25 | #N/A | 142.25 | 0 | 1.21 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8442 | 6.8579 | 6.8442 |  |  |  | -0.013816 | -7.5963e-05 | -0.013816
6.852 | 10 | 6.853 | d | Colorimeter probe problems with drift for entire day's run | 10 | 136 | 200 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 1.226 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8396 | 6.8589 | 6.8452 |  |  |  | -0.013437 | 0.0059019 | -0.0078398
6.877 | 10 | 6.878 | d | Colorimeter probe problems with drift for entire day's run | 10 | 137 | 175 | 2092 | 144.28 | 144.1993 | 144.1993 | -0.080655 | 1.231 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8642 | 6.8819 | 6.8681 |  |  |  | -0.013781 | 0.0038997 | -0.009888
6.959 | 10 | 6.96 | d | Colorimeter probe problems with drift for entire day's run | 10 | 138 | 150 | 3121 | 144.21 | 144.2162 | 144.2162 | 0.0062065 | 1.245 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9461 | 6.9597 | 6.9458 |  |  |  | -0.013907 | -0.00026711 | -0.01421
6.903 | 10 | 6.904 | d | Colorimeter probe problems with drift for entire day's run | 10 | 139 | 125 | 3189 | 133.68 | 133.587 | 133.587 | -0.093034 | 1.144 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.89 | 6.9087 | 6.8948 |  |  |  | -0.014024 | 0.004691 | -0.0091523
6.997 | 10 | 6.998 | d | Colorimeter probe problems with drift for entire day's run | 10 | 140 | 100 | 3118 | 141.72 | 140.7994 | 140.7994 | -0.92059 | 1.23 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9843 | 7.0447 | 7.0306 |  |  |  | -0.013719 | 0.046718 | 0.032604
7.126 | 10 | 7.128 | d | Colorimeter probe problems with drift for entire day's run | 10 | 141 | 75 | 2073 | 142.32 | 141.6233 | 141.6233 | -0.69668 | 1.258 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1135 | 7.1634 | 7.149 |  |  |  | -0.014479 | 0.035367 | 0.021015
7.451 | 10 | 7.452 | d | Colorimeter probe problems with drift for entire day's run | 10 | 143 | 50 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 1.236 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4374 | 7.4543 | 7.4393 |  |  |  | -0.01455 | 0.0022621 | -0.012672
7.54 | 10 | 7.541 | d | Colorimeter probe problems with drift for entire day's run | 10 | 144 | 35 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.241 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5259 | 7.5975 | 7.5823 |  |  |  | -0.015093 | 0.056491 | 0.04127
8.491 | 10 | 8.493 | d | Colorimeter probe problems with drift for entire day's run | 10 | 146 | 15 | 3108 | 143.28 | 143.2932 | 143.2932 | 0.013209 | 1.508 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4761 | 8.4923 | 8.4753 |  |  |  | -0.016887 | -0.00067032 | -0.017679
8.378 | 10 | 8.381 | d | Colorimeter probe problems with drift for entire day's run | 10 | 147 | 5 | 2099 | 148.99 | 148.9069 | 148.9069 | -0.08305 | 1.548 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3638 | 8.3853 | 8.3686 |  |  |  | -0.017172 | 0.00435 | -0.012444
6.873 | 10 | 6.874 | d | Colorimeter probe problems with drift for entire day's run | 10 | 126 | 1000 | 3107 | 141.57 | 141.5746 | 141.5746 | 0.0046261 | 1.207 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8604 | 6.8739 | 6.8602 |  |  |  | -0.013609 | -6.4191e-05 | -0.013836
6.876 | 10 | 6.878 | d | Colorimeter probe problems with drift for entire day's run | 10 | 133 | 300 | 3113 | 144.06 | 144.0711 | 144.0711 | 0.011129 | 1.229 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8636 | 6.8769 | 6.8631 |  |  |  | -0.014377 | -0.0011367 | -0.014914
7.449 | 10 | 7.45 | d | Colorimeter probe problems with drift for entire day's run | 10 | 144 | 35 | 3119 | 140.11 | 140.1242 | 140.1242 | 0.014162 | 1.294 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4352 | 7.4493 | 7.4344 |  |  |  | -0.014813 | -0.00065191 | -0.015575
6.896 | 10 | 6.897 | d | Colorimeter probe problems with drift for entire day's run | 10 | 9126 | 1000 | 3100 | 143.42 | 143.2848 | 143.2848 | -0.13519 | 1.227 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8834 | 6.9038 | 6.89 |  |  |  | -0.013591 | 0.0068267 | -0.007005
6.636 | 10 | 6.638 |  |  | 11 | 149 | 1000 | 2017 | 144.62 | 144.8189 | 144.8189 | 0.19891 | 1.191 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.6242 | 6.6283 | 6.615 |  |  |  | -0.013773 | -0.0097185 | -0.022999
6.775 | 10 | 6.777 |  |  | 11 | 150 | 750 | 2060 | 144.76 | 143.9012 | 143.9012 | -0.85877 | 1.217 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.7629 | 6.8175 | 6.8039 |  |  |  | -0.014054 | 0.040534 | 0.026875
6.903 | 10 | 6.904 |  |  | 11 | 151 | 500 | 3055 | 140.97 | 140.9856 | 140.9856 | 0.01556 | 1.207 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.89 | 6.9031 | 6.8892 |  |  |  | -0.013989 | -0.00093021 | -0.014761
0 | 10 | 7.202 |  |  | 11 | 152 | 350 | 3203 | 137.1 | 137.025 | 137.025 | -0.074958 | 1.224 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1877 | 7.2061 | 7.1917 |  |  |  | -0.01429 | 0.0041376 | -0.0103
7.411 | 10 | 7.413 |  |  | 11 | 153 | 250 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 1.341 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.398 | 7.4116 | 7.3968 |  |  |  | -0.01501 | -0.0013914 | -0.016238
7.447 | 10 | 7.448 | c | samples # 154,155,158 bubbles on probe at end of titration. | 11 | 154 | 225 | 2061 | 143.79 | 143.1712 | 143.1712 | -0.61879 | 1.328 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4334 | 7.481 | 7.466 |  |  |  | -0.014589 | 0.03298 | 0.017994
7.493 | 10 | 7.495 | c | samples # 154,155,158 bubbles on probe at end of titration. | 11 | 155 | 200 | 2027 | 142.91 | 141.8814 | 141.8814 | -1.0286 | 1.328 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4798 | 7.55 | 7.5348 |  |  |  | -0.015166 | 0.054962 | 0.039838
7.561 | 10 | 7.563 |  |  | 11 | 157 | 150 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 1.329 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5476 | 7.562 | 7.5468 |  |  |  | -0.015374 | -0.0010031 | -0.016151
7.598 | 10 | 7.6 |  | samples # 154,155,158 bubbles on probe at end of titration. | 11 | 158 | 125 | 2032 | 143.48 | 143.193 | 143.193 | -0.28696 | 1.352 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.585 | 7.6156 | 7.6004 |  |  |  | -0.015044 | 0.015627 | 0.00037208
7.684 | 10 | 7.685 | c |  | 11 | 159 | 100 | 3075 | 135.88 | 135.884 | 135.884 | 0.004022 | 1.294 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6701 | 7.6853 | 7.6699 |  |  |  | -0.014895 | 0.00027076 | -0.015125
7.664 | 10 | 7.666 |  |  | 11 | 160 | 75 | 3094 | 139.97 | 139.9812 | 139.9812 | 0.011182 | 1.33 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6508 | 7.6655 | 7.6502 |  |  |  | -0.01522 | -0.00048465 | -0.01584
7.714 | 10 | 7.716 |  |  | 11 | 161 | 50 | 2015 | 142.5 | 142.6118 | 142.6118 | 0.11175 | 1.363 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.7003 | 7.7096 | 7.6942 |  |  |  | -0.015712 | -0.0063889 | -0.021832
7.888 | 10 | 7.89 |  |  | 11 | 162 | 35 | 3058 | 138.29 | 138.3074 | 138.3074 | 0.017436 | 1.352 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.8738 | 7.8886 | 7.8728 |  |  |  | -0.016204 | -0.0014092 | -0.017211
8.218 | 10 | 8.219 |  |  | 11 | 163 | 25 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.444 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.2032 | 8.2178 | 8.2014 |  |  |  | -0.015828 | -0.0011683 | -0.017628
8.178 | 10 | 8.18 |  |  | 11 | 164 | 15 | 2049 | 144.09 | 144.1547 | 144.1547 | 0.064699 | 1.461 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.1641 | 8.1767 | 8.1603 |  |  |  | -0.015948 | -0.0032865 | -0.019664
8.031 | 10 | 8.033 |  |  | 11 | 165 | 5 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 1.369 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.0168 | 8.0308 | 8.0147 |  |  |  | -0.01622 | -0.0022481 | -0.018335
6.596 | 10 | 6.598 |  |  | 11 | 149 | 1000 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 1.172 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.5845 | 6.5987 | 6.5855 |  |  |  | -0.013514 | 0.00070821 | -0.012513
6.871 | 10 | 6.872 |  |  | 11 | 151 | 500 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 1.159 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8584 | 6.9157 | 6.9018 |  |  |  | -0.013593 | 0.043657 | 0.0298
7.502 | 10 | 7.504 |  |  | 11 | 155 | 200 | 2063 | 142.85 | 142.864 | 142.864 | 0.014046 | 1.329 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4887 | 7.503 | 7.4879 |  |  |  | -0.015319 | -0.0010356 | -0.016066
0 | 10 | 7.95 |  |  | 11 | 9149 | 1000 | 2014 | 143.89 | 143.9159 | 143.9159 | 0.025906 | 1.418 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.9339 | 7.9484 | 7.9325 |  |  |  | -0.016077 | -0.0016045 | -0.017525
5.229 | 10 | 5.23 |  |  | 13 | 192 | 585 | 3063 | 144.51 | 144.5139 | 144.5139 | 0.0038689 | 0.939 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.2194 | 5.2297 | 5.2193 |  |  |  | -0.010601 | -0.00025909 | -0.010742
5.49 | 10 | 5.491 |  |  | 13 | 196 | 500 | 3151 | 143.12 | 143.1505 | 143.1505 | 0.030486 | 0.976 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.4799 | 5.4897 | 5.4787 |  |  |  | -0.011137 | -0.0013168 | -0.01232
6.127 | 10 | 6.128 |  |  | 13 | 197 | 350 | 3171 | 143.06 | 143.0623 | 143.0623 | 0.0023427 | 1.088 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.1153 | 6.1275 | 6.1152 |  |  |  | -0.012724 | -0.00054652 | -0.012826
6.434 | 10 | 6.435 |  |  | 13 | 199 | 250 | 3194 | 134.59 | 134.541 | 134.541 | -0.048979 | 1.074 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.4217 | 6.437 | 6.4241 |  |  |  | -0.013264 | 0.0020091 | -0.010891
6.658 | 10 | 6.659 |  |  | 13 | 200 | 225 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 1.105 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.6458 | 6.7083 | 6.6948 |  |  |  | -0.01325 | 0.04929 | 0.035847
6.168 | 10 | 6.169 |  |  | 13 | 201 | 200 | 3034 | 138.61 | 138.6146 | 138.6146 | 0.0046251 | 1.061 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.1569 | 6.169 | 6.1567 |  |  |  | -0.012111 | 4.3122e-05 | -0.01232
6.131 | 10 | 6.132 |  |  | 13 | 202 | 175 | 3180 | 145.95 | 145.9067 | 145.9067 | -0.043342 | 1.111 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.1199 | 6.134 | 6.1217 |  |  |  | -0.012099 | 0.0020359 | -0.010256
6.583 | 10 | 6.584 |  |  | 13 | 203 | 150 | 3174 | 141.29 | 141.3182 | 141.3182 | 0.028153 | 1.154 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.5708 | 6.5827 | 6.5695 |  |  |  | -0.013209 | -0.0013467 | -0.014536
7.37 | 10 | 7.371 |  |  | 13 | 204 | 125 | 3164 | 141.91 | 141.9355 | 141.9355 | 0.025539 | 1.297 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3566 | 7.37 | 7.3553 |  |  |  | -0.014373 | -0.00095093 | -0.015715
7.601 | 10 | 7.603 |  |  | 13 | 205 | 100 | 3160 | 144.9 | 144.9088 | 144.9088 | 0.0087513 | 1.366 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5877 | 7.6025 | 7.5872 |  |  |  | -0.015299 | -0.00053506 | -0.015764
8.045 | 10 | 8.047 |  |  | 13 | 207 | 75 | 3190 | 131.85 | 131.8213 | 131.8213 | -0.028737 | 1.314 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.031 | 8.0489 | 8.0327 |  |  |  | -0.016037 | 0.0018648 | -0.014259
8.194 | 10 | 8.196 |  |  | 13 | 208 | 50 | 3182 | 136.7 | 135.614 | 135.614 | -1.086 | 1.388 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.1798 | 8.2629 | 8.2463 |  |  |  | -0.016163 | 0.066873 | 0.050322
8.497 | 10 | 8.499 |  |  | 13 | 209 | 35 | 3191 | 135.01 | 134.9518 | 134.9518 | -0.05824 | 1.421 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4816 | 8.5023 | 8.4853 |  |  |  | -0.017409 | 0.0033363 | -0.013694
8.871 | 10 | 8.873 |  |  | 13 | 210 | 25 | 2047 | 143.62 | 142.7708 | 142.7708 | -0.84922 | 1.579 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.8555 | 8.9268 | 8.9089 |  |  |  | -0.017494 | 0.053805 | 0.035927
9.156 | 10 | 9.159 |  |  | 13 | 211 | 15 | 3177 | 143.13 | 143.1669 | 143.1669 | 0.036864 | 1.624 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.1405 | 9.1564 | 9.1381 |  |  |  | -0.018511 | -0.0025611 | -0.020898
9.207 | 10 | 9.209 |  |  | 13 | 213 | 5 | 3096 | 142.53 | 142.5586 | 142.5586 | 0.028595 | 1.626 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.1909 | 9.2074 | 9.189 |  |  |  | -0.018138 | -0.0015684 | -0.020007
5.571 | 10 | 5.572 |  |  | 13 | 192 | 585 | 3092 | 144.37 | #N/A | 144.37 | 0 | 0.999 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.5606 | 5.5717 | 5.5606 |  |  |  | -0.011438 | -0.00027056 | -0.011438
5.492 | 10 | 5.493 |  |  | 13 | 196 | 500 | 3187 | 131.58 | 131.4873 | 131.4873 | -0.092657 | 0.897 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.4818 | 5.4967 | 5.4857 |  |  |  | -0.011227 | 0.0037152 | -0.0073044
7.574 | 10 | 7.575 |  |  | 13 | 205 | 100 | 3097 | 143.2 | 143.1879 | 143.1879 | -0.012078 | 1.345 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5605 | 7.5763 | 7.5611 |  |  |  | -0.014531 | 0.0012925 | -0.013884
6.318 | 10 | 6.32 | c | Triplicate sample # 9192 had bubble on probe. | 13 | 9192 | 585 | 3098 | 144.59 | 144.5645 | 144.5645 | -0.025475 | 1.134 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.3069 | 6.3207 | 6.308 |  |  |  | -0.013116 | 0.00067665 | -0.011989
7.074 | 10 | 7.075 |  |  | 15 | 234 | 247 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 1.193 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0607 | 7.1196 | 7.1053 |  |  |  | -0.014327 | 0.044611 | 0.030346
7.094 | 10 | 7.096 |  |  | 15 | 235 | 225 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 1.26 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0815 | 7.0968 | 7.0826 |  |  |  | -0.014515 | 0.00077852 | -0.013439
7.091 | 10 | 7.093 |  |  | 15 | 236 | 200 | 3058 | 138.29 | 138.3074 | 138.3074 | 0.017436 | 1.216 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0781 | 7.0915 | 7.0772 |  |  |  | -0.014851 | -0.0015491 | -0.015757
7.252 | 10 | 7.253 |  |  | 15 | 237 | 175 | 3100 | 143.42 | 143.2848 | 143.2848 | -0.13519 | 1.29 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.2386 | 7.2601 | 7.2455 |  |  |  | -0.014389 | 0.0070811 | -0.0074631
7.349 | 10 | 7.351 |  |  | 15 | 238 | 150 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.292 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.336 | 7.3491 | 7.3344 |  |  |  | -0.015003 | -0.0018898 | -0.016612
7.411 | 10 | 7.413 |  |  | 15 | 239 | 125 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 1.341 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.398 | 7.4116 | 7.3968 |  |  |  | -0.01501 | -0.0013914 | -0.016238
7.463 | 10 | 7.465 |  |  | 15 | 240 | 100 | 2032 | 143.48 | 143.193 | 143.193 | -0.28696 | 1.328 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4497 | 7.4798 | 7.4648 |  |  |  | -0.015301 | 0.014823 | -0.00016048
7.676 | 10 | 7.678 |  |  | 15 | 241 | 75 | 2060 | 144.76 | 143.9012 | 143.9012 | -0.85877 | 1.378 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6622 | 7.724 | 7.7085 |  |  |  | -0.015836 | 0.046007 | 0.030535
7.925 | 10 | 7.927 |  | Sample #243 had bubble on probe, sample #246,247 had bugs in sample. | 15 | 243 | 50 | 3094 | 139.97 | 139.9812 | 139.9812 | 0.011182 | 1.375 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.9108 | 7.9261 | 7.9102 |  |  |  | -0.01616 | -0.00092461 | -0.016801
8.415 | 10 | 8.417 | c |  | 15 | 244 | 35 | 2049 | 144.09 | 144.1547 | 144.1547 | 0.064699 | 1.503 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3997 | 8.4128 | 8.3959 |  |  |  | -0.017263 | -0.0042369 | -0.021086
8.703 | 10 | 8.705 |  |  | 15 | 245 | 25 | 2061 | 143.79 | 143.1712 | 143.1712 | -0.61879 | 1.551 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.6874 | 8.743 | 8.7255 |  |  |  | -0.017567 | 0.038022 | 0.020512
8.613 | 10 | 8.615 | c | Sample #243 had bubble on probe, sample #246,247 had bugs in sample. | 15 | 246 | 15 | 3203 | 137.1 | 137.025 | 137.025 | -0.074958 | 1.463 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.5983 | 8.6203 | 8.603 |  |  |  | -0.016741 | 0.0052982 | -0.011968
8.595 | 10 | 8.597 | c | Sample #243 had bubble on probe, sample #246,247 had bugs in sample. | 15 | 247 | 5 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 1.51 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.5803 | 8.5967 | 8.5794 |  |  |  | -0.016678 | -0.00034424 | -0.017562
7.129 | 10 | 7.13 |  |  | 15 | 234 | 247 | 2015 | 142.5 | 142.6118 | 142.6118 | 0.11175 | 1.26 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1158 | 7.1244 | 7.1101 |  |  |  | -0.014242 | -0.0056245 | -0.019897
7.386 | 10 | 7.387 |  |  | 15 | 238 | 150 | 3075 | 135.88 | 135.884 | 135.884 | 0.004022 | 1.244 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3723 | 7.3869 | 7.3721 |  |  |  | -0.014678 | -0.0001001 | -0.014899
8.375 | 10 | 8.377 |  |  | 15 | 244 | 35 | 2063 | 142.85 | 142.864 | 142.864 | 0.014046 | 1.483 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3605 | 8.3764 | 8.3596 |  |  |  | -0.016534 | -0.00059026 | -0.017367
7.159 | 10 | 7.161 | c |  | 15 | 9234 | 247 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 1.221 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1462 | 7.1587 | 7.1443 |  |  |  | -0.014799 | -0.0023411 | -0.016684
0 | 10 | 6.368 |  |  | 16 | 250 | 250 | 3193 | 131.71 | 131.6281 | 131.6281 | -0.081924 | 1.04 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.3553 | 6.3721 | 6.3593 |  |  |  | -0.012682 | 0.0041052 | -0.0086653
0 | 10 | 6.481 |  |  | 16 | 251 | 225 | 3198 | 136.59 | 136.5025 | 136.5025 | -0.087489 | 1.098 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.4685 | 6.4857 | 6.4727 |  |  |  | -0.012508 | 0.0046957 | -0.0083009
0 | 10 | 6.699 |  |  | 16 | 252 | 200 | 3199 | 135.79 | 134.8688 | 134.8688 | -0.92122 | 1.128 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.686 | 6.7458 | 6.7323 |  |  |  | -0.01304 | 0.046833 | 0.033316
0 | 10 | 6.791 |  |  | 16 | 253 | 175 | 2047 | 143.62 | 142.7708 | 142.7708 | -0.84922 | 1.21 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.778 | 6.8326 | 6.8189 |  |  |  | -0.013025 | 0.041553 | 0.027864
0 | 10 | 6.889 |  |  | 16 | 254 | 150 | 3177 | 143.13 | 143.1669 | 143.1669 | 0.036864 | 1.223 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.875 | 6.887 | 6.8732 |  |  |  | -0.014046 | -0.0020432 | -0.015841
0 | 10 | 7.081 |  |  | 16 | 255 | 125 | 3166 | 145.72 | 145.7412 | 145.7412 | 0.02119 | 1.28 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0673 | 7.0804 | 7.0662 |  |  |  | -0.01371 | -0.00056762 | -0.014752
0 | 10 | 7.509 |  |  | 16 | 256 | 100 | 3186 | 130.19 | #N/A | 130.19 | 0 | 1.211 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4943 | 7.5093 | 7.4943 |  |  |  | -0.014701 | 0.00034385 | -0.014701
0 | 10 | 7.197 | c | Sample #257 had leaky burette. | 16 | 257 | 75 | 3196 | 132.43 | 132.3992 | 132.3992 | -0.030816 | 1.181 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1822 | 7.1983 | 7.1839 |  |  |  | -0.014804 | 0.001316 | -0.013107
0 | 10 | 7.741 |  |  | 16 | 258 | 50 | 2059 | 142.45 | 142.4248 | 142.4248 | -0.025198 | 1.367 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.7257 | 7.7426 | 7.7271 |  |  |  | -0.015263 | 0.001633 | -0.013876
0 | 10 | 8.316 |  |  | 16 | 259 | 35 | 3174 | 141.29 | 141.3182 | 141.3182 | 0.028153 | 1.456 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.2995 | 8.3145 | 8.2979 |  |  |  | -0.016457 | -0.0014802 | -0.018134
0 | 10 | 8.845 |  |  | 16 | 260 | 25 | 3194 | 134.59 | 134.541 | 134.541 | -0.048979 | 1.474 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.8272 | 8.8482 | 8.8304 |  |  |  | -0.017821 | 0.0031625 | -0.01456
0 | 10 | 8.501 |  |  | 16 | 261 | 15 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 1.409 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4843 | 8.5641 | 8.547 |  |  |  | -0.016714 | 0.06312 | 0.045966
0 | 10 | 8.47 |  |  | 16 | 262 | 5 | 3096 | 142.53 | 142.5586 | 142.5586 | 0.028595 | 1.496 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4533 | 8.4685 | 8.4515 |  |  |  | -0.016736 | -0.0014947 | -0.018456
0 | 10 | 6.327 |  |  | 16 | 250 | 250 | 3188 | 132.43 | 131.6134 | 131.6134 | -0.81662 | 1.039 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.3141 | 6.3667 | 6.3539 |  |  |  | -0.012878 | 0.039664 | 0.026904
0 | 10 | 6.501 |  |  | 16 | 252 | 200 | 3161 | 145.03 | 145.03 | 145.03 | 0 | 1.17 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.4882 | 6.5012 | 6.4882 |  |  |  | -0.012829 | 0.00019676 | -0.012829
0 | 10 | 7.742 |  |  | 16 | 258 | 50 | 3171 | 143.06 | 143.0623 | 143.0623 | 0.0023427 | 1.373 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.7262 | 7.7416 | 7.7261 |  |  |  | -0.015757 | -0.0003778 | -0.015885
6.357 | 10 | 6.358 |  |  | 16 | 9250 | 250 | 3178 | 143.1 | 143.1174 | 143.1174 | 0.01743 | 1.129 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.3452 | 6.3572 | 6.3444 |  |  |  | -0.01277 | -0.00081535 | -0.013554
4.402 | 10 | 4.403 |  |  | 17 | 263 | 1000 | 3051 | 140.62 | 140.6477 | 140.6477 | 0.027665 | 0.77 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.3938 | 4.4017 | 4.3929 |  |  |  | -0.0092223 | -0.001271 | -0.010099
4.201 | 10 | 4.201 |  |  | 17 | 264 | 750 | 3130 | 144.6 | 143.5616 | 143.5616 | -1.0384 | 0.756 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.1929 | 4.2321 | 4.2236 |  |  |  | -0.0081343 | 0.03111 | 0.022622
5.83 | 10 | 5.831 |  |  | 17 | 265 | 500 | 3043 | 135.85 | 135.8721 | 135.8721 | 0.022116 | 0.983 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.8192 | 5.8299 | 5.8182 |  |  |  | -0.011801 | -0.0010776 | -0.012763
6.187 | 10 | 6.188 |  |  | 17 | 266 | 350 | 3004 | 134.98 | #N/A | 134.98 | 0 | 1.036 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.1751 | 6.1875 | 6.1751 |  |  |  | -0.012944 | -0.00054348 | -0.012944
6.023 | 10 | 6.024 |  |  | 17 | 267 | 250 | 3201 | 137.27 | 137.208 | 137.208 | -0.061964 | 1.026 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.0116 | 6.0264 | 6.0143 |  |  |  | -0.012427 | 0.0024062 | -0.0096718
3.898 | 10 | 3.898 |  |  | 17 | 268 | 225 | 3040 | 132.53 | 132.5551 | 132.5551 | 0.025086 | 0.643 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.8903 | 3.8974 | 3.8896 |  |  |  | -0.0076853 | -0.00061198 | -0.0084328
4.206 | 10 | 4.206 |  |  | 17 | 269 | 200 | 3069 | 138.94 | 138.9873 | 138.9873 | 0.047286 | 0.727 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.1973 | 4.2043 | 4.1959 |  |  |  | -0.0086897 | -0.0017051 | -0.010139
4.374 | 10 | 4.374 |  |  | 17 | 270 | 175 | 3018 | 143.72 | 143.7228 | 143.7228 | 0.0028485 | 0.782 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.3652 | 4.3739 | 4.3651 |  |  |  | -0.008818 | -0.00013404 | -0.0089058
4.649 | 10 | 4.649 |  |  | 17 | 271 | 150 | 3067 | 141.52 | 141.5236 | 141.5236 | 0.0035992 | 0.818 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.6398 | 4.649 | 4.6396 |  |  |  | -0.0092496 | -4.7044e-05 | -0.0093693
6.678 | 10 | 6.679 |  |  | 17 | 272 | 125 | 3202 | 135 | 134.9339 | 134.9339 | -0.066067 | 1.118 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.6657 | 6.6824 | 6.669 |  |  |  | -0.013277 | 0.0034262 | -0.0099642
3.801 | 10 | 3.801 |  |  | 17 | 273 | 100 | 3047 | 135.46 | 135.5022 | 135.5022 | 0.042208 | 0.641 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.793 | 3.7994 | 3.7918 |  |  |  | -0.0080426 | -0.0016176 | -0.0092418
6.915 | 10 | 6.916 |  |  | 17 | 274 | 75 | 3007 | 142.23 | 142.1231 | 142.1231 | -0.10691 | 1.22 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.902 | 6.9212 | 6.9073 |  |  |  | -0.01398 | 0.0051523 | -0.0087143
5.994 | 10 | 5.995 |  |  | 17 | 275 | 50 | 3082 | 140.71 | #N/A | 140.71 | 0 | 1.047 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.9832 | 5.9952 | 5.9832 |  |  |  | -0.0118 | 0.0002154 | -0.0118
4.465 | 10 | 4.465 |  |  | 17 | 276 | 35 | 3061 | 143.14 | 143.1344 | 143.1344 | -0.0056157 | 0.795 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.4566 | 4.4657 | 4.4567 |  |  |  | -0.0084387 | 0.0006941 | -0.0082614
8.051 | 10 | 8.053 |  |  | 17 | 277 | 25 | 3016 | 144.46 | 143.662 | 143.662 | -0.79803 | 1.442 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.0365 | 8.098 | 8.0818 |  |  |  | -0.016494 | 0.044998 | 0.028778
4.597 | 10 | 4.597 |  |  | 17 | 278 | 15 | 3200 | 134.59 | 134.5035 | 134.5035 | -0.086466 | 0.769 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.5876 | 4.5998 | 4.5906 |  |  |  | -0.0094133 | 0.0028056 | -0.0064196
8.943 | 10 | 8.946 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination? | 17 | 279 | 5 | 3084 | 144.62 | 144.6756 | 144.6756 | 0.055643 | 1.603 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.9276 | 8.942 | 8.9241 |  |  |  | -0.018409 | -0.0039833 | -0.021891
4.363 | 10 | 4.364 |  |  | 17 | 263 | 1000 | 3207 | 139.67 | 138.4785 | 138.4785 | -1.1915 | 0.758 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.3546 | 4.4014 | 4.3926 |  |  |  | -0.0094032 | 0.037441 | 0.028613
5.765 | 10 | 5.766 |  |  | 17 | 265 | 500 | 3183 | 144.28 | #N/A | 144.28 | 0 | 1.033 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.7546 | 5.7662 | 5.7546 |  |  |  | -0.011383 | 0.00017334 | -0.011383
5.898 | 10 | 5.898 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination? | 17 | 279 | 5 | 3006 | 143.66 | 143.7048 | 143.7048 | 0.044798 | 1.052 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.8867 | 5.8967 | 5.8849 |  |  |  | -0.011254 | -0.0012975 | -0.013115
4.221 | 10 | 4.221 |  |  | 17 | 9263 | 1000 | 3181 | 145.81 | 144.858 | 144.858 | -0.95198 | 0.766 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.213 | 4.2496 | 4.2411 |  |  |  | -0.0079683 | 0.02863 | 0.020106
4.471 | 10 | 4.472 |  |  | 18 | 280 | 1000 | 3009 | 140.81 | 140.664 | 140.664 | -0.14596 | 0.783 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.4624 | 4.4761 | 4.4671 |  |  |  | -0.0095626 | 0.0041113 | -0.0048655
3.893 | 10 | 3.893 |  |  | 18 | 281 | 750 | 3124 | 144.6 | 144.6411 | 144.6411 | 0.041072 | 0.701 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.8853 | 3.892 | 3.8842 |  |  |  | -0.0076653 | -0.00097614 | -0.008784
3.409 | 10 | 3.408 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 282 | 500 | 3048 | 137.92 | #N/A | 137.92 | 0 | 0.586 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.4017 | 3.4085 | 3.4017 |  |  |  | -0.006336 | 0.00050607 | -0.006336
4.458 | 10 | 4.458 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks | 18 | 283 | 350 | 3019 | 137.46 | 137.4714 | 137.4714 | 0.011449 | 0.762 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.4492 | 4.4578 | 4.4488 |  |  |  | -0.0088142 | -0.00024959 | -0.0091902
3.215 | 10 | 3.214 |  | before analysis. Duplicate sample had less precipitate in flask than others. | 18 | 284 | 250 | 3033 | 139.43 | 139.4588 | 139.4588 | 0.028758 | 0.559 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.2076 | 3.2134 | 3.207 |  |  |  | -0.0063605 | -0.00058005 | -0.0070316
3.151 | 10 | 3.151 |  |  | 18 | 285 | 225 | 3116 | 142.18 | 141.0862 | 141.0862 | -1.0938 | 0.559 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.1447 | 3.1758 | 3.1694 |  |  |  | -0.0062868 | 0.02482 | 0.018444
3.287 | 10 | 3.286 |  |  | 18 | 286 | 200 | 3024 | 145.15 | 145.1762 | 145.1762 | 0.026229 | 0.595 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.28 | 3.286 | 3.2794 |  |  |  | -0.0060121 | -1.7224e-05 | -0.0066129
3.634 | 10 | 3.634 |  |  | 18 | 287 | 175 | 3107 | 141.57 | 141.5746 | 141.5746 | 0.0046261 | 0.641 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.6269 | 3.6341 | 3.6268 |  |  |  | -0.0070881 | 8.4212e-05 | -0.0072083
3.867 | 10 | 3.867 |  |  | 18 | 288 | 150 | 3123 | 142.25 | #N/A | 142.25 | 0 | 0.685 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.8595 | 3.8672 | 3.8595 |  |  |  | -0.007526 | 0.0002327 | -0.007526
4.457 | 10 | 4.457 |  |  | 18 | 289 | 125 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 0.799 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.448 | 4.4844 | 4.4755 |  |  |  | -0.0090408 | 0.027445 | 0.018452
4.688 | 10 | 4.688 |  |  | 18 | 290 | 100 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 0.834 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.6789 | 4.6873 | 4.6779 |  |  |  | -0.0091265 | -0.00070287 | -0.010102
4.274 | 10 | 4.274 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 291 | 75 | 3121 | 144.21 | 144.2162 | 144.2162 | 0.0062065 | 0.767 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.266 | 4.2744 | 4.2659 |  |  |  | -0.0079608 | 0.00042576 | -0.008147
3.919 | 10 | 3.919 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 292 | 50 | 3189 | 133.68 | 133.587 | 133.587 | -0.093034 | 0.652 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.9108 | 3.9215 | 3.9136 |  |  |  | -0.0081641 | 0.0024697 | -0.005399
5.45 | 10 | 5.451 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 293 | 35 | 2099 | 148.99 | 148.9069 | 148.9069 | -0.08305 | 1.009 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.44 | 5.454 | 5.4431 |  |  |  | -0.010966 | 0.0030411 | -0.0078901
7.051 | 10 | 7.053 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 294 | 25 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.161 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0383 | 7.1053 | 7.0911 |  |  |  | -0.014653 | 0.052295 | 0.038058
4.374 | 10 | 4.374 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 295 | 15 | 3118 | 141.72 | 140.7994 | 140.7994 | -0.92059 | 0.771 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.3649 | 4.4027 | 4.3938 |  |  |  | -0.0091073 | 0.028673 | 0.019843
6.978 | 10 | 6.98 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks | 18 | 296 | 5 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 1.158 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9658 | 6.9815 | 6.9675 |  |  |  | -0.014244 | 0.0015036 | -0.012485
2.573 | 10 | 2.572 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 280 | 1000 | 3003 | 143.29 | #N/A | 143.29 | 0 | 0.461 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 2.567 | 2.5721 | 2.567 |  |  |  | -0.0050371 | 0.00013119 | -0.0050371
4.235 | 10 | 4.235 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 287 | 175 | 3137 | 142.52 | 141.6896 | 141.6896 | -0.83038 | 0.751 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.2266 | 4.2602 | 4.2517 |  |  |  | -0.008442 | 0.025227 | 0.016683
4.952 | 10 | 4.953 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 18 | 295 | 15 | 2073 | 142.32 | 141.6233 | 141.6233 | -0.69668 | 0.876 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.9429 | 4.9775 | 4.9675 |  |  |  | -0.010127 | 0.024516 | 0.014536
4.47 | 10 | 4.47 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks | 18 | 9280 | 1000 | 3012 | 143.88 | 143.9013 | 143.9013 | 0.021345 | 0.8 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.4614 | 4.4697 | 4.4607 |  |  |  | -0.0085834 | -0.00029118 | -0.0092546
4.55 | 10 | 4.55 |  | before analysis. Duplicate sample had less precipitate in flask than others. | 21 | 322 | 1000 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 0.757 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.5411 | 4.5839 | 4.5747 |  |  |  | -0.0088896 | 0.033853 | 0.024659
5.053 | 10 | 5.053 |  |  | 21 | 324 | 750 | 3096 | 142.53 | 142.5586 | 142.5586 | 0.028595 | 0.895 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.0433 | 5.0524 | 5.0423 |  |  |  | -0.0097106 | -0.00060764 | -0.010737
5.475 | 10 | 5.475 |  |  | 21 | 326 | 500 | 3188 | 132.43 | 131.6134 | 131.6134 | -0.81662 | 0.9 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.4644 | 5.5099 | 5.4988 |  |  |  | -0.010612 | 0.034863 | 0.023817
5.539 | 10 | 5.54 |  |  | 21 | 327 | 350 | 3174 | 141.29 | 141.3182 | 141.3182 | 0.028153 | 0.972 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.529 | 5.5389 | 5.5278 |  |  |  | -0.011039 | -0.0010542 | -0.012157
5.739 | 10 | 5.74 |  |  | 21 | 329 | 250 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 0.98 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.7286 | 5.7386 | 5.7271 |  |  |  | -0.01143 | -0.0014385 | -0.012941
5.893 | 10 | 5.894 |  |  | 21 | 330 | 225 | 3198 | 136.59 | 136.5025 | 136.5025 | -0.087489 | 0.999 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.882 | 5.8976 | 5.8858 |  |  |  | -0.012009 | 0.003638 | -0.0081825
6.158 | 10 | 6.159 |  |  | 21 | 331 | 200 | 3193 | 131.71 | 131.6281 | 131.6281 | -0.081924 | 1.006 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.1463 | 6.1626 | 6.1502 |  |  |  | -0.012684 | 0.0035516 | -0.0087997
6.524 | 10 | 6.525 |  |  | 21 | 332 | 175 | 3186 | 130.19 | #N/A | 130.19 | 0 | 1.053 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.5115 | 6.5246 | 6.5115 |  |  |  | -0.013464 | -0.00038826 | -0.013464
7.441 | 10 | 7.443 |  |  | 21 | 333 | 150 | 3166 | 145.72 | 145.7412 | 145.7412 | 0.02119 | 1.345 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4279 | 7.4417 | 7.4268 |  |  |  | -0.015097 | -0.0012848 | -0.016192
7.581 | 10 | 7.583 |  |  | 21 | 334 | 125 | 3196 | 132.43 | 132.3992 | 132.3992 | -0.030816 | 1.244 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5673 | 7.5843 | 7.5691 |  |  |  | -0.015673 | 0.0013102 | -0.013884
7.688 | 10 | 7.69 |  |  | 21 | 335 | 100 | 2059 | 142.45 | 142.4248 | 142.4248 | -0.025198 | 1.358 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6746 | 7.6914 | 7.676 |  |  |  | -0.015356 | 0.0014279 | -0.013979
7.572 | 10 | 7.574 |  |  | 21 | 337 | 75 | 3161 | 145.03 | 145.03 | 145.03 | 0 | 1.362 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5585 | 7.5737 | 7.5585 |  |  |  | -0.015494 | -0.00032312 | -0.015494
7.671 | 10 | 7.673 |  |  | 21 | 338 | 50 | 3199 | 135.79 | 134.8688 | 134.8688 | -0.92122 | 1.291 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6574 | 7.726 | 7.7105 |  |  |  | -0.015614 | 0.052955 | 0.037477
8.07 | 10 | 8.072 | c | Sample #339 had small bubble on probe at end of titration. | 21 | 339 | 35 | 2047 | 143.62 | 142.7708 | 142.7708 | -0.84922 | 1.437 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.056 | 8.1209 | 8.1046 |  |  |  | -0.015978 | 0.048887 | 0.032621
8.339 | 10 | 8.341 |  |  | 21 | 340 | 25 | 3180 | 145.95 | 145.9067 | 145.9067 | -0.043342 | 1.509 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3244 | 8.3437 | 8.3269 |  |  |  | -0.016563 | 0.0026546 | -0.014056
8.67 | 10 | 8.672 |  |  | 21 | 341 | 15 | 3063 | 144.51 | 144.5139 | 144.5139 | 0.0038689 | 1.553 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.6547 | 8.6719 | 8.6545 |  |  |  | -0.017269 | -0.00013597 | -0.017504
8.82 | 10 | 8.822 |  |  | 21 | 343 | 5 | 3164 | 141.91 | 141.9355 | 141.9355 | 0.025539 | 1.551 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.8042 | 8.8202 | 8.8026 |  |  |  | -0.017832 | -0.0017745 | -0.019439
4.513 | 10 | 4.513 |  |  | 21 | 322 | 1000 | 3171 | 143.06 | 143.0623 | 143.0623 | 0.0023427 | 0.803 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.5043 | 4.5133 | 4.5042 |  |  |  | -0.0086911 | 0.00028476 | -0.0087659
5.467 | 10 | 5.468 |  |  | 21 | 326 | 500 | 3177 | 143.13 | 143.1669 | 143.1669 | 0.036864 | 0.972 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.4569 | 5.4664 | 5.4555 |  |  |  | -0.011124 | -0.0015919 | -0.012549
8.574 | 10 | 8.577 |  |  | 21 | 340 | 15 | 3190 | 131.85 | 131.8213 | 131.8213 | -0.028737 | 1.4 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.559 | 8.5781 | 8.5609 |  |  |  | -0.017954 | 0.0011232 | -0.016059
4.675 | 10 | 4.675 |  |  | 21 | 9322 | 1000 | 3194 | 134.59 | 134.541 | 134.541 | -0.048979 | 0.782 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.6658 | 4.6769 | 4.6675 |  |  |  | -0.0092364 | 0.0018671 | -0.0075123
4.57 | 10 | 4.57 |  |  | 22 | 344 | 1000 | 3183 | 144.28 | #N/A | 144.28 | 0 | 0.82 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.561 | 4.5701 | 4.561 |  |  |  | -0.0090451 | 0.00011901 | -0.0090451
3.607 | 10 | 3.607 |  |  | 22 | 345 | 750 | 3051 | 140.62 | 140.6477 | 140.6477 | 0.027665 | 0.632 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.6 | 3.6065 | 3.5993 |  |  |  | -0.007 | -0.00048078 | -0.0077183
3.786 | 10 | 3.786 |  |  | 22 | 346 | 500 | 3069 | 138.94 | 138.9873 | 138.9873 | 0.047286 | 0.655 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.7781 | 3.7844 | 3.7768 |  |  |  | -0.0079154 | -0.001626 | -0.0092195
2.964 | 10 | 2.963 |  |  | 22 | 347 | 350 | 3130 | 144.6 | 143.5616 | 143.5616 | -1.0384 | 0.535 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 2.9572 | 2.9848 | 2.9788 |  |  |  | -0.0058497 | 0.021836 | 0.015842
3.64 | 10 | 3.663 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 348 | 250 | 3178 | 143.1 | 143.1174 | 143.1174 | 0.01743 | 0.653 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.6554 | 3.6623 | 3.6549 |  |  |  | -0.0076051 | -0.00070791 | -0.0080566
3.244 | 10 | 3.243 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 349 | 225 | 3202 | 135 | 134.9339 | 134.9339 | -0.066067 | 0.546 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.2365 | 3.2447 | 3.2382 |  |  |  | -0.0064555 | 0.0016679 | -0.004847
3.592 | 10 | 3.592 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 350 | 200 | 3040 | 132.53 | 132.5551 | 132.5551 | 0.025086 | 0.593 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.5849 | 3.5914 | 3.5842 |  |  |  | -0.0071108 | -0.0005908 | -0.0077997
3.885 | 10 | 3.885 |  |  | 22 | 351 | 175 | 3043 | 135.85 | 135.8721 | 135.8721 | 0.022116 | 0.657 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.8772 | 3.8844 | 3.8766 |  |  |  | -0.0077822 | -0.00062862 | -0.0084228
3.367 | 10 | 3.366 |  |  | 22 | 352 | 150 | 3067 | 141.52 | 141.5236 | 141.5236 | 0.0035992 | 0.594 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.3596 | 3.3663 | 3.3595 |  |  |  | -0.006389 | 0.00028127 | -0.0064757
4.016 | 10 | 4.016 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 353 | 125 | 3019 | 137.46 | 137.4714 | 137.4714 | 0.011449 | 0.687 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.0077 | 4.0154 | 4.0074 |  |  |  | -0.0082788 | -0.00056151 | -0.0086175
3.255 | 10 | 3.254 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 354 | 100 | 3124 | 144.6 | 144.6411 | 144.6411 | 0.041072 | 0.587 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.2479 | 3.2535 | 3.247 |  |  |  | -0.0060931 | -0.00049748 | -0.0070283
3.76 | 10 | 3.759 |  |  | 22 | 355 | 75 | 3007 | 142.23 | 142.1231 | 142.1231 | -0.10691 | 0.666 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.752 | 3.7624 | 3.7549 |  |  |  | -0.007009 | 0.0034026 | -0.0041465
3.366 | 10 | 3.365 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 356 | 50 | 3207 | 139.67 | 138.4785 | 138.4785 | -1.1915 | 0.586 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.3584 | 3.3946 | 3.3877 |  |  |  | -0.0065764 | 0.029557 | 0.022743
3.519 | 10 | 3.518 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 357 | 35 | 3200 | 134.59 | 134.5035 | 134.5035 | -0.086466 | 0.59 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.5112 | 3.5205 | 3.5134 |  |  |  | -0.0068487 | 0.0025092 | -0.0045575
4.197 | 10 | 4.196 |  |  | 22 | 358 | 25 | 3033 | 139.43 | 139.4588 | 139.4588 | 0.028758 | 0.728 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.1881 | 4.1957 | 4.1873 |  |  |  | -0.0078532 | -0.00031329 | -0.0087294
4.043 | 10 | 4.043 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 359 | 15 | 3016 | 144.46 | 143.662 | 143.662 | -0.79803 | 0.727 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.0347 | 4.0656 | 4.0574 |  |  |  | -0.0083259 | 0.022558 | 0.014403
4.24 | 10 | 4.24 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 22 | 360 | 5 | 3084 | 144.62 | 144.6756 | 144.6756 | 0.055643 | 0.763 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.2314 | 4.2383 | 4.2298 |  |  |  | -0.0085875 | -0.0017374 | -0.010238
3.732 | 10 | 3.732 |  |  | 22 | 344 | 1000 | 3201 | 137.27 | 137.208 | 137.208 | -0.061964 | 0.638 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.7245 | 3.7337 | 3.7262 |  |  |  | -0.0074782 | 0.0017213 | -0.0057713
3.76 | 10 | 3.76 |  |  | 22 | 351 | 175 | 3004 | 134.98 | #N/A | 134.98 | 0 | 0.632 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 3.7527 | 3.7602 | 3.7527 |  |  |  | -0.0073154 | 0.00023063 | -0.0073154
4.635 | 10 | 4.635 |  |  | 22 | 360 | 5 | 3006 | 143.66 | 143.7048 | 143.7048 | 0.044798 | 0.828 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.6259 | 4.6338 | 4.6245 |  |  |  | -0.0090547 | -0.0012256 | -0.010517
4.52 | 10 | 4.52 |  |  | 22 | 9344 | 1000 | 3018 | 143.72 | 143.7228 | 143.7228 | 0.0028485 | 0.808 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.5115 | 4.5204 | 4.5114 |  |  |  | -0.008537 | 0.00043716 | -0.0086277
4.562 | 10 | 4.562 |  |  | 23 | 361 | 1000 | 3118 | 141.72 | 140.7994 | 140.7994 | -0.92059 | 0.804 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.5532 | 4.5926 | 4.5834 |  |  |  | -0.0087853 | 0.030624 | 0.021414
5.059 | 10 | 5.059 |  |  | 23 | 362 | 750 | 3137 | 142.52 | 141.6896 | 141.6896 | -0.83038 | 0.896 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.0493 | 5.0895 | 5.0793 |  |  |  | -0.0096775 | 0.030541 | 0.020338
5.431 | 10 | 5.432 |  |  | 23 | 363 | 500 | 3121 | 144.21 | 144.2162 | 144.2162 | 0.0062065 | 0.973 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.421 | 5.4317 | 5.4208 |  |  |  | -0.010959 | -0.00030799 | -0.011195
5.471 | 10 | 5.472 |  |  | 23 | 365 | 350 | 3189 | 133.68 | 133.587 | 133.587 | -0.093034 | 0.908 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.461 | 5.4758 | 5.4648 |  |  |  | -0.011042 | 0.0037962 | -0.0071812
5.961 | 10 | 5.962 |  |  | 23 | 367 | 250 | 2099 | 148.99 | 148.9069 | 148.9069 | -0.08305 | 1.103 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.9499 | 5.9653 | 5.9533 |  |  |  | -0.012065 | 0.0032526 | -0.008701
6.387 | 10 | 6.388 |  |  | 23 | 368 | 225 | 2073 | 142.32 | 141.6233 | 141.6233 | -0.69668 | 1.128 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.3748 | 6.4195 | 6.4066 |  |  |  | -0.013181 | 0.031491 | 0.018627
6.98 | 10 | 6.982 |  |  | 23 | 369 | 200 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 1.239 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.9678 | 6.9804 | 6.9664 |  |  |  | -0.014182 | -0.0016492 | -0.015634
7.406 | 10 | 7.408 |  |  | 23 | 370 | 175 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 1.324 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3932 | 7.4538 | 7.4389 |  |  |  | -0.014811 | 0.045819 | 0.030887
6.882 | 10 | 6.883 |  |  | 23 | 371 | 150 | 3009 | 140.81 | 140.664 | 140.664 | -0.14596 | 1.202 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.8692 | 6.8903 | 6.8765 |  |  |  | -0.013768 | 0.0072672 | -0.0065378
7.362 | 10 | 7.363 |  |  | 23 | 372 | 125 | 3119 | 140.11 | 140.1242 | 140.1242 | 0.014162 | 1.279 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3486 | 7.3626 | 7.3478 |  |  |  | -0.014412 | -0.00041535 | -0.015165
7.323 | 10 | 7.324 |  |  | 23 | 373 | 100 | 2064 | 144.23 | 144.2453 | 144.2453 | 0.015303 | 1.31 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3095 | 7.3234 | 7.3087 |  |  |  | -0.014493 | -0.00060861 | -0.015279
7.426 | 10 | 7.428 |  |  | 23 | 374 | 75 | 2074 | 141.39 | 141.3715 | 141.3715 | -0.018512 | 1.302 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.4127 | 7.4285 | 7.4137 |  |  |  | -0.015327 | 0.00053872 | -0.014343
7.683 | 10 | 7.685 |  |  | 23 | 375 | 50 | 3108 | 143.28 | 143.2932 | 143.2932 | 0.013209 | 1.365 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6691 | 7.6837 | 7.6683 |  |  |  | -0.015938 | -0.0012632 | -0.016655
8.377 | 10 | 8.379 |  |  | 23 | 376 | 35 | 3113 | 144.06 | 144.0711 | 144.0711 | 0.011129 | 1.496 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.3622 | 8.3783 | 8.3616 |  |  |  | -0.016779 | -0.00065307 | -0.017434
8.444 | 10 | 8.447 | c | Sample #377 had leaky stopper... maybe DMQ or air in sample. | 23 | 377 | 25 | 2007 | 143.68 | 144.1373 | 144.1373 | 0.45734 | 1.504 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.4297 | 8.4194 | 8.4025 |  |  |  | -0.017328 | -0.027588 | -0.044451
8.73 | 10 | 8.733 |  |  | 23 | 378 | 15 | 2086 | 143.16 | 143.1706 | 143.1706 | 0.01058 | 1.549 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.7149 | 8.7317 | 8.7143 |  |  |  | -0.018092 | -0.0012577 | -0.018745
9.306 | 10 | 9.308 |  |  | 23 | 379 | 5 | 2087 | 143.61 | 143.5166 | 143.5166 | -0.093392 | 1.656 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.2897 | 9.3145 | 9.2958 |  |  |  | -0.018316 | 0.0064679 | -0.012185
4.61 | 10 | 4.61 |  |  | 23 | 361 | 1000 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 0.761 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.6005 | 4.6443 | 4.635 |  |  |  | -0.0094545 | 0.034315 | 0.025
5.525 | 10 | 5.525 |  |  | 23 | 365 | 350 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 0.918 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.5144 | 5.5269 | 5.5158 |  |  |  | -0.01061 | 0.0018623 | -0.0092172
7.254 | 10 | 7.256 |  |  | 23 | 372 | 125 | 3104 | 143.15 | 143.1468 | 143.1468 | -0.0032473 | 1.288 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.2412 | 7.2559 | 7.2413 |  |  |  | -0.014841 | -0.00013836 | -0.014674
9.317 | 10 | 9.319 |  |  | 23 | 379 | 5 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 1.665 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.301 | 9.3272 | 9.3086 |  |  |  | -0.018049 | 0.0082408 | -0.010438
4.557 | 10 | 4.557 |  |  | 23 | 9361 | 1000 | 3003 | 143.29 | #N/A | 143.29 | 0 | 0.812 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.5478 | 4.5569 | 4.5478 |  |  |  | -0.0092336 | -9.5774e-05 | -0.0092336
7.067 | 10 | 7.069 |  |  | 25 | 399 | 150 | 2092 | 144.28 | 144.1993 | 144.1993 | -0.080655 | 1.265 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.0548 | 7.0729 | 7.0588 |  |  |  | -0.014244 | 0.0039274 | -0.010242
5.219 | 10 | 5.219 |  |  | 25 | 400 | 125 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 0.946 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.2089 | 5.2185 | 5.208 |  |  |  | -0.010145 | -0.0005492 | -0.01101
7.372 | 10 | 7.373 |  |  | 25 | 401 | 100 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.296 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.3588 | 7.372 | 7.3572 |  |  |  | -0.014182 | -0.0010287 | -0.015797
7.729 | 10 | 7.73 |  |  | 25 | 402 | 75 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 1.303 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.7151 | 7.7795 | 7.7639 |  |  |  | -0.014938 | 0.04946 | 0.033876
8.788 | 10 | 8.791 |  |  | 25 | 403 | 50 | 3098 | 144.59 | 144.5645 | 144.5645 | -0.025475 | 1.575 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.7729 | 8.7921 | 8.7745 |  |  |  | -0.018103 | 0.0010723 | -0.016536
8.537 | 10 | 8.539 |  |  | 25 | 404 | 35 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 1.515 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.5217 | 8.5401 | 8.5229 |  |  |  | -0.017349 | 0.0010504 | -0.016054
9.858 | 10 | 9.838 |  |  | 25 | 405 | 25 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 1.727 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.8184 | 9.8371 | 9.8174 |  |  |  | -0.019584 | -0.00089686 | -0.020595
9.811 | 10 | 9.814 |  |  | 25 | 406 | 15 | 2015 | 142.5 | 142.6118 | 142.6118 | 0.11175 | 1.732 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.7944 | 9.8062 | 9.7866 |  |  |  | -0.01962 | -0.0077681 | -0.027405
9.512 | 10 | 9.515 |  |  | 25 | 407 | 5 | 2032 | 143.48 | 143.193 | 143.193 | -0.28696 | 1.691 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.4955 | 9.5339 | 9.5148 |  |  |  | -0.019529 | 0.018861 | -0.00023046
7.129 | 10 | 7.13 |  |  | 25 | 400 | 125 | 3126 | 145.41 | 145.4256 | 145.4256 | 0.01563 | 1.286 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1159 | 7.1294 | 7.1152 |  |  |  | -0.014074 | -0.00056696 | -0.01485
7.928 | 10 | 7.93 |  |  | 25 | 403 | 50 | 3075 | 135.88 | 135.884 | 135.884 | 0.004022 | 1.335 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.9143 | 7.9299 | 7.914 |  |  |  | -0.015713 | -6.513e-05 | -0.01595
7.164 | 10 | 7.166 |  |  | 25 | 9400 | 125 | 3135 | 140.35 | 140.392 | 140.392 | 0.041968 | 1.247 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.1514 | 7.1636 | 7.1492 |  |  |  | -0.014583 | -0.0024001 | -0.016752
4.496 | 10 | 4.496 |  |  | 26 | 408 | 1000 | 3100 | 143.42 | 143.2848 | 143.2848 | -0.13519 | 0.802 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.4872 | 4.5005 | 4.4915 |  |  |  | -0.0087954 | 0.0045233 | -0.0045018
4.901 | 10 | 4.902 |  |  | 26 | 410 | 750 | 3203 | 137.1 | 137.025 | 137.025 | -0.074958 | 0.835 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.8919 | 4.9044 | 4.8946 |  |  |  | -0.010121 | 0.0024285 | -0.0074056
5.486 | 10 | 5.486 |  |  | 26 | 412 | 500 | 3058 | 138.29 | 138.3074 | 138.3074 | 0.017436 | 0.942 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.4752 | 5.4854 | 5.4745 |  |  |  | -0.010847 | -0.0005514 | -0.011547
5.578 | 10 | 5.579 |  |  | 26 | 414 | 350 | 2060 | 144.76 | 143.9012 | 143.9012 | -0.85877 | 1.003 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.5677 | 5.6127 | 5.6014 |  |  |  | -0.011288 | 0.033657 | 0.022408
5.735 | 10 | 5.736 |  |  | 26 | 415 | 250 | 2061 | 143.79 | 143.1712 | 143.1712 | -0.61879 | 1.024 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.7239 | 5.7605 | 5.749 |  |  |  | -0.012107 | 0.024528 | 0.012983
5.84 | 10 | 5.841 |  |  | 26 | 417 | 225 | 2049 | 144.09 | 144.1547 | 144.1547 | 0.064699 | 1.045 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 5.8297 | 5.8387 | 5.827 |  |  |  | -0.011349 | -0.0023014 | -0.014003
6.037 | 10 | 6.038 |  |  | 26 | 418 | 200 | 2017 | 144.62 | 144.8189 | 144.8189 | 0.19891 | 1.084 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.026 | 6.0297 | 6.0176 |  |  |  | -0.011977 | -0.0082861 | -0.020369
6.39 | 10 | 6.391 |  |  | 26 | 419 | 175 | 2044 | 140.75 | 140.9981 | 140.9981 | 0.2481 | 1.116 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.378 | 6.3794 | 6.3666 |  |  |  | -0.013007 | -0.011608 | -0.024392
7.254 | 10 | 7.256 |  |  | 26 | 420 | 150 | 2085 | 143.59 | 143.5397 | 143.5397 | -0.050266 | 1.292 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.2412 | 7.2583 | 7.2438 |  |  |  | -0.014818 | 0.0022945 | -0.012246
7.62 | 10 | 7.622 |  |  | 26 | 421 | 125 | 2084 | 144.44 | 144.5316 | 144.5316 | 0.091584 | 1.365 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6066 | 7.617 | 7.6017 |  |  |  | -0.015393 | -0.0050228 | -0.020281
7.663 | 10 | 7.665 |  |  | 26 | 422 | 100 | 2027 | 142.91 | 141.8814 | 141.8814 | -1.0286 | 1.358 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6496 | 7.7213 | 7.7058 |  |  |  | -0.01541 | 0.056309 | 0.040842
7.338 | 10 | 7.34 |  |  | 26 | 423 | 75 | 2068 | 141.97 | 141.988 | 141.988 | 0.017984 | 1.292 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.325 | 7.3388 | 7.3241 |  |  |  | -0.015009 | -0.0012482 | -0.01595
7.665 | 10 | 7.667 |  |  | 26 | 425 | 50 | 3057 | 138.5 | 138.4973 | 138.4973 | -0.0027015 | 1.316 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.6514 | 7.6669 | 7.6515 |  |  |  | -0.015606 | -9.6025e-05 | -0.015455
7.881 | 10 | 7.883 |  |  | 26 | 426 | 35 | 3142 | 144.08 | 143.9688 | 143.9688 | -0.11125 | 1.408 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.8672 | 7.8892 | 7.8734 |  |  |  | -0.015806 | 0.0061609 | -0.0096413
8.138 | 10 | 8.139 |  |  | 26 | 427 | 25 | 2033 | 143.92 | 143.5593 | 143.5593 | -0.36073 | 1.452 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 8.1233 | 8.1603 | 8.144 |  |  |  | -0.015733 | 0.021312 | 0.0049672
9.238 | 10 | 9.241 |  |  | 26 | 429 | 5 | 2056 | 142.14 | 142.0985 | 142.0985 | -0.041454 | 1.627 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 9.2221 | 9.2434 | 9.2249 |  |  |  | -0.018871 | 0.0023692 | -0.016142
4.487 | 10 | 4.487 |  |  | 26 | 408 | 1000 | 3094 | 139.97 | 139.9812 | 139.9812 | 0.011182 | 0.781 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.478 | 4.4867 | 4.4777 |  |  |  | -0.0089523 | -0.00031715 | -0.0093152
6.031 | 10 | 6.032 |  |  | 26 | 418 | 200 | 2089 | 145.42 | 145.2288 | 145.2288 | -0.19118 | 1.089 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 6.0202 | 6.0403 | 6.0282 |  |  |  | -0.011792 | 0.0083479 | -0.0037565
7.546 | 10 | 7.548 |  |  | 26 | 421 | 125 | 3055 | 140.97 | 140.9856 | 140.9856 | 0.01556 | 1.319 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 7.5326 | 7.5469 | 7.5318 |  |  |  | -0.015387 | -0.0011116 | -0.01623
4.645 | 10 | 4.645 |  |  | 26 | 9408 | 1000 | 2063 | 142.85 | 142.864 | 142.864 | 0.014046 | 0.825 |  |  | 0.004 | 0.714 | 9.98 | 0.010133 |  | 4.6356 | 4.6444 | 4.6351 |  |  |  | -0.0094347 | -0.0005839 | -0.0098969
4.49 | 10 | 4.492 |  |  | 27 | 431 | 1000 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 0.793 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.4833 | 4.4914 | 4.4824 |  |  |  | -0.008652 | -0.00062874 | -0.0096358
4.982 | 10 | 4.985 |  |  | 27 | 433 | 750 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 0.905 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.9752 | 4.9844 | 4.9744 |  |  |  | -0.0098139 | -0.00064773 | -0.01064
5.397 | 10 | 5.4 |  |  | 27 | 435 | 500 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 0.952 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.3897 | 5.3999 | 5.3891 |  |  |  | -0.010327 | -5.7678e-05 | -0.010882
5.535 | 10 | 5.538 |  |  | 27 | 437 | 350 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 0.986 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5269 | 5.5388 | 5.5277 |  |  |  | -0.011144 | 0.00079783 | -0.010304
5.623 | 10 | 5.627 |  |  | 27 | 438 | 250 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 1.009 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.6156 | 5.6315 | 5.6202 |  |  |  | -0.011384 | 0.0044986 | -0.0067883
5.745 | 10 | 5.748 |  |  | 27 | 440 | 225 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 0.956 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7368 | 5.7497 | 5.7382 |  |  |  | -0.011238 | 0.0017361 | -0.0097892
6.03 | 10 | 6.034 | c | Sample #441 had bugs in sample. | 27 | 441 | 200 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 1.081 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.0222 | 6.0716 | 6.0594 |  |  |  | -0.011827 | 0.037564 | 0.025397
6.722 | 10 | 6.726 |  |  | 27 | 442 | 175 | 3104 | 143.15 | 143.1468 | 143.1468 | -0.0032473 | 1.196 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.7128 | 6.7264 | 6.7129 |  |  |  | -0.013228 | 0.0004032 | -0.013074
6.849 | 10 | 6.854 |  |  | 27 | 443 | 150 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 1.218 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.8403 | 6.8526 | 6.8389 |  |  |  | -0.01372 | -0.0014164 | -0.015146
7.201 | 10 | 7.206 |  |  | 27 | 444 | 125 | 2074 | 141.39 | 141.3715 | 141.3715 | -0.018512 | 1.265 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1917 | 7.2071 | 7.1927 |  |  |  | -0.014283 | 0.0011113 | -0.013327
7.387 | 10 | 7.392 |  |  | 27 | 445 | 100 | 3003 | 143.29 | #N/A | 143.29 | 0 | 1.315 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.3768 | 7.3916 | 7.3768 |  |  |  | -0.015192 | -0.00038429 | -0.015192
7.551 | 10 | 7.556 |  |  | 27 | 446 | 75 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.245 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.5405 | 7.6123 | 7.597 |  |  |  | -0.015473 | 0.05625 | 0.041
7.906 | 10 | 7.912 |  |  | 27 | 448 | 50 | 2007 | 143.68 | 144.1373 | 144.1373 | 0.45734 | 1.411 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.8961 | 7.8865 | 7.8707 |  |  |  | -0.015928 | -0.025537 | -0.041334
8.122 | 10 | 8.127 |  |  | 27 | 449 | 35 | 3009 | 140.81 | 140.664 | 140.664 | -0.14596 | 1.42 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.111 | 8.1358 | 8.1195 |  |  |  | -0.01604 | 0.0087937 | -0.0075024
8.44 | 10 | 8.446 |  |  | 27 | 450 | 25 | 3189 | 133.68 | 133.587 | 133.587 | -0.093034 | 1.4 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.4292 | 8.4521 | 8.4352 |  |  |  | -0.016808 | 0.006082 | -0.010848
8.815 | 10 | 8.822 |  |  | 27 | 451 | 15 | 3018 | 143.72 | 143.7228 | 143.7228 | 0.0028485 | 1.573 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.8041 | 8.8216 | 8.8039 |  |  |  | -0.017889 | -0.00039891 | -0.018066
9.183 | 10 | 9.19 |  |  | 27 | 452 | 5 | 3201 | 137.27 | 137.208 | 137.208 | -0.061964 | 1.564 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.1709 | 9.1935 | 9.1751 |  |  |  | -0.01907 | 0.0035455 | -0.014867
4.492 | 10 | 4.494 |  |  | 27 | 431 | 5 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 0.761 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.4849 | 4.5223 | 4.5133 |  |  |  | -0.0091039 | 0.028343 | 0.019272
5.766 | 10 | 5.77 |  |  | 27 | 440 | 225 | 2092 | 144.28 | 144.1993 | 144.1993 | -0.080655 | 1.035 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7584 | 5.7732 | 5.7616 |  |  |  | -0.011629 | 0.0032073 | -0.0083631
4.498 | 10 | 4.498 |  |  | 27 | 9431 | 1000 | 3098 | 144.59 | 144.5645 | 144.5645 | -0.025475 | 0.81 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.4893 | 4.4991 | 4.4901 |  |  |  | -0.0086945 | 0.0011298 | -0.0078923
6.136 | 10 | 6.138 |  |  | 29 | 454 | 150 | 3142 | 144.08 | 143.9688 | 143.9688 | -0.11125 | 1.099 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.1252 | 6.1423 | 6.13 |  |  |  | -0.012823 | 0.0042855 | -0.0080231
6.799 | 10 | 6.8 |  |  | 29 | 456 | 125 | 2027 | 142.91 | 141.8814 | 141.8814 | -1.0286 | 1.207 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.7864 | 6.85 | 6.8363 |  |  |  | -0.013631 | 0.049998 | 0.036273
7.065 | 10 | 7.066 |  |  | 29 | 457 | 100 | 2033 | 143.92 | 143.5593 | 143.5593 | -0.36073 | 1.263 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0523 | 7.0845 | 7.0703 |  |  |  | -0.013711 | 0.018453 | 0.0042603
7.375 | 10 | 7.376 |  |  | 29 | 458 | 75 | 3057 | 138.5 | 138.4973 | 138.4973 | -0.0027015 | 1.268 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.3615 | 7.3764 | 7.3616 |  |  |  | -0.014517 | 0.00040654 | -0.014371
7.69 | 10 | 7.691 |  |  | 29 | 459 | 50 | 2085 | 143.59 | 143.5397 | 143.5397 | -0.050266 | 1.371 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6761 | 7.6943 | 7.6789 |  |  |  | -0.014872 | 0.0032663 | -0.012146
8.098 | 10 | 8.1 |  |  | 29 | 460 | 30 | 2089 | 145.42 | 145.2288 | 145.2288 | -0.19118 | 1.462 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.0834 | 8.1105 | 8.0942 |  |  |  | -0.016556 | 0.010479 | -0.0057658
9.027 | 10 | 9.029 |  |  | 29 | 461 | 20 | 2084 | 144.44 | 144.5316 | 144.5316 | 0.091584 | 1.618 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.0112 | 9.0235 | 9.0054 |  |  |  | -0.017818 | -0.0055373 | -0.023608
9.853 | 10 | 9.856 |  |  | 29 | 462 | 10 | 2068 | 141.97 | 141.988 | 141.988 | 0.017984 | 1.735 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.8358 | 9.8543 | 9.8346 |  |  |  | -0.020166 | -0.0016969 | -0.02143
9.83 | 10 | 9.833 |  |  | 29 | 463 | 5 | 2044 | 140.75 | 140.9981 | 140.9981 | 0.2481 | 1.716 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.8133 | 9.8154 | 9.7958 |  |  |  | -0.019727 | -0.017587 | -0.037242
9.886 | 10 | 9.889 |  |  | 29 | 462 | 10 | 2056 | 142.14 | 142.0985 | 142.0985 | -0.041454 | 1.743 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.8694 | 9.8921 | 9.8723 |  |  |  | -0.01964 | 0.0030892 | -0.016719
2.776 | 10 | 2.775 |  |  | 30 | 464 | 2350 | 3100 | 143.42 | 143.2848 | 143.2848 | -0.13519 | 0.498 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.7696 | 2.7779 | 2.7723 |  |  |  | -0.0053796 | 0.0028503 | -0.0027294
2.782 | 10 | 2.781 |  |  | 30 | 466 | 2250 | 3094 | 139.97 | 139.9812 | 139.9812 | 0.011182 | 0.487 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.7754 | 2.7807 | 2.7752 |  |  |  | -0.0056122 | -0.00025097 | -0.0058371
2.828 | 10 | 2.827 |  |  | 30 | 467 | 2000 | 3058 | 138.29 | 138.3074 | 138.3074 | 0.017436 | 0.489 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.8213 | 2.8266 | 2.8209 |  |  |  | -0.0057154 | -0.0003981 | -0.0060763
2.961 | 10 | 2.96 |  |  | 30 | 468 | 1750 | 2091 | 143.74 | 143.57 | 143.57 | -0.16996 | 0.532 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.9544 | 2.9639 | 2.9579 |  |  |  | -0.0056156 | 0.0038829 | -0.0020689
3.275 | 10 | 3.275 |  |  | 30 | 469 | 1500 | 3135 | 140.35 | 140.392 | 140.392 | 0.041968 | 0.574 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.2685 | 3.2741 | 3.2675 |  |  |  | -0.0064805 | -0.00089888 | -0.0074717
3.649 | 10 | 3.649 |  |  | 30 | 470 | 1300 | 3203 | 137.1 | 137.025 | 137.025 | -0.074958 | 0.624 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.6419 | 3.6512 | 3.6439 |  |  |  | -0.0071389 | 0.0022105 | -0.0051171
3.938 | 10 | 3.939 |  |  | 30 | 471 | 1200 | 2061 | 143.79 | 143.1712 | 143.1712 | -0.61879 | 0.706 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.9306 | 3.9557 | 3.9478 |  |  |  | -0.0084457 | 0.016718 | 0.0087829
4.231 | 10 | 4.231 |  |  | 30 | 472 | 1100 | 3075 | 135.88 | 135.884 | 135.884 | 0.004022 | 0.716 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.2223 | 4.2306 | 4.2221 |  |  |  | -0.008738 | -0.00037817 | -0.0088648
4.48 | 10 | 4.48 |  |  | 30 | 473 | 1000 | 2015 | 142.5 | 142.6118 | 142.6118 | 0.11175 | 0.795 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.4711 | 4.4765 | 4.4675 |  |  |  | -0.0089299 | -0.0035063 | -0.012483
4.66 | 10 | 4.66 |  |  | 30 | 474 | 900 | 2049 | 144.09 | 144.1547 | 144.1547 | 0.064699 | 0.836 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.6508 | 4.658 | 4.6487 |  |  |  | -0.0091852 | -0.001962 | -0.011302
4.845 | 10 | 4.845 |  |  | 30 | 475 | 800 | 3126 | 145.41 | 145.4256 | 145.4256 | 0.01563 | 0.877 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.8357 | 4.8449 | 4.8351 |  |  |  | -0.0093317 | -0.00014525 | -0.0098587
5.018 | 10 | 5.019 |  |  | 30 | 476 | 700 | 2032 | 143.48 | 143.193 | 143.193 | -0.28696 | 0.896 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.0086 | 5.0288 | 5.0188 |  |  |  | -0.010425 | 0.0098361 | -0.00024566
5.181 | 10 | 5.181 |  |  | 30 | 477 | 600 | 3183 | 144.28 | #N/A | 144.28 | 0 | 0.93 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.1707 | 5.1811 | 5.1707 |  |  |  | -0.010295 | 9.1022e-05 | -0.010295
5.314 | 10 | 5.314 |  |  | 30 | 478 | 500 | 3137 | 142.52 | 141.6896 | 141.6896 | -0.83038 | 0.942 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.3035 | 5.3457 | 5.335 |  |  |  | -0.010529 | 0.031713 | 0.020997
5.383 | 10 | 5.384 |  |  | 30 | 479 | 400 | 2099 | 148.99 | 148.9069 | 148.9069 | -0.08305 | 0.998 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.3734 | 5.3872 | 5.3764 |  |  |  | -0.010591 | 0.003244 | -0.0075536
5.4 | 10 | 5.4 |  |  | 30 | 480 | 350 | 3043 | 135.85 | 135.8721 | 135.8721 | 0.022116 | 0.912 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.3893 | 5.3992 | 5.3884 |  |  |  | -0.010727 | -0.00079324 | -0.011617
5.438 | 10 | 5.439 |  |  | 30 | 482 | 275 | 3007 | 142.23 | 142.1231 | 142.1231 | -0.10691 | 0.962 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.428 | 5.4431 | 5.4322 |  |  |  | -0.010989 | 0.0040627 | -0.0068477
5.451 | 10 | 5.451 |  |  | 30 | 483 | 250 | 3200 | 134.59 | 134.5035 | 134.5035 | -0.086466 | 0.912 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4405 | 5.455 | 5.444 |  |  |  | -0.010513 | 0.0039733 | -0.0069624
5.514 | 10 | 5.515 |  |  | 30 | 484 | 225 | 3033 | 139.43 | 139.4588 | 139.4588 | 0.028758 | 0.956 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5038 | 5.5137 | 5.5027 |  |  |  | -0.011165 | -0.0012639 | -0.012316
5.513 | 10 | 5.513 |  |  | 30 | 486 | 200 | 3004 | 134.98 | #N/A | 134.98 | 0 | 0.925 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5024 | 5.5134 | 5.5024 |  |  |  | -0.010621 | 0.00043116 | -0.010621
2.825 | 10 | 2.824 |  |  | 30 | 466 | 2250 | 3055 | 140.97 | 140.9856 | 140.9856 | 0.01556 | 0.498 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.8184 | 2.8238 | 2.8181 |  |  |  | -0.005552 | -0.00019544 | -0.0058675
4.504 | 10 | 4.504 |  |  | 30 | 473 | 1000 | 2060 | 144.76 | 143.9012 | 143.9012 | -0.85877 | 0.812 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.4951 | 4.5314 | 4.5223 |  |  |  | -0.0088844 | 0.027406 | 0.01832
5.441 | 10 | 5.442 |  |  | 30 | 481 | 300 | 3118 | 141.72 | 140.7994 | 140.7994 | -0.92059 | 0.959 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4307 | 5.4777 | 5.4667 |  |  |  | -0.011274 | 0.035725 | 0.024745
2.825 | 10 | 2.825 |  |  | 30 | 9466 | 2250 | 2017 | 144.62 | 144.8189 | 144.8189 | 0.19891 | 0.511 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.8189 | 2.8206 | 2.815 |  |  |  | -0.0060979 | -0.0043588 | -0.010024
2.746 | 10 | 2.746 |  |  | 31 | 488 | 2045 | 3187 | 131.58 | 131.4873 | 131.4873 | -0.092657 | 0.452 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.74 | 2.7475 | 2.742 |  |  |  | -0.0059998 | 0.001482 | -0.0040392
2.791 | 10 | 2.79 |  |  | 31 | 489 | 2000 | 3092 | 144.37 | #N/A | 144.37 | 0 | 0.504 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.7847 | 2.7903 | 2.7847 |  |  |  | -0.0053009 | 0.0003036 | -0.0053009
2.813 | 10 | 2.813 |  |  | 31 | 490 | 1750 | 3191 | 135.01 | 134.9518 | 134.9518 | -0.05824 | 0.475 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.807 | 2.8139 | 2.8083 |  |  |  | -0.0059594 | 0.00092367 | -0.0047297
3.155 | 10 | 3.154 |  |  | 31 | 491 | 1500 | 3151 | 143.12 | 143.1505 | 143.1505 | 0.030486 | 0.564 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.1479 | 3.1536 | 3.1473 |  |  |  | -0.0060654 | -0.00041408 | -0.0067453
3.72 | 10 | 3.72 |  |  | 31 | 492 | 1250 | 3034 | 138.61 | 138.6146 | 138.6146 | 0.0046251 | 0.643 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.7124 | 3.7197 | 3.7122 |  |  |  | -0.0076405 | -0.00030191 | -0.0077662
4.393 | 10 | 4.393 |  |  | 31 | 493 | 1000 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 0.732 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.3842 | 4.4255 | 4.4166 |  |  |  | -0.0087654 | 0.032501 | 0.023624
4.908 | 10 | 4.909 |  |  | 31 | 494 | 750 | 3012 | 143.88 | 143.9013 | 143.9013 | 0.021345 | 0.879 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.899 | 4.9081 | 4.8983 |  |  |  | -0.0099598 | -0.0008565 | -0.010697
5.439 | 10 | 5.439 |  |  | 31 | 495 | 500 | 3186 | 130.19 | #N/A | 130.19 | 0 | 0.88 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4284 | 5.4393 | 5.4284 |  |  |  | -0.010557 | 0.00034845 | -0.010557
5.482 | 10 | 5.483 |  |  | 31 | 496 | 400 | 3160 | 144.9 | 144.9088 | 144.9088 | 0.0087513 | 0.988 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4715 | 5.4821 | 5.4711 |  |  |  | -0.011522 | -0.00086917 | -0.011857
5.531 | 10 | 5.532 |  |  | 31 | 497 | 350 | 3182 | 136.7 | 135.614 | 135.614 | -1.086 | 0.94 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5208 | 5.5768 | 5.5657 |  |  |  | -0.011205 | 0.044846 | 0.033667
5.566 | 10 | 5.567 |  |  | 31 | 498 | 300 | 2014 | 143.89 | 143.9159 | 143.9159 | 0.025906 | 0.996 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5553 | 5.5655 | 5.5543 |  |  |  | -0.011678 | -0.0015367 | -0.012692
5.699 | 10 | 5.7 |  |  | 31 | 499 | 250 | 3096 | 142.53 | 142.5586 | 142.5586 | 0.028595 | 1.01 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.6884 | 5.6987 | 5.6873 |  |  |  | -0.011584 | -0.0013196 | -0.012741
5.842 | 10 | 5.843 |  |  | 31 | 500 | 225 | 3193 | 131.71 | 131.6281 | 131.6281 | -0.081924 | 0.956 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8314 | 5.8468 | 5.8351 |  |  |  | -0.011591 | 0.0038145 | -0.0079053
6.306 | 10 | 6.307 |  |  | 31 | 501 | 200 | 3198 | 136.59 | 136.5025 | 136.5025 | -0.087489 | 1.07 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.2945 | 6.3112 | 6.2986 |  |  |  | -0.012535 | 0.0042069 | -0.0084408
7.006 | 10 | 7.007 |  |  | 31 | 502 | 175 | 3199 | 135.79 | 134.8688 | 134.8688 | -0.92122 | 1.181 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.9928 | 7.0554 | 7.0413 |  |  |  | -0.014228 | 0.048392 | 0.034256
6.941 | 10 | 6.943 |  |  | 31 | 503 | 150 | 3188 | 132.43 | 131.6134 | 131.6134 | -0.81662 | 1.141 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.9287 | 6.9864 | 6.9724 |  |  |  | -0.014299 | 0.043354 | 0.029355
7.15 | 10 | 7.151 |  |  | 31 | 504 | 125 | 3177 | 143.13 | 143.1669 | 143.1669 | 0.036864 | 1.271 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1369 | 7.1494 | 7.135 |  |  |  | -0.014095 | -0.001636 | -0.015959
7.115 | 10 | 7.117 |  |  | 31 | 505 | 100 | 3166 | 145.72 | 145.7412 | 145.7412 | 0.02119 | 1.288 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1025 | 7.1157 | 7.1014 |  |  |  | -0.014518 | -0.0013097 | -0.015565
7.491 | 10 | 7.493 |  |  | 31 | 506 | 75 | 2059 | 142.45 | 142.4248 | 142.4248 | -0.025198 | 1.325 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4776 | 7.494 | 7.479 |  |  |  | -0.015375 | 0.00097882 | -0.014033
7.852 | 10 | 7.854 |  |  | 31 | 507 | 50 | 3161 | 145.03 | 145.03 | 145.03 | 0 | 1.414 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.8382 | 7.854 | 7.8382 |  |  |  | -0.015753 | -2.127e-05 | -0.015753
8.174 | 10 | 8.176 |  |  | 31 | 508 | 35 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 1.395 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.1592 | 8.1734 | 8.157 |  |  |  | -0.016837 | -0.0026176 | -0.018989
9.957 | 10 | 9.96 |  |  | 31 | 509 | 20 | 3180 | 145.95 | 145.9067 | 145.9067 | -0.043342 | 1.803 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.9401 | 9.963 | 9.943 |  |  |  | -0.019944 | 0.0029996 | -0.01695
9.471 | 10 | 9.474 |  |  | 31 | 510 | 5 | 3171 | 143.06 | 143.0623 | 143.0623 | 0.0023427 | 1.681 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.455 | 9.4738 | 9.4548 |  |  |  | -0.019012 | -0.00019712 | -0.019169
2.712 | 10 | 2.712 |  |  | 31 | 489 | 2000 | 3097 | 143.2 | 143.1879 | 143.1879 | -0.012078 | 0.486 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.7063 | 2.7119 | 2.7065 |  |  |  | -0.0057398 | -6.0339e-05 | -0.0055083
5.601 | 10 | 5.602 |  |  | 31 | 498 | 300 | 3174 | 141.29 | 141.3182 | 141.3182 | 0.028153 | 0.984 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5904 | 5.6005 | 5.5893 |  |  |  | -0.011585 | -0.0014891 | -0.012715
7.07 | 10 | 7.071 |  |  | 31 | 505 | 100 | 3196 | 132.43 | 132.3992 | 132.3992 | -0.030816 | 1.162 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0569 | 7.0728 | 7.0586 |  |  |  | -0.014088 | 0.0017513 | -0.01242
2.755 | 10 | 2.743 |  |  | 31 | 9489 | 2000 | 2047 | 143.62 | 142.7708 | 142.7708 | -0.84922 | 0.493 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.7376 | 2.7597 | 2.7541 |  |  |  | -0.0054054 | 0.016653 | 0.011109
2.894 | 10 | 2.893 |  |  | 32 | 511 | 1473 | 3048 | 137.92 | #N/A | 137.92 | 0 | 0.499 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 2.8876 | 2.8934 | 2.8876 |  |  |  | -0.0054482 | 0.00036351 | -0.0054482
3.556 | 10 | 3.556 |  |  | 32 | 512 | 1250 | 3181 | 145.81 | 144.858 | 144.858 | -0.95198 | 0.647 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.5486 | 3.5795 | 3.5723 |  |  |  | -0.0073546 | 0.023476 | 0.016293
4.447 | 10 | 4.447 |  |  | 32 | 513 | 1000 | 3107 | 141.57 | 141.5746 | 141.5746 | 0.0046261 | 0.784 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.4381 | 4.4469 | 4.438 |  |  |  | -0.0088981 | -0.00012714 | -0.0090452
4.798 | 10 | 4.798 |  |  | 32 | 514 | 750 | 3024 | 145.15 | 145.1762 | 145.1762 | 0.026229 | 0.867 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.7888 | 4.7976 | 4.7879 |  |  |  | -0.0091768 | -0.0004352 | -0.010054
5.465 | 10 | 5.465 |  |  | 32 | 515 | 500 | 3084 | 144.62 | 144.6756 | 144.6756 | 0.055643 | 0.983 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4543 | 5.4631 | 5.4522 |  |  |  | -0.010698 | -0.0018747 | -0.012825
5.565 | 10 | 5.566 |  |  | 32 | 516 | 400 | 3019 | 137.46 | 137.4714 | 137.4714 | 0.011449 | 0.951 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5545 | 5.5652 | 5.554 |  |  |  | -0.011515 | -0.00082911 | -0.011985
5.653 | 10 | 5.654 |  |  | 32 | 517 | 350 | 3082 | 140.71 | #N/A | 140.71 | 0 | 0.989 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.6425 | 5.6538 | 5.6425 |  |  |  | -0.011505 | -0.00017272 | -0.011505
5.612 | 10 | 5.613 |  |  | 32 | 518 | 300 | 3061 | 143.14 | 143.1344 | 143.1344 | -0.0056157 | 0.999 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.6018 | 5.6132 | 5.602 |  |  |  | -0.011231 | 0.00024228 | -0.011008
5.75 | 10 | 5.751 |  |  | 32 | 519 | 250 | 3047 | 135.46 | 135.5022 | 135.5022 | 0.042208 | 0.968 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7392 | 5.7489 | 5.7373 |  |  |  | -0.011843 | -0.0021344 | -0.013658
5.835 | 10 | 5.836 |  |  | 32 | 520 | 225 | 3016 | 144.46 | 143.662 | 143.662 | -0.79803 | 1.048 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8238 | 5.8683 | 5.8566 |  |  |  | -0.012238 | 0.03233 | 0.020569
6.237 | 10 | 6.238 |  |  | 32 | 521 | 200 | 3124 | 144.6 | 144.6411 | 144.6411 | 0.041072 | 1.121 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.2257 | 6.2364 | 6.2239 |  |  |  | -0.012305 | -0.001601 | -0.014098
6.908 | 10 | 6.91 |  |  | 32 | 522 | 175 | 3178 | 143.1 | 143.1174 | 143.1174 | 0.01743 | 1.228 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.8957 | 6.9087 | 6.8949 |  |  |  | -0.014253 | -0.0012636 | -0.015105
7.081 | 10 | 7.082 |  |  | 32 | 523 | 150 | 3194 | 134.59 | 134.541 | 134.541 | -0.048979 | 1.183 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0681 | 7.0849 | 7.0707 |  |  |  | -0.013928 | 0.0028792 | -0.011316
7.269 | 10 | 7.271 |  |  | 32 | 524 | 125 | 3202 | 135 | 134.9339 | 134.9339 | -0.066067 | 1.218 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2558 | 7.274 | 7.2594 |  |  |  | -0.01516 | 0.0030194 | -0.011554
7.345 | 10 | 7.347 |  |  | 32 | 525 | 100 | 3069 | 138.94 | 138.9873 | 138.9873 | 0.047286 | 1.267 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.332 | 7.3442 | 7.3295 |  |  |  | -0.014985 | -0.0028027 | -0.017516
8.108 | 10 | 8.11 |  |  | 32 | 526 | 75 | 3051 | 140.62 | 140.6477 | 140.6477 | 0.027665 | 1.415 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.0934 | 8.108 | 8.0917 |  |  |  | -0.016646 | -0.00202 | -0.018261
8.24 | 10 | 8.242 |  |  | 32 | 527 | 50 | 2063 | 142.85 | 142.864 | 142.864 | 0.014046 | 1.461 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.2253 | 8.241 | 8.2245 |  |  |  | -0.016716 | -0.0010298 | -0.017536
9.609 | 10 | 9.612 |  |  | 32 | 528 | 35 | 3108 | 143.28 | 143.2932 | 143.2932 | 0.013209 | 1.708 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.5924 | 9.6108 | 9.5916 |  |  |  | -0.019551 | -0.0012024 | -0.020448
10.484 | 10 | 10.487 |  |  | 32 | 529 | 20 | 3006 | 143.66 | 143.7048 | 143.7048 | 0.044798 | 1.868 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 10.4661 | 10.4838 | 10.4628 |  |  |  | -0.020873 | -0.0031899 | -0.024182
9.395 | 10 | 9.397 |  |  | 32 | 530 | 5 | 3121 | 144.21 | 144.2162 | 144.2162 | 0.0062065 | 1.681 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.3785 | 9.3969 | 9.3781 |  |  |  | -0.018471 | -6.233e-05 | -0.01888
3.527 | 10 | 3.526 |  |  | 32 | 512 | 1250 | 3123 | 142.25 | #N/A | 142.25 | 0 | 0.626 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.5195 | 3.5266 | 3.5195 |  |  |  | -0.0065129 | 0.00056451 | -0.0065129
5.603 | 10 | 5.604 |  |  | 32 | 516 | 400 | 3207 | 139.67 | 138.4785 | 138.4785 | -1.1915 | 0.973 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5926 | 5.6527 | 5.6414 |  |  |  | -0.011428 | 0.048727 | 0.037397
6.276 | 10 | 6.277 |  |  | 32 | 521 | 200 | 3130 | 144.6 | 143.5616 | 143.5616 | -1.0384 | 1.128 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.2648 | 6.3234 | 6.3107 |  |  |  | -0.012215 | 0.04641 | 0.033739
3.568 | 10 | 3.568 |  |  | 32 | 9512 | 1250 | 3116 | 142.18 | 141.0862 | 141.0862 | -1.0938 | 0.633 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.561 | 3.5962 | 3.589 |  |  |  | -0.0069907 | 0.028231 | 0.021014
3.397 | 10 | 3.397 |  |  | 33 | 531 | 1042 | 3104 | 143.15 | 143.1468 | 143.1468 | -0.0032473 | 0.607 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.3899 | 3.3968 | 3.3899 |  |  |  | -0.0071447 | -0.00024908 | -0.0070667
3.534 | 10 | 3.534 |  |  | 33 | 532 | 1000 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 0.636 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.527 | 3.537 | 3.5299 |  |  |  | -0.0070242 | 0.0029598 | -0.0041381
5.729 | 10 | 5.73 |  |  | 33 | 535 | 400 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 1.02 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7186 | 5.731 | 5.7195 |  |  |  | -0.011369 | 0.00098581 | -0.0105
3.451 | 10 | 3.451 |  |  | 33 | 532 | 1000 | 2074 | 141.39 | 141.3715 | 141.3715 | -0.018512 | 0.609 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.4441 | 3.4515 | 3.4445 |  |  |  | -0.0069171 | 0.00046767 | -0.0064597
5.758 | 10 | 5.759 |  |  | 33 | 535 | 400 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 1.032 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7476 | 5.7948 | 5.7832 |  |  |  | -0.01136 | 0.03578 | 0.024166
3.455 | 10 | 3.455 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others. | 33 | 9532 | 1000 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 0.617 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 3.448 | 3.4542 | 3.4473 |  |  |  | -0.0070185 | -0.00080455 | -0.007737
5.747 | 10 | 5.748 |  |  | 33 | 9535 | 400 | 3018 | 143.72 | 143.7228 | 143.7228 | 0.0028485 | 1.027 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7362 | 5.7476 | 5.7361 |  |  |  | -0.011827 | -0.00042287 | -0.011942
7.324 | 10 | 7.326 |  |  | 34 | 558 | 200 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 1.327 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.3109 | 7.3244 | 7.3097 |  |  |  | -0.015059 | -0.0016008 | -0.016273
7.672 | 10 | 7.674 |  |  | 34 | 562 | 100 | 2087 | 143.61 | 143.5166 | 143.5166 | -0.093392 | 1.368 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6582 | 7.6786 | 7.6632 |  |  |  | -0.015826 | 0.0046088 | -0.010772
7.263 | 10 | 7.242 |  |  | 34 | 558 | 200 | 2007 | 143.68 | 144.1373 | 144.1373 | 0.45734 | 1.292 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2272 | 7.2184 | 7.204 |  |  |  | -0.01477 | -0.023563 | -0.038024
7.639 | 10 | 7.641 |  |  | 34 | 562 | 100 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.259 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6257 | 7.6983 | 7.6829 |  |  |  | -0.01526 | 0.057273 | 0.041851
7.272 | 10 | 7.273 |  |  | 34 | 9558 | 200 | 3038 | 141.75 | 141.7644 | 141.7644 | 0.014395 | 1.28 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2587 | 7.2725 | 7.2579 |  |  |  | -0.014336 | -0.00051479 | -0.015084
7.642 | 10 | 7.644 |  |  | 34 | 9562 | 100 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.345 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6285 | 7.6421 | 7.6268 |  |  |  | -0.015503 | -0.0018685 | -0.017177
5.604 | 10 | 5.604 |  |  | 35 | 570 | 500 | 3189 | 133.68 | 133.587 | 133.587 | -0.093034 | 0.931 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.593 | 5.6082 | 5.5969 |  |  |  | -0.011015 | 0.0041811 | -0.0070611
5.726 | 10 | 5.726 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others. | 35 | 570 | 500 | 3209 | 136.03 | 135.1873 | 135.1873 | -0.84268 | 0.968 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7147 | 5.7625 | 5.7509 |  |  |  | -0.01125 | 0.036457 | 0.024907
5.61 | 10 | 5.611 |  |  | 35 | 9570 | 500 | 3164 | 141.91 | 141.9355 | 141.9355 | 0.025539 | 0.99 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.5998 | 5.61 | 5.5988 |  |  |  | -0.011209 | -0.00098626 | -0.012231
0 | 10 | 7.456 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate | 36 | 571 | 150 | 2014 | 143.89 | 143.9159 | 143.9159 | 0.025906 | 1.332 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.441 | 7.4546 | 7.4397 |  |  |  | -0.014978 | -0.0014031 | -0.016336
7.491 | 10 | 7.492 |  | by 3.0ml/L DO value... contamination? | 36 | 572 | 150 | 3092 | 144.37 | #N/A | 144.37 | 0 | 1.343 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4775 | 7.4925 | 7.4775 |  |  |  | -0.014539 | 0.00046958 | -0.014539
7.498 | 10 | 7.499 |  |  | 36 | 573 | 125 | 3187 | 131.58 | 131.4873 | 131.4873 | -0.092657 | 1.224 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4842 | 7.5046 | 7.4896 |  |  |  | -0.014785 | 0.0056059 | -0.0094296
7.648 | 10 | 7.649 |  |  | 36 | 574 | 100 | 3186 | 130.19 | #N/A | 130.19 | 0 | 1.235 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6337 | 7.649 | 7.6337 |  |  |  | -0.0153 | 2.4887e-05 | -0.0153
7.668 | 10 | 7.67 |  |  | 36 | 575 | 95 | 2017 | 144.62 | 144.8189 | 144.8189 | 0.19891 | 1.377 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6542 | 7.6589 | 7.6435 |  |  |  | -0.015808 | -0.011127 | -0.026469
7.824 | 10 | 7.826 |  |  | 36 | 576 | 75 | 3034 | 138.61 | 138.6146 | 138.6146 | 0.0046251 | 1.346 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.8102 | 7.8256 | 7.81 |  |  |  | -0.015774 | -0.00036182 | -0.016038
8.298 | 10 | 8.3 |  |  | 36 | 578 | 50 | 3151 | 143.12 | 143.1505 | 143.1505 | 0.030486 | 1.474 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.2829 | 8.2977 | 8.2811 |  |  |  | -0.017096 | -0.0022656 | -0.018885
8.959 | 10 | 8.961 |  |  | 36 | 579 | 30 | 3191 | 135.01 | 134.9518 | 134.9518 | -0.05824 | 1.5 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.9436 | 8.9655 | 8.9475 |  |  |  | -0.017405 | 0.0044694 | -0.013487
9.279 | 10 | 9.281 |  |  | 36 | 580 | 21 | 3182 | 136.7 | 135.614 | 135.614 | -1.086 | 1.573 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.2629 | 9.357 | 9.3382 |  |  |  | -0.018056 | 0.075972 | 0.057232
9.54 | 10 | 9.543 |  |  | 36 | 581 | 15 | 3188 | 132.43 | 131.6134 | 131.6134 | -0.81662 | 1.566 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.5235 | 9.6027 | 9.5835 |  |  |  | -0.019544 | 0.05969 | 0.040458
7.491 | 10 | 7.493 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate | 36 | 582 | 10 | 3160 | 144.9 | 144.9088 | 144.9088 | 0.0087513 | 1.348 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4776 | 7.4921 | 7.4771 |  |  |  | -0.01541 | -0.00085946 | -0.015868
7.502 | 10 | 7.479 |  | by 3.0ml/L DO value... contamination? | 36 | 572 | 150 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 1.242 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4646 | 7.5349 | 7.5198 |  |  |  | -0.014359 | 0.055883 | 0.040788
8.326 | 10 | 8.328 |  |  | 36 | 578 | 50 | 3012 | 143.88 | 143.9013 | 143.9013 | 0.021345 | 1.487 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.3115 | 8.3269 | 8.3102 |  |  |  | -0.016501 | -0.0010734 | -0.017751
9.779 | 10 | 9.781 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate | 36 | 582 | 10 | 3198 | 136.59 | 136.5025 | 136.5025 | -0.087489 | 1.656 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.7616 | 9.7875 | 9.7679 |  |  |  | -0.019407 | 0.0065426 | -0.013058
5.77 | 10 | 5.771 |  | by 3.0ml/L DO value... contamination? | 38 | 612 | 250 | 2086 | 143.16 | 143.1706 | 143.1706 | 0.01058 | 1.027 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7589 | 5.7701 | 5.7585 |  |  |  | -0.012071 | -0.00093799 | -0.012502
7.125 | 10 | 7.127 |  |  | 38 | 619 | 75 | 3113 | 144.06 | 144.0711 | 144.0711 | 0.011129 | 1.275 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1126 | 7.1263 | 7.112 |  |  |  | -0.014395 | -0.00067569 | -0.014952
5.744 | 10 | 5.745 |  |  | 38 | 612 | 250 | 3047 | 135.46 | 135.5022 | 135.5022 | 0.042208 | 0.967 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7332 | 5.7429 | 5.7314 |  |  |  | -0.01181 | -0.0021112 | -0.013622
7.113 | 10 | 7.114 |  |  | 38 | 619 | 75 | 3048 | 137.92 | #N/A | 137.92 | 0 | 1.218 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1 | 7.1142 | 7.1 |  |  |  | -0.014039 | 0.00021433 | -0.014039
5.784 | 10 | 5.785 |  |  | 39 | 9612 | 250 | 3107 | 141.57 | 141.5746 | 141.5746 | 0.0046261 | 1.018 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7732 | 5.7846 | 5.773 |  |  |  | -0.011814 | -0.00041132 | -0.012005
7.11 | 10 | 7.111 |  |  | 39 | 9619 | 75 | 3024 | 145.15 | 145.1762 | 145.1762 | 0.026229 | 1.282 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0974 | 7.1103 | 7.0961 |  |  |  | -0.013614 | -0.0006694 | -0.014914
4.632 | 10 | 4.632 |  |  | 39 | 625 | 763 | 3181 | 145.81 | 144.858 | 144.858 | -0.95198 | 0.841 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.6229 | 4.663 | 4.6537 |  |  |  | -0.0091237 | 0.031032 | 0.021682
5.488 | 10 | 5.489 |  |  | 39 | 628 | 500 | 3061 | 143.14 | 143.1344 | 143.1344 | -0.0056157 | 0.977 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4776 | 5.4889 | 5.4779 |  |  |  | -0.011356 | -0.00013594 | -0.011138
5.701 | 10 | 5.701 |  |  | 39 | 631 | 250 | 3119 | 140.11 | 140.1242 | 140.1242 | 0.014162 | 0.993 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.6901 | 5.7009 | 5.6895 |  |  |  | -0.010929 | -8.5566e-05 | -0.011512
6.1 | 10 | 6.101 |  |  | 39 | 633 | 200 | 2047 | 143.62 | 142.7708 | 142.7708 | -0.84922 | 1.089 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.0888 | 6.1379 | 6.1256 |  |  |  | -0.012156 | 0.036875 | 0.024575
6.764 | 10 | 6.765 |  |  | 39 | 634 | 150 | 3016 | 144.46 | 143.662 | 143.662 | -0.79803 | 1.214 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.7517 | 6.8033 | 6.7897 |  |  |  | -0.01334 | 0.038325 | 0.024694
7.387 | 10 | 7.389 |  |  | 39 | 635 | 100 | 3207 | 139.67 | 138.4785 | 138.4785 | -1.1915 | 1.281 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.3741 | 7.4534 | 7.4385 |  |  |  | -0.014884 | 0.064426 | 0.049494
7.691 | 10 | 7.693 |  |  | 39 | 638 | 75 | 3130 | 144.6 | 143.5616 | 143.5616 | -1.0384 | 1.381 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6776 | 7.7494 | 7.7339 |  |  |  | -0.015398 | 0.056443 | 0.04092
9.106 | 10 | 9.109 |  |  | 39 | 642 | 50 | 2073 | 142.32 | 141.6233 | 141.6233 | -0.69668 | 1.608 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.0906 | 9.1543 | 9.1359 |  |  |  | -0.018424 | 0.045268 | 0.026935
9.216 | 10 | 9.218 |  |  | 39 | 645 | 25 | 3194 | 134.59 | 134.541 | 134.541 | -0.048979 | 1.538 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.2001 | 9.222 | 9.2035 |  |  |  | -0.017853 | 0.004017 | -0.014453
8.766 | 10 | 8.768 |  |  | 39 | 648 | 5 | 3069 | 138.94 | 138.9873 | 138.9873 | 0.047286 | 1.511 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.7509 | 8.7654 | 8.7479 |  |  |  | -0.017109 | -0.0025738 | -0.020129
4.647 | 10 | 4.647 |  |  | 39 | 625 | 763 | 3019 | 137.46 | 137.4714 | 137.4714 | 0.011449 | 0.795 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.6374 | 4.6463 | 4.637 |  |  |  | -0.0095768 | -0.00065094 | -0.0099687
6.757 | 10 | 6.758 |  |  | 39 | 634 | 150 | 3124 | 144.6 | 144.6411 | 144.6411 | 0.041072 | 1.214 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.745 | 6.7566 | 6.7431 |  |  |  | -0.012969 | -0.001374 | -0.014911
9.095 | 10 | 9.098 |  |  | 39 | 642 | 50 | 3108 | 143.28 | 143.2932 | 143.2932 | 0.013209 | 1.617 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.0795 | 9.0969 | 9.0787 |  |  |  | -0.018467 | -0.0010977 | -0.019316
4.721 | 10 | 4.721 |  |  | 39 | 9625 | 763 | 3178 | 143.1 | 143.1174 | 143.1174 | 0.01743 | 0.841 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.7117 | 4.7205 | 4.7111 |  |  |  | -0.0093356 | -0.00045238 | -0.0099176
7.249 | 10 | 7.251 |  |  | 40 | 652 | 150 | 3116 | 142.18 | 141.0862 | 141.0862 | -1.0938 | 1.28 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2364 | 7.3079 | 7.2933 |  |  |  | -0.014602 | 0.056947 | 0.042306
7.771 | 10 | 7.773 |  |  | 40 | 655 | 50 | 2074 | 141.39 | 141.3715 | 141.3715 | -0.018512 | 1.364 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.7573 | 7.7739 | 7.7583 |  |  |  | -0.01571 | 0.00089294 | -0.014679
7.258 | 10 | 7.259 |  |  | 40 | 652 | 150 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 1.3 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2451 | 7.2656 | 7.251 |  |  |  | -0.013908 | 0.0065758 | -0.0079793
7.774 | 10 | 7.765 |  |  | 40 | 655 | 50 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 1.38 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.7492 | 7.7659 | 7.7504 |  |  |  | -0.015816 | 0.00091735 | -0.014639
7.279 | 10 | 7.281 |  |  | 40 | 9652 | 150 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 1.303 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.266 | 7.3256 | 7.3109 |  |  |  | -0.015022 | 0.044565 | 0.02989
7.78 | 10 | 7.782 |  |  | 40 | 9655 | 50 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 1.382 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.766 | 7.7799 | 7.7643 |  |  |  | -0.016035 | -0.0020692 | -0.017653
6.086 | 10 | 6.087 |  |  | 41 | 674 | 300 | 3121 | 144.21 | 144.2162 | 144.2162 | 0.0062065 | 1.091 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.0748 | 6.0867 | 6.0745 |  |  |  | -0.012219 | -0.00028646 | -0.012484
6.124 | 10 | 6.124 |  |  | 41 | 677 | 200 | 3181 | 145.81 | 144.858 | 144.858 | -0.95198 | 1.11 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.1124 | 6.1655 | 6.1531 |  |  |  | -0.011597 | 0.041489 | 0.029135
6.564 | 10 | 6.565 |  |  | 41 | 678 | 175 | 3194 | 134.59 | 134.541 | 134.541 | -0.048979 | 1.097 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.5516 | 6.5672 | 6.554 |  |  |  | -0.013431 | 0.0021501 | -0.01101
6.927 | 10 | 6.929 |  |  | 41 | 679 | 150 | 3123 | 142.25 | #N/A | 142.25 | 0 | 1.224 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.9148 | 6.9287 | 6.9148 |  |  |  | -0.014172 | -0.0002907 | -0.014172
7.232 | 10 | 7.233 |  |  | 41 | 680 | 125 | 3084 | 144.62 | 144.6756 | 144.6756 | 0.055643 | 1.299 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2187 | 7.2303 | 7.2159 |  |  |  | -0.014319 | -0.0026502 | -0.017135
7.139 | 10 | 7.141 |  |  | 41 | 681 | 100 | 2064 | 144.23 | 144.2453 | 144.2453 | 0.015303 | 1.279 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1265 | 7.14 | 7.1257 |  |  |  | -0.014501 | -0.00096399 | -0.015268
7.019 | 10 | 7.021 |  |  | 41 | 682 | 75 | 3061 | 143.14 | 143.1344 | 143.1344 | -0.0056157 | 1.248 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0066 | 7.021 | 7.0069 |  |  |  | -0.014367 | -2.2447e-05 | -0.014088
7.298 | 10 | 7.3 |  |  | 41 | 683 | 50 | 3107 | 141.57 | 141.5746 | 141.5746 | 0.0046261 | 1.283 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.2851 | 7.2995 | 7.2849 |  |  |  | -0.014859 | -0.00047674 | -0.0151
9.635 | 10 | 9.638 |  |  | 41 | 684 | 35 | 2073 | 142.32 | 141.6233 | 141.6233 | -0.69668 | 1.701 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.6184 | 9.6857 | 9.6663 |  |  |  | -0.01965 | 0.047739 | 0.028343
10.56 | 10 | 10.563 |  |  | 41 | 685 | 25 | 3130 | 144.6 | 143.5616 | 143.5616 | -1.0384 | 1.894 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 10.5423 | 10.641 | 10.6197 |  |  |  | -0.020673 | 0.077964 | 0.056658
10.361 | 10 | 10.364 |  |  | 41 | 686 | 15 | 3119 | 140.11 | 140.1242 | 140.1242 | 0.014162 | 1.8 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 10.3431 | 10.3628 | 10.342 |  |  |  | -0.020925 | -0.0012349 | -0.021985
6.13 | 10 | 6.131 |  |  | 41 | 674 | 300 | 3202 | 135 | 134.9339 | 134.9339 | -0.066067 | 1.028 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.1182 | 6.1336 | 6.1213 |  |  |  | -0.012753 | 0.00258 | -0.0097127
6.6 | 10 | 6.601 |  |  | 41 | 678 | 175 | 3069 | 138.94 | 138.9873 | 138.9873 | 0.047286 | 1.139 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.5877 | 6.5986 | 6.5854 |  |  |  | -0.013314 | -0.0023655 | -0.015588
7.179 | 10 | 7.181 |  |  | 41 | 681 | 100 | 3207 | 139.67 | 138.4785 | 138.4785 | -1.1915 | 1.245 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.1659 | 7.243 | 7.2284 |  |  |  | -0.015116 | 0.061954 | 0.047444
6.115 | 10 | 6.116 |  |  | 41 | 9674 | 300 | 3006 | 143.66 | 143.7048 | 143.7048 | 0.044798 | 1.092 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.104 | 6.1143 | 6.1021 |  |  |  | -0.012012 | -0.0016889 | -0.013942
6.802 | 10 | 6.803 |  |  | 42 | 693 | 175 | 3097 | 143.2 | 143.1879 | 143.1879 | -0.012078 | 1.21 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.7893 | 6.8036 | 6.7899 |  |  |  | -0.01365 | 0.0005616 | -0.01307
6.826 | 10 | 6.828 |  |  | 42 | 694 | 150 | 3174 | 141.29 | 141.3182 | 141.3182 | 0.028153 | 1.198 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.8138 | 6.8261 | 6.8125 |  |  |  | -0.014156 | -0.0018559 | -0.015533
6.905 | 10 | 6.906 |  |  | 42 | 695 | 125 | 3160 | 144.9 | 144.9088 | 144.9088 | 0.0087513 | 1.243 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.8925 | 6.9059 | 6.8921 |  |  |  | -0.013526 | -0.00011229 | -0.013948
6.451 | 10 | 6.452 |  |  | 42 | 696 | 100 | 3199 | 135.79 | 134.8688 | 134.8688 | -0.92122 | 1.088 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.4392 | 6.4969 | 6.4839 |  |  |  | -0.012762 | 0.044903 | 0.031884
6.675 | 10 | 6.654 |  |  | 42 | 697 | 75 | 3177 | 143.13 | 143.1669 | 143.1669 | 0.036864 | 1.183 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.6404 | 6.652 | 6.6386 |  |  |  | -0.013628 | -0.0020344 | -0.015362
7.027 | 10 | 7.029 |  |  | 42 | 698 | 50 | 3161 | 145.03 | 145.03 | 145.03 | 0 | 1.266 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0143 | 7.0283 | 7.0143 |  |  |  | -0.014739 | -0.00065809 | -0.014739
9.008 | 10 | 9.01 |  |  | 42 | 699 | 35 | 3096 | 142.53 | 142.5586 | 142.5586 | 0.028595 | 1.593 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.992 | 9.0082 | 8.9902 |  |  |  | -0.018006 | -0.0017948 | -0.019835
10.794 | 10 | 10.798 |  |  | 42 | 700 | 25 | 3171 | 143.06 | 143.0623 | 143.0623 | 0.0023427 | 1.915 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 10.776 | 10.7974 | 10.7758 |  |  |  | -0.02203 | -0.00058955 | -0.022208
10.305 | 10 | 10.308 |  |  | 42 | 701 | 15 | 3012 | 143.88 | 143.9013 | 143.9013 | 0.021345 | 1.839 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 10.2871 | 10.3062 | 10.2856 |  |  |  | -0.020867 | -0.001778 | -0.022414
6.815 | 10 | 6.816 |  |  | 42 | 694 | 150 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.2 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.8023 | 6.8145 | 6.8008 |  |  |  | -0.013675 | -0.0015139 | -0.015167
6.479 | 10 | 6.48 |  |  | 42 | 696 | 100 | 3052 | 137.55 | 137.5858 | 137.5858 | 0.03576 | 1.107 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.4672 | 6.4785 | 6.4655 |  |  |  | -0.01275 | -0.0014739 | -0.014456
6.798 | 10 | 6.799 |  |  | 42 | 9694 | 150 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.121 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.7858 | 6.8503 | 6.8366 |  |  |  | -0.013216 | 0.051331 | 0.037604
5.402 | 10 | 5.402 |  |  | 45 | 728 | 435 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 0.969 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.3916 | 5.4069 | 5.396 |  |  |  | -0.010367 | 0.0048832 | -0.0059545
4.989 | 10 | 4.99 |  |  | 45 | 729 | 350 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 0.889 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.9797 | 4.9905 | 4.9805 |  |  |  | -0.010265 | 0.00049684 | -0.0095082
4.976 | 10 | 4.976 |  |  | 45 | 729 | 350 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 0.886 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.9663 | 4.9753 | 4.9653 |  |  |  | -0.0096686 | -0.00072897 | -0.010704
5.859 | 10 | 5.86 |  |  | 50 | 807 | 350 | 3109 | 144.13 | 143.2569 | 143.2569 | -0.87312 | 1.05 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8485 | 5.8965 | 5.8846 |  |  |  | -0.011511 | 0.036456 | 0.024639
5.799 | 10 | 5.799 |  |  | 50 | 809 | 250 | 3116 | 142.18 | 141.0862 | 141.0862 | -1.0938 | 1.025 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7878 | 5.8451 | 5.8333 |  |  |  | -0.011171 | 0.04606 | 0.034346
5.777 | 10 | 5.778 |  |  | 50 | 810 | 225 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 1.048 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7667 | 5.7773 | 5.7657 |  |  |  | -0.011314 | -0.00069345 | -0.012272
5.843 | 10 | 5.843 |  |  | 50 | 811 | 200 | 3018 | 143.72 | 143.7228 | 143.7228 | 0.0028485 | 1.044 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8317 | 5.8433 | 5.8316 |  |  |  | -0.011305 | 0.0002885 | -0.011422
6.024 | 10 | 6.025 |  |  | 50 | 812 | 175 | 3104 | 143.15 | 143.1468 | 143.1468 | -0.0032473 | 1.072 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.0132 | 6.0254 | 6.0133 |  |  |  | -0.01179 | 0.00042377 | -0.011651
6.101 | 10 | 6.102 |  |  | 50 | 813 | 150 | 3063 | 144.51 | 144.5139 | 144.5139 | 0.0038689 | 1.096 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.0899 | 6.102 | 6.0898 |  |  |  | -0.012068 | -5.3554e-06 | -0.012233
6.455 | 10 | 6.456 |  |  | 50 | 814 | 125 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 1.073 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.4434 | 6.458 | 6.445 |  |  |  | -0.01261 | 0.0019588 | -0.010983
6.314 | 10 | 6.315 |  |  | 50 | 815 | 100 | 2060 | 144.76 | 143.9012 | 143.9012 | -0.85877 | 1.136 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.3024 | 6.3533 | 6.3405 |  |  |  | -0.012613 | 0.038259 | 0.025529
6.486 | 10 | 6.487 |  |  | 50 | 816 | 75 | 3055 | 140.97 | 140.9856 | 140.9856 | 0.01556 | 1.136 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.4743 | 6.4865 | 6.4735 |  |  |  | -0.012733 | -0.00046064 | -0.013458
7.448 | 10 | 7.449 |  |  | 50 | 818 | 35 | 3201 | 137.27 | 137.208 | 137.208 | -0.061964 | 1.269 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4343 | 7.4526 | 7.4377 |  |  |  | -0.014693 | 0.0036448 | -0.011286
7.808 | 10 | 7.809 |  |  | 50 | 819 | 25 | 3003 | 143.29 | #N/A | 143.29 | 0 | 1.389 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.7939 | 7.8095 | 7.7939 |  |  |  | -0.015125 | 0.00051802 | -0.015125
8.369 | 10 | 8.371 |  |  | 50 | 820 | 15 | 3033 | 139.43 | 139.4588 | 139.4588 | 0.028758 | 1.448 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.3546 | 8.3697 | 8.3529 |  |  |  | -0.016353 | -0.0013365 | -0.018101
9.753 | 10 | 9.756 | c | Sample #821 had bugs in sample. Duplicate sample had small bubble on probe | 50 | 821 | 5 | 3004 | 134.98 | #N/A | 134.98 | 0 | 1.632 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.7361 | 9.7556 | 9.7361 |  |  |  | -0.019941 | -0.00040377 | -0.019941
5.773 | 10 | 5.774 | c | Sample #821 had bugs in sample. Duplicate sample had small bubble on probe | 50 | 807 | 350 | 2051 | 143.08 | 143.1094 | 143.1094 | 0.029407 | 1.027 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7622 | 5.7726 | 5.761 |  |  |  | -0.011805 | -0.0014366 | -0.013006
6.039 | 10 | 6.04 |  | at end of titration. | 50 | 812 | 175 | 3190 | 131.85 | 131.8213 | 131.8213 | -0.028737 | 0.989 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.0275 | 6.0409 | 6.0288 |  |  |  | -0.012503 | 0.00093938 | -0.011169
7.479 | 10 | 7.48 |  |  | 50 | 818 | 35 | 2092 | 144.28 | 144.1993 | 144.1993 | -0.080655 | 1.34 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.4654 | 7.4846 | 7.4696 |  |  |  | -0.0146 | 0.0046277 | -0.010365
5.758 | 10 | 5.758 |  |  | 50 | 9807 | 350 | 2069 | 143.18 | 143.1585 | 143.1585 | -0.021454 | 1.025 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.7468 | 5.7592 | 5.7477 |  |  |  | -0.011167 | 0.0012487 | -0.010294
4.874 | 10 | 4.874 |  |  | 52 | 823 | 192 | 3182 | 136.7 | 135.614 | 135.614 | -1.086 | 0.829 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.8646 | 4.914 | 4.9041 |  |  |  | -0.0094116 | 0.039981 | 0.030127
4.869 | 10 | 4.87 |  |  | 52 | 824 | 175 | 3164 | 141.91 | 141.9355 | 141.9355 | 0.025539 | 0.86 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.8599 | 4.8688 | 4.859 |  |  |  | -0.01012 | -0.001245 | -0.011007
5.13 | 10 | 5.118 |  |  | 52 | 825 | 150 | 3048 | 137.92 | #N/A | 137.92 | 0 | 0.878 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.108 | 5.1183 | 5.108 |  |  |  | -0.0099989 | 0.00026263 | -0.0099989
5.447 | 10 | 5.448 |  |  | 52 | 827 | 125 | 3113 | 144.06 | 144.0711 | 144.0711 | 0.011129 | 0.976 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4366 | 5.4471 | 5.4361 |  |  |  | -0.011435 | -0.000943 | -0.011861
5.89 | 10 | 5.89 |  |  | 52 | 829 | 100 | 3188 | 132.43 | 131.6134 | 131.6134 | -0.81662 | 0.969 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8786 | 5.9275 | 5.9156 |  |  |  | -0.011412 | 0.037507 | 0.025626
6.387 | 10 | 6.389 |  |  | 52 | 830 | 75 | 3151 | 143.12 | 143.1505 | 143.1505 | 0.030486 | 1.136 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.3756 | 6.3871 | 6.3743 |  |  |  | -0.01337 | -0.0019493 | -0.014748
7.871 | 10 | 7.873 | c | Sample #834 had bugs in sample. | 52 | 834 | 35 | 3047 | 135.46 | 135.5022 | 135.5022 | 0.042208 | 1.323 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.8573 | 7.8706 | 7.8548 |  |  |  | -0.015666 | -0.0023836 | -0.01815
9.12 | 10 | 9.123 |  |  | 52 | 835 | 25 | 2086 | 143.16 | 143.1706 | 143.1706 | 0.01058 | 1.62 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.1042 | 9.1218 | 9.1035 |  |  |  | -0.018825 | -0.0012393 | -0.019507
9.491 | 10 | 9.494 |  |  | 52 | 837 | 15 | 3199 | 135.79 | 134.8688 | 134.8688 | -0.92122 | 1.598 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.4747 | 9.5596 | 9.5404 |  |  |  | -0.019252 | 0.065584 | 0.046439
9.499 | 10 | 9.501 |  |  | 52 | 838 | 5 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 1.576 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.482 | 9.5712 | 9.5521 |  |  |  | -0.018995 | 0.070224 | 0.051056
4.882 | 10 | 4.882 |  |  | 52 | 824 | 175 | 3034 | 138.61 | 138.6146 | 138.6146 | 0.0046251 | 0.842 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 4.8724 | 4.882 | 4.8722 |  |  |  | -0.0096469 | -2.2996e-05 | -0.0098118
6.387 | 10 | 6.388 |  |  | 52 | 830 | 75 | 3191 | 135.01 | 134.9518 | 134.9518 | -0.05824 | 1.071 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.3752 | 6.3908 | 6.378 |  |  |  | -0.012777 | 0.0028224 | -0.0099848
9.452 | 10 | 9.454 |  |  | 52 | 838 | 5 | 3193 | 131.71 | 131.6281 | 131.6281 | -0.081924 | 1.543 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.4351 | 9.46 | 9.4411 |  |  |  | -0.018882 | 0.006027 | -0.012919
5.097 | 10 | 5.098 |  |  | 52 | 9824 | 175 | 3198 | 136.59 | 136.5025 | 136.5025 | -0.087489 | 0.866 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.0875 | 5.101 | 5.0908 |  |  |  | -0.010522 | 0.003015 | -0.0072123
5.469 | 10 | 5.47 |  |  | 54 | 861 | 137 | 2088 | 144.21 | 144.0937 | 144.0937 | -0.11628 | 0.981 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.4588 | 5.4743 | 5.4633 |  |  |  | -0.011172 | 0.0042678 | -0.0067047
5.788 | 10 | 5.789 |  |  | 54 | 862 | 100 | 3063 | 144.51 | 144.5139 | 144.5139 | 0.0038689 | 1.04 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.777 | 5.7885 | 5.7769 |  |  |  | -0.011984 | -0.00053988 | -0.012141
6.335 | 10 | 6.336 |  |  | 54 | 863 | 75 | 2074 | 141.39 | 141.3715 | 141.3715 | -0.018512 | 1.113 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.3234 | 6.3369 | 6.3242 |  |  |  | -0.012637 | 0.00090122 | -0.011797
7.064 | 10 | 7.065 |  |  | 54 | 864 | 50 | 3104 | 143.15 | 143.1468 | 143.1468 | -0.0032473 | 1.256 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0513 | 7.0656 | 7.0514 |  |  |  | -0.013731 | 0.0005868 | -0.013568
8.323 | 10 | 8.325 | c | Samples #865, 866, 867 had bugs in sample. | 54 | 865 | 35 | 3033 | 139.43 | 139.4588 | 139.4588 | 0.028758 | 1.44 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.3083 | 8.3232 | 8.3066 |  |  |  | -0.016707 | -0.0017743 | -0.018446
9.503 | 10 | 9.505 | c | Samples #865, 866, 867 had bugs in sample. | 54 | 866 | 25 | 2025 | 145.87 | 145.8939 | 145.8939 | 0.023894 | 1.72 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.4862 | 9.5036 | 9.4846 |  |  |  | -0.018819 | -0.0013628 | -0.020394
9.49 | 10 | 9.492 | c | Samples #865, 866, 867 had bugs in sample. | 54 | 867 | 15 | 3003 | 143.29 | #N/A | 143.29 | 0 | 1.687 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.4734 | 9.4924 | 9.4734 |  |  |  | -0.018587 | 0.00042193 | -0.018587
9.387 | 10 | 9.39 |  |  | 54 | 868 | 5 | 3004 | 134.98 | #N/A | 134.98 | 0 | 1.571 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.3708 | 9.3896 | 9.3708 |  |  |  | -0.019223 | -0.00041801 | -0.019223
5.825 | 10 | 5.826 |  |  | 54 | 862 | 100 | 2092 | 144.28 | 144.1993 | 144.1993 | -0.080655 | 1.045 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8143 | 5.8293 | 5.8176 |  |  |  | -0.011661 | 0.0033194 | -0.0083631
5.85 | 10 | 5.851 |  |  | 54 | 9862 | 100 | 3206 | 133.85 | 133.8167 | 133.8167 | -0.033286 | 0.973 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 5.8394 | 5.8526 | 5.8409 |  |  |  | -0.011565 | 0.0016402 | -0.010091
8.054 | 10 | 8.056 |  |  | 55 | 870 | 102 | 3196 | 132.43 | 132.3992 | 132.3992 | -0.030816 | 1.323 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.0399 | 8.0579 | 8.0418 |  |  |  | -0.016134 | 0.0019083 | -0.014234
8.344 | 10 | 8.345 |  |  | 55 | 871 | 75 | 3182 | 136.7 | 135.614 | 135.614 | -1.086 | 1.415 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.3289 | 8.4134 | 8.3966 |  |  |  | -0.016115 | 0.068433 | 0.051581
8.331 | 10 | 8.334 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and | 55 | 873 | 50 | 3151 | 143.12 | 143.1505 | 143.1505 | 0.030486 | 1.48 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.3168 | 8.3317 | 8.315 |  |  |  | -0.017239 | -0.0023482 | -0.019036
9.068 | 10 | 9.071 |  | bubbles on probe at end of titration. | 55 | 875 | 35 | 3012 | 143.88 | 143.9013 | 143.9013 | 0.021345 | 1.619 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.0524 | 9.0692 | 9.051 |  |  |  | -0.018638 | -0.0018376 | -0.02
9.28 | 10 | 9.282 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and | 55 | 876 | 25 | 3161 | 145.03 | 145.03 | 145.03 | 0 | 1.67 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.2635 | 9.2821 | 9.2635 |  |  |  | -0.018481 | 0.00010729 | -0.018481
9.456 | 10 | 9.459 |  | bubbles on probe at end of titration. | 55 | 878 | 15 | 3096 | 142.53 | 142.5586 | 142.5586 | 0.028595 | 1.672 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.4396 | 9.4567 | 9.4377 |  |  |  | -0.019352 | -0.0023344 | -0.021272
9.645 | 10 | 9.648 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and | 55 | 879 | 5 | 3160 | 144.9 | 144.9088 | 144.9088 | 0.0087513 | 1.734 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.6286 | 9.6473 | 9.628 |  |  |  | -0.019411 | -0.00068239 | -0.020001
8.32 | 10 | 8.322 |  | bubbles on probe at end of titration. | 55 | 873 | 50 | 3047 | 135.46 | 135.5022 | 135.5022 | 0.042208 | 1.398 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.3048 | 8.3189 | 8.3022 |  |  |  | -0.017164 | -0.0031263 | -0.01979
9.257 | 10 | 9.259 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and | 55 | 876 | 25 | 3028 | 141.76 | 141.7907 | 141.7907 | 0.030674 | 1.628 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.241 | 9.2575 | 9.2389 |  |  |  | -0.018045 | -0.0015331 | -0.020072
9.506 | 10 | 9.497 |  | bubbles on probe at end of titration. | 55 | 878 | 15 | 3174 | 141.29 | 141.3182 | 141.3182 | 0.028153 | 1.664 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.4779 | 9.495 | 9.476 |  |  |  | -0.019053 | -0.0019535 | -0.020968
8.525 | 10 | 8.527 |  |  | 55 | 9873 | 50 | 3034 | 138.61 | 138.6146 | 138.6146 | 0.0046251 | 1.466 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.5097 | 8.5265 | 8.5094 |  |  |  | -0.01728 | -0.00049011 | -0.017568
8.518 | 10 | 8.52 |  |  | 58 | 904 | 104 | 2017 | 144.62 | 144.8189 | 144.8189 | 0.19891 | 1.529 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.5029 | 8.5081 | 8.491 |  |  |  | -0.01712 | -0.011923 | -0.028963
8.457 | 10 | 8.458 | c | Sample #905 had small bubble on probe at end of titration. | 58 | 905 | 75 | 3048 | 137.92 | #N/A | 137.92 | 0 | 1.447 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.4416 | 8.4585 | 8.4416 |  |  |  | -0.016396 | 0.00054652 | -0.016396
8.609 | 10 | 8.611 |  |  | 58 | 906 | 50 | 3157 | 133.84 | 132.8731 | 132.8731 | -0.96686 | 1.429 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.5941 | 8.675 | 8.6576 |  |  |  | -0.016877 | 0.06399 | 0.046614
9.107 | 10 | 9.11 |  |  | 58 | 907 | 35 | 3193 | 131.71 | 131.6281 | 131.6281 | -0.081924 | 1.487 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.0913 | 9.1153 | 9.0971 |  |  |  | -0.018677 | 0.0053253 | -0.012932
9.08 | 10 | 9.083 |  |  | 58 | 908 | 25 | 2086 | 143.16 | 143.1706 | 143.1706 | 0.01058 | 1.613 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.0647 | 9.0822 | 9.064 |  |  |  | -0.018313 | -0.00080402 | -0.018992
9.234 | 10 | 9.237 |  |  | 58 | 909 | 15 | 3188 | 132.43 | 131.6134 | 131.6134 | -0.81662 | 1.516 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 9.2182 | 9.2949 | 9.2763 |  |  |  | -0.018809 | 0.057886 | 0.03927
9.01 | 10 | 9.012 |  |  | 58 | 910 | 5 | 3180 | 145.95 | 145.9067 | 145.9067 | -0.043342 | 1.632 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.9941 | 9.0149 | 8.9968 |  |  |  | -0.017897 | 0.0028656 | -0.015188
8.49 | 10 | 8.492 |  |  | 58 | 905 | 75 | 3199 | 135.79 | 134.8688 | 134.8688 | -0.92122 | 1.43 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.4748 | 8.5507 | 8.5336 |  |  |  | -0.017185 | 0.058701 | 0.041574
8.672 | 10 | 8.674 |  |  | 58 | 906 | 50 | 3191 | 135.01 | 134.9518 | 134.9518 | -0.05824 | 1.452 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.6562 | 8.6774 | 8.66 |  |  |  | -0.017775 | 0.003397 | -0.013983
8.467 | 10 | 8.469 |  |  | 58 | 9905 | 75 | 3113 | 144.06 | 144.0711 | 144.0711 | 0.011129 | 1.514 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.4523 | 8.4686 | 8.4517 |  |  |  | -0.016684 | -0.00038509 | -0.017346
8.254 | 10 | 8.256 |  |  | 59 | 912 | 100 | 3003 | 143.29 | #N/A | 143.29 | 0 | 1.468 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.2391 | 8.2557 | 8.2391 |  |  |  | -0.016878 | -0.00034302 | -0.016878
8.899 | 10 | 8.902 |  |  | 59 | 915 | 35 | 3063 | 144.51 | 144.5139 | 144.5139 | 0.0038689 | 1.596 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.8838 | 8.9014 | 8.8836 |  |  |  | -0.018175 | -0.00058992 | -0.018417
8.253 | 10 | 8.255 |  |  | 59 | 912 | 100 | 3033 | 139.43 | 139.4588 | 139.4588 | 0.028758 | 1.428 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.2388 | 8.2536 | 8.237 |  |  |  | -0.016239 | -0.0014311 | -0.017963
8.902 | 10 | 8.904 |  |  | 59 | 915 | 35 | 3198 | 136.59 | 136.5025 | 136.5025 | -0.087489 | 1.508 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.8859 | 8.9096 | 8.8917 |  |  |  | -0.018064 | 0.0055602 | -0.012284
8.913 | 10 | 8.916 |  |  | 59 | 9915 | 35 | 3004 | 134.98 | #N/A | 134.98 | 0 | 1.492 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 8.8977 | 8.9156 | 8.8977 |  |  |  | -0.018293 | -0.00043646 | -0.018293
6.917 | 10 | 6.918 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 941 | 210 | 3051 | 140.62 | 140.6477 | 140.6477 | 0.027665 | 1.208 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 6.9042 | 6.9167 | 6.9028 |  |  |  | -0.013776 | -0.0012957 | -0.015154
7.091 | 10 | 7.092 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 942 | 200 | 3204 | 132.83 | 131.8575 | 131.8575 | -0.97252 | 1.169 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0779 | 7.1453 | 7.131 |  |  |  | -0.014057 | 0.053267 | 0.03895
7.093 | 10 | 7.095 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 943 | 175 | 3177 | 143.13 | 143.1669 | 143.1669 | 0.036864 | 1.261 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0805 | 7.0928 | 7.0786 |  |  |  | -0.014519 | -0.0021585 | -0.016368
7.709 | 10 | 7.71 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 949 | 35 | 3006 | 143.66 | 143.7048 | 143.7048 | 0.044798 | 1.375 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.6948 | 7.7078 | 7.6924 |  |  |  | -0.01518 | -0.0021732 | -0.017613
7.09 | 10 | 7.091 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 943 | 175 | 3097 | 143.2 | 143.1879 | 143.1879 | -0.012078 | 1.261 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.077 | 7.0918 | 7.0776 |  |  |  | -0.014029 | 0.00078359 | -0.013424
7.702 | 10 | 7.703 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 949 | 35 | 3171 | 143.06 | 143.0623 | 143.0623 | 0.0023427 | 1.368 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.688 | 7.7033 | 7.6879 |  |  |  | -0.014967 | 0.00033633 | -0.015094
7.107 | 10 | 7.109 | c | Colorimeter drifting exceeding 100% for entire file. | 62 | 9943 | 175 | 2063 | 142.85 | 142.864 | 142.864 | 0.014046 | 1.261 |  |  | 0.004 | 0.715 | 9.98 | 0.010134 |  | 7.0946 | 7.1081 | 7.0938 |  |  |  | -0.014444 | -0.00091091 | -0.015151
[26 empty] | MAX | -0.0050371 | 0.077964 | 0.057232
[26 empty] | MIN | -0.02203 | -0.027588 | -0.044451

## Reformatting - step1
2 March 2017: Recalculated oxygen being reformated for merging into chemistry data.
Replicate | ID Replicate | Cast | Sample_ Number | Pressure | Oxy (mL/L) BOTH KIO3 STD vol and flask vol changes | Flag, Nov 2007 | Comment, Nov 2007 | Secondary | Flag, Nov 2007 | Comment, Nov 2007 | Triplicate | Flag, Nov 2007 | Comment, Nov 2007
|  | 3 | 16 | 75 | 7.8931 | d | Colorimeter probe problems with drift for entire day's run | 7.9063 | d | Colorimeter probe problems with drift for entire day's run | 7.8786 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 3 | 16 | 75 | 7.9063 | d | Colorimeter probe problems with drift for entire day's run
3 | 0 | 3 | 16 | 75 | 7.8786 | d | Colorimeter probe problems with drift for entire day's run
4 | 0 | 3 | 16 | 75 | 7.8311 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 3 | 18 | 50 | 8.4986 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 3 | 18 | 50 | 8.4104 | d | Colorimeter probe problems with drift for entire day's run | 8.4104 | d | Colorimeter probe problems with drift for entire day's run | 8.3073 | d | Colorimeter probe problems with drift for entire day's run
3 | 0 | 3 | 18 | 50 | 8.3073 | d | Colorimeter probe problems with drift for entire day's run
4 | 0 | 3 | 18 | 50 | 8.0019 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 3 | 20 | 25 | 8.7134 | d | Colorimeter probe problems with drift for entire day's run | 8.3627 | d | Colorimeter probe problems with drift for entire day's run | 8.498 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 3 | 20 | 25 | 8.3627 | d | Colorimeter probe problems with drift for entire day's run
3 | 0 | 3 | 20 | 25 | 8.498 | d | Colorimeter probe problems with drift for entire day's run
4 | 0 | 3 | 20 | 25 | 8.3087 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 3 | 22 | 5 | 7.3353 | d | Colorimeter probe problems with drift for entire day's run | 7.8255 | d | Colorimeter probe problems with drift for entire day's run | 7.032 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 3 | 22 | 5 | 7.8255 | d | Colorimeter probe problems with drift for entire day's run
3 | 0 | 3 | 22 | 5 | 7.032 | d | Colorimeter probe problems with drift for entire day's run
4 | 0 | 3 | 22 | 5 | 7.0015 | d | Colorimeter probe problems with drift for entire day's run
| 23 | 5 | 45 | 214 | 7.5716 | d | Colorimeter probe problems with drift for entire day's run
| 3 | 5 | 48 | 200 | 7.6139 | d | Colorimeter probe problems with drift for entire day's run | 7.8188 | d | Colorimeter probe problems with drift for entire day's run | 7.3791 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 5 | 48 | 200 | 7.8188 | d | Colorimeter probe problems with drift for entire day's run
3 | 0 | 5 | 48 | 200 | 7.3791 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 5 | 49 | 150 | 7.6711 | d | Colorimeter probe problems with drift for entire day's run | 7.9293 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 5 | 49 | 150 | 7.9293 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 5 | 50 | 125 | 7.549 | d | Colorimeter probe problems with drift for entire day's run | 8.0135 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 5 | 50 | 125 | 8.0135 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 5 | 51 | 100 | 8.2914 | d | Colorimeter probe problems with drift for entire day's run
| 3 | 5 | 54 | 75 | 8.1226 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 5 | 55 | 50 | 7.9362 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 5 | 56 | 35 | 7.7574 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 5 | 57 | 25 | 8.4056 | d | Colorimeter probe problems with drift for entire day's run
| 4 | 5 | 61 | 5 | 10.0441 | d | Colorimeter probe problems with drift for entire day's run
| 17 | 7 | 78 | 2000 | 6.7089 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 7 | 79 | 1500 | 7.0498 | d | Colorimeter probe problems with drift for entire day's run
| 22 | 8 | 101 | 3459 | 7.0208 | d | Colorimeter probe problems with drift for entire day's run | 6.9979 | d | Colorimeter probe problems with drift for entire day's run | 7.0285 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 8 | 101 | 3459 | 6.9979 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 102 | 3000 | 5.8909 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 103 | 2500 | 6.3467 | d | Colorimeter probe problems with drift for entire day's run | 6.6389 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 8 | 103 | 2500 | 6.6389 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 8 | 105 | 1000 | 6.8913 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 8 | 107 | 750 | 7.0676 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 8 | 109 | 500 | 7.0934 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 110 | 400 | 6.9364 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 111 | 300 | 6.9819 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 112 | 250 | 6.9825 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 8 | 114 | 200 | 6.9575 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 115 | 150 | 7.1029 | d | Colorimeter probe problems with drift for entire day's run | 7.0874 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 8 | 115 | 150 | 7.0874 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 116 | 125 | 7.3031 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 117 | 100 | 7.39 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 8 | 119 | 75 | 7.3522 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 120 | 50 | 7.3634 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 121 | 35 | 6.8814 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 122 | 25 | 7.6298 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 8 | 123 | 5 | 7.5377 | d | Colorimeter probe problems with drift for entire day's run
| 3 | 10 | 126 | 1000 | 7.0099 | d | Colorimeter probe problems with drift for entire day's run | 6.8602 | d | Colorimeter probe problems with drift for entire day's run | 6.89 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 10 | 126 | 1000 | 6.8602 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 10 | 128 | 750 | 6.9344 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 10 | 130 | 500 | 6.8956 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 10 | 132 | 350 | 7.059 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 133 | 300 | 7.0262 | d | Colorimeter probe problems with drift for entire day's run | 6.8631 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 10 | 133 | 300 | 6.8631 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 134 | 250 | 6.8442 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 10 | 136 | 200 | 6.8452 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 137 | 175 | 6.8681 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 138 | 150 | 6.9458 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 139 | 125 | 6.8948 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 140 | 100 | 7.0306 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 141 | 75 | 7.149 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 10 | 143 | 50 | 7.4393 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 144 | 35 | 7.5823 | d | Colorimeter probe problems with drift for entire day's run | 7.4344 | d | Colorimeter probe problems with drift for entire day's run
2 | 0 | 10 | 144 | 35 | 7.4344 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 10 | 146 | 15 | 8.4753 | d | Colorimeter probe problems with drift for entire day's run
| 1 | 10 | 147 | 5 | 8.3686 | d | Colorimeter probe problems with drift for entire day's run
| 2 | 11 | 149 | 1000 | 6.615 |  |  | 6.5855 |  |  | 7.9325
2 | 0 | 11 | 149 | 1000 | 6.5855
| 1 | 11 | 150 | 750 | 6.8039
| 1 | 11 | 151 | 500 | 6.8892
2 | 0 | 11 | 151 | 500 | 6.9018
| 1 | 11 | 152 | 350 | 7.1917
| 1 | 11 | 153 | 250 | 7.3968
| 1 | 11 | 154 | 225 | 7.466 | c | samples # 154,155,158 bubbles on probe at end of titration.
| 1 | 11 | 155 | 200 | 7.5348 | c | samples # 154,155,158 bubbles on probe at end of titration. | 7.4879
2 | 0 | 11 | 155 | 200 | 7.4879
| 2 | 11 | 157 | 150 | 7.5468
| 1 | 11 | 158 | 125 | 7.6004 |  | samples # 154,155,158 bubbles on probe at end of titration.
| 1 | 11 | 159 | 100 | 7.6699 | c
| 1 | 11 | 160 | 75 | 7.6502
| 1 | 11 | 161 | 50 | 7.6942
| 1 | 11 | 162 | 35 | 7.8728
| 1 | 11 | 163 | 25 | 8.2014
| 1 | 11 | 164 | 15 | 8.1603
| 1 | 11 | 165 | 5 | 8.0147
| 27 | 13 | 192 | 585 | 5.2193 |  |  | 5.5606 |  |  | 6.308 | c | Triplicate sample # 9192 had bubble on probe.
2 | 0 | 13 | 192 | 585 | 5.5606
| 4 | 13 | 196 | 500 | 5.4787
2 | 0 | 13 | 196 | 500 | 5.4857
| 1 | 13 | 197 | 350 | 6.1152
| 2 | 13 | 199 | 250 | 6.4241
| 1 | 13 | 200 | 225 | 6.6948
| 1 | 13 | 201 | 200 | 6.1567
| 1 | 13 | 202 | 175 | 6.1217
| 1 | 13 | 203 | 150 | 6.5695
| 1 | 13 | 204 | 125 | 7.3553
| 1 | 13 | 205 | 100 | 7.5872 |  |  | 7.5611
2 | 0 | 13 | 205 | 100 | 7.5611
| 2 | 13 | 207 | 75 | 8.0327
| 1 | 13 | 208 | 50 | 8.2463
| 1 | 13 | 209 | 35 | 8.4853
| 1 | 13 | 210 | 25 | 8.9089
| 1 | 13 | 211 | 15 | 9.1381
| 2 | 13 | 213 | 5 | 9.189
| 21 | 15 | 234 | 247 | 7.1053 |  |  | 7.1101 |  |  | 7.1443 | c
2 | 0 | 15 | 234 | 247 | 7.1101
| 1 | 15 | 235 | 225 | 7.0826
| 1 | 15 | 236 | 200 | 7.0772
| 1 | 15 | 237 | 175 | 7.2455
| 1 | 15 | 238 | 150 | 7.3344 |  |  | 7.3721
2 | 0 | 15 | 238 | 150 | 7.3721
| 1 | 15 | 239 | 125 | 7.3968
| 1 | 15 | 240 | 100 | 7.4648
| 1 | 15 | 241 | 75 | 7.7085
| 2 | 15 | 243 | 50 | 7.9102 |  | Sample #243 had bubble on probe, sample #246,247 had bugs in sample.
| 1 | 15 | 244 | 35 | 8.3959 | c |  | 8.3596
2 | 0 | 15 | 244 | 35 | 8.3596
| 1 | 15 | 245 | 25 | 8.7255
| 1 | 15 | 246 | 15 | 8.603 | c | Sample #243 had bubble on probe, sample #246,247 had bugs in sample.
| 1 | 15 | 247 | 5 | 8.5794 | c | Sample #243 had bubble on probe, sample #246,247 had bugs in sample.
| 3 | 16 | 250 | 250 | 6.3593 |  |  | 6.3539 |  |  | 6.3444
2 | 0 | 16 | 250 | 250 | 6.3539
| 1 | 16 | 251 | 225 | 6.4727
| 1 | 16 | 252 | 200 | 6.7323 |  |  | 6.4882
2 | 0 | 16 | 252 | 200 | 6.4882
| 1 | 16 | 253 | 175 | 6.8189
| 1 | 16 | 254 | 150 | 6.8732
| 1 | 16 | 255 | 125 | 7.0662
| 1 | 16 | 256 | 100 | 7.4943
| 1 | 16 | 257 | 75 | 7.1839 | c | Sample #257 had leaky burette.
| 1 | 16 | 258 | 50 | 7.7271 |  |  | 7.7261
2 | 0 | 16 | 258 | 50 | 7.7261
| 1 | 16 | 259 | 35 | 8.2979
| 1 | 16 | 260 | 25 | 8.8304
| 1 | 16 | 261 | 15 | 8.547
| 1 | 16 | 262 | 5 | 8.4515
| 1 | 17 | 263 | 1000 | 4.3929 |  |  | 4.3926 |  |  | 4.2411
2 | 0 | 17 | 263 | 1000 | 4.3926
| 1 | 17 | 264 | 750 | 4.2236
| 1 | 17 | 265 | 500 | 5.8182 |  |  | 5.7546
2 | 0 | 17 | 265 | 500 | 5.7546
| 1 | 17 | 266 | 350 | 6.1751
| 1 | 17 | 267 | 250 | 6.0143
| 1 | 17 | 268 | 225 | 3.8896
| 1 | 17 | 269 | 200 | 4.1959
| 1 | 17 | 270 | 175 | 4.3651
| 1 | 17 | 271 | 150 | 4.6396
| 1 | 17 | 272 | 125 | 6.669
| 1 | 17 | 273 | 100 | 3.7918
| 1 | 17 | 274 | 75 | 6.9073
| 1 | 17 | 275 | 50 | 5.9832
| 1 | 17 | 276 | 35 | 4.4567
| 1 | 17 | 277 | 25 | 8.0818
| 1 | 17 | 278 | 15 | 4.5906
| 1 | 17 | 279 | 5 | 8.9241 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination? | 5.8849 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination?
2 | 0 | 17 | 279 | 5 | 5.8849 | c | Sample #279 and duplicate deviate by 3.0ml/L DO value... contamination?
| 1 | 18 | 280 | 1000 | 4.4671 |  |  | 2.567 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 4.4607 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks
2 | 0 | 18 | 280 | 1000 | 2.567 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 281 | 750 | 3.8842
| 1 | 18 | 282 | 500 | 3.4017 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 283 | 350 | 4.4488 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks
| 1 | 18 | 284 | 250 | 3.207 |  | before analysis. Duplicate sample had less precipitate in flask than others.
| 1 | 18 | 285 | 225 | 3.1694
| 1 | 18 | 286 | 200 | 3.2794
| 1 | 18 | 287 | 175 | 3.6268 |  |  | 4.2517 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
2 | 0 | 18 | 287 | 175 | 4.2517 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 288 | 150 | 3.8595
| 1 | 18 | 289 | 125 | 4.4755
| 1 | 18 | 290 | 100 | 4.6779
| 1 | 18 | 291 | 75 | 4.2659 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 292 | 50 | 3.9136 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 293 | 35 | 5.4431 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 294 | 25 | 7.0911 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 295 | 15 | 4.3938 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle. | 4.9675 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
2 | 0 | 18 | 295 | 15 | 4.9675 | c | Sample #282 niskin bottle closed without wait. Sample #293 had leaky niskin bottle.
| 1 | 18 | 296 | 5 | 6.9675 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks
| 26 | 21 | 322 | 1000 | 4.5747 |  | before analysis. Duplicate sample had less precipitate in flask than others. | 4.5042 |  |  | 4.6675
2 | 0 | 21 | 322 | 1000 | 4.5042
| 2 | 21 | 324 | 750 | 5.0423
| 2 | 21 | 326 | 500 | 5.4988 |  |  | 5.4555
2 | 0 | 21 | 326 | 500 | 5.4555
| 1 | 21 | 327 | 350 | 5.5278
| 2 | 21 | 329 | 250 | 5.7271
| 1 | 21 | 330 | 225 | 5.8858
| 1 | 21 | 331 | 200 | 6.1502
| 1 | 21 | 332 | 175 | 6.5115
| 1 | 21 | 333 | 150 | 7.4268
| 1 | 21 | 334 | 125 | 7.5691
| 1 | 21 | 335 | 100 | 7.676
| 2 | 21 | 337 | 75 | 7.5585
| 1 | 21 | 338 | 50 | 7.7105
| 1 | 21 | 339 | 35 | 8.1046 | c | Sample #339 had small bubble on probe at end of titration.
| 1 | 21 | 340 | 25 | 8.3269 |  |  | 8.5609
2 | 0 | 21 | 340 | 15 | 8.5609
| 1 | 21 | 341 | 15 | 8.6545
| 2 | 21 | 343 | 5 | 8.8026
| 1 | 22 | 344 | 1000 | 4.561 |  |  | 3.7262 |  |  | 4.5114
2 | 0 | 22 | 344 | 1000 | 3.7262
| 1 | 22 | 345 | 750 | 3.5993
| 1 | 22 | 346 | 500 | 3.7768
| 1 | 22 | 347 | 350 | 2.9788
| 1 | 22 | 348 | 250 | 3.6549 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 349 | 225 | 3.2382 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 350 | 200 | 3.5842 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 351 | 175 | 3.8766 |  |  | 3.7527
2 | 0 | 22 | 351 | 175 | 3.7527
| 1 | 22 | 352 | 150 | 3.3595
| 1 | 22 | 353 | 125 | 4.0074 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 354 | 100 | 3.247 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 355 | 75 | 3.7549
| 1 | 22 | 356 | 50 | 3.3877 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 357 | 35 | 3.5134 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 358 | 25 | 4.1873
| 1 | 22 | 359 | 15 | 4.0574 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis.
| 1 | 22 | 360 | 5 | 4.2298 | c | Samples #348, 349, 350, 353, 354, 356, 357, 359 had tiny bubbles in flask at time of analysis. | 4.6245
2 | 0 | 22 | 360 | 5 | 4.6245
| 1 | 23 | 361 | 1000 | 4.5834 |  |  | 4.635
2 | 0 | 23 | 361 | 1000 | 4.635
| 1 | 23 | 362 | 750 | 5.0793
| 1 | 23 | 363 | 500 | 5.4208
| 2 | 23 | 365 | 350 | 5.4648 |  |  | 5.5158
2 | 0 | 23 | 365 | 350 | 5.5158
| 2 | 23 | 367 | 250 | 5.9533
| 1 | 23 | 368 | 225 | 6.4066
| 1 | 23 | 369 | 200 | 6.9664
| 1 | 23 | 370 | 175 | 7.4389
| 1 | 23 | 371 | 150 | 6.8765
| 1 | 23 | 372 | 125 | 7.3478 |  |  | 7.2413
2 | 0 | 23 | 372 | 125 | 7.2413
| 1 | 23 | 373 | 100 | 7.3087
| 1 | 23 | 374 | 75 | 7.4137
| 1 | 23 | 375 | 50 | 7.6683
| 1 | 23 | 376 | 35 | 8.3616
| 1 | 23 | 377 | 25 | 8.4025 | c | Sample #377 had leaky stopper... maybe DMQ or air in sample.
| 1 | 23 | 378 | 15 | 8.7143
| 1 | 23 | 379 | 5 | 9.2958 |  |  | 9.3086
2 | 0 | 23 | 379 | 5 | 9.3086
| 20 | 25 | 399 | 150 | 7.0588
| 1 | 25 | 400 | 125 | 5.208 |  |  | 7.1152 |  |  | 7.1492
2 | 0 | 25 | 400 | 125 | 7.1152
| 1 | 25 | 401 | 100 | 7.3572
| 1 | 25 | 402 | 75 | 7.7639
| 1 | 25 | 403 | 50 | 8.7745 |  |  | 7.914
2 | 0 | 25 | 403 | 50 | 7.914
| 1 | 25 | 404 | 35 | 8.5229
| 1 | 25 | 405 | 25 | 9.8174
| 1 | 25 | 406 | 15 | 9.7866
| 1 | 25 | 407 | 5 | 9.5148
| 1 | 26 | 408 | 1000 | 4.4915 |  |  | 4.4777 |  |  | 4.6351
2 | 0 | 26 | 408 | 1000 | 4.4777
| 2 | 26 | 410 | 750 | 4.8946
| 2 | 26 | 412 | 500 | 5.4745
| 2 | 26 | 414 | 350 | 5.6014
| 1 | 26 | 415 | 250 | 5.749
| 2 | 26 | 417 | 225 | 5.827
| 1 | 26 | 418 | 200 | 6.0176 |  |  | 6.0282
2 | 0 | 26 | 418 | 200 | 6.0282
| 1 | 26 | 419 | 175 | 6.3666
| 1 | 26 | 420 | 150 | 7.2438
| 1 | 26 | 421 | 125 | 7.6017 |  |  | 7.5318
2 | 0 | 26 | 421 | 125 | 7.5318
| 1 | 26 | 422 | 100 | 7.7058
| 1 | 26 | 423 | 75 | 7.3241
| 2 | 26 | 425 | 50 | 7.6515
| 1 | 26 | 426 | 35 | 7.8734
| 1 | 26 | 427 | 25 | 8.144
| 2 | 26 | 429 | 5 | 9.2249
| 2 | 27 | 431 | 1000 | 4.4824 |  |  | 4.5133 |  |  | 4.4901
2 | 0 | 27 | 431 | 5 | 4.5133
| 2 | 27 | 433 | 750 | 4.9744
| 2 | 27 | 435 | 500 | 5.3891
| 2 | 27 | 437 | 350 | 5.5277
| 1 | 27 | 438 | 250 | 5.6202
| 2 | 27 | 440 | 225 | 5.7382 |  |  | 5.7616
2 | 0 | 27 | 440 | 225 | 5.7616
| 1 | 27 | 441 | 200 | 6.0594 | c | Sample #441 had bugs in sample.
| 1 | 27 | 442 | 175 | 6.7129
| 1 | 27 | 443 | 150 | 6.8389
| 1 | 27 | 444 | 125 | 7.1927
| 1 | 27 | 445 | 100 | 7.3768
| 1 | 27 | 446 | 75 | 7.597
| 2 | 27 | 448 | 50 | 7.8707
| 1 | 27 | 449 | 35 | 8.1195
| 1 | 27 | 450 | 25 | 8.4352
| 1 | 27 | 451 | 15 | 8.8039
| 1 | 27 | 452 | 5 | 9.1751
| 2 | 29 | 454 | 150 | 6.13
| 2 | 29 | 456 | 125 | 6.8363
| 1 | 29 | 457 | 100 | 7.0703
| 1 | 29 | 458 | 75 | 7.3616
| 1 | 29 | 459 | 50 | 7.6789
| 1 | 29 | 460 | 30 | 8.0942
| 1 | 29 | 461 | 20 | 9.0054
| 1 | 29 | 462 | 10 | 9.8346 |  |  | 9.8723
2 | 0 | 29 | 462 | 10 | 9.8723
| 1 | 29 | 463 | 5 | 9.7958
| 1 | 30 | 464 | 2350 | 2.7723
| 2 | 30 | 466 | 2250 | 2.7752 |  |  | 2.8181 |  |  | 2.815
2 | 0 | 30 | 466 | 2250 | 2.8181
| 1 | 30 | 467 | 2000 | 2.8209
| 1 | 30 | 468 | 1750 | 2.9579
| 1 | 30 | 469 | 1500 | 3.2675
| 1 | 30 | 470 | 1300 | 3.6439
| 1 | 30 | 471 | 1200 | 3.9478
| 1 | 30 | 472 | 1100 | 4.2221
| 1 | 30 | 473 | 1000 | 4.4675 |  |  | 4.5223
2 | 0 | 30 | 473 | 1000 | 4.5223
| 1 | 30 | 474 | 900 | 4.6487
| 1 | 30 | 475 | 800 | 4.8351
| 1 | 30 | 476 | 700 | 5.0188
| 1 | 30 | 477 | 600 | 5.1707
| 1 | 30 | 478 | 500 | 5.335
| 1 | 30 | 479 | 400 | 5.3764
| 1 | 30 | 480 | 350 | 5.3884
| 1 | 30 | 481 | 300 | 5.4667
| 1 | 30 | 482 | 275 | 5.4322
| 1 | 30 | 483 | 250 | 5.444
| 1 | 30 | 484 | 225 | 5.5027
| 2 | 30 | 486 | 200 | 5.5024
| 2 | 31 | 488 | 2045 | 2.742
| 1 | 31 | 489 | 2000 | 2.7847 |  |  | 2.7065 |  |  | 2.7541
2 | 0 | 31 | 489 | 2000 | 2.7065
| 1 | 31 | 490 | 1750 | 2.8083
| 1 | 31 | 491 | 1500 | 3.1473
| 1 | 31 | 492 | 1250 | 3.7122
| 1 | 31 | 493 | 1000 | 4.4166
| 1 | 31 | 494 | 750 | 4.8983
| 1 | 31 | 495 | 500 | 5.4284
| 1 | 31 | 496 | 400 | 5.4711
| 1 | 31 | 497 | 350 | 5.5657
| 1 | 31 | 498 | 300 | 5.5543 |  |  | 5.5893
2 | 0 | 31 | 498 | 300 | 5.5893
| 1 | 31 | 499 | 250 | 5.6873
| 1 | 31 | 500 | 225 | 5.8351
| 1 | 31 | 501 | 200 | 6.2986
| 1 | 31 | 502 | 175 | 7.0413
| 1 | 31 | 503 | 150 | 6.9724
| 1 | 31 | 504 | 125 | 7.135
| 1 | 31 | 505 | 100 | 7.1014 |  |  | 7.0586
2 | 0 | 31 | 505 | 100 | 7.0586
| 1 | 31 | 506 | 75 | 7.479
| 1 | 31 | 507 | 50 | 7.8382
| 1 | 31 | 508 | 35 | 8.157
| 1 | 31 | 509 | 20 | 9.943
| 1 | 31 | 510 | 5 | 9.4548
| 1 | 32 | 511 | 1473 | 2.8876
| 1 | 32 | 512 | 1250 | 3.5723 |  |  | 3.5195 |  |  | 3.589
2 | 0 | 32 | 512 | 1250 | 3.5195
| 1 | 32 | 513 | 1000 | 4.438
| 1 | 32 | 514 | 750 | 4.7879
| 1 | 32 | 515 | 500 | 5.4522
| 1 | 32 | 516 | 400 | 5.554 |  |  | 5.6414
2 | 0 | 32 | 516 | 400 | 5.6414
| 1 | 32 | 517 | 350 | 5.6425
| 1 | 32 | 518 | 300 | 5.602
| 1 | 32 | 519 | 250 | 5.7373
| 1 | 32 | 520 | 225 | 5.8566
| 1 | 32 | 521 | 200 | 6.2239 |  |  | 6.3107
2 | 0 | 32 | 521 | 200 | 6.3107
| 1 | 32 | 522 | 175 | 6.8949
| 1 | 32 | 523 | 150 | 7.0707
| 1 | 32 | 524 | 125 | 7.2594
| 1 | 32 | 525 | 100 | 7.3295
| 1 | 32 | 526 | 75 | 8.0917
| 1 | 32 | 527 | 50 | 8.2245
| 1 | 32 | 528 | 35 | 9.5916
| 1 | 32 | 529 | 20 | 10.4628
| 1 | 32 | 530 | 5 | 9.3781
| 1 | 33 | 531 | 1042 | 3.3899
| 1 | 33 | 532 | 1000 | 3.5299 |  |  | 3.4445 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others. | 3.4473 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others.
2 | 0 | 33 | 532 | 1000 | 3.4445
| 3 | 33 | 535 | 400 | 5.7195 |  |  | 5.7832 |  |  | 5.7361
2 | 0 | 33 | 535 | 400 | 5.7832
| 23 | 34 | 558 | 200 | 7.3097 |  |  | 7.204 |  |  | 7.2579
2 | 0 | 34 | 558 | 200 | 7.204
| 4 | 34 | 562 | 100 | 7.6632 |  |  | 7.6829 |  |  | 7.6268
2 | 0 | 34 | 562 | 100 | 7.6829
| 8 | 35 | 570 | 500 | 5.5969 |  |  | 5.7509 |  |  | 5.5988
2 | 0 | 35 | 570 | 500 | 5.7509
| 1 | 36 | 571 | 150 | 7.4397 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination?
| 1 | 36 | 572 | 150 | 7.4775 |  |  | 7.5198
2 | 0 | 36 | 572 | 150 | 7.5198
| 1 | 36 | 573 | 125 | 7.4896
| 1 | 36 | 574 | 100 | 7.6337
| 1 | 36 | 575 | 95 | 7.6435
| 1 | 36 | 576 | 75 | 7.81
| 2 | 36 | 578 | 50 | 8.2811 |  |  | 8.3102
2 | 0 | 36 | 578 | 50 | 8.3102
| 1 | 36 | 579 | 30 | 8.9475
| 1 | 36 | 580 | 21 | 9.3382
| 1 | 36 | 581 | 15 | 9.5835
| 1 | 36 | 582 | 10 | 7.4771 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination? | 9.7679 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate by 3.0ml/L DO value... contamination?
2 | 0 | 36 | 582 | 10 | 9.7679 | c | Sample #571 had 2 volumes added together. Sample #582 and duplicate deviate
| 30 | 38 | 612 | 250 | 5.7585 |  |  | 5.7314 |  |  | 5.773
2 | 0 | 38 | 612 | 250 | 5.7314
| 7 | 38 | 619 | 75 | 7.112 |  |  | 7.1 |  |  | 7.0961
2 | 0 | 38 | 619 | 75 | 7.1
| 6 | 39 | 625 | 763 | 4.6537 |  |  | 4.637 |  |  | 4.7111
2 | 0 | 39 | 625 | 763 | 4.637
| 3 | 39 | 628 | 500 | 5.4779
| 3 | 39 | 631 | 250 | 5.6895
| 2 | 39 | 633 | 200 | 6.1256
| 1 | 39 | 634 | 150 | 6.7897 |  |  | 6.7431
2 | 0 | 39 | 634 | 150 | 6.7431
| 1 | 39 | 635 | 100 | 7.4385
| 3 | 39 | 638 | 75 | 7.7339
| 4 | 39 | 642 | 50 | 9.1359 |  |  | 9.0787
2 | 0 | 39 | 642 | 50 | 9.0787
| 3 | 39 | 645 | 25 | 9.2035
| 3 | 39 | 648 | 5 | 8.7479
| 4 | 40 | 652 | 150 | 7.2933 |  |  | 7.251 |  |  | 7.3109
2 | 0 | 40 | 652 | 150 | 7.251
| 3 | 40 | 655 | 50 | 7.7583 |  |  | 7.7504 |  |  | 7.7643
2 | 0 | 40 | 655 | 50 | 7.7504
| 19 | 41 | 674 | 300 | 6.0745 |  |  | 6.1213 |  |  | 6.1021
2 | 0 | 41 | 674 | 300 | 6.1213
| 3 | 41 | 677 | 200 | 6.1531
| 1 | 41 | 678 | 175 | 6.554 |  |  | 6.5854
2 | 0 | 41 | 678 | 175 | 6.5854
| 1 | 41 | 679 | 150 | 6.9148
| 1 | 41 | 680 | 125 | 7.2159
| 1 | 41 | 681 | 100 | 7.1257 |  |  | 7.2284
2 | 0 | 41 | 681 | 100 | 7.2284
| 1 | 41 | 682 | 75 | 7.0069
| 1 | 41 | 683 | 50 | 7.2849
| 1 | 41 | 684 | 35 | 9.6663
| 1 | 41 | 685 | 25 | 10.6197
| 1 | 41 | 686 | 15 | 10.342
| 7 | 42 | 693 | 175 | 6.7899
| 1 | 42 | 694 | 150 | 6.8125 |  |  | 6.8008 |  |  | 6.8366
2 | 0 | 42 | 694 | 150 | 6.8008
| 1 | 42 | 695 | 125 | 6.8921
| 1 | 42 | 696 | 100 | 6.4839 |  |  | 6.4655
2 | 0 | 42 | 696 | 100 | 6.4655
| 1 | 42 | 697 | 75 | 6.6386
| 1 | 42 | 698 | 50 | 7.0143
| 1 | 42 | 699 | 35 | 8.9902
| 1 | 42 | 700 | 25 | 10.7758
| 1 | 42 | 701 | 15 | 10.2856
| 27 | 45 | 728 | 435 | 5.396
| 1 | 45 | 729 | 350 | 4.9805 |  |  | 4.9653
2 | 0 | 45 | 729 | 350 | 4.9653
| 78 | 50 | 807 | 350 | 5.8846 |  |  | 5.761 |  |  | 5.7477
2 | 0 | 50 | 807 | 350 | 5.761
| 2 | 50 | 809 | 250 | 5.8333
| 1 | 50 | 810 | 225 | 5.7657
| 1 | 50 | 811 | 200 | 5.8316
| 1 | 50 | 812 | 175 | 6.0133 |  |  | 6.0288
2 | 0 | 50 | 812 | 175 | 6.0288
| 1 | 50 | 813 | 150 | 6.0898
| 1 | 50 | 814 | 125 | 6.445
| 1 | 50 | 815 | 100 | 6.3405
| 1 | 50 | 816 | 75 | 6.4735
| 2 | 50 | 818 | 35 | 7.4377 |  |  | 7.4696
2 | 0 | 50 | 818 | 35 | 7.4696
| 1 | 50 | 819 | 25 | 7.7939
| 1 | 50 | 820 | 15 | 8.3529
| 1 | 50 | 821 | 5 | 9.7361 | c | Sample #821 had bugs in sample. Duplicate sample had small bubble on probe at end of titration.
| 2 | 52 | 823 | 192 | 4.9041
| 1 | 52 | 824 | 175 | 4.859 |  |  | 4.8722 |  |  | 5.0908
2 | 0 | 52 | 824 | 175 | 4.8722
| 1 | 52 | 825 | 150 | 5.108
| 2 | 52 | 827 | 125 | 5.4361
| 2 | 52 | 829 | 100 | 5.9156
| 1 | 52 | 830 | 75 | 6.3743 |  |  | 6.378
2 | 0 | 52 | 830 | 75 | 6.378
| 4 | 52 | 834 | 35 | 7.8548 | c | Sample #834 had bugs in sample.
| 1 | 52 | 835 | 25 | 9.1035
| 2 | 52 | 837 | 15 | 9.5404
| 1 | 52 | 838 | 5 | 9.5521 |  |  | 9.4411
2 | 0 | 52 | 838 | 5 | 9.4411
| 23 | 54 | 861 | 137 | 5.4633
| 1 | 54 | 862 | 100 | 5.7769 |  |  | 5.8176 |  |  | 5.8409
2 | 0 | 54 | 862 | 100 | 5.8176
| 1 | 54 | 863 | 75 | 6.3242
| 1 | 54 | 864 | 50 | 7.0514
| 1 | 54 | 865 | 35 | 8.3066 | c | Samples #865, 866, 867 had bugs in sample.
| 1 | 54 | 866 | 25 | 9.4846 | c | Samples #865, 866, 867 had bugs in sample.
| 1 | 54 | 867 | 15 | 9.4734 | c | Samples #865, 866, 867 had bugs in sample.
| 1 | 54 | 868 | 5 | 9.3708
| 2 | 55 | 870 | 102 | 8.0418
| 1 | 55 | 871 | 75 | 8.3966
| 2 | 55 | 873 | 50 | 8.315 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration. | 8.3022 |  |  | 8.5094
2 | 0 | 55 | 873 | 50 | 8.3022
| 2 | 55 | 875 | 35 | 9.051
| 1 | 55 | 876 | 25 | 9.2635 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration. | 9.2389 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration.
2 | 0 | 55 | 876 | 25 | 9.2389 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration.
| 2 | 55 | 878 | 15 | 9.4377 |  |  | 9.476
2 | 0 | 55 | 878 | 15 | 9.476
| 1 | 55 | 879 | 5 | 9.628 | c | Samples #873, 876, 879, and duplicate sample #0876 had bugs in samples and bubbles on probe at end of titration.
| 25 | 58 | 904 | 104 | 8.491
| 1 | 58 | 905 | 75 | 8.4416 | c | Sample #905 had small bubble on probe at end of titration. | 8.5336 |  |  | 8.4517
2 | 0 | 58 | 905 | 75 | 8.5336
| 1 | 58 | 906 | 50 | 8.6576 |  |  | 8.66
2 | 0 | 58 | 906 | 50 | 8.66
| 1 | 58 | 907 | 35 | 9.0971
| 1 | 58 | 908 | 25 | 9.064
| 1 | 58 | 909 | 15 | 9.2763
| 1 | 58 | 910 | 5 | 8.9968
| 2 | 59 | 912 | 100 | 8.2391 |  |  | 8.237
2 | 0 | 59 | 912 | 100 | 8.237
| 3 | 59 | 915 | 35 | 8.8836 |  |  | 8.8917 |  |  | 8.8977
2 | 0 | 59 | 915 | 35 | 8.8917
| 26 | 62 | 941 | 210 | 6.9028 | c | Colorimeter drifting exceeding 100% for entire file.
| 1 | 62 | 942 | 200 | 7.131 | c | Colorimeter drifting exceeding 100% for entire file.
| 1 | 62 | 943 | 175 | 7.0786 | c | Colorimeter drifting exceeding 100% for entire file. | 7.0776 | c | Colorimeter drifting exceeding 100% for entire file. | 7.0938 | c | Colorimeter drifting exceeding 100% for entire file.
2 | 0 | 62 | 943 | 175 | 7.0776 | c | Colorimeter drifting exceeding 100% for entire file.
| 6 | 62 | 949 | 35 | 7.6924 | c | Colorimeter drifting exceeding 100% for entire file. | 7.6879 | c | Colorimeter drifting exceeding 100% for entire file.
2 | 0 | 62 | 949 | 35 | 7.6879 | c | Colorimeter drifting exceeding 100% for entire file.
| 8152 | 8 | 9101 | 3459 | 7.0285 | d | Colorimeter probe problems with drift for entire day's run
| 25 | 10 | 9126 | 1000 | 6.89 | d | Colorimeter probe problems with drift for entire day's run
| 23 | 11 | 9149 | 1000 | 7.9325
| 43 | 13 | 9192 | 585 | 6.308 | c | Triplicate sample # 9192 had bubble on probe.
| 42 | 15 | 9234 | 247 | 7.1443 | c
| 16 | 16 | 9250 | 250 | 6.3444
| 13 | 17 | 9263 | 1000 | 4.2411
| 17 | 18 | 9280 | 1000 | 4.4607 |  | Samples #291, 292, 294, 295 and duplicate samples #0287, 0295 had bubbles noticed in flasks
| 42 | 21 | 9322 | 1000 | 4.6675
| 22 | 22 | 9344 | 1000 | 4.5114
| 17 | 23 | 9361 | 1000 | 4.5478
| 39 | 25 | 9400 | 125 | 7.1492
| 8 | 26 | 9408 | 1000 | 4.6351
| 23 | 27 | 9431 | 1000 | 4.4901
| 35 | 30 | 9466 | 2250 | 2.815
| 23 | 31 | 9489 | 2000 | 2.7541
| 23 | 32 | 9512 | 1250 | 3.589
| 20 | 33 | 9532 | 1000 | 3.4473 | c | Triplicate sample #9532 had lighter color precipitate in flask than the others.
| 3 | 33 | 9535 | 400 | 5.7361
| 23 | 34 | 9558 | 200 | 7.2579
| 4 | 34 | 9562 | 100 | 7.6268
| 8 | 35 | 9570 | 500 | 5.5988
| 42 | 39 | 9612 | 250 | 5.773
| 7 | 39 | 9619 | 75 | 7.0961
| 6 | 39 | 9625 | 763 | 4.7111
| 27 | 40 | 9652 | 150 | 7.3109
| 3 | 40 | 9655 | 50 | 7.7643
| 19 | 41 | 9674 | 300 | 6.1021
| 20 | 42 | 9694 | 150 | 6.8366
| 113 | 50 | 9807 | 350 | 5.7477
| 17 | 52 | 9824 | 175 | 5.0908
| 38 | 54 | 9862 | 100 | 5.8409
| 11 | 55 | 9873 | 50 | 8.5094
| 32 | 58 | 9905 | 75 | 8.4517
| 10 | 59 | 9915 | 35 | 8.8977
| 28 | 62 | 9943 | 175 | 7.0938 | c | Colorimeter drifting exceeding 100% for entire file.