#### Nutrients

(Linda White – IOS)

Sampling

Water samples for nutrient determination were collected into acid-washed polystyrene test tubes after the tube and cap had been rinsed three times with the sample water.  If analysis could be performed within 24 hours the samples were stored at 4 °C, if not they were frozen at -20 °C.

Analysis and Results

Nutrients (nitrate + nitrite, silicate and orthophosphate) were analyzed by Linda White onboard ship using a three channel Technicon Auto Analyzer, following the methods described by Barwell-Clarke and Whitney (1996).

Surfactants for all three nutrients were not located after unpacking the nutrient supplies.  Ian Wrohan had set up his filtration equipment in the nutrient lab when the Louis left Halifax.  This year I relocated to the main U shaped counters in the Main Lab area.  This area is not as convenient to sinks, the Nanopure water system and is a high traffic area.  Ship’s water tested August 14, 2007 contained 0.05 µm/L nitrate + nitrite and 0.90 µm/L silicate and zero ammonium.

Palmolive dish detergent was an effective surfactant for nitrate and silicate but produced white precipitates in the platen tubing, flow cell and waste lines for phosphate.  Samples were frozen in glass for phosphate determination back at IOS.

Terry Whitledge of the University of Alaska, Fairbanks was contacted and he kindly sent Brij 35 and sodium lauryl sulfate surfactants to Point Barrow, Alaska where the Canadian Coast Guard Crew picked them up, August 3, 2007.  Jenny Hutchings and her husband Pat assisted in making this possible.

Sodium lauryl sulfate worked very well for silicate as Terry’s method described.  It does not work well with phosphate; it too, produced a white precipitate clogging the lines.  August 12th, I decided to freeze all phosphate samples in glass starting from Cast #43 for analysis at IOS.  Brij 35 is the usual surfactant for nitrate and Brij blocks colour development in the phosphate chemistry.

Reagents were prepared onboard using water from a NANOpure system that produced 17 to 18 mega ohm-cm resistance Type I reagent grade water.  The system was supplied with ship’s distilled water.  A 3.2% weight-to-volume solution of sodium chloride (Sigma, after 2003) was prepared daily and used to rinse the system between samples and to prepare working standards.  Pump tubing was changed after approximately 500 samples. One cadmium column was used for all samples unless noted below.  The Auto Analyzer was cleaned every other day as follows; rinsed with 3N NaOH first and then 10% HCl for approximately 5 minutes and rinsed with DMQ for over 20 minutes after all reagents and salt were disconnected at the end of the day.  Data were logged both by analog (chart) and digitally using the IOS “Newget” program.

- Standards and blanks

NANOpure water was analyzed daily before connecting the reagents and analyzing the initial standards and after the last set of standards to establish the baseline and record the purity of the reagents.  A set of working standards (low, medium and high) were prepared from the stock standard solution, using freshly prepared 3.2% sodium chloride solution.  The stock solutions were prepared from: Potassium nitrate, Fluka #60414, Sodium silicofluoride, Fluka #71596 and Dihydrogen potassium phosphate, Fluka #60218.  The working standards were analyzed at the start and close of each day or, if more than 60 samples were to be analyzed in a day, standards were also run mid-day or after three hours.  Concentrations of the standards were selected to bracket the expected nutrient levels in the samples.  A medium standard for each nutrient was analyzed between stations consisting of 12 to 27 samples and as an unknown sample followed by two zero standards.

Standards purchased from Wako (0 µm/L and 20 µm/L nitrate and 0 µm/L and 50 µm/L silicate) and Reference Samples (RS) purchased from KANSO (AS and AT) were analyzed at the end of each day.  See below for specific details.  An onboard reference sample was collected at sea, stored at 4 °C in the dark, and analyzed daily to provide an operational check.

The order of the sample analysis was from the surface to depth and sample peaks that appeared to be out of order were re-analyzed.  Duplicate samples were collected approximately every 10 samples.  One sample from every cast was collected in triplicate with two samples analyzed the day of sampling and the third sample analyzed the following day to verify the day-to-day calibrations.  The results of the replicate and standards comparisons are listed below.

The turbidity of surface samples where salinity is less than 27 PSU were analyzed through the phosphate channel with no reagents being added to the sample.  When the nitrate level in surface samples was the same or slightly lower than the 3.2% sodium chloride solution it was reported as zero.

Table 1.  Quality control and assurance for nutrient samples.

| Nutrient | Nitrate + Nitrite mmol/m3 | Silicate mmol/m3 | Phosphate mmol/m3 | Phosphate mmol/m3 |
| --- | --- | --- | --- | --- |
| Sample Replicates |  |  | Fresh | Frozen |
| sp | 0.09 | 0.15 | 0.03 | 0.03 |
| No. of duplicates | 131 | 133 | 36 | 43 |
| Range | 0.0 – 16.6 | 2.2 – 39.6 | 0.21 – 1.87 | 0.40 – 2.12 |
| Medium check standard |  |  |  |  |
| Calibrated value | 16.0 | 40.0 | 2.00 | 2.00 |
| Average and std dev | 16.1 ± 0.1 | 40.0 ± 0.2 | 1.99 ± 0.02 | 2.00 ± 0.02 |
| No. of duplicates | 47 | 47 | 16 | 19 |
| WAKO | 20 | 50 |  |  |
| Wako Standard | 20.1 ± 0.1 | 50.2 ± 0.17 | n/a | n/a |
| No. of duplicates | 15 | 4 |  |  |
| KANSO LOW | 0.1 | 1.84 | 0.05 | AY 0.51 Opened earlier |
| Japanese Reference Sample: AS | 0.07 ± 0.04 | 1.84 ± 0.13 | 0.04 ± 0.01 | 0.57 ± 0.03 |
| No. of duplicates | 12 | 12 | 9 |  |
| KANSO HIGH | 7.42 | 18.2 | 0.54 |  |
| Japanese Reference Sample: AT | 7.57 ± 0.06 | 18.4 ± 0.19 | 0.57 ± 0.01 |  |
| No. of duplicates | 12 | 10 | 9 |  |
| CB29b #283 shipboard reference | 13.0 ± 0.2 | 8.6 ± 0.1 | 0.94 ± 0.01 | 0.93 ± 0.02 |
| Range | 13.0 – 11.4 | 8.1 – 7.9 | 0.96 – 0.93 | 0.91 – 0.96 |
| n | 26 | 24 | 11 | 6 |
| Detection Limit | 0.09 | 0.13 |  |  |
| n | 10 | 10 |  |  |

2007: Cruise Specific Details

- Palmolive dish detergent

- July 31rst, #198 - #221 (N, Si and P),

- August 2, phosphate repeats plus Cast #16, Cast #10, Cast #9, Cast #2.

- Sodium Lauryl Sulfate:

August 3rd, tested phosphate with Sodium Lauryl sulfate(SLS) Cast# 9, Cast #1, and Cast #2.

- Sodium Lauryl Sulfate and Brij 35

August 4th, Nitrate is now with Brij, Silicate and phosphate with Sodium Lauryl Sulfate Casts # 15, 21, 23, 24, 4, 2, 20, and 14.

August 5th,,6th and 7th continue with SLS for silicate and phosphate Casts # 17, 30, 28, 26, 32, 31., Casts #25, 27, 37, 39, 29, Casts 36, 37, 40, 41.

August 8th, the following samples had been frozen and then thawed to analyze: Cast #3, 5, 11, 4.

- August 10th, Cast 43, 44, 45, 47, 48, 50.

August 12th, began freezing sample to be analyzed at IOS and continued use of Brij for nitrate and SLS for silicate with good results.

Sodium chloride was purchased from Sigma.  Wako 0µm/l and 20µm/l nitrate standards, a Wako 50µm/l silicate standard and KANSO AS (low) and AT (high) Reference Samples (RS) were analyzed at the end of each day though I ran out of Silicate early in the trip.  All KANSO reference samples, newly opened, were 3 years old.

At the end of each day’s run the shipboard reference sample, collected at CB 28b #283 from Cast #16  from 1,200 m depth on August 1, 2007 was analyzed throughout the course of the trip.

Duplicate samples were drawn for each station.  Several triplicates were taken and two analyzed one day and the third the next day.

Samples with salinity less than 27 were analyzed for phosphate turbidity and no phosphate samples required a turbidity correction. When the nitrate level in surface samples is the same or slightly lower than the 3.2% sodium chloride solution it is reported as zero.

At station CB18 quadruplets were taken for plastic and duplicate in glass.  2 sets will be analyzed for nitrate and silicate on day one, the triplicate on day 2 and the 4th set will be analyzed at IOS.  Both sets in glass will be analyzed for phosphate at IOS.

Samples collected at the last 5 stations were frozen because the Autoanalyser was dismantled and packed for return to IOS.  Frozen replicate samples were analyzed back at IOS to check on the use of Palmolive detergent and sodium lauryl sulfate as surfactants.