## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2007-26

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG – JdF		Party Chief: Masson D.

Platform: Vector		Date: April 2, 2007 – April 7, 2007

Processed by: Germaine Gatien

Date of Processing: 24 May 2007 – 28 August 2007

Number of original CTD casts: 74	Number of CTD casts processed: 74

Number of original rosette casts: 22	Number of rosette casts processed: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea/Seatech transmissometer (#1005DR), a PAR sensor (#4656) and an Altimeter (#1117). A SeaBird dissolved oxygen sensor (#1117) and a Seapoint Fluorometer (#2845) with a 10X cable were mounted on the pumps. The deck unit # was not recorded. The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log book and rosette sheets were available. There were no details entered about the equipment used, so there is some possibility of error in the serial numbers of the external sensors since they are based only on the entries in the configuration file.

The CTD was started deeper than usual for the Vector. The calibration for this pressure sensor has drifted over the years and an offset of +1.2db should be entered in the configuration file.

The stops for bottles were generally very short, sometimes less than 10s. The standard deviations in CTD salinity were high for many of the outliers in the bottle comparison, probably due to the high natural variability in this region and insufficient stop time for shed wake activity to dissipate. Those conditions would also mean that the DO sensors would often not have had sufficient time to equilibrate. These factors limit the quality of bottle comparisons and the recalibrations based on them.

The format was wrong for the quality flag channel of the dissolved oxygen files; it should be 2 columns wide, not 3. This is an error that occurs frequently.

The surface oxygen saturation (from the DO sensor on the CTD) was very high in Saanich Inlet and Malaspina Strait (~135% and 150%). High values were found during both the downcasts and upcasts and a bottle confirms the values in Saanich Inlet, so this is assumed to due to intense biological activity.

The SBE dissolved oxygen data in the CTD files should be considered:

•	±0.6ml/l from    0 -  50m

•	±0.2ml/l from   50 - 125m

•	±0.1ml/l below 125m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained. There were no details about the instruments used. The configuration file was used to determine the equipment used, but that can be subject to error especially for the external sensors.

The nutrient, chlorophyll and salinity data were each obtained in spreadsheet format. Each spreadsheet was edited to remove information not needed for the chemistry files. There were no quality flags for the salinity but comments were entered. Individual dissolved oxygen bottle files were obtained; flag channels and comments were included but the format was wrong for the quality flags.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The CTD calibrations were checked, offsets were added to the pressure and PAR sensor co-efficients and the file saved as 2007-26-CTD.con.

##### Conversion of Raw Data

The files were converted to CNV files. A few files were checked and the surface pressures are a little higher than usual for the Vector, but during cast #72 pressures at the end of the cast were ~0.5db with “in- water” values for conductivity. Had the surface offset not been applied the pressure would be <0. So the pressure offset applied does appear to be appropriate.

The descent rate was very steady for some casts, and none are really noisy.

The log book indicates that bottles were fired without stops, but this was not the case. Bottle stops appear to have been quite short; some lasted 30s but many were much shorter.

The temperature and conductivity channels are reasonably close with upcasts significantly noisier than downcasts.

Transmissivity shows reasonable agreement between downcast and upcast with most deep spikes in the upcasts.

The dissolved oxygen looks as expected with upcast and downcast offset vertically and some detail in the traces; there is some hysteresis but it is not severe.

The fluorescence looks ok with dark values of about 0.07ug/l.

The altimetry is looks reliable in the bottom 30db.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format. The event numbers were in the headers so CLEAN was not run; the extensions were changed to BOT. All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Four casts with a steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried. A choice of (0.03, 9) was best for the primary for 2 casts and for the other 2 there was no significant difference presumably because there was little temperature variation. For the secondary the best choice was to leave the data alone. CELLTM was run using (0.03, 9) for the primary channel only.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 270 | 0.0008 | XN ~0 | XN ~0 | Moderate, steady |
| 32 | 330 | 0.0008 | ~0.0002 | 0.0012 | High, steady |
| 42 | 240 | 0.0004 | ~0.0002 | 0.001 | Moderate, steady |
| 74 | 345 | 0.0008 | ~0.00015 | 0.0007 | High, steady |

The differences were all very noisy with many large spikes. When the primary and secondary were plotted together it is clear that the primary is noisier than the secondary. There was no significant pressure dependence in the salinity differences, but none of the casts are very deep.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

A header check shows no problems except that the platform name is wrong. That will be corrected later.

The header summary was checked against the log book entries and the only error was the station name for cast #62 which was changed from 12 to 11 in the header in the CLN file. There was no rosette file for that cast. The time in the header differed from that in the log book, but the log is clearly wrong and the header time looks right.

Track plots were produced and added to the end of this report.

The average surface pressure is ~3.0db which is a little high for the Vector. As mentioned earlier examination of upcast data suggests the pressure offset used is suitable since without it there would be some negative pressure values for records that clearly were taken while the CTD was in the water. It is possible that given the 1.2db error, the operators thought the CTD was shallower than it really was. No adjustment will be made to the pressure.

The altimeter values were exported to a spreadsheet; a few casts were examined and the algorithm was found to have worked well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Three problems were noted:

- The analyst noted that there was a sample marked as #173 whereas the rosette sheets indicate sampling was from #172. In fact, the rosette sheets indicate there was a salinity sample #174, but not #172 or 173. Looking at values of salinity compared to the CTD it seems likely that the label on the bottle is correct and it is from the bottom of cast #48.

- For cast #57 the surface bottle was not fired and no samples were taken so that line was removed from the ADDSAMP file.

- For cast #67 the rosette sheet shows 11 bottles fired while there are only 10 in the rosette file. The rosette sheet shows sampling of oxygen, nutrients and chlorophyll but no oxygen or nutrient samples were reported by analysts. There was a chlorophyll sample labelled as being from that bottle; this is assumed to be an error. That line was removed from the ADDSAMP file.

The file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data was provided in EXCEL format with a comment channel but no flag channel. The analyst was not available so “c” flags were added to two samples based on the comments. This should be reviewed after COMPARE. In adding event numbers to the spreadsheet  a number of inconsistencies were found between the salinity file and rosette log sheet, with some samples not noted at all on the sheets and some noted as from different bottles than on the sheets. Channel names were changed to standard format and the file was saved as 2007-26-sal.csv. The files were converted to individual SAL files.

The extracted chlorophyll data was in spreadsheet format and included a flag channel and comments. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The station name for sample #297 was corrected. As mentioned above, the sample given as #286 (cast 367) was changed to #285 since the bottle corresponding to #286 was not fired; the value was flagged “c” and a note of explanation put in the header. The resulting file is 2007-26-chl.csv. That spreadsheet was converted to individual files. The following comments from the analysis sheet were added to the comments that will later be added to the headers of the bottle files: “Average of two samples is reported unless stated otherwise. Variability is assessed as the % (std dev/mean*100).”  (Output: CHL)

The dissolved oxygen data were received in individual files with flag channel and comments. The flagged values had format errors; those were fixed.

The nutrient data were provided in spreadsheet format with flag channel and comments. Channel names were changed to standard format and the file was saved as 2007-26-nuts.csv. The data were converted to individual NUT files.

The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4.)

The MRG4 files were then merged with SAMAVG files (MRG) and then put through CLEAN to remove SeaBird headers and comments from the secondary files. (MRGCLN)

##### 11. COMPARE

Salinity

COMPARE was run and there is a lot of scatter in the results for both pairs of sensors but some of the outliers vary from one pair of sensors to the other. There are several explanations for this. The standard deviation in the CTD data is high for many of the outliers probably because the stops for bottles were generally short, the primary salinity is quite noisy and there is a lot of variability in this region. The two most extreme outliers come from stops that were less than 10s. When just samples with low standard deviations in the CTD are used and a few other outliers excluded the CTD is higher than the bottles by <0.001 for both channels. So either channel is usable.

There is no obvious time or pressure dependence, but the noise level is so high that this is not significant.

Two samples had been flagged earlier. The one for cast 41 was based on the lack of a liner; this flag will be left as “c” since it is a mild outlier. The one for cast #48 was based on confusion of sample #s.  It will also be left as “c” since there is some concern, but the comparison with the CTD suggests that it has been identified correctly.

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD and excluding outliers, the fit was:

DOX_BOT = 1.0274 * DOX_CTD + 0.023

Most of the outliers were already flagged by the analyst. The other outliers were either very shallow and/or had very short stops for bottles so no further flags will be assigned. (While the samples are not flagged, they may be affected by the short bottle stops since they may contain water brought from below by shed wakes.)

The fit is reasonably close to the results of 2007-11 from late March when the relationship was:

DOX_BOT = 1.0101 * DOX_CTD + 0.1109

The fit against pressure was reasonably tight, but not as good as against DO. The stops were very short so the fit may include some effects of time response errors as well as calibration errors.

The fit versus time shows no significant temporal drift when the same outliers are excluded as for the other fits.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. When all data is included the fluorescence is about 56% of the titrated chlorophyll values with an offset of about 0.20; setting the offset to 0 made little difference. None of the samples look like severe outliers. When the samples were broken down into different geographical regions there was little change and when only samples with CHL<1 were plotted the fluorescence was ~75% of the CHL. The ratio of CHL/FL showed some variation through the cruise with the highest ratios at the beginning of the cruise in Saanich Inlet and Haro Strait and the lowest at the mouth of Juan de Fuca Strait and in Johnstone Strait.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors were used for 2007-11 in March 2007. The bottle comparisons indicated that the primary salinity was low by 0.0015 and the secondary was low by 0.0004.

- This DO sensor is a new one that has been used only three times before for 2006-35, 2007-01 and 2007-11. The results of 2006-35 were considered unreliable because the stops were too short for the sensor to equilibrate.

For 2007-01 the fit was found to be

CTD_BOT = 0.9893 DOX_CTD -0.0142

with an offset of -0.047ml/l applied as a second calibration for the early casts and -0.156ml/l for later casts. Low oxygen water was sampled and an anoxic plug was used prior to use which probably accounted for most of the drift during the cruise.

For 2007-11 the fit was

CTD_BOT = 1.0101 DOX_CTD +0.1109

with a pressure-dependent second recalibration of

CTD_DOX (Corrected) = CTD_DOX + 0.00009 *Pressure - 0.041

There were once again problems with short stops that probably led to the complex recalibration.

Historic ranges – There are only 2 minor excursions from the historic ranges with salinity a little high in one case and a little low in another. These excursions are not considered indicative of instrument calibration problems.

##### 13. SHIFT

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were on the order of 1s which is the shift that has been used in most other cruises. All data was shifted by +24 records and a few casts were checked afterwards and found to be shifted appropriately. (Output: SHFFL)

Dissolved Oxygen

Based on the recent results of previous uses of this instrument, tests were run using settings from +80 to +140 records on 5 casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +120 records (5s).

Conductivity

Tests were run on a few casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be +0.5 for the primary sensor which is the same result as found during 2007-11. The secondary conductivity looks best with a setting of -0.7s where as -0.5 was used during 2007-11.

SHIFT was run using +0.5 and -0.7s for the primary and secondary conductivity channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: none

##### DETAILED EDITING

The salinity values from both pairs of sensors are close to the bottles. During 2007-11 when the same sensors were used the secondary salinity was closest to the bottles and the primary sensors were a little noisier in T-S space. So the secondary sensors were selected for archiving.

CTDEDIT was used to clean salinity.  On-screen plots of descent rate and pump status were used to aid in distinguishing real variability from instrumental noise.

The following casts required no editing: 3, 17, 19, 29.

The following casts required fairly heavy editing: 6, 7, 8, 10.

All other casts required only lightly editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

Salinity will not be recalibrated.

From section 11 we have the following equation for recalibration of DOX:

DOX_BOT = 1.0274 * DOX_CTD + 0.023

File 2007-26-recal1.ccf was prepared to apply the above correction to the DO channel in the SAM and MRG files. COMPARE was rerun for dissolved oxygen and when the same points were left out of the comparison as in the first run, the average difference is -0.001ml/l indicating that the recalibration worked properly; there is little variation with pressure and time. (See 2007-26-dox-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2007-26-recal1.ccf. (Output:COR1)

The COR1 files were clipped to 120db and set aside to be processed later for Angelica Peña. A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time, and the comparison with titrated samples those due to drift in calibration. But there remains an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged (0.25db bins) & thinned to the usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX, and outliers were excluded. A simple offset is often found to be appropriate at this stage, but was not for this data in which there was significant variation both with pressure and with DO value. This may be because some stops were quite short making errors higher near the surface where the gradients are higher. (See 2007-26-dox-comp3.xls.) File 2007-26-recal2.ccf was used to apply the following correction to the averaged files:

CTD_DOX (Corrected) = CTD_DOX + 0.0002 *Pressure - 0.137

COMPARE was rerun and the results indicate the above recalibration was effective. (See 2007-26-dox-comp4.xls.)

File 2007-26-recal2.ccf was then applied to the AVG files, but not to the bottle files since this error does not apply to data collected while stopped. (Output: COR2)

##### Special Fluorometer Processing

All COR1 files were put through CLIP to produce files with data to 120db only for the use of A. Peña.

Those files were then bin-averaged (1/4db bins), recalibrated suing 2007-26-recal2.ccf and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data was put through a median filter with fixed width 11, before bin-averaging.

The recalibrated CTD bottle files (SAMCOR1) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names, fix the platform name in the headers and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered

•	±0.6ml/l from    0 -  50m

•	±0.2ml/l from   50 - 125m

•	±0.1ml/l below 125m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~75% to 150% with the highest values in Saanich Inlet and Malaspina Stait. The lowest values were in Juan de Fuca Strait and Johnstone Strait. For Saanich Inlet a surface bottle confirms that the dissolved oxygen values are ok, and where high DO values are seen in downcasts they are also high in upcasts; so the high saturation values are assumed to be due to intense biological activity.

As a final check on the files a track plot, cross-reference listing and HEADER CHECK were run. Errors were found in the station name of a few casts; these were fixed.

##### Final Bottle Files

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE, Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units, to fix the platform name in the headers and to add a comment about quality flags including the following comments on extracted chlorophyll analysis:

Average of two samples is reported unless stated otherwise.

Variability is assessed as the % (std dev/mean*100) and analysis methods.

Plots were made of DO versus pressure and SBE:DO and Titrated DO versus CTD salinity. Outliers were checked and most had already been flagged. The following samples were flagged “c” because they were outliers in those plots and in profile plots:

- Cast #45, sample #171

- Cast #52, sample #198

A few other outliers were not flagged either because the standard deviation in the CTD DO data was high, or because they were very close to the surface where the CTD data may be unreliable. Where plot outliers corresponded to values already flagged, notes to that effect were added to the comments from the analyst.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### 22. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the CTD sensors.

Particulars from logs

1. Stopped early to avoid anoxic sampling?

7. Duplicates run on DO for Windows 98 / XP comparison.

38. High Fluorescence near the surface.

45. Niskin #2 fired early

46. Fraser plume

57. No surface sample

66. Error in labelling (written as 65 should be 66).

71. System left on at end for a few minutes.

# Institute of Ocean Sciences  CRUISE SUMMARY

| Cruise ID#: 2007-26 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 2 April 2007 End: 7 April 2007 |  |  |  |  |  |  |
| Location: SoG/JdF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Masson D. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550         Cruise ID#:		2007-26

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 23/12/06 | Factory “ |  |  |
| Conductivity | 1766 | 23/12/06 | “ |  |  |
| Secondary Temp. | 2374 | 21/12/06 | “ |  |  |
| Secondary Cond. | 1729 | 23/12/06 | “ |  |  |
| Transmissometer | 1005 | 28/12/06 | Factory |  |  |
| SBE 43 DO sensor | 1117 | 17/10/06 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |
| PAR | 4656 | 11/02/03 |  |  |  |
| Fluorometer | 2845 |  | IOS |  |  |
| Pressure Sensor | 75636 | 04/06/1999 | Factory |  |  |