## Overview
| Overview | - all instruments & sensors used in Arctic
*NEEDS TO BE COMPLETED & MAINTAINED !!
Inventory listed as Arctic or Melling to distinguish equipment purchased and/or maintained by Humfrey Melling
Inventory | Main | Type | S/N | date new | status (6 Nov 2006) | description | 2007 | max depth (m)
Arctic | CTD | SBE-9 | 0756 | 2004-03-01 |  | backup unit for LSSL | LSSL secondary
Arctic | CTD | SBE-9 | 0724 | 2003-03-01 | 3 bulkhead connectors require replacement | primary unit for LSSL | LSSL
Arctic | CTD | SBE-25 | 0293 |  | problem with startup - sent to Seabird for servicing | main unit for use on Laurier with 12-bottle Carousel sampler. Max pressure varies with Psensor (now 1024m), housing 3400db ? | SWL
Arctic | CTD | SBE-25 | 0415 | 2007-03-01 |  |  | Nahidik | 6800 but 1000m p sensor
IOS | CTD | SBE-25 |  | 2007-03-01 |  |  | SWL secondary
Melling | CTD | SBE-25 | 0307 |  |  | modified by Melling/Lindsay for through-ice use. Does not work with pylon? Max pressure (sensor )?? | Larsen
Arcitc | CTD | SBE-19 | 1030 |  |  | Only deep rated SBE19 in our inventory. Upgraded from 600m to 3400m max however pressure resolution is low (1.5db?) | LSSL/Nahidik, depending on whether Jamstec brings on on LSSL
Arctic | CTD | SBE-19 | 2688 |  |  | max 600m depth, used on lakes and Nahidik | Nahidik (spare, in TSG)
Acoustics | CTD | SBE-19 | 0904 |  |  | used 1998 at Arrow Lake, now (2006) in E-lab (rm.1248) Svein Vagle's inventory. Likely we can use it if its available. 600m max depth. | Larsen | 600
UBC | CTD | SBE-19 | 3112 |  | data interface problems, may need repair | UBC-Ingram - on loan to SEDNA Antarctic. 600m max depth | Sedna, being replaced
from Gower | CTD | SBE-19plus | 4132 |  | IOS-Gower, indefinite loan, modified from max 100 to max 600m depth | IOS-Gower, indefinite loan, modified from max 100 to max 600m depth | LSSL bio | 600
Arctic | CTD | SBE-19plus | 4057 |  |  | SBE 16+ converted to SBE19plus. Max depth 600m? (may have wrong sn…might be the 4056 SBE 16 that underwent conversion) | Lakes
Arctic | CTD | SBE-16plus | 4056 |  |  | $3000 would convert to SBE19plus | Svein's mooring in Bellot St.?
UBC? Arctic? | CTD | XR420 | 10855 | 2005 |  | UBC-Ingram - whale-survey-2005. Small, light instruments with inductive conductivity. Good for biological work. Max depth housing =740m, Psensor max??
Arctic | Pylon | SBE32 | 3245464-0616 | 2007-03-01 |  | 12 position Nahidik | Nahidik
Arctic | Pylon | SBE32 |  | 2007-03-01 |  | modified compact, up to 12l bottles | LSSL bio
Arctic | Pylon | SBE32 | 3215407-0172 | 2000-03-01 |  | 12 position (SWL) - factory repaired summer 2005 after 1st leg SWL | SWL
Arctic | Pylon | SBE32 | 3231679-0452 | 2003-03-01 | need post-cruise servicing to clean out lanolin from latches | 24 position (LSL) 3 bulkhead connectors at base (confirmed this is 2003 system 5/2006) | LSSL
Arctic | Pylon | SBE32 | 3235152-0498 | 2004-03-01 | need post-cruise servicing to clean out lanolin from latches | 24 position 1 bulkhead connector at base (confirmed this is 2004 system 5/2006)(LSL) | LSSL
Arctic | Deck Unit | SBE11 |  | 2004-03-01 |  | deck unit for SBE9+ | LSSL
Arctic | Deck Unit | SBE11 |  | 2003-03-01 |  | deck unit for SBE9+ | LSSL
Arctic | Deck Unit | SBE-33 | 3315407-0028 | 1997-03-24 |  | deck unit (never used? found in e-lab by Ron L.) has PAR | SWL
Arctic | Deck Unit | SBE-33 | 0132 | 2007-03-01 |  | has PAR | Nahidik
IOS | Deck Unit | SBE-33 |  |  |  |  | SWL backup
Arctic | Deck Unit | SBE-33 | 0058 |  | Check NMEA | deck unit with SBE25#0293 CHECK NMEA option! No PAR | Nahidik backup
Melling | Deck Unit | SBE-33 | 0082 |  | needs repair | deck unit (bought by Humfrey for use with modified SBE25-0307) No PAR
Arctic | PDIM | SBE | 0051 |  |  | underwater power-data-interface-module for live data via deck-unit used with SBE25#0293
Melling | PDIM | SBE | ?? |  |  | " " - built into SBE-25 0307 pressure housing
Arctic | AutoFireModule | SBE AFM | 23 |  | needs testing and possibly repair | Use with SBE25-0293 to close bottles on rosette without conducting wire.
Arctic | AutoFireModule | SBE AFM |  | 2007-03-01 |  |  | LSSL bio
Arctic | Pressure | SBE-29 | 0436 |  |  | 2006: deep sensor (spare for SWL SBE-25). Max depth 3400m.
Melling | Pressure | SBE-29 | 0571 | 2006 |  | 2006: 1450psia spare bought by Humfrey. (Is this approx 1000m?)
| Temp | SBE-03 | 2913
Melling | Temp | SBE-03 | 4044 |  |  | 2006: spare for HM (Nares Strait)
| Temp | SBE-03 | 2926
Arctic | Temp | SBE-03 | 4115 |  |  | primary unit for CTD25#0293
| Temp | SBE-03 | 4109 |  |  | 2006: spare for SWL (SBE-25)
Arctic | Temp | SBE-03 | P4322 | 2003-03-01 |  | primary unit for CTD9 #0724
Arctic | Temp | SBE-03 | P4239 | 2003-03-01 |  | secondary unit for CTD9 #0724
| Temp | SBE-03 | 1086
Arctic | Temp | SBE-03 | 4397 | 2004-03-01 |  | primary unit for CTD9 #0756
Arctic | Temp | SBE-03 | 4402 | 2004-03-01 |  | secondary unit for CTD9 #0756
| Cond | SBE-04 | 2232 |  |  | 2006: spare for HM (Nares Strait)
| Cond | SBE-04 | 2582
Arctic | Cond | SBE-04 | 2607 |  |  | primary unit for CTD25#0293
| Cond | SBE-04 | 2676 |  |  | 2006: spare for SWL (SBE-25)
Arctic | Cond | SBE-04 | 2809 | 2003-03-01 |  | primary unit for CTD9 #0724
Arctic | Cond | SBE-04 | 2810 | 2003-03-01 |  | secondary unit for CTD9 #0724
| Cond | SBE-04 | 2913
| Cond | SBE-04 | 0771
Arctic | Cond | SBE-04 | 2992 | 2004-03-01 |  | primary unit for CTD9 #0756
Arctic | Cond | SBE-04 | 2984 | 2004-03-01 |  | secondary unit for CTD9 #0756
Arctic | Oxygen | SBE-43 | 1202 | 2007-03-01 |  | wet-pluggable bulkhead connector | SWL
Arctic | Oxygen | SBE-43 | 1115 | 2006-10-01 |  | wet-pluggable bulkhead connector | LSSL (backup)
Arctic | Oxygen | SBE-43 | 615 | 2004-03-01 |  |  | Nahidik
Arctic | Oxygen | SBE-43 | 435 | 2003-03-01 |  | primary unit for CTD9 #0724 LSSL | LSSL
Arctic | Oxygen | SBE-43 | 52 |  |  | used on SBE19+ onNahidik
Arctic | pH meter | 18-I | 180232 |  |  | Not often used…given to Keith Johnson in water properties 11/2006
Arctic | Altimeter | Datasonics PSA-916D | 1161 | 2005-03-01 |  | primary unit for CTD9 #0724 LSSL | LSSL
Arctic | Altimeter | Datasonics PSA-916D | 40853 | 2007-02-01 |  | available for Laurier system | SWL
Arctic | Altimeter | Datasonics PSA-916D | 41098 | 2007-03-01 |  |  | Nahidik | 6000
Arctic | Transmissometer | Wetlabs | CST-DR |  |  |  | Lakes? | 6000
UBC | Transmissometer | Wetlabs | CST-DR |  |  |  | LSSL | 6000
UBC | Transmissometer | Wetlabs | CST-DR |  |  |  | SWL | 6000
Arctic | Transmissometer | Wetlabs | CST-662DR | 2007-03-01 | needs new cable | primary unit for CTD9 #0724 LSSL | LSSL | 6000
Arctic | Transmissometer | Wetlabs | CST-1047DR | 2007-03-01 |  |  | Nahidik | 6000
Arctic | Transmissometer | Seatech 25cm | 93367 |  | needs testing/repair/replacement. Showing large hysteriesis | used with SBE25 Laurier? | SWL
Arctic | Transmissometer | Seatech 10cm | 0163 |  | needs testing | used with SBE19+ | Nahidik
Arctic | Fluorometer | Seapoint | SCF 2841 | 2006-10-01 |  |  | SWL (tsg) | 6000
Arctic | Fluorometer | Seapoint | SCF 2859 | 2007-03-01 |  |  | Nahidik | 6000
Arctic | Fluorometer | Seapoint | SCF 2569 |  |  | primary unit for CTD9 #0724 LSSL | LSSL | 6000
Arctic | Fluorometer | Seapoint | SCF 2336 |  |  | for SBE25 #0293 SWL | SWL | 6000
Arctic | Fluorometer | Seapoint | SCF 2745 |  |  | for SBE19P #4132 (lakes) | LSSL | 6000
Arctic? | Fluorometer | Wetlabs | Wetstar WS3S521P |  | working? | Currently in Antarctica | lakes?
Arctic | Fluorometer | Seatech | S93 |  | needs testing…if we want to keep | Power hungry, heavy monster of a fluorometer. Is battery pack location known?
Arctic | Fluorometer | Seapoint | SCF ???? | 2006-05-01 |  | Purchased for Nahidik work | Nahidik | 6000
Arctic | Optical Backscatter | Seapoint |  |  |  | Used on Nahidik 2006
Arctic | PAR | Biospherical QSP2200PD | 20273 | 2007-03-01 |  | underwater | Nahidik | 2000
Arctic | PAR | Biospherical QSR2100 | 10350 | 2007-03-01 |  | Surface PAR, RS232 output | Nahidik | 2000
Arctic | PAR | Biospherical QSP2200PD |  | 2007-03-01 |  | underwater, log amplifier integrated | LSSL | 2000
Arctic | PAR | Biospherical QSR2100 |  | 2007-03-01 |  | Surface PAR | LSSL | 2000
UBC | PAR | Biospherical QSP2200PD |  | 2007-03-01 |  | underwater | SWL | 2000
UBC | PAR | Biospherical QSR2100 |  | 2007-03-01 |  | Surface PAR | SWL | 2000
Melling | Pump | SBE-5T | 3373 |  |  | 2006: spare for HM (Nares Strait)
Cage
Cage
Still have, but no longer in use…
Arctic | CTD | MCTD | 1598 |  | loan to P.Vallarta, Mexico (Univ. Guadelahare ) | …until further notice (data quality limited).
Arctic | CTD | FSI - ICTD | 1329 |  | storage… | no longer used: data quality problems (slow response of inductive cond to rapid Temp changes
Arctic | CTD | Applied Microsystems
Arctic | Deck Unit | General Oceanic | 1015 SCI
| Oxygen | 23-01-Y | 230326 |  |  | not used, discontinued, should be surplused
Arctic | Transmissometer | Seatech | 192D |  | Not good, need to surplus | Not good, need to surplus
Used at one time, but not our inventory…
Water Properties | CTD | SBE-9 | 0443
Water Properties | Pylon | SBE33 | ? | ? | Water Properties (S.Romaine) believes there is an extra pylon for small Carousel | Water Properties (S.Romaine) believes there is an extra pylon for small Carousel
| CTD | SBE-19 | 1031 |  | might still be in our inventory?? | lost at sea in Aug 1998 (during SHEBA)
Arctic | CTD | MCTD | 1534 |  |  | (2004?) returned to SFI for repair: C-cell broke in transport, too expensive to fix: no longer available.

## SBE C-sensors
SBE04-
Conductivity | date: | 2000-12-22 | 2002-02-15 | 2002-03-07 | 2002-12-31 | 2003-01-24
# 04-2232 | g= | -10.9583 | -10.9697 | -10.4419 | -10.4431 | -10.417
| h= | 1.5495 | 1.5528 | 1.4776 | 1.4777 | 1.4733
| i= | -0.0031019 | -0.003917 | -0.0034982 | -0.003469 | -0.0031073
| j= | 0.00027151 | 0.00034012 | 0.00029422 | 0.00029279 | 0.00026962
| Cpcor= | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  | 0.0006
Conductivity | date: | 2001-03-20 |  |  |  |  | 2004-12-16
# 04-2582 | g= | -10.4002 |  |  |  |  | -10.4027
| h= | 1.5713 |  |  |  |  | 1.5718
| i= | 0.0002858 |  |  |  |  | 0.00024741
| j= | 7.3728e-05 |  |  |  |  | 7.7726e-05
| CPcor= | -9.57e-08 |  |  |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  |  |  | 3.25e-06
drift PSU/month |  | new |  |  |  |  | -0.0001
Conductivity | date: | 2001-04-11 |  |  | 2002-12-31 | 2004-05-14 | 2004-12-10 | 2005-01-20 | 2006-02-10 | 2007-01-05 | 2007-12-04
# 04-2607 | g= | -10.7508 |  |  | -10.7616 | -10.7617 | -10.7616 | -9.1256 | -9.1348 | -9.1299 | -9.1354
| h= | 1.5007 |  |  | 1.5037 | 1.5034 | 1.5035 | 1.2734 | 1.276 | 1.2743 | 1.2765
| i= | -0.0023592 |  |  | -0.0031131 | -0.0029381 | -0.0029996 | -0.001792 | -0.0024126 | -0.0019715 | -0.0025967
| j= | 0.00025292 |  |  | 0.00030162 | 0.00028378 | 0.00029037 | 0.00018425 | 0.00022655 | 0.00019818 | 0.00024295
| Cpcor= | -9.57e-08 |  |  | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  | 0 | -0.0001 | 0 | -0.0001 | -0.0004 | -0.0001 | 0.0001
Conductivity | date: | 2002-03-14 |  |  | 2002-12-31
# 04-2676 | g= | -10.6862 |  |  | -10.677
| h= | 1.4924 |  |  | 1.4889
| i= | -0.0070427 |  |  | -0.0059867
| j= | 0.00063624 |  |  | 0.00054955
| Cpcor= | -9.57e-08 |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06
drift since last cal: |  | 0 |  |  | +0.00010
post-cruise cal:
Conductivity | date: | 2003-02-21 |  |  | 2004-05-25 | 2004-10-05 | 2006-02-10 | 2007-01-09
# 04-2809 | g= | -9.9553 |  |  | -9.961 | -9.9605 | -9.9672 | -9.9552
| h= | 1.4118 |  |  | 1.4134 | 1.4134 | 1.4147 | 1.4116
| i= | -0.0002185 |  |  | -0.00058677 | -0.00060551 | -0.00091874 | -0.00013431
| j= | 9.6919e-05 |  |  | 0.00012396 | 0.00012563 | 0.00014512 | 9.0548e-05
| Cpcor= | -9.57e-08 |  |  | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | new |  |  | -0.0004 | 0.0002 | 0.0001 | -0.0001
Conductivity | date: | 2003-02-21 |  |  | 2004-05-25 | 2004-10-05 | 2006-02-10 | 2007-01-09
# 04-2810 | g= | -9.8663 |  |  | -9.8736 | -9.8721 | -9.8788 | -9.8695
| h= | 1.4576 |  |  | 1.4595 | 1.4592 | 1.4608 | 1.4582
| i= | -0.00065958 |  |  | -0.001096 | -0.0010272 | -0.0014372 | -0.00072556
| j= | 0.00013121 |  |  | 0.00016409 | 0.00016068 | 0.00018624 | 0.00013489
| Cpcor= | -9.57e-08 |  |  | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CTcor= | 3.25e-06 |  |  | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  | -0.0004 | 0 | 0.0001 | 0.0001
Conductivity | date: | 2001-06-04
# 04-2913 | g= | -9.8663
| h= | 1.4576
| i= | -0.00065958
| j= | 0.00013121
| Cpcor= | -9.57e-08
| CTcor= | 3.25e-06
drift since last cal: | PSU/mon | 0
Conductivity | date: |  |  | 2002-05-06
# 04-0771 | g= |  |  | -4.2696
(JAMSTEC?) | h= |  |  | 0.50542
| i= |  |  | 0.0002447
| j= |  |  | 1.0228e-05
| Cpcor= |  |  | -9.57e-08
| CTcor= |  |  | 3.25e-06
drift since last cal: | PSU/mon | 0 |  | -0.00105
Conductivity | date: | 11-Mar-04 NEW |  |  |  | 2005-01-20
# 04-2992 | g= | -9.9454 |  |  |  | -9.9409
| h= | 1.372 |  |  |  | 1.37
| i= | 0.00033338 |  |  |  | 0.00093495
| j= | 0.00047816 |  |  |  | 1.3188e-05
| Cpcor= | -9.57e-08 |  |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  |  | 3.25e-06
drift since last cal: | PSU/mon | 0 |  |  |  | -0.0014
Conductivity | date: | 11-Mar-04 NEW |  |  |  | 2005-01-26
# 04-2984 | g= | -10.1869 |  |  |  | -10.1896
| h= | 1.3835 |  |  |  | 1.3841
| i= | 0.00021555 |  |  |  | 6.7445e-05
| j= | 5.7268e-05 |  |  |  | 6.9622e-05
| Cpcor= | -9.57e-08 |  |  |  | -9.57e-08
| CTcor= | 3.25e-06 |  |  |  | 3.25e-06
drift since last cal: | PSU/mon | 0 (new) |  |  |  | -0.0004
Conductivity | date: | 2007-04-25
# 04-2754 | g= | -10.5843
IOS pool, | h= | 1.6106
for SBE25-0334 | i= | -0.0011046
Arctic-2007 | j= | 0.00017875
| Cpcor= | -9.57e-08
| CTcor= | 3.25e-06
drift since last cal: | PSU/mon | unknown - check w/ Doug Andersen
Note on C41: BvH: on new SBE9/11 Primary Conductivity NEW FROM FACTORY (IS THIS SAME ACCURACY AS SENSORS FOR SBE-25?)
Note on C50: BvH: on new SBE9/11 Secondary Conductivity NEW FROM FACTORY (IS THIS SAME ACCURACY AS SENSORS FOR SBE-25?)
Note on D5: Bon: these coeffs are for the post-cruise conductivity sensor calibration before electrode repair!!
Note on E5: Bon: new C-cell, was replaced under warranty (defective electrodes)
Note on F5: Bon: these coeffs are for post-cruise conductivity sensor calibration before electrode repair!!
Note on G5: Bon: C-cell had cracks, was replaced and then re-calibrated at our expense
Note on H23: BvH: POST CRUISE calibration
Note on H41: BvH: Calibration after modifications
Note on H50: BvH: Calibration after modifications
Note on I23: BvH: Calibration after modifications

## SBE 21+38
SBE21 | S/N 3274
T: | calibr | 2006-01-15 | 2007-12-05
| G | 0.0041949
| H | 0.00062806
| I | 1.9915e-05
| J | 1.4304e-06
| Fo | 1000
| Slope | 1
| Offset | 0
C: | calibr | 2006-01-15 | 2007-12-04
| G | -4.1916 | -9.1354
| H | 0.49525 | 1.2765
| I | -0.00033458 | -0.0025967
| J | 4.3595e-05 | 0.00024295
| CTcor | 3.25e-06 | 3.25e-06
| CPcor | -9.57e-08 | -9.57e-08
| Slope | 1
| Offset | 0
SBE38 | S/N 0271
T: | calibr |  | 2007-12-05
| a0 |  | -0.00021868
| a1 |  | 0.00030691
| a2 |  | -4.4042e-06
| a3 |  | 2.0411e-07
| drift (˚C/yr) | new | -0.00032

## SBE T-sensors
SBE03-
Temperature | date: | 2001-06-04 |  |  |  | 2004-12-10
# 03-2913 | g= | 0.0043476 |  |  |  | 0.004348
(coef for ITS-90) | h= | 0.00064083 |  |  |  | 0.00064166
| i= | 2.2647e-05 |  |  |  | 2.3223e-05
| j= | 2.104e-06 |  |  |  | 2.2355e-06
| fo= | 1000 |  |  |  | 1000
drift degrC/yr: |  | new |  |  |  | 0.00111
Temperature | date: | 2001-01-01 | 2002-02-15 | 2002-12-31
# 03-4044 | g= | 0.0043426 | 0.0043427 | 0.0043426
coeffs for | h= | 0.00064252 | 0.00064268 | 0.00064264
ITS-90 | i= | 2.2433e-05 | 2.2553e-05 | 2.2535e-05
| j= | 1.9681e-06 | 1.9947e-06 | 1.9932e-06
| fo= | 1000 | 1000 | 1000
drift degrC/yr: |  | 0 | -0.0007 | +0.00047
Temperature | date: | 2000-06-24
# 03-2926 | g= | 0.0043698
| h= | 0.00064799
| i= | 2.3381e-05
| j= | 2.1484e-06
| fo= | 1000
Temperature | date: | 2001-03-21 |  | 2002-12-31 | 2004-05-15 | 2004-12-10 | 2006-02-24 | 2007-01-04 | 2007-12-07
# 03-4115 | g= | 0.0043583 |  | 0.0043583 | 0.0043583 | 0.0043584 | 0.0043583 | 0.0043584 | 0.0043583
| h= | 0.00064523 |  | 0.00064527 | 0.00064536 | 0.00064564 | 0.00064537 | 0.00064549 | 0.00064536
| i= | 2.2859e-05 |  | 2.29e-05 | 2.2954e-05 | 2.314e-05 | 2.2941e-05 | 2.3035e-05 | 2.2942e-05
| j= | 1.9982e-06 |  | 2.0089e-06 | 2.0203e-06 | 2.0616e-06 | 2.0116e-06 | 2.0378e-06 | 2.0158e-06
| fo= | 1000 |  | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
drift degrC/yr: |  |  |  | +0.00003 | +0.00017 | +0.00026 | -0.00032 | -0.00042 | -0.00021
Temperature | date: | 2002-03-14 |  | 2002-12-31
# 03-4109 | g= | 0.0043846 |  | 0.004384
| h= | 0.00064987 |  | 0.00064875
| i= | 2.3482e-05 |  | 2.2729e-05
| j= | 1.9523e-06 |  | 1.7861e-06
| fo= | 1000 |  | 1000
drift degrC/yr: |  |  |  | +0.00102
Temperature | date: | 2003-03-08 |  | 2004-05-25 | 2004-10-05 | 2006-02-24 | 2007-01-05
# 03P4322 | g= | 0.0044195 |  | 0.0044195 | 0.0044193 | 0.0044193 | 0.0044193
(ITS-90 coeffs) | h= | 0.00065479 |  | 0.00065467 | 0.00065433 | 0.00065437 | 0.00065436
| i= | 2.4111e-05 |  | 2.4027e-05 | 2.3804e-05 | 2.3817e-05 | 2.3817e-05
| j= | 1.9761e-06 |  | 1.9572e-06 | 1.9079e-06 | 1.9073e-06 | 1.9107e-06
| fo= | 1000 |  | 1000 | 1000 | 1000 | 1000
drift degrC/yr: |  |  |  | -0.00015 | -0.00018 | -0.00032 | -0.00043
Temperature | date: | 2003-03-07 |  | 2004-05-25 | 2004-10-05 | 2006-02-24 | 2006-12-30
# 03P4239 | g= | 0.0044146 |  | 0.0044146 | 0.0044146 | 0.0044146 | 0.0044146
| h= | 0.00065408 |  | 0.00065401 | 0.00065399 | 0.00065398 | 0.00065396
| i= | 2.3742e-05 |  | 2.3682e-05 | 2.3676e-05 | 2.3653e-05 | 2.365e-05
| j= | 1.9085e-06 |  | 1.8943e-06 | 1.8944e-06 | 1.8848e-06 | 1.8884e-06
| fo= | 1000 |  | 1000 | 1000 | 1000 | 1000
drift degrC/yr: |  |  |  | -0.00018 | -0.00062 | -0.00022 | -0.00049
Temperature | date: |  |  | 2002-05-04
# 03-1086 | g= |  |  | 0.0048156
| h= |  |  | 0.00067502
| i= |  |  | 2.6032e-05
| j= |  |  | 2.0502e-06
| fo= |  |  | 1000
drift degrC/yr: |  |  |  | -0.0003
Temperature | date: | 12/03/2004 (NEW) |  |  | 2005-01-20
# 03-4397 | g= | 0.0043828 |  |  | 0.0043828
| h= | 0.00064745 |  |  | 0.00064752
| i= | 2.2363e-05 |  |  | 2.24e-05
| j= | 1.7932e-06 |  |  | 1.7995e-06
| fo= | 1000 |  |  | 1000
drift since last cal | degrC/yr: |  |  |  | 8e-05
Temperature | date: | 06/03/2004 (NEW) |  |  | 2005-01-20
# 03-4402 | g= | 0.0043692 |  |  | 0.0043692
| h= | 0.00064929 |  |  | 0.00064934
| i= | 2.2614e-05 |  |  | 2.2638e-05
| j= | 1.7781e-06 |  |  | 1.7822e-06
| fo= | 1000 |  |  | 1000
drift degrC/yr: |  |  |  |  | 8e-05
Temperature | date: | 14-Mar-07 (NEW) | 2007-12-18
# 03-4444 | g= | 0.0043388 | 0.0043381
| h= | 0.00063768 | 0.00063633
| i= | 2.157e-05 | 2.0655e-05
| j= | 1.9188e-06 | 1.7164e-06
| fo= | 1000 | 1000
drift degrC/yr: |  |  | 0.0014
Temperature | date: | 2007-04-25
# 03-2668 | g= | 0.0043564
IOS pool, | h= | 0.00064885
for SBE25-0334 | i= | 2.4647e-05
Arctic-2007 | j= | 2.4568e-06
| fo= | 1000
drift degrC/yr: |  | unknown - check w/ Doug Andersen
Note on C42: BvH: on new SBE9/11 Primary Temperature HIGHER ACCURACY THAN SBE-25 SENSORS?
Note on C50: BvH: on new SBE9/11 Primary Temperature HIGHER ACCURACY THAN SBE-25 SENSORS?

## SBE P-sensors
SBE-Psensors |  |  | see individual CTD's for SBE-19's
# 290442 | date: | 2000-11-15 | 2002-02-21
max=1024m | A0= | 787.8446 | 787.0138
| A1= | -0.20726 | -0.20721
| A2= | 2.7948e-07 | 2.6257e-07
| TC= |  | -140
#195334 | date: | 2001-04-27
max=1024m | A0= | 795.367
| A1= | -0.20741
| A2= | 2.8042e-07
TC=
# 290436 | date: | 2001-01-19 | 2002-02-11
max=6800m | A0= | 5285.741 | 5286.5012
| A1= | -2.7596 | -2.7587
| A2= | -1.205e-06 | -1.2983e-06
| TC= |  | 19
# 63507 | date: | 1996-01-11
10,000 m? | C1 | -44354.68
on SBE9-0443 | C2 | -0.18457
| C3 | 0.011927
| D1 | 0.039799
| D2 | 0
| T1 | 29.9357
| T2 | -0.00033362
| T3 | 3.7592e-06
| T4 | 1.5653e-09
| T5 | 0
| AD590M | 0.012833
| AD590B | -9.295
# 290464 | date: | 2001/2 ? | 2002-12-27 | 2006-03-23
max= 1024 m | A0= |  | 796.2542 | 796.5379
(1500 psia) | A1= |  | -0.20736 | -0.20729
| A2= |  | 2.5193e-07 | 2.4589e-07
| TC= |  | -264 | -264
# 290571 | date: | 2006-04-12
max= 1000 m | PA0= | 920.3101
(1450 psia) | PA1= | -0.25184
S/N 8531 | PA2= | -1.4623e-08
| TCV= | -6
# 90559 | date: | 2002-10-29
max= 6800 m? | C1 | -44678.87
(fixed) | C2 | -0.628
on SBE9-0724 | C3 | 0.014524
| D1 | 0.037286
| D2 | 0
| T1 | 30.1586
| T2 | -0.00042482
| T3 | 3.8582e-06
| T4 | 4.3914e-09
| T5 | 0
| Slope | 1
| Offset | 0
| AD590M | 0.012854
| AD590B | -9.035
# 91164 | date: | 2004-03-18
max= 6800 m? | C1 | -46369.51
max= 10,000psi | C2 | -0.074006
on SBE9-0756 | C3 | 0.014342
| D1 | 0.034424
| D2 | 0
| T1 | 30.4004
| T2 | -0.0003058
| T3 | 4.4282e-06
| T4 | 1.9912e-09
| T5 | 0
| Slope | ?
| Offset | ?
| AD590M | ?
| AD590B | ?
# 290544 | date: | 2005-05-13
max= 1000 m | PA0= | 925.434
(1450 psia) | PA1= | -0.25292
S/N 5707 | PA2= | -9.5868e-08
| TCV= | -20
Note on A70: BvH: on SBE9/11 system
Note on C46: BvH: spare for SBE25-0307 bought by Humfrey
Note on C52: BvH: on new SBE9/11, not interchangeable
Note on C86: BvH: from IOS pool on spare SBE25-0334 (SWL 2007)

## SBE-05 pumps
Serial Number | New |  | Comment
# 053373
# 053164 |  |  | on SBE9-0443 for Arctic-2002
# 051604 |  | IOS/OSAP | on SBE9-0443 for Arctic-2002
# 053610 | 2003-03-01 |  | on new SBE9-0724
# 053615 | 2003-03-01 |  | on new SBE9-0724
#053869 | 2004-03-01 |  | on new SBE9-0756
#053871 | 2004-03-01 |  | on new SBE9-0756
? | 2005-03-01 |  | new stronger pump bought for SBE 19+
# 054051 | ? |  | on SBE19plus S/N 4132

## SBE9-0724
SBE9 - 0724 |  |  | NEW March 2003 from IOS for Arctic use
date | Use | temperature | conductivity | pump | temperature | conductivity | pump | pressure | D.O. | transmissometer | fluorometer | turbidity | Bottom | Altimeter | Water Sampler | Deck Unit | Comments
|  | primary | primary | primary | secondary | secondary | secondary |  | sensor | Wetlabs | Seapoint | sensor | Contact Warning | sensor
2003-03-31 | Set-up as received March 2003: | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | 10133 |  | none | 3235152-??? | 11P31679-0649 | original setup
2003-07-15 | Set-up for August 2003 LSL (2003-21): | Same as original configuration??
2004-07-25 | Set-up for August 2004 LSL (2004-16 S Zimmermann, Canada Basin): | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | none | yes | Datasonics PSO-916T#640 | 3235152-??? | 11P31679-0649 | original setup with addition of WHOI altimeter. Only CTD problem was leak on oxygen sensor cable at the CTD end. Bulkhead connector and cable required cleaning.
2004-09-25 | Set-up for Sept-2004 LSL (2004-33 Jane Eert, circum Somerset Island): | 4397 | 2992 | 3610 | 4402 | 2984 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | none | yes | none | 3235152-??? | 11P31679-0649 | As original but T&C pairs from NEW 9/11 (SBE9-0756) system. Oxygen sensor found to have torn membrane in post cruise calibration. Problems with initial termination - kept blowing bulkhead connector fuse until new connectors put on sea-cable termination. Post-cruise Doug Sieberg replaced CTD's bulkhead connectors for the sea cable and the aux channel used with oxygen sensor.
2005-06-03 | Set-up for August 2005 LSL (2005-04 S Zimmermann, Canada Basin): | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 43-0435 | CST-662DR | SCF 2569 | none | yes | Datasonics PSA-916D #1161 | 3235152-??? And 3231679-0452 | 11P31679-0649 and 11P35152-0680 | Original setup now with first use of Arctic Altimeter. Post cruise cal shows oxygen sensor has manufacturers defect in membrane. Primary pylon replaced mid-cruise with secondary pylon. CTD to pylon cable shows corrosion at pylon end. Communication errors with pylons, data spikes and sticking values in Seasave throughout cruise. See problem list in CTD 9#724 directory. New winch this year. It is now permanently on the LSSL.
2006-06-01 | Set-up for August 2006 LSL (2006-18 S Zimmermann, Canada Basin): | 4322 | 2809 | 3610 | 4239 | 2810 | 3615 | 90559 | 3 total: 43-0435, JAMSTEC, Seabird loaner | 2 total: CST-662DR; Jen Jackson | SCF 2569 | none | yes | Datasonics PSA-916D #1161 | 3235152-??? And 3231679-0452 | 11P31679-0649 and 11P35152-0680 | Communication errors with pylon did not occur this year (due to new seacable, new termination, new pylon cable?). However sticky pylon pins were a problem, partly due to lanolin buildup and partly due to bad pull angle between pin and bottle cap. Pylon caps were swapped out and cleaned to deal with this. Dual oxygen sensors on at start of cruise, the seabird loaner failed, then the Jamstec sensor was put on however a leak at the sensors bulkhead connector could not be fixed so was also removed. Two transmissometers were put on. Jen is looking at the transmissometer data and was interested in comparison to help look at historic data.

## SBE9-0756
SBE9 - 0756 |  |  | NEW March 2004 'for Arctic Use'
date | Use | temperature | conductivity | pump | temperature | conductivity | pump | pressure | D.O. | transmissometer | fluorometer | turbidity | Bottom | Altimeter | Water Sampler | Deck Unit | Comments
|  | primary | primary | primary | secondary | secondary | secondary |  | sensor | Wetlabs | Seapoint | sensor | Contact Warning | sensor
2004-03-31 | Set-up as received March 2004: | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | SBE43-615 | none | none | none | none | none | 3231679-0452 | 11P35152-0680 | original setup
2004-07-25 | Set-up for August 2004 LSL (2004-16 S Zimmermann, Canada Basin) as backup. Test cast only, not used afterwards. | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | none | none | none | none | none | none |  |  | did not have to use
2005-06-03 | Set-up for August 2005 LSL (2005-04 S Zimmermann, Canada Basin) as backup. Used to help troubleshoot problems in primary system. Secondary pylon used for last half of cruise. | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | none | none | none | none | none | none |  |  | pylon needs inspection
2006-06-01 | Set-up for August 2006 LSL (2006-18 S Zimmermann, Canada Basin) as backup. Secondary pylon head used for last half of cruise. | 4397 | 2984 | yes | 4402 | 2992 | yes | 91164 | none | none | none | none | none | none |  |  | pylon 'head' needs cleaning

## SBE9-0443 (IOS)
SBE9 - 0443 | on loan from IOS/OSAP |  | for 2002-20, -21, -22, -23 (LAURIER Leg1-3 & LOUIS STLAURENT)
date | temperature | conductivity | pump | temperature | conductivity | pump | pressure | transmissometer | fluorometer | Comments
| primary | primary | primary | secondary | secondary | secondary
1996-07-31 | 2106 | 1748 | 1602 | 2062 | 1764 | 1604 | 63507 |  |  | original setup
---
2002-06-15 | 4044 | 2232 | 3164 | 4109 | 2676 | 1604 | 63507 | SeaTech# 139 | SeaPoint #2336 | also Touchdown Switch
Laurier leg1,2,3
SeaTech# 192D
deep casts LSL
Note on I13: BvH: used deep transm from OSAP for 2002-23, but it died after several casts, so swapped #139 back for shallower casts.

## SBE25-0293
SBE25-0293 | check sensors ratings for |  |  | press rating
# 2524864-0293 | max.depth |  |  |  |  | LogAmp for PAR: M=-0.770231, B=-3.530844
housing: | 3400 m |  | ARCTIC-2001
sensors | Temperature | Conductivity | Pump: | Pressure | P-range | Transmissom | Fluorometer | DO | PAR
type | SBE-03 | SBE-04 | SBE5-01 | SBE-29 |  | SeaTech | Seapoint | SBE-43
press rating | [5 empty] | 2000m | 6000m
ARCTIC-2001 | # 034044 | # 042232 | #053000 | # 290442 | 1024m | # 139 | # 2336
|  |  |  | # 290436 | 6800m
ARCTIC-2002 | #034115 | #042607 | #053100? | # 290464 | 1024m
|  |  |  | # 290436 | 6800m
ARCTIC-2003 | #034115 | #042607 |  | # 290464 | 1024m | # 139 | # 2336 | # 430052
ARCTIC-2004 | #034115 | #042607 |  | # 290464 | 1024m |  | # 2336 | # 430052
30x (5mg/l)
ARCTIC-2005 | #034115 | #042607 |  | # 290464 | 1024m | SeaTech #139 | # 2336 | # 430615
30x (5mg/l) ?
ARCTIC-2006 | #034115 | #042607 |  | # 290464 | 1024m | SeaTech #139 | # 2336 | # 430615 |  | touchdown
SWL spares: | #034109 | #042676 |  | # 290436 | 6800m |  | 3x (alt:10x) |  |  | (on V1)
ARCTIC-2007 | #034115 | #042607 |  | # 290464 | 1024m | CST-1052DR | # 2336 | # 430615 | QSP-20280 | Altimeter
SWL spares: | #034109 | #042676 |  | # 290436 | 6800m |  | 3x (alt:10x) |  | QSR-20281 | (on V1)
Note on G27: BvH: new 2007
Note on J27: BvH: new 2007, UBC Grant Ingram
Note on J28: BvH: new 2007, UBC Grant Ingram

## SBE25-0307
SBE25-0307 | purchased 2001
#2525901-0307 |  |  | spares:
max.depth=?? | Temperature | # 2926 | #4044
| Conductivity | # 2607 | #2232
| Pressure | # 195334 | max 1024 dbar
| Pump: SBE5-01 | #053000 | #3373
THIS SBE-25 EXTENSIVELY MODIFIED BY MELLING/LINDSAY:
no longer standard modular but internal pump,
sensors in-line for through-ice deployments
built-in PDIM (still requires internal batteries)

## SBE25-0415
SBE25-0293 | check sensors ratings for |  |  | press rating
# 2524864-0415 | max.depth
housing: | 6800m |  | ARCTIC-2007
sensors | Temperature | Conductivity | Pump: | Pressure | P-range | Transmissom | Fluorometer | DO | Altimeter
type | SBE-03F | SBE-04C | SBE-05T | SBE-29 |  | Wetlabs | Seapoint | SBE-43 | PSA-916D
press rating | [5 empty] | 6000m | 6000m |  | 6000m
ARCTIC-2007 | #034115 | #042607 |  | # 290603 | 1024m | CST-1047DR | # 2589 | # 430202 | 41098
spares: |  |  |  | # 290530 | 1024m

## SBE25-0334
SBE25-0334 | check sensors ratings for |  |  |  | press rating
IOS/OSD POOL | max.depth
housing: | 6800m |  | ARCTIC-2007
sensors | Temperature | Conductivity | Pump: | Pressure | P-range | Transmissom | Fluorometer | DO | Altimeter
type | SBE-03 | SBE-04C | SBE-05T | SBE-29 |  | Wetlabs | Seapoint | SBE-43 | PSA-916D
press rating | [5 empty] | 6000m | 6000m |  | 6000m
ARCTIC-2007 | 2668 | 2754 | 1990 | 0544 | 1024m | CST-1047DR | # 2589 | 0997 | 41098
spares:

## SBE19-2688
SBE19-2688 |  |  |  | repaired pump #SBE5-050478
#1920437-2688
Temperature | date: | 1999-02-02 | 1999-12-16 | 2001-03-31 | 2002-02-20 | 2004-12-15
# 2688 | G= | 0.004238 | 0.004237 | 0.0042425 | 0.0042349 | 0.0042364
| H= | 0.000618 | 0.00061518 | 0.00062824 | 0.00061001 | 0.00061361
| I= | 1.0711e-05 | 8.2355e-06 | 1.8565e-05 | 4.1161e-06 | 7.0516e-06
| J= | -4.4827e-07 | -1.1436e-06 | 1.5603e-06 | -2.2012e-06 | -1.4298e-06
| F0= | 1000 | 1000 | 1000 | 1000 | 1000
| drift/month = | none, new | 0.00218 | 0.00359 | -0.00331 | 0.00041
Conductivity | date: | 1999-02-02 | 1999-12-16 | 2001-03-31 | 2002-02-20 | 2004-12-15
# 2688 | G | -4.2223 | -0.0042219 | -4.2266 | -4.2313 | -4.2255
| H | 0.5052 | 0.50487 | 0.50557 | 0.50701 | 0.50544
| I | 0.00068681 | 0.00078607 | 0.00066489 | 0.00039076 | 0.00074027
| J | -2.1004e-06 | -7.713e-06 | -1.9877e-06 | 1.176e-05 | -6.4719e-06
| CP-corr | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08 | -9.57e-08
| CT-corr | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06 | 3.25e-06
| drift PSU/month = | none, new | 0.0006 | 0.0001 | -0.0003 | -0.0002
Pressure | date: | 1999-03-29 | 1999-12-28 | 2001-04-04 | 2002-02-23 | 2005-01-04
# 192787 | PA0 | 500.4465 | 500.8254 | 501.396 | 501.4606 | 502.1478
(1000psia) | PA1 | -0.13016 | -0.13019 | -0.13018 | -0.13015 | -0.13018
| PA2 | 1.3081e-07 | 1.2742e-07 | 1.2591e-07 | 1.2897e-07 | 1.1719e-07
T-compens | TC= |  |  |  | -70 | -70
note: 6-pin IO conn
damaged at Seabird
replaced by factory
battery bulkh re-glued
2005:00
2005: used with
DO SBE430052
Fluo SCF 2745

## SBE19-1030
SBE19-1030 |  | (pump: SBE5-01) | [8 empty] | mod: 1998/05 new housing, new press 3K psi?
# 195891-1030 |  | #050478 | [8 empty] | mod sept-98: | new 5Kpsia pressure sensor
Please see bottom of spreadsheet for notes on when CTD was used.
Factory cal
Temperature | cal: | 1992-02-20 | 1993-06-25 | 1994-03-16 | 1995-05-31 | 1995-10-31 |  |  |  | 1997-05-15 | date: | 1998-08-28 | date: | 1999-12-16 | 2001-03-23 | 2002-02-20
# 1030 | a= | 0.0036742 | 0.0036816 | 0.0036734 | 0.0036834 | 0.0036835 |  |  | G= | 0.0041813 | G= | 0.0041787 | G= | 0.0041786 | 0.0041809 | 0.0041783
| b= | 0.00058077 | 0.00058052 | 0.00058055 | 0.00057941 | 0.00058064 |  |  | H= | 0.00059305 | H= | 0.00058661 | H= | 0.00058602 | 0.00059202 | 0.00058514
| c= | 9.2078e-06 | 5.7046e-06 | 9.2139e-06 | 1.823e-06 | 7.4723e-06 |  |  | I= | 5.5332e-06 | I= | 3.3133e-07 | I= | -2.9748e-07 | 4.6952e-06 | -1.0628e-06
| d= | -5.995e-07 | -5.4104e-07 | -4.1614e-07 | -9.4383e-06 | -2.7471e-06 |  |  | J= | -1.2749e-06 | J= | -2.6466e-06 | J= | -2.8314e-06 | -1.5024e-06 | -3.0535e-06
| fo= | 2371.56 | 2341.4 | 2374.56 | 2333.47 | 2371.56 |  |  | Fo= | 1000 | Fo= | 1000 | Fo= | 1000 | 1000 | 1000
drift (degC/yr) | = | none (new) |  | 0.0015 | [7 empty] | 0.00134 |  | 0.00143 | -0.0038 | 0.00232
factory cal:
Conductivity | cal: | 1992-02-20 | 1993-06-25 | 1994-03-16 | 1995-05-31 | 1995-10-31 |  | 1997-05-15 |  | 1997-05-15 |  | 1998-08-28 |  | 1999-12-16 | 2001-03-23 | 2002-02-20
# 1030 | a= | 0.043492 | bad cal? | 0.0234 | -0.063445 | 0.29952 | G= | -4.1731 | G= | -4.1731 | G= | -4.1815 | G= | -4.1816 | -4.1846 | -4.1928
| b= | 0.45132 |  | 0.50216 | 0.75248 | -0.31391 | H= | 0.5052 | H= | 0.49813 | H= | 0.49927 | H= | 0.49933 | 0.49976 | 0.50151
| c= | -4.1579 |  | -4.187 | -12.3762 | 12.9003 | I= | 0.00068681 | I= | 0.0011809 | I= | 0.0011068 | I= | 0.0011022 | 0.0010729 | 0.00076572
| d= | 0.00052781 |  | 0.00038087 | -0.19057 | 0.28918 | J= | -2.1004e-06 | J= | -2.7492e-05 | J= | -2.4442e-05 | J= | -2.4174e-05 | -2.1087e-05 | -7.2566e-06
| m= | 2.1 |  | 3.3 | 2.3 | 2.3 | Cpcorr | -9.57e-08 | Cpcorr | -9.57e-08 | Cpcorr | -9.57e-08 | Cpcorr | -9.57e-08 | -9.57e-08 | -9.57e-08
[7 empty] | CTcorr | 3.25e-06 | CTcorr | 3.25e-06 | CTcorr | 3.25e-06 | CTcorr | 3.25e-06 | 3.25e-06 | 3.25e-06
drift (psu/mon) | = | none, new |  | 0.0001 | [7 empty] | 0.0026 |  | 0.002 | 0.0018 | 0.0001
Pressure | cal: | 1992-02-27 |  | 1994-03-09 |  |  |  |  | date: | 1997-05-12
# 154898 | A0 | 148.8465 |  | 148.925 |  |  |  |  | A0 | 143.0902
(300 psia) | A1 | -0.039069 |  | -0.039071 |  |  |  |  | A1 | -0.039057
| A2 | 3.4771e-08 |  | 3.2421e-08 |  |  |  |  | A2 | 3.3418e-08
Pressure | cal: | [8 empty] | 1997-05-19
# 177965 | A0 | [8 empty] | 499.8601
(1000 psia) | A1 | [8 empty] | -0.13045
| A2 | [8 empty] | 1.9117e-07
Pressure | cal: | [9 empty] | date: | 1998-06-04
# 184833 | A0 | [9 empty] | A0 | 1509.2799
(3000 psia) | A1 | [9 empty] | A1 | -0.39093
| A2 | [9 empty] | A2 | -3.0934e-09
Pressure | cal: | [9 empty] | date: | 1998-08-26 | date: | 1999-12-27 | 2001-03-27 | 2002-02-27
# 189759 | A0 | [9 empty] | A0 | 2501.659 | A0 | 2502.7156 | 2502.824 | 2502.9772
(5000 psi) | A1 | [9 empty] | A1 | -0.65204 | A1 | -0.65194 | -0.65197 | -0.65202
| A2 | [9 empty] | A2 | -5.9724e-08 | A2 | -1.6646e-08 | -1.751e-08 | -3.8302e-09
T-compensation | TC | [14 empty] | -221
SERVICE: |  |  |  | Factory cal. T,C,P | Factory cal. T,C | [6 empty] | Calibr T, C, P |  | Factory cal. T,C,P | Factory cal. T,C,P
[5 empty] | check & cal 300psi | [6 empty] | replaced 3Kpsi with 5Kpsi |  | replaced 6-pin conn | replaced 6-pin I/O
[5 empty] | install & cal 1000psi | [6 empty] | Pump SBE5 #050478: |  | on pump 050478 | and 6-pin AUX conn
[5 empty] | found spreadsheet: | [6 empty] | replaced 6-pin conn &
[5 empty] | 300psi test (IOS?) | [6 empty] | new pump interface cable
1995-05-26
88.5604
-0.026882
1.9057e-08
-2.3418e-12
USE: |  | Date | Cruise | User | Comments
|  | 2004-08-01 | LSSL 2004-16 | S Zimmermann | Used with Par sensor, attached to bongo-net wire for 100m deep casts. Trouble with pressure equilibrating.

## SBE19-0904
SBE19-03
#195178-0904
Temperature | date: | 1998-01-17 |  | 2002-10-22 | 2005-02-24
# 0904 | A+ | 0.0035994 | G= |  | 0.0042377
| B= | 0.00058105 | H= |  | 0.00059764
| C= | 9.706e-06 | I= |  | 5.0796e-06
| D= | -1.9463e-06 | J= |  | -1.582e-06
| F0= | 2948.267 | Fo= |  | 1000
drift/month =
Conductivity | date: | 1998-01-17 |  |  | 2005-02-24
# 0904 | A | 0.052404 | G= |  | -4.276
| B | 0.41199 | H= |  | 0.51014
| C | -3.9154 | I= |  | 0.0013208
| D | -0.00016657 | J= |  | -3.5836e-05
| M | 2.1 | CT-corr |  | 3.25e-06
| CP-corr | 0 | CP-corr |  | -9.57e-08
CT-corr
| drift PSU/month = | none, new
Pressure | date: | 1998-02-02 |  |  | 2005-02-26
# 163246 | A0 | 500.4465 | A0 |  | 247.9882
(1000psia) | A1 | -0.13016 | A1 |  | -0.065244
| A2 | 1.3081e-07 | A2 |  | 5.8996e-08
T-compens |  |  | Offset |  | 0
TC=

## SBE19+4132
SBE19P #4132 |  | This SBE19plus is on loan from Jim Gower.
19P28301-4132 |  | We have replaced the 100m pressure sensor with a 600m (1000psia) and C,T,P sensor calibrations Jan 2005: see 4132-05.CON
Temperature | date: | 2005-01-04
# 4132 | A0 | 0.0011941
| A1 | 0.00028429
| A2 | -3.7353e-06
| A3 | 3.3708e-07
| Slope | 1
| Offset | 0
Conductivity | date: | 2005-01-04
# 4132 | G | -0.96451
| H | 0.13547
| I | -0.00042622
| J | 4.8554e-05
| CTcor | 3.25e-06
| CPcor | -9.57e-08
| Slope | 1
| Offset | 0
Pressure | date: | 2004-12-30
# 4132 | PA0 | -0.68787
(1000psia) | PA1 | 0.0044295
| PA2 | -1.3836e-11
| PTEMPA0 | -72.2228
| PTEMPA1 | 48.4525
| PTEMPA2 | 0.076849
| PTCA0 | 520412.8
| PTCA1 | -13.8036
| PTCA2 | 0.3222
| PTCB0 | 24.711
| PTCB1 | 0.0016
| PTCB2 | 0
| Offset | 0
Pump: 5T | flow rate?
S/N 054051
2006: setup for
Quesnel Lake
DO SBE430052
Fluo SCF 2745

## SBE16+4056
SBE 16plus SEACAT
S/N | 16P25902-4056
1) Count, Temperature
Serial number | 4056
Calibrated on | 20-Apr-01
A0 | 0.0012433
A1 | 0.00026616
A2 | -6.0888e-07
A3 | 1.7815e-07
Slope | 1
Offset | 0
2) Frequency, Conductivity
Serial number | 4056
Calibrated on | 20-Apr-01
G | -0.96832
H | 0.13551
I | -0.00033778
J | 4.3232e-05
CTcor | 3.25e-06
CPcor | -9.57e-08
Slope | 1
Offset | 0
3) Count, Pressure, Strain Gauge
Serial number | 4056
Calibrated on | 25-Apr-01
PA0 | 0.11494
PA1 | 0.0015959
PA2 | 6.7871e-12
PTEMPA0 | -65.5077
PTEMPA1 | 52.4302
PTEMPA2 | -0.58815
PTCA0 | 525308.1
PTCA1 | 4.6884
PTCA2 | -0.068187
PTCB0 | 24.2797
PTCB1 | -0.00045
PTCB2 | 0
Offset | 0

## SBE16+4057
SBE 16plus SEACAT |  | CONVERTED TO SBE16plus PROFILER 12 June 2006
S/N | 16P25902-4057 |  | MAX DEPTH 600 m | memory 8192K
zero cond freq= 2693.4 Hz
1) Count, Temperature |  |  |  | Pump SBE 5T # 054398 2000 rpm
Serial number | 4057 | post-cruise |  | Final cal: | ITS-90
Calibrated on | 20-Apr-01 | Calibr: | 25-May-06 | Calibr: | 2-Jun-06
A0 | 0.0012408 | A0= | 0.0012686 | A0= | 0.0013251
A1 | 0.00026678 | A1= | 0.0002564 | A1= | 0.00023526
A2 | -5.7575e-07 | A2= | 7.1559e-07 | A2= | 3.3438e-06
A3 | 1.7983e-07 | A3= | 1.2632e-07 | A3= | 1.7612e-08
Slope | 1 | Cpcor= | -9.57e-08
Offset | 0 | Ctcor= | 3.25e-06
|  | drift | -0.00048 | drift since Apr'01 | -0.00071
2) Frequency, Conductivity
Serial number | 4057 | post-cruise |  | Final cal:
Calibrated on | 20-Apr-01 | Calibr: | 25-May-06 | Calibr: | 2-Jun-06
G | -0.98226 | g= | -0.98073 | g= | -0.98118
H | 0.1364 | h= | 0.1358 | h= | 0.13594
I | -0.00051315 | i= | -0.00034029 | i= | -0.00037167
J | 5.6495e-05 | j= | 4.3835e-05 | j= | 4.5773e-05
CTcor | 3.25e-06 | Cpcor= | -9.57e-08 | Cpcor= | -9.57e-08
CPcor | -9.57e-08 | Ctcor= | 3.25e-06 | Ctcor= | 3.25e-06
Slope | 1 | drift/month= | -0.0001 |  | -0.0001
Offset | 0
3) Count, Pressure, Strain Gauge
Serial number | 4057 | new P: | 870psia #9877
Calibrated on | 25-Apr-01 |  |  | Calibr: | 2-Jun-06
PA0= | 0.024922 |  |  | PA0= | -0.0009201
PA1= | 0.0015895 |  |  | PA1= | 0.0026662
PA2= | 6.8826e-12 |  |  | PA2= | -7.6053e-12
PTEMPA0= | -70.2803 |  |  | PTEMPA0= | -61.5162
PTEMPA1= | 49.3556 |  |  | PTEMPA1= | 51.9853
PTEMPA2 = | -0.040851 |  |  | PTEMPA2= | 0.022318
PTCA0= | 524442.6 |  |  | PTCA0= | 519298.6
PTCA1= | 3.036 |  |  | PTCA1= | -4.146
PTCA2= | -0.097682 |  |  | PTCA2= | 0.19222
PTCB0= | 24.3084 |  |  | PTCB0= | 2.4964
PTCB1= | -0.000925 |  |  | PTCB1= | -0.00085
PTCB2= | 0 |  |  | PTCB2= | 0
Offset | 0

## SBE19-3112
On loan from Grant Ingram (UBC)
SBE19-3112 |  |  | 2005/6: SEDNA
1925666-3112 |  |  | (Antarctic mission)
Temperature | date: | 2001-03-22
# 3112 | G= | 0.004195
| H= | 0.00061401
| I= | 1.2935e-05
| J= | -4.9472e-07
| F0= | 1000
| drift/month = | none (new)
Conductivity | date: | 2001-03-22
# 3112 | G | -4.1164
| H | 0.49289
| I | 0.00053015
| J | 7.859e-06
| CT-corr | 3.25e-06
| CP-corr | -9.57e-08
| drift PSU/month = | none, new
Pressure | date: | 2001-03-23
# 165700 | PA0 | 149.5054
(300 PSIA) | PA1 | -0.039268
Max.200m | PA2 | 4.0474e-85
T-compens | TC= | 87
Pump
SBE 5T # 261
Fluorometer | cal: | 2004-05-14
WS3S-512P | CWO | 0.068
(on SEDNA'05/06) | SF | 14
Note on B33: BvH: Clean Water Offset at 24 degr C (Wetlabs term)
Note on B34: BvH: Scale Factor in microgram/l/V at 24 C

## SBE19-1294
SEABIRD SEACAT CALIBRATIONS (from IOS spreadsheet by Minkley)
serial number 1294
date | 1998-04-29 | 1999-06-17
temperature
g | 0.0042368 | 0.0042388
h | 0.00059866 | 0.00060264
I | 5.2131e-06 | 7.9157e-06
j | -1.4243e-06 | -8.0729e-07
f0 | 1000 | 1000
Conductivity
g | -4.0406 | -4.0336
h | 0.48217 | 0.48061
I | 0.0012708 | 0.0015497
j | -3.7073e-05 | -4.9584e-05
Cpcor | -9.57e-08 | -9.57e-08
Ctcor | 3.25e-06 | 3.25e-06
Pressure (quadratic Fit)
A0 | 496.4681 | 496.5765
A1 | -0.13028 | -0.13028
A2 | 1.2446e-07 | 1.2761e-07
TC | -44 | -44

## SBE-21 TS
SBE-21 Thermo Salinograph
S/N 3274 |  | purchased 2006 | used on SLW July - October 2006
|  | new | post-cruise | after re-platinizing
Conductivity: | Calibr.: | 2006-01-15 | 2007-02-22 | 2007-03-06
| g |  | -4.1947 | -4.1898
| h |  | 0.49527 | 0.49461
| i |  | -0.00026254 | -0.00021167
| j |  | 3.993e-05 | 3.6979e-05
| Cpcorr |  | -9.57e-08 | -9.57e-08
| Ctcorr |  | 3.25e-06 | 3.25e-06
| drift(PSU/month) |  | -0.0017 | 0.0003
|  |  | post-cruise | after cleaning
Temperature: | Calibr.: | 2006-01-15 | 2007-02-22 | 2007-03-06
| g |  | 0.0041941 | 0.0041944
| h |  | 0.00062609 | 0.00062685
| i |  | 1.8278e-05 | 1.8864e-05
| j |  | 9.834 | 1.1277
| f0 |  | 1000 | 1000
| drift(deg/yr) |  | -0.00125 | -0.00231
Remarks: |  |  | firm ware upgraded
T,C post-cruise calibrations
clean & replatinized
O-ring replaced
Pressure tested and
system check & diagnostic
CHECK--> |  |  | SBE38 remote-T not post-cruise recalibrated
S/N SBE3842376-0271

## XR-420-10855
XR-420-CTD | mfg: | RBR
| owner: | Grant Ingram (UBC)
C-sensor repaired Dec 2005 |  | new coefficients:
logger conductivity =a+b*VR+c*VRI\2+d*VRI\3
| a | 0.028863
| b | 124.4691
| c | 0
| d | 0
max depth?

## SBE37-0071
SBE-37 # 0459 |  | Humfrey's?
Temperature: | Cal. date: | 2002-11-26 | 2004-12-14
#0459 | a0 | (Lindsay?) | -6.3299e-05
| a1 |  | 0.00028965
| a2 |  | -3.5177e-06
| a3 |  | 1.8365e-07
drift since prev cal | (deg.C/yr) |  | -1e-05
Conductivity: | Cal. date: | 2002-11-26 | 2004-12-14
# 0459 | g |  | -1.0478
| h |  | 0.15201
| i |  | -0.00044553
| j |  | 5.6951e-05
| Cpcor |  | -9.57e-08
| Ctcor |  | 3.25e-06
| WBOTC |  | -8.079e-05
drift since prev cal | (PSU/mon) |  | -0.0005
RTC freq | Cal. date: | 2002-11-26 | 2004-12-14
ITS-90 scale | rtca0 |  | 0.99997
| rtca1 |  | 1.7725e-06
| rtca2 |  | -2.8914e-08

## SBE43 D.O.
ordered May 2001 - delivered end June - used in sea trials - see remarks
# 430052 | Cal. Date: | 2001-06-12 | 2001-08-06 |  | 2005-10-07
|  |  | repaired! |  | repaired!
Coefficients: | Soc= | 0.3557 | 0.3159 |  | 0.3322
| Boc= | 0.0138 | 0.0019 |  | 0
| Offset | -0.6129 | -0.6129 |  | -0.5987
| Tcor= | 0.0013 | 0.0003 |  | 0.0007
| Pcor= | 0.00015 | 0.00015 |  | 0.000135
|  | 2001-07-01 |  | Aug,2004 | Sept. 2005
Services performed: |  | large hysteresis in initial tests: returned unit from Dutch Harbor to factory for repair | factory replaced torn membrane, | returned to factory for repair (?), replaced bulkhead connector. Bad data in 2004-13 cruise!! | returned for repair: SPAR assembly replaced, no post-cruise cal, new cal above.
#430615 | New date: | 2004-03-09 |  |  | 2006-02-08 |  | 2007-01-10
Came with SBE9-0756
Coefficients: | Soc= | 0.3916 |  |  | 0.4037 |  | 0.3955
| Boc= | 0 |  |  | 0 |  | 0
| Offset | -0.4643 |  |  | -0.4752 |  | -0.4797
| Tcor= | 0.0011 |  |  | 0.0013 |  | 0.0013
| Pcor= | 0.000135 |  |  | 0.000135 |  | 0.000135
#43-0435 | New date: | 2003-03-18 |  | 2005-03-14 | 2006-02-16 | 2006-06-06 | 2007-01-25
Came with SBE9-0724
Coefficients: | Soc= | 0.3581 |  | 0.4159 | 0.3563 | 0.3938 | 0.3874
| Boc= | 0 |  | 0 | 0 | 0 | 0
| Voffset | -0.4804 |  | -0.4812 | -0.4925 | -0.495 | -0.4935
| Tcor= | 0.0015 |  | -0.0006 | 0.0009 | 0.0016 | 0.001
| Pcor= | 0.000135 |  | 0.000135 | 0.000135 | 0.000135 | 0.000135
| Tau= | 0 |  | 0
Comments |  |  |  | Torn membrane replaced prior to calibration | SBE replaced defective membrane (no charge) | SBE will replaced defective SPAR assembly(no charge)
#430071 | New date: | 2001-06-01 | Factory loaner used for Leg 1 of 2001-21
Coefficients: | Soc= | 0.4273
| Boc= | 0.0201
| Offset | -0.6275
| Tcor= | 0.0014
| Pcor= | 0.00015
#431202 | New date: | 21-Mar-07 (NEW)
Coefficients: | Soc= | 0.3428
| Boc= | 0
| Voffset | -0.519
| Tcor= | 0.0014
| Pcor= | 0.000135
| Tau= | 0
#431115 | New date: | 2006-10-17
Coefficients: | Soc= | 0.3663
| Boc= | 0
| Voffset | -0.5053
| Tcor= | 0.0015
| Pcor= | 0.000135
Tau=
Comments: | with wet-plug connector (& cable) !
2007: from IOS pool, on SWL spare CTD SBE25-0334
#430997 | Cal date: | 2006-03-29
Coefficients: | Soc= | 0.4146
| Boc= | 0
| Voffset | -0.4831
| Tcor= | 0.0011
| Pcor= | 0.000135
| Tau= | 0

## SBE23-0326 DO
SBE-23 YSI Dissolved Oxygen sensor
# 230326 | Cal. Date: | 1992-02-28 | 1993-03-30 | 1993-03-31 | 2001-04-10
|  |  | post-cruise | new diaphram | new diaphram
Sensor current | m= | 1.053e-05 | 1.053e-05 | 1.053e-05 | 1.053e-05
| b= | -1.053e-09 | -1.053e-09 | -1.053e-09 | -1.053e-09
Comp.Temp. | k= | 6.4851 | 6.4851 | 6.4851 | 6.4852
| c= | -2.17 | -2.17 | -2.17 | -2.17
Seasoft coeffs: | Soc= | 0.0652 | 0.3254 | 0.078 | 0.0783
( Oxfit ) | Boc= | -0.0517 | -0.4451 | -0.036 | -0.0156
| tcor= | -0.033 | -0.033 | -0.033 | -0.033
| pcor= | 0.00015 | 0.00015 | 0.00015 | 0.00015
| tau= | 2 | 2 | 2 | 2
| wt= | 0.85 | 0.85 | 0.85 | 0.85
| barom(mB)= | 1013.13 | 1013.037 | 1005.349 | 1014.015
| Twater(C)= | 23.34 | 21.6 | 21.85 | 4.495
| Tcomp(C)= | 23.35 | 21.73 | 21.97 | 4.263
| Isat(uA)= | 33.937 | 7.633 | 26.629 | 15.015
| Iair(uA)= | 35.074 | 7.37 | 27.293 | 25.955
| Izero(uA)= | 0.793 | 1.368 | 0.462 | 0.199

## Fluorometers
FLUOROMETERS
Unit |  | Settings | Comments
SeaTech | new: | 1994 ?
S/N S93 | unknown | set by internal jumper | April 2006 BVH: The Seatech S/N 93 is that big anchor in the E-lab: weighs a ton, very power-hungry, requires opening pressure case and soldering a shunt on a circuit board to switch range, only used with the FSI ICTD...
SeaPoint | new: | 2001-03-01
S/N 2336 | gain: | 30 x, 0-5 ug/L | for SBE25-0293 on SWL
max.depth | offset: | 0
6000m |  |  | 2002: Note: a Y-cable was used in 2002 was modified 10 Aug 2002 on LSL to float (open) pin# 5 & 6 Fluoro range was 30X (i.e. pin 5&6 open)
SeaPoint | new: | 2003 ?
S/N 2745 | gain: |  | on SBE19P-4132 (Apr'06)
max.depth | offset:
6000m
SeaPoint | new: | 2004 ? | SEAPOINT GAIN is controlled by cable: setting pin #5 and/or #6 either open (=5V) or short-to-ground (=0V) see manual.
S/N 2569 | gain: | 30 x, 0-5 ug/L
max.depth | offset: | 0
6000m
SeaPoint | new: | 2006 ? (Bill W) | for use on NAHIDIK
S/N 2805 | gain: | 30 x, 0-5 ug/L
max.depth | offset: | 0
6000m
SeaPoint | new: | 2006-10-01 | new Fall 2006
S/N 2841 | gain: | 30 x, 0-5 ug/L
max.depth | offset: | 0
6000m
Wetlabs | date: |  | DGS modified the SBE19-1030 to power this unit, was used in 2005 for surveys of Okanagan Lake,
Wetstar Chla | Vblank= | 0.068 | Sept-'02: shipped to SEDNA for use on SBE19-3112 (UBC loan, Seabird checked that it was modified to take Wetstar fluorom)
WS3S521P | Scale fact= | 14
April 2006 BVH: The old Wetlabs Wetstar is on loan to the SEDNA Antarctic for use with the Ingram SBE19 until they return (in about another year). They had serious problems with that CTD and have bought another one. I asked them by email about their science program or CTD data but have not had any replies recently.
Note on C10: BvH: RANGE SETTING: Y-cable used in 2002 on SBE9/11 connected Fluorometer AND Transmissom; modified 10 Aug 2002 to float pin 5&6 (30X)
Note on C21: BvH: GAIN is controlled by cable !
Note on C26: BvH: GAIN is controlled by cable !
Note on C37: BvH: convert fluorescense into chl-a concentration

## Transmissometers
TRANSMISSOMETERS=SeaTech |  |  |  | (added April 2002 from separate spreadsheet)
serial | cal.date | V-air | lab | V-ref | blank | Comments | Guildline: |  | Seabird:
number | mm/dd/yy | voltage | cal. |  | voltage |  | Offset | Slope | M= | B=
Wet-Labs: |  | A | B | 90.2 | Z | 600 nm,25cm path,6000m
CST-662DR | 2003-03-20 | 4.86 | 4.86 | 4.772 | 0.054 | Tcal=19.9 Trans%= |  |  | 18.9019 | -1.0207
CST-1047DR | 2007-02-07 | 4.767 | 4.676 | 4.648 | 0.057 |  | (new) |  | 19.6672 | -1.121 | (Bill for Nahidik)
CST-1052DR | 2007-04-27 | 4.741 | 4.741 | 4.6434 | 0.058 |  | (new) |  | 19.3984 | -1.1251 | (for SWL)
|  | Factory | Cal. |  | Cal | Pmax=? Pathlength? | Guildline: |  | Seabird:
| mm/dd/yy | V-air | V-air |  | V-block | comments | Offset | Slope | M= | B=
SeaTech: |  | A | B |  | Z
23 | 1980-08-06 | 4.79 |  | 0.003 |  | factory
23 | 1995-07-12 |  | 4.439 |  | 0.003 | JL
26 | 1980-08-06 | 4.79 |  | -0.007 |  | factory
26 | 1995-07-12 |  | 4.3 |  | 0.004 | field, JL
28 | 1980-08-06 | 4.74 |  | -0.007 |  | factory
28 |  |  | 4.63 |  | 0.002 | JL | -2.0475 | 102.3349
78 | 1981-06-03 | 4.77 |  | -0.001 |  | factory
78 | 1988-02-02 |  | 4.576 |  | 0.0017 | lab | -2.0848 | 104.2041
78 | 1989-03-20 |  | 4.5468 |  | 0.0007 | lab | -2.0982 | 104.8943
78 | 1990-03-16 |  | 4.5455 |  | 0.0008 | lab | -2.0988 | 104.9222
78 | 1992-05-11 |  | 4.5262 |  | 0.0011 | lab | -2.1077 | 105.3632
78 | 1994-04-01 |  |  |  |  | failed at 400m
81 | 1981-06-03 | 4.81 |  | -0.002 |  | factory
81 | 1988-02-02 | 4.8101 |  | 0.0007 |  | new factory value after repair
81 | 1989-03-20 |  | 4.8578 |  | 0.0008
81 | 1989-08-22 |  |  |  |  | hit by Russian, meter and trans not returned
Seabird CON file:
109 | 1982-05-05 | 4.748 | 4.748 | 0 | 0 | factory (Bon) |  |  | M= | B=
109 | 1997-10-30 |  | 4.392 |  | 0.001 | to SHEBA (on sheet from DGS) |  |  | 21.6261 | -0.021626
130 | 1982-11-12 | 4.734 | 0.001 |  |  | EDR 004547A
| 1990-04-09 |  | 4.622 |  | 0.0015
132 | 1982-10-23 | 4.725 |  | 0.002 |  | factory
132 | 1995-07-01 |  | 4.31 |  | 0.003 | JL | -2.1926 | 109.563
136 | 1993-10-06 | 4.75 |  | 0 |  | factory |  |  | Seabird CON file:
139 | 1982-12-02 | 4.74 | 4.74 | 0.003 | 0.003 | factory |  |  | M= | B=
139 | 1993-08-31 |  | 4.74 |  | 0.009 | RP |  |  | 20.0254 | -0.18023
139 | 1998-09-17 |  | 4.434 |  | 0.004 | Bon |  |  | 21.386 | -0.085544
139 | 2000-01-15 |  | 4.562 |  | 0.006 | Doug |  |  | 20.7946 | -0.12477
139 | 2001-04-23 |  | 4.554 |  | 0.006 | Doug |  |  | 20.8311 | -0.12499
139 | 1905-06-29 |  |  |  |  | Mike? |  |  | #DIV/0! | #DIV/0!
163 | guess only: | 4.746 |  | 0.002 |  | need factory values!
163 | 2000-01-15 |  | 4.564 |  | 0.001 | Doug | from where? ---> |  | 20.7945 | -0.12477
163 | fomula test: |  | 4.564 |  | 0.001 |  |  |  | 20.7933 | -0.020793
163 | 2006-07-12 |  | 4.656 |  | 0 | Ron Lindsay |  |  | 20.378 | 0
175 | 1985-09-24 | 4.803 |  | 0 |  | factory
175 | 1988-08-01 |  |  |  |  | smashed by fishing boat
176 | 1985-09-25 | 4.728 |  | -0.001 |  | factory
176 | 1988-01-01 |  |  |  |  | flooded
197 | 1985-10-08 | 4.749 | 4.749 | 0 | 0 | factory - EDR 007333A
197 | 1988-02-02 |  | 4.6697 |  | 0.0007 | side rods look slightly bent | -2.034 | 101.6839
197 | 1988-09-07 |  | 4.6452 |  | 0.0012 | lab | -2.0447 | 102.21
197 | 1988-11-15 |  | 4.6336 |  | 0.0011 | lab | -2.0498 | 102.468
197 | 1988-11-22 |  | 4.637 |  | 0.0009 | lab | -2.0483 | 102.3969
197 | 1989-03-20 |  | 4.6283 |  | 0.0008 | lab | -2.0522 | 102.5915
197 | 1990-03-16 |  | 4.5724 |  | 0.0009 | lab | -2.0772 | 103.8436
197 | 1990-09-25 |  | 4.5476 |  | 0.0015 | lab | -2.0886 | 104.3974
197 | 1992-05-07 |  | 4.4863 |  | 0.0018 | lab | -2.1171 | 105.8175
197 | 1994-04-04 |  | 4.39 |  | 0.002 | JL | -2.1636 | 108.1344
197 | 1996-02-08 |  | 4.209 |  | 0.002 | lab, BGM | -2.2566 | 112.7845
197 | 1996-10-24 | 4.73 | 4.73 | -0.001 | -0.001 | factory calibration | -2 | 100
197 | 1997-02-07 |  | 4.74 |  | 0.005 | lab, Doug | 1.9958 | 99.6892
231 | 1992-02-17 | 4.734 |  | 0.0001 |  | factory
231
231 | 1993-03-29 |  | 4.661 |  | -0.0004 | lab, JL | -2.0313 | 101.5743
231 | 1994-08-13 |  | 4.61 |  |  | JL
231 | 1995-04-29 |  | 4.56 |  | 0 | field, JL | -2.0763 | 103.8158
231 | 1996-02-08 |  | 4.511 |  | 0.012 | lab, BGM, unstable with large zero offset | -2.0989 | 104.6916
240 | 1992-12-29 | 4.674 |  | 0 |  | factory
241 | 1986-12-31 | 4.754 |  | 0.001 |  | factory
241 | 1992-02-17 |  | 4.669 |  | 0.0025 | lab, JL, sent to Russia.
242 | 1987-01-01 | 4.765 | 4.765 | 0.001 |  | factory 25cm path |  |  | 19.9958 | 0
242 | 1993-06-21 |  | 4.82 |  | 0.0006 |  |  |  | 19.7701 | -0.011862
242 | 1994/08/00 |  | 4.76 |  | 0 |  |  |  | 20.0168 | 0
242 | 1995-01-23 |  | 4.76 |  | 0 | Ron Cooke (?) |  |  | 20.0168 | 0
242 | 96? note found… | [7 empty] | 22.58 | -0.027
254 | 1987-06-23 | 4.755 | 4.755 | 0.003 |  | factory
254 | 1989-03-20 |  | 4.754 |  | 0.0031 | lab | -2.0004 | 99.959
254 | 1990-03-16 |  | 4.7603 |  | 0.0029 | lab | -1.9978 | 99.8307
254 | 1990-09-30 |  | 4.7462 |  | 0.0018 | lab | -2.0037 | 100.1493
254 | 1992-05-11 |  | 4.7516 |  | 0.0022 | lab | -2.0014 | 100.0275
254 | 1993-10-16 |  | 4.73 |  | -0.002 | lab | -2.0106 | 100.5688
254 | 1996-02-08 |  | 4.708 |  | 0.004 | lab, BGM | -2.02 | 100.9175
571 | 1993-03-04 | 4.673 |  | 0 |  | factory
572 | 1993-03-04 | 4.674 |  | 0 |  | factory
572 | 1994-08-13 |  | 4.58 |  |  | field, JL | -2.041 | 102.0524
598 | 1994-03-21 | 4.662 | 4.662 | 0.002 | 0.002 | factory
598 | 1994-04-02 |  | 4.65 |  | 0.001 | lab, JL | -2.0052 | 100.238 | 20.0473 | -0.020047
598 | 1995-04-02 |  | 4.63 |  | 0.002 | JL, date is a guess only!!! | -2.0138 | 100.6509 | 20.1383 | -0.040277
598 | 1996-02-13 |  | 4.622 |  | 0.004 | lab, BGM | -2.0173 | 100.7847 | 20.1819 | -0.080728
598 | 1999-06-21 |  | 3.86 |  | 0.007 | lab,DGS | -2.4155 | 120.6081 | 24.1889 | -0.16932
598 | 2000-01-15 |  | 4.029 |  | 0.002 | Doug | -2.3142 | 115.6648 | 23.1438 | -0.046288
598 | 2001-01-16 |  | 3.922 |  | 0.002 | Doug | -2.3774 | 118.8204 | 23.7755 | -0.047551
182D | 1992-07-07 | 4.672 | 4.672 | 0 |  | factory, 5000 metres | -2 | 100
182D | 1994-04-06 |  | 4.5829 |  | 0.0005 | lab | -2.0389 | 101.934
182D | 1996-02-05 |  | 4.33 |  | 0 | lab, BGM, 28448,-103.11, WOCE ctd | -2.158 | 107.8984 | 28448 | -103 | 0.32937 | 0.0031978
182D | 1996-02-08 |  | 4.323 |  | 0.001 | lab, BGM | -2.1615 | 108.0515
182D | 1996-06-10 |  | 4.332 |  | 0 | lab, BGM | -2.157 | 107.8486
182D | 1996-07-01 |  | 4.305 |  | 0.0002 | lab, Doug | -2.1705 | 108.5206
182D | 1996-08-07 |  | 4.29 |  | 0 | lab, bgm, 28136, -137 | -2.1781 | 108.9044 | 28136 | -137 | 0.4424 | 0.0032292
182D | 1996-11-13 |  | 4.682 |  | -0.001 | factory calibration
182D | [5 empty] | factory rebuild
mm/dd/yy
192D | 1993-02-18 | 4.675 |  | 0 |  | factory, 5000 metres
192D | 1994-04-06 | 4.6967 | 4.6967 | 0.0003 |  | new factory value, 18/2/93 value not aged.
192D | 1994-06-30 |  | 4.6612 |  | 0 | JL | -2.0152 | 100.7616
192D | 1995-02-05 |  | 4.59 |  | 0 |  | -2.0465 | 102.3246
192D | 1995-04-27 |  | 4.559 |  | 0.0038 | 29936, -108 | -2.0604 | 102.9421
192D | 1995-07-01 |  | 4.52 |  | 0.001 | field, JL | -2.0782 | 103.8885
192D | 1996-02-05 |  | 4.433 |  | 0 | lab, BGM, 29132,-99.34,WOCE CTD |  |  | 29132 | -99 | 0.30922 | 0.0031234
192D | 1996-02-08 |  | 4.4478 |  | 0.0018 | lab, BGM | -2.1119 | 105.558
192D | 1996-06-06 |  | 4.258 |  | 0.002 | lab, BGM | -2.2061 | 110.2588
192D | 1996-07-18 |  | 4.36 |  | 0.002 | lab, Doug | -2.1544 | 107.6794
192D | 1996-08-07 |  | 4.36 |  | 0.001 | lab, bgm, 28652, -129, woce ctd | -2.1544 | 107.7009 | 28652 | -129 | 0.40912 | 0.0031715
192D | 1996-10-31 |  | 4.66 |  | 0 | factory checkup and calibration | -2 | 100
V=(A/B*(X-Z)
V=corrected output voltage
A=air calibration, factory
B air calibration, present
X=data value (output voltage measured in water
Z=zero offset with light path blocked
Note on E5: BvH: Tw - transm in pure water for 660 nm = 90.2 (to derive M,B)
For Guildline, slope = -2.0*(A/B)
offset = 20.0*(A/B*(5.00-zero offset)
For SBE-19: M = 20(A0 - Y0)/(A1 - Y1)
B = -M Y1
A0 & A1 are the original & present AIR CALIBRATION voltages
Y0 & Y1 are the original & present BLOCKED PATH voltages
For Guildline, slope = -2.0*(A/B)
offset = 20.0*(A/B*(5.00-zero offset)
For SBE-19: M = 20(A0 - Y0)/(A1 - Y1)
B = -M Y1
A0 & A1 are the original & present AIR CALIBRATION voltages
Y0 & Y1 are the original & present BLOCKED PATH voltages

## Turbidity
Turbidity Meter
Manufact: | Seapoint
S/N: | # 10133
range: | selected | range select by grounding Pin#5 &/or 6 (same as fluorometer)!
by cable !
max depth?
Manufact: | Seapoint | new 03/23/07
S/N: | # 11074
range: | selected | range select by grounding Pin#5 &/or 6 (same as fluorometer)!
by cable !
max depth?

## PAR
PAR SENSORS |  | Notes: | * annual calibr recommended (to NIST std)
*clean collector frequently with alcohol
* calibr performed with customer cable when available
model | S/N | cal.date | Std.lamp | Lamp date | Excitation voltage VDC | Calibr. Factor |  | Calibr. Factor
[6 empty] | quanta/cm2.s per Volt |  | uEinstein/cm2.s per Volt
QSP-2300 | 70123 | 2007-03-13 | 91537 | 2005-04-12 |  | Dry: | 3170000000000 | 5.26e-06
Nahidik | [5 empty] | Wet: | 3330000000000 | 5.53e-06
[6 empty] | V/(quanta/cm2sec) |  | V/(uE/cm2sec)
QSR-2200 | 20279 | 2007-03-13 | 91537 | 2005-04-12 |  | Dry: | 1.05e-17 | 6.35
quanta/cm2.s per Volt
QSP-2200 | 20280 | 2007-04-06 | 91537 | 2005-04-12 | 6 to 18 | Dry: | 56497175141242896 |  | cal volts= | 6 | VDC
Laurier | [5 empty] | Wet: | 95238095238095296 |  | probe current= | 4 | mA
[9 empty] | probe output illuminated= | 157.6 | mV
[9 empty] | probe output dark= | 1.8 | mV
[9 empty] | probe net response= | 155.9 | mV
[9 empty] | corrected lamp output in air: | 8810000000000000 | quanta/cm2.sec
[9 empty] | output corrected for immersion in water: | 14830000000000000 | quanta/cm2.sec
[6 empty] | V/(quanta/cm2sec) |  | V/(uE/cm2sec)
QSR-2200 | 20281 | 2007-04-09 | 91537 | 2005-04-12 | 6 to 18 | Dry: | 9.3093e-18 | 5.606 | cal volts= | 6 | VDC
[9 empty] | probe current= | 4 | mA
[9 empty] | probe output illuminated= | 83.7 | mV
| Surface light (uEinstein/m2.sec) = volts * conversion factor | [7 empty] | probe output dark= | 1.7 | mV
| Conversion factor = Output in air / probe net response | [7 empty] | probe net response= | 82 | mV
| Corrected PAR = 100 * ratio multiplier * underwater light/ surface light | [7 empty] | corrected lamp output in air= | 8810000000000000 | quanta/cm2.sec
[9 empty] | or = | 0.01463 | uEinstein/cm2.s
[9 empty] | CONVERSION FACTOR= | 1784.1463 | uEinstein/m2.s/V

## Altimeter
Altimeter
New: | 2005-03-31 | Used with SBE911+ on LSSL 2005-04
Manufact: | Benthos
Model: | Datasonics PSA-916D
S/N: | 1161
range: | 0-6000m
New: | 2007-02-12
Manufact: | Benthos
Model: | Datasonics PSA-916D
S/N: | 40853
range: | 0-6000m
New: | 2/12/2007 ?
Manufact: | Benthos
Model: | Datasonics PSA-916D
S/N: | 41098 | Scl= | 15
range: | 0-6000m | Offs= | 0
Note on C22: BvH: Scale factor =FSvolt*300/FSrange

## ICTD-1329
ICTD CALIBRATION INFO
S/N = 1329
VER = 2.4
CDATE = 21MAY96
A01 = 8.287852e-02
A501 = 3.743052e+01
A1001 = 7.482841e+01
B01 = 1.179096e-05
B501 = 1.909808e-05
B1001 = 3.347727e-05
C01 = 3.858695e-10
C501 = 1.037280e-08
C1001 = 1.091776e-08
A1 = 1.336248e-02
B1 = 9.487742e-01
C1 = -5.103178e-05
D1 = 4.515775e-07
A02 = 3.592097e+01
A502 = 1.567107e+01
A1002 = -4.478627e+00
B02 = -1.651400e-05
B502 = -1.823800e-05
B1002 = -2.638600e-05
C02 = -1.925820e-08
C502 = -1.977630e-08
C1002 = -1.400120e-08
A2 = -5.716996e-04
B2 = 1.001717e+00
C2 = -1.631159e-04
D2 = 3.604622e-06
A03 = -6.070145e-01
A503 = 1.444939e+01
A1003 = 2.955527e+01
B03 = 1.614800e-04
B503 = 2.103200e-04
B1003 = 8.756400e-05
C03 = -2.800560e-08
C503 = 1.952680e-07
C1003 = 1.781560e-07
A3 = -1.995102e-02
B3 = 1.050521e+00
C3 = -4.940318e-03
D3 = 1.108796e-04
A04 = 3.554498e+01
A504 = 1.530000e+01
A1004 = -4.763292e+00
B04 = -6.646200e-05
B504 = -1.238300e-05
B1004 = -3.009600e-05
C04 = -8.802370e-09
C504 = -1.024190e-08
C1004 = -8.208591e-09
A4 = 0.000000e+00
B4 = 1.000000e+00
C4 = 0.000000e+00
D4 = 0.000000e+00
A05 = 6.951330e+03
A505 = 3.515259e+03
A1005 = 6.139834e+01
B05 = -1.024400e-02
B505 = 5.197400e-04
B1005 = 1.160800e-02
C05 = 3.284700e-06
C505 = 7.119200e-07
C1005 = 5.978300e-06
A5 = 8.020587e-01
B5 = 1.000720e+00
C5 = -3.663700e-07
D5 = 3.187000e-11
PLTC = 4.000000e+02
TRTC = 1.000000e+02
XOFF = 0.000000e+00
XGAN = 1.000000e+00
CH1 = on, CH2 = on, CH3 = on, CH4 = off, CH5 = on, EXTD = off
DC1=on, DC2=off, DC3=on, DC4=on, DC5=off, DC6=off, DC7=off, DC8=off, Alt=off
Bytes/Scan = 16, Max Bytes/Scan allowed = 122

## MCTD-1357
WHOSE INSTRUMENT??
CALIBRATION INFO
original MCTD
S/N = 1357
VER = E.2
CDATE = 26JUN95
A01 = 7.604567e+01
A501 = 3.805856e+01
A1001 = 1.528660e-01
B01 = 1.195773e-05
B501 = 5.502455e-06
B1001 = 2.855086e-06
C01 = 2.815857e-09
C501 = 1.756495e-09
C1001 = 2.061360e-09
A1 = 9.251564e-03
B1 = 9.508290e-01
C1 = -8.058304e-05
D1 = 6.863835e-07
A02 = 3.914346e-01
A502 = 1.548280e+01
A1002 = 3.048322e+01
B02 = -3.294100e-06
B502 = 3.247900e-06
B1002 = 6.598900e-06
C02 = -9.607980e-09
C502 = -2.778820e-09
C1002 = 1.582310e-10
A2 = -2.134496e-02
B2 = 1.048333e+00
C2 = -4.658864e-03
D2 = 1.033809e-04
A03 = 4.556631e+00
A503 = 1.540935e+02
A1003 = 2.992850e+02
B03 = 8.327920e-04
B503 = -2.285100e-05
B1003 = -6.088090e-04
C03 = -3.486200e-07
C503 = -3.085400e-07
C1003 = -3.163300e-07
A3 = -2.507200e-01
B3 = 1.011834e+00
C3 = -9.197901e-05
D3 = 1.778900e-07
XOFF = 0.000000e+00
XGAN = 1.000000e+00

## MCTD-1534
Pmax = ???
CALIBRATION INFO | CALIBRATION INFO | Date: 21 Mar 1996 | Date: 17 May 1996 | Date: 15 Mar 1999
original MCTD | upgraded MCTD | new 1400 dBAR transducer | 7000 dBAR TRANSDUCER | with Psensor for 7000m | RETURNED TO FSI FOR CALIBRATION: Spring 2004
(later modified to S/N 1534) | (new electronics) | Pressure Calibration | Pressure Calibration |  | damaged by FEDEX (broken-off C-sensor)
S/N = 1344 | S/N = 1533 | S/N = 1534 | S/N = 1534 | S/N = 1534 | repair estimate ~US$ 4000.-
VER = E.2 | VER = H.7 | VER = H.8 | VER = H.8 | VER = I.0 | Asked FSI to return unit without repair for spare parts of other MCTD
CDATE = 15JUN95 | CDATE = 3APR96 | CDATE = 21MAY96 | CDATE = 17MAY96 | CDATE = 17FEB99
A01 = 7.619018e+01
A01 = 7.481707e+01 | A01 = 7.687679e+01 | A03 = -4.503574e+00 | A03 = 3.063934e+01 | A501 = 3.819031e+01
A501 = 3.740881e+01 | A501 = 3.846718e+01 | A503 = 7.940275e+02 | A503 = 3.755900e+03 | A1001 = 2.476492e-01
A1001 = 1.280668e-01 | A1001 = 2.628167e-01 | A1003 = 1.564395e+03 | A1003 = 7.257228e+03 | B01 = 7.901935e-06
B01 = 1.478333e-05 | B01 = 1.196468e-05 | B03 = 3.125000e-04 | B03 = -3.371500e-03 | B501 = 1.822803e-06
B501 = 5.057728e-06 | B501 = 5.900900e-06 | B503 = 4.421800e-03 | B503 = 3.095300e-02 | B1001 = -7.228598e-07
B1001 = 2.013860e-06 | B1001 = 6.645934e-06 | B1003 = 7.693600e-03 | B1003 = 5.611130e-02 | C01 = 2.239277e-09
C01 = 1.412444e-09 | C01 = -5.000634e-09 | C03 = 1.738800e-07 | C03 = 2.287600e-06 | C501 = 1.266642e-09
C501 = 4.282004e-10 | C501 = -3.119247e-09 | C503 = 9.459300e-07 | C503 = 8.474700e-06 | C1001 = 7.272759e-10
C1001 = 7.991649e-10 | C1001 = -3.288075e-09 | C1003 = 1.536800e-06 | C1003 = 1.314800e-05 | A1 = 2.321624e-02
A1 = 8.645451e-03 | A1 = 7.837162e-03 | A3 = -2.187000e-01 | A3 = -2.513260e+00 | B1 = 9.398204e-01
B1 = 9.586888e-01 | B1 = 9.228891e-01 | B3 = 1.001102e+00 | B3 = 1.004783e+00 | C1 = -3.802821e-05
C1 = -9.121145e-05 | C1 = -7.983520e-05 | C3 = -1.605100e-06 | C3 = -2.033500e-06 | D1 = 3.344590e-07
D1 = 7.672685e-07 | D1 = 6.977060e-07 | D3 = 6.579800e-10 | D3 = 1.925600e-10 | A02 = 6.706564e-01
A02 = 6.443921e-01 | A02 = -4.330795e-03 |  |  | A502 = 1.591774e+01
A502 = 1.588781e+01 | A502 = 1.527188e+01 |  |  | A1002 = 3.108448e+01
A1002 = 3.106772e+01 | A1002 = 3.054178e+01 |  |  | B02 = -2.317800e-05
B02 = -1.903600e-05 | B02 = -1.825400e-05 |  |  | B502 = -2.056000e-05
B502 = -1.631500e-05 | B502 = -1.721800e-05 |  |  | B1002 = -1.702900e-05
B1002 = -1.194100e-05 | B1002 = -1.483700e-05 |  |  | C02 = -2.821210e-10
C02 = -8.786180e-10 | C02 = 8.165930e-10 |  |  | C502 = -1.180450e-10
C502 = -4.897190e-10 | C502 = 1.496950e-09 |  |  | C1002 = -1.890760e-09
C1002 = 6.559269e-11 | C1002 = 1.955320e-09 |  |  | A2 = -2.016589e-02
A2 = -2.136674e-02 | A2 = -1.911568e-02 |  |  | B2 = 1.047491e+00
B2 = 1.049017e+00 | B2 = 1.047338e+00 |  |  | C2 = -4.637278e-03
C2 = -4.711407e-03 | C2 = -4.578197e-03 |  |  | D2 = 1.035693e-04
D2 = 1.043942e-04 | D2 = 1.018532e-04 |  |  | A03 = 3.063934e+01
A03 = 5.089466e+02 | A03 = 5.244594e+01 |  |  | A503 = 3.755900e+03
A503 = 4.173482e+03 | A503 = 8.477444e+02 |  |  | A1003 = 7.257228e+03
A1003 = 7.684304e+03 | A1003 = 1.621784e+03 |  |  | B03 = -3.371500e-03
B03 = -1.265490e-02 | B03 = -2.416600e-03 |  |  | B503 = 3.095300e-02
B503 = -6.529770e-02 | B503 = 1.287500e-03 |  |  | B1003 = 5.611130e-02
B1003 = -1.163910e-01 | B1003 = 3.583330e-03 |  |  | C03 = 2.287600e-06
C03 = -3.709300e-06 | C03 = 1.036400e-06 |  |  | C503 = 8.474700e-06
C503 = -2.094900e-05 | C503 = 1.774100e-06 |  |  | C1003 = 1.314800e-05
C1003 = -3.664200e-05 | C1003 = 2.349800e-06 |  |  | A3 = -2.513260e+00
A3 = 5.685000e+00 | A3 = 3.113400e-01 |  |  | B3 = 1.004783e+00
B3 = 1.008536e+00 | B3 = 1.003375e+00 |  |  | C3 = -2.033500e-06
C3 = -3.697700e-06 | C3 = -6.469000e-06 |  |  | D3 = 1.925600e-10
D3 = 3.451000e-10 | D3 = 2.844000e-09 |  |  | XOFF = 0.000000e+00
XOFF = 0.000000e+00 | XOFF = 0.000000e+00 |  |  | XGAN = 1.000000e+00
XGAN = 1.000000e+00 | XGAN = 1.000000e+00
Open Mode

## MCTD-1598
Pressure Calibration
3400?? dBAR TRANSDUCER
(5000 PSIA? ...not on FSI pr.cal printout)
MCTD
MCTD S/N = 1598
VER = ??
CDATE = 29SEP98
A03 = 168.069911
A503 = 1811.033911
A1003 = 3461.290496
B03 = -10428.738977D-06
B503 = 84312.562567D-07
B1003 = 23003.272361D-06
C03 = 38760.727267D-11
C503 = 27665.705167D-10
C1003 = 52381.759251D-10
(--different page in printout: )
A3 = -20.072216E-02
B3 = 10.125706E-01
C3 = -11.101592E-06
D3 = 21.629247E-10

## SBE19-1031 LostAtSheba
SBE19-1031 |  | Pump: SBE5-01
LOST at SHEBA - 1998/08 (with transmissometer and borrowed pump… )
box with manual, download cable, pump, etc. should have been returned to FWI?!
Temperature | date:
# 1031 | a=
b=
c=
d=
fo=
Conductivity | date:
# 1031 | a=
b=
c=
d=
m=
Pressure | date:
# 154898 | A0
(300 psia) | A1
A2
Pressure | date:
# 177965 | A0
(1000 psia) | A1
A2
Pressure | date:
# 184833 | A0
(3000 psia) | A1
A2
Pressure | date:
# 189759 | A0
(5000 psi) | A1
A2
SERVICE: