## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2007-36

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Project: Broughton Salmon and Sea Lice

Area: Broughton Archipelago/Knight Inlet

Chief Scientist: Hargreaves B.

Platform: Clupea

Date: 31 March 2007 – 5 April 2007

Processed by: Germaine Gatien

Date of Processing: 21 December 2007 – ??

Number of original CTD casts: 86

Number of casts processed: 85

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19+ SEACAT CTD (S/N 4345) was used; it was mounted with a SBE 43 dissolved oxygen sensor, (S/N 0766).

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available. The times in the log book are in Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

A study of another SBE 19+ CTD during 2007-11 suggests that a soak of 2-3 minutes may lead to better salinity data, especially if the instrument experiences large temperature changes as it goes into the water. For this cruise the CTD soaked for less than a minute for most of the casts, though the first cast had a very long soak which is a good idea after long storage. It is likely that there was not a large difference between air and sea temperatures at this time of year, so the potential error in salinity is likely to be small.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard so were adjusted. The proper names were obvious from the original names.
The dates in the log book are in Pacific Daylight Time.

Configuration file 4345.con contains the calibration coefficients used in acquisition of this data. This is the first known processing of data using these sensors on a profiling instrument. All coefficients were checked and errors in the conductivity were corrected and the file was then saved as 2007-36-ctd.con.

- Results of inter-comparison study of an SBE19 and an SBE 911+ during 2007-11

- Leaving alignment of the temperature and conductivity to later in the processing is a better idea. The tests are easier to run in IOSSHELL and alignment can be done using the SHIFT routine.

- The SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature. This may matter most for the first cast after being stored.

- Conversion of Raw Data

- The raw data were converted using conversion file 2007-36-ctd.con. A few casts were checked and no problems were found. The pressure looks good with conductivity rising sharply between 0.2db and 0.45db during the downcasts and going down rapidly between 0.3db and -0.05db of the upcast.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to add details.

The time zone given in the headers is UTC, but the times actually agree with the times in the log which are in PDT.

8. Checking Headers

Track plots and header checks turned up no problems.

A header summary was produced and details were checked against the log records; errors in positions and station names were found and corrected. The dates of the first 2 casts were wrong.

Track plots were added to the end of this report.

The surface check was run. The average surface pressure is -0.05db, but all the surface salinity values are extremely low, so some of these data are probably from above the surface. The pressure offset applied looks appropriate and no further correction is needed.

ADD TIME CHANNEL was used to add 7 hours to the start times so they are in UTC.

9. Test Plots

During 2007-11 tests were done comparing an SBE19 with an SBE911+. It was concluded that the SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature. This may matter most for the first cast after being stored.

The soak times for this cruise were generally <l minute, but for the first cast the soak lasted about 2.5 minutes.

Profiles were plotted for all casts and up and down traces compared. There is an offset between upcasts and downcasts and upcasts are generally smoother as expected due to the sensors measuring in the wake of the CTD. No significant problems were noted. While there are some doubts about the first 5m of the downcasts for casts with short soak periods, there is no obvious evidence of upcast data being better.

10. SHIFT

Temperature

Generally the temperature needs aligning for the SBE19 with a setting of +0.5s given by SeaBird as a typical advancement relative to the other channels. Tests were done on 5 casts to see which setting made up and downcasts look most alike on a T-S plot and +2 or +2.5 records looked best, depending on the cast. So +2.25 was selected which is equivalent to about +0.56s.

Dissolved Oxygen

There is no signal in the dissolved oxygen voltage channel; all values are 3.1409. This data will not be processed further and should be removed later.

11. DELETE

The data from the first 4 seconds is unlikely to be useful as the sensors have not had time to equilibrate, but that data is likely to be selected by DELETE. So to reduce the editing needed, CLIP was used to remove the first 20 records from all casts.

Then DELETE was run using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There only warnings in the Delete log were that for casts #4 and #79  there is no data below 10db; the data files before DELETE was run contain maximum pressures of <10db for both those casts, so this is not a problem.

12. CTDEDIT

Only light editing was required.

CTDEDIT was used to remove bad surface records for most casts; some bottom records were removed and light cleaning of salinity was required for some casts. Notes were entered in the header for all casts that were edited.

The following casts required no editing: 5, 8, 11, 16, 18, 21, 25, 38, 58, 62, 63, 64,

For the following casts the downcast data looked bad so upcast data files were prepared and examined. The SHFT files were put through REVERSE and then DELETE to produce upcast files.

- The downcast data for #9 looked bad until about 30db; probably the pump did not come on until that point, or something was plugging the intake until pressure cleared it. The upcast data looks fine, so it was selected and edited lightly.

- The temperature and salinity data for #66 looks very different during downcast and upcast sections with a shift of 2 salinity units at the bottom; the upcast looks more like nearby casts but still looks bad with little variation. The cast was shallow (<30db) so if there was an obstruction the CTD probably did not go deep enough to clear it. This cast will not be processed further.

For a few casts (ex. #56) there were unstable features that were seen in both downcast and upcast; they do not appear to be due to instrumental noise, so were left unedited as they are probably real.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and a little more editing was done.

14. Intercomparisons

COMPARE – There were no calibration samples.

Previous Use of CTD – None.

Historic ranges – No local climatology is available.

15. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

There was no salinity calibration sampling; no recalibration was applied.

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

The standards check routine was run and no further problems found.

16. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2007-36

Dates: 	Start:	31 March 2007	    End:	5 April 2007

Location: Broughton Archipelago/ Knight Inlet                 Vessel:    CLUPEA

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19+ SEACAT | 4345 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/19 SEACAT / 4345

Cruise ID#:	2007-36

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4345 | 29 Nov. 2005 | Factory |  |  |
| Conductivity | 4345 | 29 Nov. 2005 | Factory |  |  |
| Pressure Sensor | 4345 | 06 Dec. 2005 | Factory |  |  |
| SBE43 DO sensor | 0766 | 4 Jan. 2007 | Factory |  |  |