## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2007-37

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Project: Broughton Salmon and Sea Lice

Area: Broughton Archipelago/Knight Inlet

Chief Scientist: Hargreaves B.

Platform: Clupea

Date: 24 April 2007 – 3 May 2007

Processed by: Germaine Gatien

Date of Processing: 2 January 2008 – 11 January 2008

Number of original CTD casts: 113

Number of casts processed: 36

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19+ SEACAT CTD (S/N 4345) was used; it was mounted with a SBE 43 dissolved oxygen sensor, (S/N 0766).

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available. The times in the log book are in Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

The bottom depth in the headers is taken from the log sheets, and should be taken to be the bottom depth at the beginning of the cast. Often sampling went deeper than the bottom depth, presumably because of drift into deeper water during the cast.

The only useful CTD data from this cruise are from casts 87 to 119, plus the upcast sections of casts 122, 130 and 138. All other casts have bad data, with unbelievable temperature and salinity values; this may be due to a corroded bulkhead connector. Similar problems were noted during 2007-38 which followed, and to a limited extent in 2007-39.

There was no signal in the dissolved oxygen channel; either the sensor was not actually present or was not connected properly.

A study of another SBE 19+ CTD during 2007-11 suggests that a soak of 2-3 minutes may lead to better salinity data, especially if the instrument experiences large temperature changes as it goes into the water. For this cruise the CTD was soaked for less than a minute for most of the casts, though the first cast had a very long soak, which is a good idea after long storage. It is likely that there was not a large difference between air and sea temperatures at this time of year, so the potential error in salinity is likely to be small.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform and gradients large. These errors are likely to be much larger than calibration errors. For this cruise, the large errors are only likely near the surface where errors may be as high as 0.05units.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard so were adjusted. The proper names were obvious from the original names.
The dates in the log book are in Pacific Daylight Time.

Configuration file 4345.con contains the calibration coefficients used in acquisition of these data. Cruises 2007-36 -37, -38 and -39 were the first ones using these sensors on a profiling instrument and there was no calibration sampling.

All coefficients were checked and errors in the conductivity were corrected and the file was then saved as 2007-37-ctd.con.

- Results of inter-comparison study of an SBE19 and an SBE 911+ during 2007-11

- Leaving alignment of the temperature and conductivity to later in the processing is a better idea. The tests are easier to run in IOSSHELL and alignment can be done using the SHIFT routine.

- The SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature. This may matter most for the first cast after being stored.

- Conversion of Raw Data

- The raw data were converted using conversion file 2007-37-ctd.con. A few casts were checked and no problems were found. The pressure looks good with conductivity rising sharply between 0.3db and 0.6db during the downcasts and going down rapidly between 0db and -0.3db of the upcast.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to add details.

The time zone given in the headers is UTC, but the times actually agree with the times in the log which are in PDT.

8. Checking Headers

Track plots and header checks turned up a few errors which were corrected.

A cross-reference listing was checked against the log records and a few more errors were found and corrected.

Track plots were added to the end of this report.

The surface check was run. The average surface pressure is -0.06db, but all the surface salinity values are extremely low, so some of these data are probably from above the surface. The pressure offset applied looks appropriate and no further correction is needed.

ADD TIME CHANNEL was used to add 7 hours to the start times so they are in UTC.

9. Test Plots

During 2007-11 tests were done comparing an SBE19 with an SBE911+. It was concluded that the SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature. This may matter most for the first cast after being stored.

The soak times for this cruise were generally <l minute, but for the first cast the soak lasted more than 3 minutes.

Profiles were plotted for all casts and up and down traces compared. There is an offset between upcasts and downcasts and upcasts are generally smoother as expected due to the sensors measuring in the wake of the CTD.

The data from cast #120 to the end mostly look bad with upcast and downcast temperature and salinity very different. There are occasional sections with closer agreement, particularly during deeper casts. Upcast files will be prepared and considered for archiving if they prove to be better than the downcasts.

10. SHIFT

Temperature

Generally the temperature needs aligning for the SBE19 with a setting of +0.5s given by SeaBird as a typical advancement relative to the other channels. For 2007-36 tests were done on 5 casts to see which setting made up and downcasts look most alike on a T-S plot and +2 or +2.5 records looked best, depending on the cast. Tests were run on 2 casts from this cruise (from early in the cruise when the data look similar for up and downcasts) and the results were the same.

SHIFT was run using +2.25 records which is equivalent to about +0.56s.

Dissolved Oxygen

There is no signal in the dissolved oxygen voltage channel; all values are 3.1409. That channel will not be processed further and should be removed later.

11. CLIP and REVERSE

The data from the first 4 seconds are unlikely to be useful as the sensors have not had time to equilibrate, but those data are likely to be selected by DELETE. So to reduce the editing needed, CLIP was used to remove the first 22 records from all casts. This is 2 more points than was done for 2007-36, since most casts had at least 2 useless points at the top for the earlier cruise.

At this point all files were put through REVERSE.

12. DELETE

DELETE was run using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

A second run of DELETE was done using the reversed files, and there were no warnings for those either.

13. CTDEDIT

As noted earlier there are many casts with data that are bad. Typically the downcast data are very clearly wrong with large swings in values, and unbelievable values. The upcasts are smoother, but suspiciously constant. Given there is no calibration sampling, it was decided to exclude any casts with bad downcasts, and no clear evidence that the problem cleared before the upcast.

There are several possible explanations for this problem. It has happened in the past that the turn-on for the pump was set to too high a conductivity value. This does not seem to be a good explanation for these data since the conductivity is not particularly low. Conductivity was fairly low for 2007-36 when no such problems were noted. It also does not explain why the setting would have changed part-way through the cruise; in fact, it occurs in the middle of a day’s sampling.

It is possible that the pump was not working properly, or something was interfering with the flow to the sensors. A clogged bleed valve or particles plugging the line might also clear at depth, but this does not explain some of the extreme values in T and S. Another suggestion is that there was corrosion in a bulkhead connector and this could lead to the sort of unbelievable values that are sometimes observed. The intermittence is consistent with the bulkhead connector explanation.

Files that will not be processed further are: 120-121, 123-129, 131-137 and 139-199.

The upcasts will be processed for: 122, 130 and 138. There is no way to confirm that the data are good, but they look reasonable. The SHFT files were put through REVERSE, then DELETE to produce upcast files.

CTDEDIT was used to remove bad surface records for most casts; some bottom records were removed and light cleaning of salinity was required for some casts. Notes were entered in the header for all casts that were edited.

All casts required only light editing except for casts #109 and #114 which required no editing.

14. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and a little more editing was done.

15. Inter-comparisons

COMPARE – There were no calibration samples.

Previous Use of CTD – The sensors were used during 2007-36, 2007-38 and 2007-39 but there was no calibration sampling from those cruises.

Historic ranges – No local climatology is available.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

There was no salinity calibration sampling; no recalibration was applied.

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

There was a dissolved oxygen sensor mounted on the CTD, but there was no

signal recorded so the channel was removed.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

d.) Groups of near-by casts were plotted together in T-S space and they look reasonably similar.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2007-37

Dates: 	Start:	24 April 2007	    End:	3 May 2007

Location: Broughton Archipelago/ Knight Inlet                 Vessel:    CLUPEA

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19+ SEACAT | 4345 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/19 SEACAT / 4345

Cruise ID#:	2007-37

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4345 | 29 Nov. 2005 | Factory |  |  |
| Conductivity | 4345 | 29 Nov. 2005 | Factory |  |  |
| Pressure Sensor | 4345 | 06 Dec. 2005 | Factory |  |  |
| SBE43 DO sensor | 0766 | 4 Jan. 2007 | Factory |  |  |