## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2007-38

Agency: PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.

Project: Broughton Salmon and Sea Lice

Area: Broughton Archipelago/Knight Inlet

Chief Scientist: Hargreaves B.

Platform: Walker Rock

Date: 23 May 2007 – 31 May 2007

Processed by: Germaine Gatien

Date of Processing: 5 January 2008 – 11 January 2008

Number of original CTD casts: 113

Number of casts processed: 35

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19+ SEACAT CTD (S/N 4345) was used; it was mounted with a SBE 43 dissolved oxygen sensor, (S/N 0766).

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available. The times in the log book are in Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

The bottom depth in the headers is taken from the log sheets, and should be taken to be the bottom depth at the beginning of the cast. Often sampling went deeper than the bottom depth, presumably because of drift into deeper water during the cast.

Only 35 casts contained useful data. All other casts have corrupted temperature and conductivity data with large variations, instabilities and unbelievable values. Upcasts were affected less than the downcasts. Data to be archived include 13 downcasts and 22 upcasts. Similar problems occurred during 2007-37 and to a lesser extent during 2007-39.

The CTD was soaked for less than a minute for most of the casts, though the first cast had a very long soak, which is a good idea after long storage. Study of this CTD during 2007-11 suggests that a longer soak (2-3 minutes) may lead to better salinity data.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform and gradients large. These errors are likely to be much larger than calibration errors. For this cruise, the large errors are only likely near the surface where errors may be as high as 0.05units.

There was no signal in the dissolved oxygen channel; either the sensor was not actually present or was not connected properly.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard so were adjusted. The proper names were obvious from the original names.
The dates in the log book are in Pacific Daylight Time.

Configuration file 4345.con contains the calibration coefficients used in acquisition of this data. Cruises 2007-36 -37, -38 and -39 were the first ones using these sensors on a profiling instrument and there was no calibration sampling.

All coefficients were checked and errors in the conductivity were corrected and the file was then saved as 2007-38-ctd.con.

- Results of inter-comparison study of an SBE19 and an SBE 911+ during 2007-11

- Leaving alignment of the temperature and conductivity to later in the processing is a better idea. The tests are easier to run in IOSSHELL and alignment can be done using the SHIFT routine.

- The SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature. This may matter most for the first cast after being stored.

- Conversion of Raw Data

- The raw data were converted using conversion file 2007-38-ctd.con. A few casts were checked and no problems were found. The pressure looks good with conductivity rising sharply between 0.3db and 0.6db during the downcasts and going down rapidly between 0db and -0.3db of the upcast.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

6. DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets.

During 2007-37 the temperature and salinity data from the second half of the cruise were bad and there is evidence of trouble with this cruise as well. To avoid wasting time in editing the headers of casts that won’t be used, that job will be done after the data has been edited. In the course of editing, decisions can be made about whether to use downcasts or upcasts or to reject the cast.

8. Test Plots

During 2007-11 tests were done comparing an SBE19 with an SBE911+. It was concluded that the SBE19 should be allowed to soak for at least 3 minutes and perhaps longer if it has been stored at a temperature that is very different from the sea temperature. This may matter most for the first cast after being stored.

The soak times for this cruise were generally about l minute but often significantly shorter.

Profiles were plotted for all casts and up and down traces compared. There are many casts with major differences between the traces that look like evidence of pump malfunction or other flow problems. In some cases the upcast data look reasonable as though whatever was causing the problem cleared due to flushing or pressure. Upcast files will be prepared since they are likely to be more useful than downcasts.

9. SHIFT

Temperature

Generally the temperature needs aligning for the SBE19 with a setting of +0.5s given by SeaBird as a typical advancement relative to the other channels. For 2007-36 and -37 tests were done to see which setting made up and downcasts look most alike on a T-S plot and +2 or +2.5 records looked best, depending on the cast, so a value of +2.25 records was used for the SHIFT routine. For this cruise there are many doubts about the data, so tests are not easy to run. For a few casts that look ok, a check was made that using +2.25 looked reasonable and it did improve the data.

SHIFT was run using +2.25 records which is equivalent to about +0.56s.

Dissolved Oxygen

There is no signal in the dissolved oxygen voltage channel; all values are 3.1409. These data will not be processed further and should be removed later.

10. CLIP and REVERSE

The data from the first 4 seconds are unlikely to be useful as the sensors have not had time to equilibrate, but those data are likely to be selected by DELETE. So to reduce the editing needed, CLIP was used to remove the first 21 records from all casts. This is 2 more points than was done for 2007-36, since most casts had at least 2 useless points at the top for the earlier cruise.

At this point all files were put through REVERSE.

11. DELETE

DELETE was run using the following parameters:

Surface Record Removal: Last Press Min      Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	       Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

A second run of DELETE was done using the reversed files, and there were no warnings for those either.

12. CTDEDIT

As noted earlier there are many casts with data that are bad. The same thing occurred during 2007-37. Typically the downcast data is marked by large unstable swings in values that suggest that the pump was not working. But there are also extreme temperature and salinity values (outside the expected range) and that is not characteristic of pump problems. The upcasts are sometimes bad too, but more often there are no large variations there; in fact, often the upcast data is suspiciously constant. Some of the upcasts look usable even when the downcast does not. The problem may be due to a corroded bulkhead connector.

The following downcast files were processed and edited: 206-210, 212-216, 218, 276 and 289.

The following upcast files were processed and edited: 211, 219, 221, 257, 258, 261, 263, 265-268, 271-273, 275, 277-281, 291 and 293.

The following files were rejected: 200-205, 217, 220, 222-256, 259, 260, 262, 264, 269, 270, 274, 282-288, 1290, 292, 294-299.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and a little more editing was done.

14. Checking Headers

The EDIT HEADERS routine was used to add the lines for positions and bottom depth and a text editor was used to add details. This was done for just the 35 casts that are to be archived.

The time zone given in the headers is UTC, but the times actually agree with the times in the log which are in PDT.

ADD TIME CHANNEL was used to add 7 hours to the start times so they are in UTC.

Track plots and header checks turned up a few errors which were corrected.

A cross-reference listing was checked against the log records and no further errors were found.

Track plots were added to the end of this report.

The surface check was run. The average surface pressure is +0.65db, which is higher than for 2007-37 and -37, but all the surface salinity values are also higher so presumably the CTD just started a little later, so that no “out of water” data was recorded. The pressure offset applied looks appropriate and no further correction is needed.

15. Inter-comparisons

COMPARE – There were no calibration samples.

Previous Use of CTD – The sensors were used during 2007-36, 2007-37 and 2007-39 but there was no calibration sampling from those cruises.

Historic ranges – No local climatology is available.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Descent_Rate and Flag.

The HEADEDIT routine was used to add a comment to the headers:

NOTE: The data in this file are from the DOWNCAST (or UPCAST as appropriate)

***********************************************************

There were problems with the temperature and salinity data for many casts

that may be due to a corroded bulkhead connector. Many cast files were

heavily corrupted and unusable.

For some the downcast section was corrupted but the upcast was better.

There was no salinity calibration sampling; no recalibration was applied.

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

d.) Because of the problems with this data special care was taken checking the final files. Groups of near-by casts were plotted together in T-S space and they look consistent with expectations.

CRUISE SUMMARY

Cruise ID#:	2007-38

Dates: 	Start:	23 May 2007	    End:	31 May 2007

Location: Broughton Archipelago/ Knight Inlet                 Vessel:    Walker Rock

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19+ SEACAT | 4345 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/19 SEACAT / 4345         Cruise ID#:	2007-38

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4345 | 29 Nov. 2005 | Factory |  |  |
| Conductivity | 4345 | 29 Nov. 2005 | Factory |  |  |
| Pressure Sensor | 4345 | 06 Dec. 2005 | Factory |  |  |
| SBE43 DO sensor | 0766 | 4 Jan. 2007 | Factory |  |  |