## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27 May 2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2007-46

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Effingham and Muchalat Inlets

Project: Effingham / Muchalat

Party Chief: Dallimore A.

Platform: Vector

Date: May 21, 2007 – May 28, 2007

Processed by: Germaine Gatien

Date of Processing: 3 October 2007 – 4 November 2007

Number of original CTD casts: 11	Number of CTD casts processed: 10

Number of original rosette casts: 11	Number of rosette casts processed: 10

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Chelsea/Seatech transmissometer (#723DR) and an Altimeter (#1233). A SeaBird dissolved oxygen sensor (#1176) and a Seapoint Fluorometer (#2356) with a 10X cable were mounted on the pumps. The deck unit was an SBE model 911 s/n 0619. The data logging computer was PAC 02588. The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

There were problems with the identification of salinity samples. Two samples could not be related to the rosette records and two were clearly mislabelled.

This was the first time this oxygen sensor was used. The calibration sampling was limited and the comparison to CTD data very noisy, but the errors small.

The SBE dissolved oxygen data in the CTD files should be considered, roughly:

•	±0.5ml/l from    0 -  50 db

•	±0.3ml/l from   50 - 175 db

•	±0.2ml/l below 175 db

The peak fluorescence was off-scale (>15ug/l) for casts 13 through 25.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained. Problems were noted with Niskin bottle #2 and with the altimeter.

Salinity data were available in spreadsheet format but had no flag column.

Dissolved oxygen files were available but lacked flags.

The nutrients were not yet available.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The CTD calibrations were checked, the offset and date of calibration was corrected for the pressure channel and the file was saved as 2007-46-CTD.con.

##### Conversion of Raw Data

The files were converted to CNV files. Plots were made of a few files.

Cast #1 contains only surface data with the pumps off and is not listed in the log as being a CTD cast, so it will not be processed further.

The temperature and conductivity channels are reasonably close during downcasts, but are frequently quite different during upcasts; the secondary conductivity has a lot of fine-scale noise in it.

The altimetry is quite noisy near the bottom for some casts, so entries will need to be checked carefully.

The fluorescence seems noisier than usual with large spikes even at depth.

Transmissivity and dissolved oxygen look as usual, with the usual hysteresis in the DO channel.

Bottle stops were all at least 30s.

Conductivity values are low at 2.5db with the pumps on, but given the area low surface salinity is likely so this is not considered evidence of errors in the pressure.

Rosette files were converted using a start time of -2s and duration of 5s.

All files were then converted to IOS HEADER format. The event numbers were missing in the headers so CLEAN was run to add those and produce BOT files. All BOT files were plotted and no significant outliers were found.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were:  Pass 1- Std Dev = 2	 Pass 2- Std Dev = 5	 Points per block = 50

##### CELLTM

Tests for CELLTM are difficult because there were no casts without stops in the upcast. The settings found appropriate when this equipment was used for 2007-26 were (0.03, 9) for the primary and no change for the secondary and for 2007-43 (0.03, 9) was applied to both channels. Tests were run on two casts using a variety of settings and (0.03, 9) does look best overall for both channels and both casts.

CELLTM was run using (0.03, 9) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Two casts were plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 26 | 350 | ~0.0004 | ~-0.0001 | ~-0.0005 | Moderate very steady |
| 27 | 350 | ~0.0004 | ~0.00005 | ~-0.0005 | Moderate, very steady |

The differences are small and there is no evidence of pressure dependence in the differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

A header check was run. It was found that the water depth had not been entered for one cast, so all were checked and the water depth was added to the headers of 3 casts.

The header summary was checked against the log book entries and no errors were found.

Track plots were produced and added to the end of this report.

The average surface pressure is ~1.6db which is a little low for the Vector. However, the salinity is suitably low for the pressures <2db, so the pressure offset looks appropriate. No adjustment will be made to the pressure.

The altimeter values were exported to a spreadsheet; all casts were examined since the altimetry was often noisy near the bottom. The algorithm was found to have worked well except for cast #27. For that cast the header was edited and a note entered to explain how the value was determined..

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There were inconsistencies between the data files, CTD Daily Log records and the rosette log sheets. The salinity samples are consistent with the rosette. he NUT results to make sure this is correct:

The file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data were in spreadsheet format but lacked a flag column. There was a comment column with no entries so it is assumed that no flags are required. There were some inconsistencies between the salinity spreadsheet and the information given on the rosette log sheets:

- Two samples were said to be sample #13 and #14 with labels EFF13 and EFF14. The log book and rosette sheets indicate that there was no CTD cast at EFF13 and at station EFF14 the salinity samples are said to be #8 and 10. The CTD salinity for the bottles labelled as samples #8 and 10 differ from the two samples by ~1 and ~0.8 salinity units. The SAMAVG files were all examined to see if there was any salinity data close to the two samples and there was not. So either there was a problem with the analyses or these samples may be from something other than the rosette – say surface samples. The two samples were removed from the spreadsheet 2007-46-sal.csv.

- For cast #25 the sample numbers given in the spreadsheet do not agree with those on the rosette sheets. The entries were edited to reflect the information from the rosette sheets, but “c” flags were entered and comments entered to show there is some possibility of confusion.

The spreadsheet file was then converted into individual SAL files.

The dissolved oxygen data were received in individual files without flag channel and comments. The rosette sheets include some comments. ADD files were created with flag and comment channels, and entries were made based on the notes on the rosette sheets. There were a few extra lines that had to be removed and based on the rosette notes changes were made to the flask number for one entry and the depth of another. There is an OXY file for cast #38 which was not a CTD cast. The only data point has value 0, so this is assumed to be an error and will not be processed further.

The nutrient data were provided in spreadsheet format with comments but no flag channel. The duplicates had not been averaged. After checking with the analyst it was determined that no flags were needed. The duplicates were averaged. Channel names were changed to standard format and the file was saved as 2007-46-nuts.csv. The data were converted to individual NUT files.

The SAL, ADD and NUTS files were merged with the CST files in three steps. (Output: MRG1, MRG3, MRG4.)

The MRG4 files were then merged with SAMAVG files (MRG) and then put through CLEAN to remove SeaBird headers and comments from the secondary files. (MRGCLN)

##### 11. COMPARE

Salinity

COMPARE was run. There is a lot of noise in the bottles from above 200db. Based on the 4 bottles from below 200db the CTD primary salinity is found to be low by an average of 0.0002 and the secondary low by 0.0005.

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD and gradually excluding outliers based on residuals, the fit was:

DOX_BOT = 1.0179 * DOX_CTD + 0.0355

The fit is not very convincing because there is not a clear division between outliers and points in the fit and dropping a single record makes a significant difference. There is only one point that is a clear outlier – when only that one is removed the fit is:

DOX_BOT = 1.0098 * DOX_CTD + 0.0395

When the differences are plotted against pressure and one outlier is excluded, the result is quite flat with corrections ranging from +0.05 at the surface to +0.064 at 350db.

When all bottles are included the dissolved oxygen is found to be low by an average of 0.035 and the median is low by 0.065ml/l; the range is from low by 0.38 to high by 0.51ml/l. When one outlier is exclude the average difference is -0.0045. These are small errors with most of the points between 0.3 and -0.3ml/l. It is clear that there is a slight bias to low values. Adding 0.0045ml/l to all values will bring the average difference ~0.0002ml/l.

There are no outliers significant enough to require the addition of flags.

Plots of DO vs Salinity show some slight outliers, but in each case the values are confirmed by bottles and reflect the unusual profiles in Muchalat Inlet with local DO maxima between 100 and 200db. The variability will limit the quality of this data as the sensor is probably not able to keep up to the changes.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors were used for 2007-24 (March 2007) but there was no calibration sampling. During 2007-43 (June 2007) the primary and secondary sensors appeared to be low by about 0.004 and 0.005 but most of the bottle stops were very short so the result was not really trusted. The primary sensors were selected for archiving for both cruises. The secondary sensors were also used for 2007-01 in January 2007 when the salinity was found to be high by 0.005 but it is believed that problems with the Autosal account for much of the difference and that it was really only high by 0.0004.

- The DO sensor was used during 2007-43. Most bottle stops were very short so the differences reflect both calibration and response time errors. The fit was:

DOX_BOT = 1.01351 * DOX_CTD + 0.0078

but the fit is expected to be different from this cruise during which the stops were long.

- The pressure sensor has been used often. An offset of +0.4db has been applied since June 2006.

Historic ranges – There were no excursions from the historic ranges where local climatology was available; most of the casts in Muchalat Inlet did fall outside the climatology below 175db, but the only information available was from a large-scale climatology, not representative of inlet conditions, so this is not considered indicative of instrumental problems.

##### 13. SHIFT

Conductivity

Tests were run on a few casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be +0.3 for the primary conductivity and +0.7s for the secondary.

SHIFT was run using +0.3 and +0.7s for the primary and secondary conductivity channels.

Dissolved Oxygen

Tests were run using settings from +40 to +140 records on 4 casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +120 records (5s).

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the fluorescence is unusually noisy and there were bottle stops confusing the issue. So the value usually used, +24 records, will be applied.

All data were shifted by +24 records. (Output: SHFFL)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: none

##### DETAILED EDITING

Based on the bottle calibration data the primary sensors look slightly better than the secondary. Looking at a few casts in detail shows that the secondary has a lot of fine-scale noise. Primary channels were selected for editing and eventual archiving.

CTDEDIT was used to clean salinity.  On-screen plots of descent rate and pump status were used to aid in distinguishing real variability from instrumental noise.

All casts required some editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

No recalibration will be applied to the salinity as it is very close to the bottles.

File 2007-46-recal1.ccf was prepared to recalibrate the dissolved oxygen by adding 0.045ml/l.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time. The comparison with titrated samples allows a correction for drift in calibration, leaving an error due to poor time response of the sensor. To analyze that, a comparison is usually made between the downcast values at the depths of bottles and the titrated DO values. For this cruise most bottles were fired after an adequate stop but not all and there was some unusual structure in the DO and Temperature profiles so even if the profiles are constant with time, small mismatches in the comparison will lead to large errors. It is also likely that there were significant changes between downcast and upcast. So this approach may not work.

The FIL files were bin-averaged (0.25db bins) & thinned to the usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX, and outliers were excluded. There was a lot of scatter. Points excluded included some with large standard deviations in the CTD data, some with short stops and some from around 135 to 150db in the areas with the large deep anomalies in temperature and dissolved oxygen. With those excluded there were no records left with high DO value. The average difference showed the CTD to be high by 0.05ml/. As long as the most severe outliers were removed this did not change much with changes in what was included in the fit. There is insufficient information to justify using a pressure-dependent or DO-dependent correct. Applying an offset of -0.05ml/l appears to be justified. (See 2007-46-dox-comp3.xls.)

File 2007-46-recal2.ccf was used to apply an offset of -0.05ml/l.to the thinned files: COMPARE was rerun and the results indicate the above recalibration was effective. (See 2007-46-dox-comp4.xls.)

File 2007-46-recal2.ccf was then applied to the AVG files, but not to the bottle files since this error does not apply to data collected while stopped. (Output: COR2)

##### Special Fluorometer Processing

All COR1 files were put through CLIP to produce files with data to 100db only for the use of A. Peña. These were then recalibrated using 2007-46-recal2.ccf.

Those files were then bin-averaged (1/4db bins) and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD1) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data was put through a median filter with fixed width 11, before bin-averaging.

There was no chlorophyll sampling so special bottle files were not prepared.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names, fix the platform name in the headers and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered, roughly:

•	±0.5ml/l from    0 -  50 db

•	±0.3ml/l from   50 - 175 db

•	±0.2ml/l below 175 db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~135% in Effingham Inlet and from 105 – 115% in Muchalat Inlet. For Effingham Inlet there were no dissolved oxygen samples taken, so the high values seen in the CTD dissolved oxygen data cannot be confirmed.

As a final check on the files a track plot, cross-reference listing and HEADER CHECK were run. No errors were found.

##### Final Bottle Files HEADER file IS READY

The MRGCLN files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE, Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### 22. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the CTD sensors.

Particulars from logs

1. Surface data only. Pumps not on. Not processed.

9. Altimeter not working. Bottle #3 did not trip.

13. Bottle #2 did not fire.

14. Fired extra bottle #2 misfired.

32. Problems with CTD display.

# Institute of Ocean Sciences  CRUISE SUMMARY

| Cruise ID#: 2007-46 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 21 May 2007 End: 28 May 2007 |  |  |  |  |  |  |
| Location: Effingham Inlet |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Dallimore A. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585         Cruise ID#:		2007-46

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 31May06 | Factory |  |  |
| Conductivity | 2128 | 22Dec06 | “ |  |  |
| Secondary Temp. | 2038 | 20Dec06 | “ |  |  |
| Secondary Cond. | 2173 | 23Dec06 | “ |  |  |
| Transmissometer | 723DR | 28/12/06 | Factory |  |  |
| SBE 43 DO sensor | 1176 | 17/10/06 | Factory |  |  |
| Altimeter | 1233 | ? | ? |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/2000 | Factory |  |  |