## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 5 Jan 2021 | ### Added HPLC Data. S.H. |
| 31 May 2008 | Recalibrated CTD and Bottle files based on post-cruise calibration; see note at end of report. |
| 19June 2008 | A CTD file for cast #17 was prepared containing unpumped channels only. |

## PROCESSING NOTES

Cruise: 2007-61

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG – JdF

Party Chief: Chandler P.

Platform: Vector

Date: 1 October 2007 – 5 October 2007

Processed by: Germaine Gatien

Date of Processing: 4 December 2007 – 13 December 2007

Number of original CTD casts: 74

Number of CTD casts processed: 72 (1 file upcast only, 1 had pumps off)

Number of original rosette casts: 21

Number of rosette casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with a Chelsea/Seatech transmissometer (#953DR), a PAR sensor (#4615), a surface PAR (#16504) and an Altimeter (#1252). A SeaBird dissolved oxygen sensor (#1119) and a Seapoint Fluorometer (#2228) with a 10X cable were mounted on the pumps. The deck unit was an SBE model 911 s/n 0619. The data logging computer was PAC02570. The salinometer was an 8400B model Autosal (S/N 68572).

# SUMMARY OF QUALITY AND CONCERNS

The rosette sheets and daily log book were in good order with useful comments included. The bottle analyses were complete before data processing began; this helped make the job go smoothly. There were a few problems in some of the bottle files, but these were easily resolved with the help of the analysts.

The quality of the bottle comparisons is limited by several factors. The casts are relatively shallow and the vertical gradients and temporal variability high in this region. The conductivity sensors had been recalibrated since last use, so there was no useful history. There has also been some evidence of possible problems with recent salinity analyses. Recalibration was applied to salinity, but that decision should be revisited when more information about these sensors becomes available, either through use on another cruise or after the next factory recalibration.

There were unusual features in temperature, fluorescence and dissolved oxygen profiles in some casts near the mouth of Juan de Fuca Strait. The dissolved oxygen values are confirmed by titrated values and there is no evidence of instrumental error.

The pumps were not turned on for cast #17. It was noted in the log that the oxygen and salinity data looked bad. Only a CHE file was prepared for this cast and no pumped channels were included in that.

The data from the first station was split into two casts: cast #1 contained all downcast data and part of the upcast data; cast #2 contained only the data from the top 22db of the upcast including the data concerning all bottles fired at this station. Only the rosette file was processed and it was renamed as 2007-61-0001.che, since it contains the all the bottle data associated with the downcast file 2007-61-0001.ctd.

The comparison of SBE dissolved oxygen data with titrated bottle samples showed less noise than usual. This is probably because the range of DO values was relatively small and the casts shallow enough that hysteresis should not be a factor. There were few severe outliers so the high quality of bottle analyses may have been a factor as well. The data should be considered roughly:

•	±0.6ml/l from 0 to 50db

•	±0.2ml/l below 50m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. The raw file names were not in standard format so had to be renamed.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained. It is noted that syringes were left off the CTD in between stations and the oxygen plug was not used. There were several notes about problems with the altimeter. Some repairs were done that improved the data but some problems continued. On the rosette log sheets there are many notes about problems with oxygen analysis.

Individual dissolved oxygen bottle files were obtained; flag channels and comments were included.

The nutrient, chlorophyll and salinity data were each obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The CTD calibrations were checked, the offset was adjusted in the pressure channel and the file was saved as 2007-61-CTD.con. Recent cruises have needed this adjustment because the pressure calibration is drifting significantly. This will be checked and can be adjusted further later.

##### Conversion of Raw Data

The files were converted to CNV files. Plots were made of a few files and indicate that temperature and conductivity channels are close during downcasts, but are frequently quite different during upcasts due to a lot of noise in all channels. The upcast and downcast traces are further apart than usual; this is true of both primary and secondary data. The altimetry, while noisy, looks usable near the bottom. The fluorescence is higher than usual at depth for some casts (for #4 it is ~0.275 below 200db) but this is probably ok in this region. Fluorescence is very odd in cast #1 with a maximum around 60db, but there is a temperature step at that level, so this may be real. Later in the cruise the fluorescence gets down to expected dark values. Transmissivity, PAR and SPAR data look fine. The descent rate is extremely noisy at the mouth of Juan de Fuca Strait, but very quiet in other areas.

It was mentioned in the log that the oxygen and salinity data looked bad for cast #17. In fact all the pumped data looks bad because the pumps were not turned on. There is a note that cast #30 looks wrong; cast #31 was a repeat. However, it is likely that #30 was just well-mixed and odd looking.

Rosette files were converted using a start time of -2s and duration of 5s.

CTD cast #1 included a full downcast and most of the upcast. A new file was started before the first bottle was fired. Only file #1 is needed for downcast data and only file #2 exists for bottles. The rosette file 2007-61-0002 was renamed as 2007-61-0001.  This should prevent confusion.

All files were then converted to IOS HEADER format and renamed as *.BOT. All BOT files were plotted and no significant outliers except for cast #17. The pumped channels should be removed later.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1    Std Dev = 2	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

A few casts with a very steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried. The noisiness of the upcast data makes this a difficult judgment. Choosing either (0.0245, 9.5) or (0.03, 7) seem best for both conductivity sensors.

CELLTM was run using (0.0245, 9.5) for both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The shaded entries are deeper than 300db. VN=very noisy descent rate and XN=extremely noisy descent rate.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 5 | 225 | -0.0005 XN | +0.0002 VN | +0.002 XN | Low, steady |
| 43 | 225 | -0.0003 VN | +0.00025 | +0.0028 | Mod, very steady |
| 51 | 400 | -0.0003 | +0.00028 | +0.003 | Mod, very steady |
| 59 | 225 | -0.0007 | +0.0002 | +0.0029 | Mod, very steady |
| 59 | 325 | -0.0005 | +0.00024 | +0.0029 | Mod, very steady |
| 63 | 225 | -0.0005 | +0.00025 | +0.003 | Mod, very steady |
| 74 | 225 | -0.0003 | +0.00027 | +0.0032 | Mod, very steady |
| 74 | 325 | -0.0003 | +0.00027 | +0.0028 XN | Mod, very steady |

Though not large, the differences in temperature and salinity were very noisy with many large spikes. The differences in conductivity were also noisy, but not as bad as in temperature. When the primary and secondary were plotted together with the differences between them, it is not clear that one channel is worse than the other. But examining features in the profiles, it looks like the spikes in differences are associated with vertical offsets between the two temperature channels. Since the descent rate is quite steady for these casts, the problem is likely irregularities in one or both pumps rather than the result of the rosette moving in and out of the vertical. The salinity differences are fairly steady if the first cast is excluded and there is no evidence of pressure dependence, though no cast was deeper than 400db.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

A header check was run and the only problem noted is that the minimum pressure is <0. One cast had a pressure of -0.26 at the end of the upcast with pumps on and with in-water values.

The average surface pressure is ~1.9db which is a little low for the Vector. There are a few cases of very low pressure that seem a little unlikely. Adding 0.3db to the pressure seems advisable.

The header summary was checked against the log book entries and it was found that no station names had been entered. These were added to the CLN files based on the log book. Track plots were produced and added to the end of this report. The track plot with the station names was shown to Diane Masson who confirmed they are correct. There was a discrepancy between the start times in the file header and the log for cast #28, but the log time looks wrong.

The altimeter values were exported to a spreadsheet. Because of comments in the log book, values from the first 12 casts were examined carefully. Casts #6, 8, 9 and 10 had very confused signals, but the value chosen by the algorithm looks correct. For cast #12 no header entry was created and this looks appropriate because most values were near zero throughout the profile. For cast #7 a value was entered that does not look reliable, so that was removed from the header. From cast #13 onwards the algorithm was found to have worked well. Casts #48 and 50 were noisy right at the bottom, but the values chosen were appropriate.

Files 2007-61-0002.CLN and -0017.CLN will not be processed beyond this point as one has no downcast data and the pumps were not working for the second.

##### BOTTLE FILE PREPARATION

Station names were added to the headers of the BOT files. Cast #17 will be processed since there was sampling at that station; the pressure from the CTD data is fine, just the pumped values are bad.

A track plot was produced to check that they were correct and no problems were found.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The ADDSAMP file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data was provided in EXCEL format with a comment and flag channel. The cast identified as #2 was changed to #4. (The 2 was based on an error on the rosette sheet.) Channel names were changed to standard format and the file was saved as 2007-61-sal.csv. The files were converted to individual SAL files.

The extracted chlorophyll data was in spreadsheet format and included a flag channel and comments. For cast #1 there was only 1 value, but the rosette sheet indicated there should be 2. The analyst found the data for the second bottle. (It had been mistakenly entered as a duplicate from the first bottle.) The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2007-61-chl.csv. That spreadsheet was converted to individual files. (Output: CHL)

The dissolved oxygen data were received in individual files with flag channel and comments. The file given as 2007-61-0002.add was changed to 2007-61-0004. A few small editing changes were made to the comments for clarity. Some problems were found later in some of the ADD files and the analyst corrected these.

The nutrient data were provided in spreadsheet format with flag channel and comments. Channel names were changed to standard format, the event # for station 59 was changed from 2 to 4, data was reordered on increasing sample number, blank values were replaced with pad values, comments were amended by the addition of “Nutrients:” and the file was saved as 2007-61nuts.csv. The data were converted to individual NUT files.

The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4.)

The MRG4 files were then merged with SAMAVG files (MRG) and then put through CLEAN to remove SeaBird headers and comments from the secondary files. (MRGCLN1)

##### 11. COMPARE

Salinity

COMPARE was run. Two serious outliers were identified:

- Cast #1, sample #1: The value looks more like water from 15db than from 22db. This was the cast during which the CTD was stopped at about 15db during the upcast and then lowered to 22db for a bottle. Perhaps the bottle closed at 15db, or for some other reason did not flush well. This salinity sample was flagged “d”.

- Cast #64, sample #237: The salinity looks like it is from near the bottom. The DO and nutrients do not look out of line, so this is not a case of a bottle that misfired. The most likely scenario is that the sample was taken from Niskin #2 as that was the bottle sampled for every other cast. Salinity sample #237 was flagged “d”.  The value was entered for sample #233 with a “c” flag and a note expressing doubt about this sample.

There are other minor outliers, but the standard deviations in the CTD salinity data are high, so it is likely that the bottles are fine.

After making the change to cast #64 COMPARE was rerun. Points were removed from the fits against pressure (based on differences and standard deviations in the CTD salinity) until the trendline was flat against pressure. The same points were excluded from both sensor pairs. The average of points in the fit show the primary to be low by 0.0058 and the secondary by 0.0030; the average of 5 points below 300db indicate the primary is low by 0.0060 and the secondary by 0.0036. There are only 10 points in the fit.

The fit against file pair number shows increasing differences with time for both salinity channels. This is complicated by the fact that the deepest casts are all later in the cruise so it is impossible to distinguish pressure dependence from time dependence.

The primary sensors have slightly more time dependence and the secondary slightly more pressure dependence, but there is sufficient scatter that the differences between the two are probably not significant. Since the secondary are closer to the bottles, it is probably best to use them for archiving.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. When all data is included the fluorescence is about 60% of the titrated chlorophyll. That goes down to about 50% when an estimated dark value (~0.1) is subtracted. For low CHL values, fluorescence is higher relative to CHL.

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD showed 1 extreme outlier. Investigation showed that 6 DO bottle values were entered incorrectly in the ADD file for cast #20. The flask numbers were out of order which may have led to the problem. It was also found that many bottles had been missed from cast #11. When these problems were corrected COMPARE was rerun. There were no severe outliers and the noise level in the differences is lower than usual. Residuals were used to exclude minor outliers and the fit found was:

DOX_BOT = 1.0263 * DOX_CTD - 0.0314

This compares with the fit found for 2007-62:

DOX_BOT = 1.0476 * DOX_CTD - 0.0066

It is believed that an anoxic plug was used during 2007-62, but not during this cruise; 2007-62 sampled a much larger range of DO values.

The fit versus time shows no significant temporal drift.

Plots were made of Dissolved Oxygen bottle data versus CTD salinity and the only problem noted are odd features in casts #9 and 11 with a minimum at about 30db and a maximum at about 35db. The same sort of feature is seen in temperature and fluorescence profiles, so it is assumed that this is not an instrumental problem. T-S is stable. The salinity is slightly lower than the historic minimum at about 50db, but temperature is entirely within the climatology.

There is less noise in the comparison than usual. This is probably the result of low DO gradients experienced during this cruise. But it could be that this sensor has a better response time than others that have been used at IOS. To investigate this, plots were made in high gradient regions to see how the sensor responded when the CTD stopped.

The figure below shows a small section of upcast data plotted against increasing scan #. As the CTD slows down a shed wake is observed in the temperature and salinity data at about scan #33825; about 75 scans later (3s) it is seen in the DO data. This delay is presumed to be due to DO transit time which is estimated in section 13 to be about 4.5s. The DO values never reach the value we would expect in the shed wake; before the sensor can fully respond a new signal is arriving.

There is a lot of variability in the temperature and salinity, so it is impossible to make a definitive statement about the response time, but something between 12 and 20s looks reasonable. Response time is gradient dependent, so this is not something we can compare easily among different data sets. At the bottom of this cast the DO varies little, so the response time is ~0. The response time is definitely better than experienced in the early days of SBE DO sensor use at IOS, but it does not seem so good as to explain the reduced scatter. So the reduced range is probably the explanation, and if the D.O. titrations were better than usual, that might be a factor.

##### Other Comparisons

Previous experience with these sensors –

- Both conductivity sensors have been recalibrated since last used.

- The DO sensor was used during 4 other cruises in 2007, of which 3 have been processed. During two cruises an anoxic plug was used.

- The pressure sensor has been used often. The offset has been increasing rapidly and was +5.9 for 2007-62 in September 2007.

Historic ranges – There are minor excursions from the historic ranges with salinity a little high near the surface west of San Juan Island and very slightly high for two casts west of Texada Island. Temperatures were all within the historic ranges. The salinity excursions are not considered indicative of instrument calibration problems as this is an area where excursions have been very common in recent years as the local climatology is out of date.

##### 13. SHIFT

Conductivity

Tests were run on a few casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.4 for the primary and -0.2 for the secondary conductivity.

SHIFT was run using -0.4 and -0.2s for the primary and secondary conductivity channels.

Dissolved Oxygen

Tests were run using settings from +90 to +140 records on 4 casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +110 records (4.6s).

Fluorescence

The method generally used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. Values found were from 1.25 to 2.5s, with most values at the lower end of that range. The vertical offset seems larger than usual in both temperature and fluorescence. There is also some suggestion that at least one pump might not have been working well especially during the upcasts. The value that has been used for almost all cruises in recent years is +24 records, but occasionally +36 has been used.

All data were shifted by +24 records and a few casts were checked afterwards and the results seem reasonable. (Output: SHFFL)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: none

##### DETAILED EDITING

Based on the bottle calibration data the secondary sensors look slightly better than the primary. Secondary channels were selected for editing and eventual archiving.

CTDEDIT was used to clean salinity and to remove corrupted data.  On-screen plots of descent rate and pump status were used to aid in distinguishing real variability from instrumental noise.

Casts #7 to 15 required heavy editing since extremely noisy descent rates caused much corruption by shed wakes. All other casts required only lightly editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

COMPARE indicates that the secondary salinity is too low. However, there have been some concerns with the performance of the Autosal in recent months. During 2007-62 the salinity analysis showed the salinity to be low by 0.002, but post-cruise factory calibration of the conductivity sensor shows it was probably by high by about 0.003. The temperature sensor was the same one as used on this cruise for the secondary channels. This could be a problem with the temperature sensor, but that is less likely than that it is an Autosal problem. Recalibration will be applied, but this should be revisited when more information becomes available for this sensor, from other cruises, or from the next recalibration.

From section 11 we have the following equation for recalibration of DOX:

DOX_BOT = 1.0263 * DOX_CTD - 0.0314

File 2007-61-recal1.ccf was prepared to apply the above correction to the DO channel, to add 0.3db to the Pressure and to add 0.003 to the Salinity:Primary channel in the SAM and MRG files.

COMPARE was rerun for salinity and dissolved oxygen and showed that the recalibrations were effective. (See 2007-61-dox-comp2.xls and 2007-61-sal-comp2.xls.)

The edited downcast files, EDT, were recalibrated using 2007-61-recal1.ccf. (Output:COR1)

The COR1 files were clipped to 150db and set aside to be processed later for Angelica Peña. A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time. The comparison with titrated samples allows a correction for drift in calibration, leaving an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged (0.25db bins) & thinned to the usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX, and outliers were excluded. The fit against DOX suggests that a simple offset fits the data reasonably well, but the plot of differences against pressure suggests a linear fit against pressure. Both were tested and the latter proved better. (See 2007-61-dox-comp3.xls.) File 2007-61-recal2.ccf was used to apply the following correction to the thinned files:

CTD_DOX (Corrected) = CTD_DOX - 0.0005*Pressure +0.1527

COMPARE was rerun and the results indicate the above recalibration was effective. (See 2007-61-dox-comp4.xls.)

File 2007-61-recal2.ccf was then applied to the AVG files, but not to the bottle files since this error does not apply to data collected while stopped. (Output: COR2)

##### Special Fluorometer Processing

The CLIP files with data from the surface to 150db were recalibrated using 2007-61-recal2.ccf.

Those files were then bin-averaged (1/4db bins)and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD1) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data was put through a median filter with fixed width 11, before bin-averaging.

The recalibrated CTD bottle files (SAMCOR1) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

Plots were made to check that the PAR sensor was on for all casts and it was.

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names, fix the platform name in the headers and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered, roughly:

•	±0.6ml/l from 0 to 50db

•	±0.2ml/l below 50m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~55% to 95% with the lowest values in Saanich Inlet and Haro. The highest values were at the mouth of Juan de Fuca Strait and in the northern half of the Strait of Georgia. Surface bottles confirm that the dissolved oxygen values are ok, and where low SBE DO values are seen in downcasts they are also low in upcasts; so the low saturation values are assumed to be real.

As a final check on the files a track plot, cross-reference listing and HEADER CHECK were run. No errors were found.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE, Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units, to fix the platform name in the headers and to add a comment about quality flags including the following comments on extracted chlorophyll analysis:

Average of two samples is reported unless stated otherwise.

Variability is assessed as the % (std dev/mean*100) and analysis methods.

The pumped channels (temperature, salinity, fluorescence and dissolved oxygen) were removed from file 2007-61-0017.che.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### 22. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the CTD sensors.

Particulars from logs

1. Separate up and down files for this cast.

6. Depth left out of header file

8. Bottle #9 did not trip.

10. Altimeter not working well - taken off for repairs.

14. Forgot to trigger bottle 13 – used a bucket of surface water.

17. Problem with Bottle #11. Repaired after this cast. Bad oxygen and salinity data on downcast. PUMPS NOT ON.

30. Temp/Sal look wrong.

31. Rerun of 30.

42. Bottle 16 not fired, bucket surface sample taken

49. Weird altimetry.

50. Altimeter flickering from 98 to bottom.

52. Cast in wake of quarry barge – mixing event at surface?

72. Altimeter does not detect bottom.

NOTE: May 31, 2008

The post-cruise calibration of the secondary conductivity sensor showed that salinity was high by ~0.0018 (measured at 3S/m) in January 2008. The drift might be linear with time or with use or might all have occurred early or late. In any case it is unlikely that the salinity was low by 0.003 at the time of this cruise as was found by the comparison with bottles. Over the past 2 years, there have been other indications that the Autosal is sometimes giving bottle salinity values that are too high leading to the CTD salinity looking low. Initial linearity tests suggest that the problem is most significant for lower salinity values, thus being particularly serious for cruises with relatively shallow sampling.

Based on these results it was decided to recalibrate all CTD salinity by subtracting 0.004 which should lead to values that are good to ±0.002. File 2007-61-recal3.ccf was prepared to subtract 0.004 from all CTD salinity data.

# Institute of Ocean Sciences  CRUISE SUMMARY

| Cruise ID#: 2007-61 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Dates: Start: 1 October 2007 End: 5 October 2007 |  |  |  |  |  |
| Location: SoG/JdF |  |  |  |  |  |
| Vessel: Vector Party Chief: Chandler P. |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

CTD Calibration Information

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 25Apr07 | Factory |  |  |
| Conductivity | 1766 | 21Aug07 | “ |  |  |
| Secondary Temp. | 4752 | 06Mar07 | “ |  |  |
| Secondary Cond. | 2399 | 25Apr07 | “ |  |  |
| Transmissometer | 953DR | 23/04/07 | Factory |  |  |
| SBE 43 DO sensor | 1119 | 10Nov06 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |
| PAR | 4615 | 15Dec00 |  |  |  |
| Surface PAR | 16504 | 2Jan2004 |  |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |