## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 5 Feb. 2019 | Bottle spreadsheet converted to searchable BOT files. |

## PROCESSING NOTES

Cruise: 2007-65

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, B.C.

Project: High Seas Salmon

Chief Scientist: Trudel M.

Platform: Viking Storm

Date: October 10, 2007 – November 5, 2007

Processed by: Germaine Gatien

Date of Processing: 13 December 2007 – 20 December 2007

Number of original CTD casts: 159 	Number of casts processed:  159

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#0404) was run with pressure sensor #0573 and was mounted with the following external sensors: Seapoint fluorometer #2356 and transmissometer #498DR. The fluorometer was pumped.

# SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard.

The CTD Salinity may be low by 0.01, but it was not recalibrated because the evidence was weak. When this instrument is next recalibrated, or if useful information becomes available from another cruise, the calibration issue should be revisited.

Transmissivity was removed from 3 casts because the signal was extremely noisy with many unbelievable zero values.

There were some inconsistencies between event numbers and file names in the data files and the log. These were resolved, but there remains a slight possibility of error for casts #8, 10 and 14.

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX. The names were not in standard format and had to be changed.

2. Preliminary Steps

The Daily Log was obtained. It is noted in the log that all times are in Pacific Daylight Time.

The file names were changed to standard format.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

The configuration file used at sea was obtained. Coefficients were checked and no errors were found.

All hex files were converted using 0404CTD.con.

A few casts were plotted. The pressure signal has steps of about 0.2db as is usual for this model CTD; the manufacturer states the resolution is 1db.

Temperature and conductivity data look reasonable.

The descent rate generally looks very noisy for some casts with some complete reversals in direction, but for others it is quite steady. The downcast descent rate tends to be higher than the upcast. The CTD was stopped at about the 10db level during the upcast for most casts.

Four cast files were missing: #1, 4, 5 and 168. Hugh Maclean was informed and found the missing files. The first three had been mislabelled.

The files named 2007-65-0007, -0013, -0071 and -0258 are identified in the log book as 2007-65-0008, -0014, -0070 and -0259 respectively.  This will be checked after conversion to IOS Headers at which point positions and water depth will be added to the headers.

4. WILDEDIT

WILDEDIT was run on all casts on pressure, temperature and conductivity channels using 2,20,25,0 for “Standard deviations for pass 1” and “Standard deviations for pass 2”, scans per block and “Keep data within this distance of the mean”.

5. WFILTER

Based on the results of many other cruises using this equipment, the SeaSoft routine WFILTER was run for all casts to apply a cosine filter, size 5, to the pressure, temperature and conductivity. This removes the steps caused by the limitations of the pressure sensor. Tests were done on the fluorescence channel to see if that should be filtered as well. The fluorescence does have a lot of fine-scale noise; the cosine filter does smooth the data somewhat, but not in the way it does for the other variables, presumably because the source of the noise is different. With an SBE911+ we apply a median filter. Tests show that a median filter, size 5, does a reasonable job on this data as well. However, it is better to do this step later, so that special unfiltered files can be prepared for the use of Angelica Peña after conversion to IOS SHELL and editing.

6. CELLTM

Tests were run using a variety of setting for CELLTM. The best results in the past have been with (α = 0.03, 1/β = 9.0) and that proved to be the best choice for this data as well.

CELLTM was run on all casts using α = 0.03, 1/β = 9.0.

7. DERIVE

Program DERIVE was run to calculate salinity.

8.  Conversion to IOS Headers

The IOSSHELL routine was used to convert Sea-Bird ASCII data to IOS Headers.

CLEAN was used to add event numbers.

A number of items are missing from the headers – positions, event numbers and water depth. Some file names and station names are wrong.

HEADER EDIT was used to add the missing lines to the headers, and a text editor was used to enter the latitude and longitude and (where available) water depth based on the log book entries.

In the course of checking these entries the following errors were found and corrected:

- 2007-65-0007 – Event #7 was a Bongo according to the log. The header time is that of event #8. But the station name looks like it is event #10. But there is another file with the right time and station name for event #10. Event # and file name were changed to #8.

- 2007-65-0013 – Event #13 was a Bongo according to the log. The header time and station name are those of event #10. Event # was changed to #10.

- 2007-65-0010 – Event #10 was a CTD cast according to the log but the header time looks like cast #14 as does the station name. Event # was changed to #14.

- 2007-65-0016 – Station name is given as JF07 but time and event number agree with the log entry for event #16, but that has station JF04. The station name was changed.

- 2007-65-0028 – Station name corrected. Obviously just a typo.

- 2007-65-0071 – Event # and file name changed to 70.

- 2007-65-0258 – Event # and file name changed to 259

A track plot was used to ensure the entries made sense.

The file list was updated with the new file names.

9. Checking Headers

Header Check was run and the speed check led to finding one position error.

A Header Summary was produced and checked against the log. A number of errors in position were found and fixed in the CLN2 files.

The Surface Check produces an average of -0.45db but the very low salinity values suggest that the CTD was out-of-water when the first values were recorded; there is no indication of a problem with the pressure calibration.

Track plots were produced and added to the end of this report.

The log indicates the time is in PDT. The headers say UTC but clearly the times are PDT. ADD TIME CHANNEL was run to add 7 hours to the times.

10. SHIFT

Conductivity

On previous cruises using this type of CTD good results were found when SHIFT was run to advance the conductivity by +0.7 records. Tests were run on 4 casts using settings from -0.5 to +1.3 records and the best results were found to be from +0.9 to +1.1 records.

SHIFT was run on all casts advancing the conductivity by +1 record.

Fluorescence (Note: it was discovered later in the processing that the fluorescence had not been shifted, so this step was run on the SHF files, then the output SHFFL files were put through DELETE with output DELFL. Those files were then merged with the EDT files, choosing the fluorescence from the DELFL files and all other channels from the EDT files.)

In previous cruises using a fluorometer mounted on a different SBE25, the fluorescence was aligned by shifting the channel by +42 records. A few tests on this data suggest that value is much too high. A shift of +24 records looks appropriate. All casts were put through SHIFT using +24records. Plots were examined and the results look satisfactory.

11. DELETE

The SHF files were put through DELETE using the following parameters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00    Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted

Sample interval taken from the header.  Pressure was not filtered.

COMMENTS ON WARNINGS: There were no warnings.

As mentioned above, the SHFFL were put through DELETE using the above settings, output DELFL.

- DETAILED EDITING

Page plots were examined on-screen and examined for spikes and instabilities and used to guide the use of CTDEDIT.  All casts required some editing except for cast #303.

Note was made of the editing details in the files.

MERGE was run to select all channels except fluorescence from the EDT files and fluorescence from the DELFL files; pressure was the merge channel.

A median filter, size 5, was applied to the fluorescence channel in the EDTFL files. Output:FIL

- Files for Angelica Peña

The EDTFL files were clipped to 120db and set aside for the use of Angelica Peña. Those files were then bin-averaged (0.5db bins), then put through REMOVE and HEADEDIT. A second set of files were produced using a size 5 filter first, then BIN AVERAGE, REMOVE and HEADEDIT.

A set of upcast files were produced for just the casts with CHL data. For those the SHFFL files were put through reverse, clipped to 15db, and then processed either with a filter or bin-averaging, or with neither.

- Inter-comparisons

Comparisons with bottle samples

- Salinity: There were many surface salinity samples. The CTD was stopped at about 9 to 11db and then it generally drifted downwards to 10 to 12db. The Niskin was probably mounted 5m above the CTD. So it is not obvious what depth to use in the comparison, but the lowest point while stopped – 5db is probably the best bet. The spreadsheet from the analyst was edited to remove information that is not needed in the comparison and to add event #s. The bin-averaged CTD files were thinned to 5db and exported to a spreadsheet 2007-65-CTD at 5db.csv. That data was then combined with the salinity bottle data. Differences were found and showed huge variability. When the casts were reduced to those that were well-mixed (based on changes in salinity compared to the values at 4db), the CTD salinity was found to be low by an average of ~0.01. The process was repeated using CTD data from 6db and the average was almost identical. As a final approach plots of the well-mixed casts were examined, the pressure at which the stop ended recorded and then the CTD data from 5 db above that depth was found in the downcast file. The averages were the same no matter which method was used, so indeed these are well-mixed casts! All these approaches were combined in file 2007-65-sal-comp.xls.

The CTD salinity appears to be low by 0.01. There have been recent doubts about the quality of the salinity comparisons, which may be due to problems with the salinometer or possibly due to poor cleaning of the sample bottles. Any problems in the flushing of the Niskin bottle may result in the water in the bottle having higher salinity than is representative of the stopping depth, though using well-mixed casts should minimize that error. However, none of the well-mixed layers is very deep. The 3 best-mixed casts show the CTD to be low by 0.005, 0.011 and 0.013.

- Fluorescence: There was extracted chlorophyll for 24 casts. Data from 5.5db of the downcasts was extracted and compared to the bottle data. The chlorophyll values were all quite low, <2.4ug/l, and most were <1. The ratio of fluorescence/chlorophyll was 0.91 if all data was included, 0.97 for the 4 values from well-mixed casts and 0.95 for the casts with CHL<1.

Sensor History

The conductivity sensor was used during 2006-34, but there was no calibration sampling. No other data has been processed that used this sensor.

Historic ranges

All temperature traces were within the historic ranges, where local climatology was available. Salinity was occasionally lower than the historic minima, but only at stations that were in inlets or near-shore; such casts are probably not well represented in the climatology.

- Recalibration

The CTD salinity is probably low, but given no deep sampling, no history for the instrument and some doubts about the salinity analysis, it seems unwise to recalibrate at this time. When the sensor is next calibrated at the factory, or when it has been used on other cruises, we can revisit this decision.

- BIN AVERAGE

The following Bin Average values were applied to the FIL files:

Bin channel = pressure 		Averaging interval = 1db

Minimum bin value =   .000	Average value will be used

Interpolated values are NOT used for empty bins.

- REMOVE

The following channels were removed from all casts: Scan_number; Conductivity:Primary, Descent Rate and Flag.

Plots were examined on-screen to see if there were any further problems and it was discovered that the transmissivity looked bad for casts 58, 141 and 146.

Transmissivity was removed from casts 58, 141 and146.

- HEADER EDIT

Header Edit was used to fix formats and channel names and to add the following comment:

Fluorescence and Transmissivity data are unedited and nominal except that some

records were removed in the processing of editing of temperature and salinity.

For the 3 casts from which transmissivity was removed, this comment was edited to reflect that fact.

Standards Check was run and Header Edit rerun until all problems were resolved.

- Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) The conductivity and pressure sensor history files were updated.

Notes from log book

287. CTD left on for an hour after end of cast.

291. Check that data is not in file for 287. (looks ok.)

299. Noted that still recording in PDT. (This was the time when official change to PST ashore.)

## Institute of Ocean Sciences

## CRUISE SUMMARY

Cruise ID#: 2007-65

Dates: Start: 10 October 2007 End: 5 November 2007

Location: N.E.Pacific

Vessel: Viking Storm

Party Chief: Trudel M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0404 | No | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 0573 | 22May06 | Factory |  |  |
| Temperature | 4700 | 29Jul06 | Factory |  |  |
| Conductivity | 3184 | 04Aug06 | Factory |  |  |
| Seapoint Fluorometer | 2356 | n/a | n/a |  |  |
| Transmissometer | 498DR | 3Oct07 | IOS |  |  |