## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2007-67

Agency: IOS, Ocean Sciences Division, Sidney, BC

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG – JdF

Party Chief: Chandler P.

Platform: Vector

Date: November 27, 2007 – December 1, 2007

Processed by: Germaine Gatien

Date of Processing: 11 January 2008 – 27 February 2008

Number of original CTD casts: 70 	Number of CTD casts processed: 70

Number of original rosette casts: 23 	Number of rosette casts processed: 20 (no sampling for 3 others)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0505) was mounted with a Chelsea/Seatech transmissometer (#953DR), a PAR sensor (#4615), a SPAR sensor (#16504), an Altimeter (#1024), a SeaBird dissolved oxygen sensor (#1117) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was an SBE model 911 (S/N 0619). The data logging computer was PAC02570. The salinometer was an 8400B model Autosal (S/N 68572). There were 24 20L calibration bottles mounted 1.4m above the CTD.

# SUMMARY OF QUALITY AND CONCERNS

It is recommended that the pressure offset be changed to 0 for this CTD.

There were a lot of severe outliers in the salinity comparison. All involved either CTD values that are too low or bottle values that are too high. Possible causes include failure to adequately clean the salinity sample bottles and intermittent problems with the salinity analysis. Poor flushing or premature firing of Niskin bottles might be a cause of low values, but both seem unlikely since some of the bottle values were lower than any recorded by the CTD. Outliers were found in casts with a very steady ascent rate (which might be expected to impede flushing) and others with a very noisy ascent rate (which would lead to a more random distribution of errors if flushing is an issue) and outliers are seen in several Niskin bottles.

Cast #7 is considered to be of lower quality than usual due to large patches of very low descent rate with frequent complete reversals of direction. DELETE will have removed many records with low descent rate but the effects are obvious in what is left, and editing the data was generally unsuccessful because it was impossible to distinguish good data from bad. Cast #29 is also considered to be of lower quality than usual below 70db. Above that it is ok.

The SBE dissolved oxygen data in the CTD files should be considered roughly:

•	±0.5ml/l from   0 - 100m

•	±0.3ml/l from 100 - 175m

•	±0.1ml/l below 175m

Duplicate dissolved oxygen samples were taken from 15 bottles. For 13 of them the differences were <2% and <0.045ml/l. For 2 pairs the differences were 3.8% and 6.9%.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Daily Log Book and Rosette Log were obtained. It is noted that the syringes were left off the CTD unless there was a steam of several hours between casts. The cast in Saanich Inlet was moved from the usual site because of glider operations going on in the inlet. There were problems with the transmissometer for two casts, though this appears to have only affected the upcast. There were also problems with bottles either not firing, or firing but being empty. Bottle #7 was the biggest problem. During casts #21 and 57 many bottles were fired for test purposes, but no sampling was done.

Individual dissolved oxygen bottle files were obtained; flag channels and comments were included. For one cast duplicate DO samples had been taken, and the ADD file contains both in separate lines, so this will have to be changed before it is merged. A study should be made of the duplicates, and a decision made on whether to average duplicates or choose one set for archiving. First the CTD data should be prepared for comparison.

The nutrient, chlorophyll and salinity data were each obtained in spreadsheet format. Each spreadsheet was edited to remove information not needed for the chemistry files.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The CTD calibrations were checked, the pressure sensor # and the calibration date for the SPAR were corrected and the file was saved as 2007-67-CTD.con.

##### Conversion of Raw Data

The files were converted to CNV files. Plots were made of a few files and indicate that temperature and conductivity channels are closer during downcasts than upcasts. The secondary conductivity has a lot of small spikes even during the downcasts. Fluorescence, PAR, transmissivity and dissolved oxygen look as usual. While there are the usual alignment and hysteresis problems in the DO channel, there is some detail in the DO trace suggesting that the response time is not too bad. The SPAR traces look a little odd, but this is probably just because there were such low values that we are mostly seeing noise. The altimetry has an obvious signal near the bottom, but there are often spikes there so care will need to be taken in checking that the header algorithm works well.

As usual for the SoG/JdeF cruises, the descent rate was very steady for many casts, but noisy for a few near the mouth of Juan de Fuca Strait. There are negative pressures in a few of the upcasts, but the pumps were off, so they are close to the surface. This should be checked further later.

Rosette files were converted using a start time of -2s and duration of 5s. Those files were then converted to IOS HEADER format. All BOT files were plotted. CTDEDIT was used to clean the primary salinity for cast #46 around 40db.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.

Parameters used were: 	Pass 1 Std Dev = 2	Pass 2 Std Dev = 5	Points per block = 50.

##### CELLTM

Four of the deeper casts with a fairly steady descent rate were studied to determine the choice of parameters for CELLTM. Settings of (0.01, 7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9) and (0.0245, 9.5) were tried. There was little difference between the choices, probably due to the fairly low range of temperature values. CELLTM was run using (0.0245, 9.5) for the primary and (0.02, 7.0) for the secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 4 | 100 280 | -0.0005 ~0 | +0.0001 +0.0002 VN | +0.0025 +0.0025 VN | Steady |
| 44 | 100 280 | -0.0004 VN -0.0002 | +0.00012 +0.00015 | +0.002 +0.0018 | Very steady |
| 49 | 100 280 | -0.0006 VN -0.0002 | +0.00016 N +0.00022 | +0.0020 N +0.0023 N | Very steady |
| 70 | 100 280 | ~0.001 XN -0.0001 N | +0.00005 VN +0.0002 | +0.001 VN +0.002 | High, very steady |

The differences were not large, but they were very noisy at times.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

A header check was run and no problems were noted. A header summary and cross-reference listings were produced.

A cross-reference list was checked against the log book entries and the only header errors found were the event # for casts #37 and 63; those were corrected in the headers. There were discrepancies in the latitude for casts #12 and 28, but the log entries for the starting positions are clearly wrong (the bottom and stop positions are correct).

Track plots were produced and added to the end of this report.

The average surface pressure is ~1.2db which is a little low for the Vector. A few casts have negative surface pressures; those have anomalously low temperatures, but the conductivity looks like they are “in-water” values. The pumps came on while the pressures were<0 for cast #66 and were not turned off until the pressure was -0.8db. For cast #68 the transmissivity and fluorescence look like they go from air to water values at about -0.7db of the downcast. The pumps were not turned on until +0.9db. It is noted in the Bridge Log that it was snowing during those two casts. For cast #69 the transmissivity suggests the surface was at about -0.6db, which is the same level at which temperatures move from negative values to about 7C˚. The conductivity looks “in-water” above that. For cast #3 the upcast pressures are negative while all values continue to look like the CTD was in water. Transmissivity begins to look like it is in air at about -0.3db. The pumps were off, so pumped channels aren’t useful. For cast #4 there are negative pressures of -0.7db with pumps on and “in-water” values. It is possible that freezing conditions have affected the CTD, but this should not affect the upcasts. It does seem clear that the pressures are too low. The last time this sensor was used it was noted that the surface pressures were a little low and that it would soon be time to adjust the offset. It looks like the time is now! So to ensure not too many data are lost in DELETE, the pressure will be recalibrated now.

File 2007-67-press-recal.ccf was prepared to add 0.8db to both the CLN and BOT files. It is suggested that when next used the offset should be set to 0 in the configuration file.

The altimeter values were exported to a spreadsheet; many casts were examined and the algorithm was found to have worked well despite a lot of noise at the bottom of some casts.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There were 3 casts with bottles fired, but no samples assigned, so those were removed from the list of bottle files to be prepared.

The ADDSAMP file was then converted to CST files to form the framework for the bottle files. Sample numbers were added to the BOT files (output: SAM) which were then bin-averaged (SAMAVG) on bottle number.

The salinity data were provided in EXCEL format with a comment and flag channel. Channel names were changed to standard format and the file was saved as 2007-67-sal.csv. The files were converted to individual SAL files.

The extracted chlorophyll data were in spreadsheet format and included a flag channel and comments. The spreadsheet was edited by changing channel names, adding an event number for each sample and adding “Extracted Chlorophyll:” before each of the comments. The resulting file is 2007-67-chl.csv. That spreadsheet was converted to individual files. The following comments from the analysis sheet were added to the comments that will later be added to the headers of the bottle files: “Average of two samples is reported unless stated otherwise. Variability is assessed as the % (std dev/mean*100).” (Output: CHL)

The dissolved oxygen data were received in individual files with flag channel and comments. There were duplicates for cast #17; the ADD file was edited to choose the first sample in each case of duplicates. The pad values were set wrong in the ADD files, so they were put through clean to change those to -99. The output files are ADDCLN. (There were DO duplicate samples for cast #33, but no sample #s were assigned and no other sampling done.)

The nutrient data were provided in spreadsheet format with flag channel and comments. Channel names were changed to standard format, data were reordered on increasing sample number, comments amended by the addition of “Nutrients:” and the file was saved as 2007-67nuts.csv. The data were converted to individual NUT files.

The SAL, CHL, ADD and NUTS files were merged with the CST files in four steps. (Output: MRG1, MRG2, MRG3, MRG4.)

The MRG4 files were then merged with SAMAVG files (MRG) and then put through CLEAN to remove SeaBird headers and comments from the secondary files. (MRGCLN2)

##### 11. COMPARE

Salinity

COMPARE was run. There was a lot of scatter. When differences >0.01 were excluded the fits against pressure were fairly flat and the excluded points were the same for both channels. The primary was found to be low by an average of 0.0067 and the secondary by 0.0049. If only bottles below 200db are used the differences are slightly higher at 0.0075 and 0.0054. The secondary was slightly flatter. There is no hint of time dependence. A second run of COMPARE used Niskin Bottle # as reference channel to see if there was a bad bottle. Most samples were from Niskin #2, with some from #3 and #6. There were outliers from all of them.

The outliers (salinity differences from -0.013 to -0.096) came from casts 1, 3, 8, 20, 39, 56, 59, 62 and 66 and from a variety of depths. The salinity analysis sheet shows 3 of the outliers being run at the beginning of the session and 2 at the end; the rest are scattered between, so this does not seem to be related to drift in the salinometer. There are many notes on the analysis sheet about cleaning and rerunning of samples, but there are no flags attached to the any of the data and no indication that the outliers were associated with more problems than other samples.

Examining plots of a few casts turned up no obvious explanations for the outliers. The standard deviations in the CTD salinity are not high for most of these outliers. Cast #1 showed little variability in the salinity; while the stop was a little short, there is nothing to suggest an error of 0.02 would be caused by that. For cast #3 there is more variability, but not enough to explain an error of almost 0.1 in salinity. There is noise in the CTD salinity for cast #20 at about the point of bottle firing, but it would only explain an error of ~0.005, not 0.06 as seen in COMPARE. Similarly for cast #66 there is enough noise in the CTD to account for an error of ~0.005, but not 0.06. The samples were not taken at the bottom, so neither bottom boundary currents nor sediment in samples will help explain the results.

There could be intermittent salinometer problems, but there is no obvious sign of that. If the Niskin bottles fired before the intended time that would lead to CTD values looking low, but that is not likely, given that the dissolved oxygen does not show the same pattern. Moreover, in a few cases the salinity in the bottle is higher than the maximum CTD value for the cast. There could have been problems in sampling, especially since the air temperature was very low. Failure to clean the salinity bottles was found to be a problem on another recent cruise, and salt crystals left behind would lead to high bottle values. Poor flushing, like premature firing, does not seem likely since the largest errors are found in both casts with very noisy ascent rates and others with a very steady ascent; the bottles that produced salinity outliers have DO values that do not stand out.

The outliers were flagged “c” in the salinity files.

Dissolved Oxygen

COMPARE was run. Plotting the differences versus CTD and excluding outliers, the fit was:

DOX_BOT = 1.0203 * DOX_CTD + 0.0055

This correction is significantly smaller than for other recent cruises using this sensor. This may be partly because of a smaller range of DO values, with no values lower than 1.7. With a smaller range, the response time is not so critical. While this correction is not supposed to address response time, if the stops are not really long it is affected to some extent. Nonetheless it is surprising to see so small a correction.

The only outliers significant enough to be flagged were as follows:

Cast #1, Sample 8, flagged “c”

Cast #11, Sample 46, flagged “c”

Cast #46, Sample 154, already flagged by analyst; a note was added that it was also an outlier in COMPARE

Cast #59, Sample 216, flagged “c”

Cast #64, Sample 250, flagged “c”

Cast #66, Sample 254, already flagged by analyst; a note was added that it was also an outlier in COMPARE.

The fit versus time does not show any significant drift.

Plots were made of Dissolved Oxygen bottle data versus CTD salinity and Bottle salinity; all outliers had already been flagged, but notes were added to the headers for the outliers in casts 11 and 46.

There were duplicate DO samples from 2 casts. For cast #17 duplicates were taken from 8 bottles at a variety of depths. The average difference was 0.024ml/l, with errors ranging from 0 to 3.8%. Only 1 pair had differences >1.3%. For cast #33 there were 7 bottles from one depth. The differences from the average of the 7 bottles ranged from 0.5% to 6.9%. Only one pair differed by more than 1.7%. So in summary, out of 15 duplicates, only 2 differed by more than 2%, or by more than 0.04ml/l.

Fluorescence versus Extracted Chlorophyll

COMPARE was run for a quick check on the data. When all data are included the fluorescence is about 104% of the titrated chlorophyll; most values are low so this is about what we expect.

##### Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensor was recalibrated shortly before this cruise so there is no history available for it. The secondary conductivity sensor was used for the latter part of 2007-02, and during 2007-28 and 2007-03. For two of those cruises there was evidence of poor flushing of bottles. It was found to be low by between 0.0043 and 0.0073, with the best calibration being from 2007-03 when it was believed to be low by about 0.0043.

- The DO sensor has been used 8 times since it was last calibrated. The comparison with bottles has shown an increasing correction with time. It is noted that the uses were all between November 2006 and June 2007, so that the changes might be related to DO range as well as time. The alignment has also drifted.

- The pressure sensor has been used 3 times since last calibrated. During the last use it was noted that the offset might be drifting. An offset of -0.8db was applied to all 3.

Historic ranges – The only excursions from the historic ranges were for salinity; it was slightly low at 130db for cast #66 and slightly high at 70db for cast #70. Temperatures were all within the ranges. There is no evidence of instrument calibration problems.

##### 13. SHIFT

Fluorescence

The method used to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the results ranged from 0.8s to 1.4s. A comparison of features in T and FL traces shows that the offsets tend to be a little higher for the upcasts than the downcasts, so that the SHIFT value should be slightly lower than found in the calculation. This is presumed to be because the pumps don’t work as efficiently on the upcast. The value that has been used for almost all cruises in recent years is +24 records (+1s) and that looks appropriate for this one as well.

All data were shifted by +24 records and a few casts were checked afterwards and the results are good.

Conductivity

Tests were run on a few casts with a steady descent rate to determine the best shift of the conductivity sensors based on reduction of instabilities in salinity without oversmoothing. The best settings proved to be -0.5 for the primary. The secondary conductivity looks best with a setting of -0.7s.

SHIFT was run using -0.5 and -0.7s for the primary and secondary conductivity channels.

Dissolved Oxygen

Tests were run using settings from +110 to +150 records on 2 casts. Judging by how the downcast vs upcast trace offset compares with that of temperature, the best choice overall appears to be to advance the DO channel by +130 records (5s). The same setting was found to be best for the past 4 cruises using this sensor. SHIFT was run using +130 records for the Oxygen:Dissolved:SBE channel for all casts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range 10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: none

##### DETAILED EDITING

Based on the bottle calibration data the secondary sensors are closer to the bottles and less pressure dependent. The secondary seem slightly noisier than the primary sensors, though given the location, these may be real variations that the primary is missing. The secondary sensors were selected for archiving.

CTDEDIT was used to clean salinity.  On-screen plots of descent rate and pump status were used to aid in distinguishing real variability from instrumental noise.

The following casts required no editing: 3

The following casts required fairly heavy editing:  8-11, 28, 39.

The following casts were edited lightly but there are areas of low quality where good data could not be distinguished from bad data:

- Cast 7: The data were heavily corrupted by shed wakes; in patches there were frequent incidents of very low descent rate or complete reversals of direction. Identification of data corrupted by shed wakes requires that the incidents be discrete, so much of this cast could not be edited. The data must be considered of lower quality than usual, except near the surface and bottom where the data are more reliable.

- Cast 29: The descent rate was very low below 70db – above 70db it looks ok.

All other casts required only light editing.

Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration and Fluorescence Filter

It was discovered after editing that the wrong files had been put through DELETE, so that the shift of DO had been skipped. Since the only editing that would have affected DO values involves deletion of bad records, the EDT files were put through REMOVE (EDTREM) to remove the DO channel. The SHFO files were then put through DELETE (SHFODEL) using the same parameters as given above; the output was then merged with the EDTREM files (matching scan #s) to produce EDX files, with correct DO data. MERGE was set to ignore any records missing from the EDTREM files.

The CTD secondary salinity was found to be low by 0.005 for this cruise and by 0.004 during 2007-03. During 2007-02 and 2007-28 the differences were larger, but there were some doubts about the comparison. The salinity will be recalibrated by adding 0.004, but this should be revisited after the sensors are next recalibrated.

From section 11 we have the following equation for recalibration of DOX:

DOX_BOT = 1.0203 * DOX_CTD + 0.0055

File 2007-67-recal1.ccf was prepared to apply the above correction to the DO channel and to add 0.004 to the Salinity:T1:C1 channel in the SAM and MRG files. COMPARE was rerun for salinity and using the same points the average difference was -0.0009 so the recalibration was effective. COMPARE was rerun for dissolved oxygen and the fit again indicates that the recalibration worked properly. (See 2007-67-sal-comp2.xls and 2007-67-dox-comp2.xls.)

The edited downcast files, EDX, were recalibrated using 2007-67-recal1.ccf. (Output:COR1)

The COR1 files were clipped to 150db and set aside to be processed later for Angelica Peña. A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There are small instabilities, but aside for cast #7, they were in areas where this is expected. As noted earlier cast #7 is considered poor. No further editing will be applied.

##### Final DO comparison and calibration

SHIFT addresses errors due to transit time. The comparison with titrated samples allows a correction for drift in calibration, leaving an error due to poor time response of the sensor. To analyze that, a comparison is made between the downcast values at the depths of bottles and the titrated DO values.

The FIL files were bin-averaged (0.5db bins) & thinned to the usual bottle levels. Then CTD DO values were compared with the upcast bottle DO values. Plots were made of differences against pressure and DOX, and outliers were excluded, including all points that were excluded from the first comparison. Averaging the differences shows the CTD DO is high by about 0.11ml/l with dependence on both pressure and DOX, with the fit against pressure looking tightest. (See 2007-67-dox-comp3.xls.)

File 2007-67-recal2.ccf was used to apply the following correction to the thinned files:

CTD_DOX (Corrected) = 0.0005 * Pressure – 0.1559

COMPARE was rerun and the results indicate the above recalibration was effective. (See 2007-67-dox-comp4.xls.)

File 2007-67-recal2.ccf was then applied to the AVG files, but not to the bottle files since this error does not apply to data collected while stopped. (Output: COR2)

##### Special Fluorometer Processing

The CLIP files with data from the surface to 150db were recalibrated using 2007-67-recal1.ccf.

Those files were then bin-averaged (1/4db bins), recalibrated using 2007-67-recal2.ccf and put through REMOVE to remove extraneous channels. HEADEDIT was run to fix formats and channel names and the final files (FCTD1) were saved in a separate directory.

A second set of files (FCTD2) were prepared in exactly the same way except that the fluorescence data were put through a median filter with fixed width 11, before bin-averaging.

The recalibrated CTD bottle files (SAMCOR1) were put through REMOVE and HEADEDIT and saved as BOF files in a separate directory.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed from the COR2 files: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Descent_Rate, Status:Pump, Altimeter and Flag channels. (Output: *.REM)

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names, fix the platform name in the headers and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files should be considered roughly:

•	±0.5ml/l from    0 - 100m

•	±0.3ml/l from  100 - 175m

•	±0.1ml/l below 175m

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The files were named CTD.

A final check on the dissolved oxygen values was done by calculating DO saturation. Plots show surface values ranged from ~60% to 90% with the highest values in Saanich Inlet, at the mouth of Juan de Fuca and in the central section of the Strait of Georgia.

Profile plots were examined on-screen to check fluorescence, PAR, dissolved oxygen and transmissivity channels. The only problem detected was that the transmissivity is 0 at the bottom of cast #8. The zero values first appear when the altimetry and descent rate indicate that the CTD was not at the bottom. The upcast data have large sections with zero values. The log notes problems during upcasts with this channel. A text editor was used to enter pad values for the transmissivity for the zero values in the CTD file at the bottom of cast #8. The CHE files were checked since the upcast problem would affect them. Problems were found with casts #8 and 11; pad values were entered in the transmissivity channel where appropriate.

As a final check on the files a track plot, cross-reference listing and HEADER CHECK were run. No errors were found.

##### Final Bottle Files

The MRGCLN2 files were put through SORT to rearrange data with increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Oxygen:Voltage:SBE, Descent_Rate, Altimeter and Flag for all casts.

CHANGE UNITS was used to derive Oxygen:Dissolved:SBE in umol/kg.

REORDER was used to get the two SBE DO channels together.

HEADER EDIT was run to fix formats and units, to fix the platform name in the headers and to add a comment about quality flags.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The files were named CHE.

##### 22. Producing final files

A cross-reference listing was produced for CTD and CHE files.

HEADER CHECK was rerun; no errors were detected.

The sensor history was updated for the CTD sensors.

Particulars from logs

6. CTD hit bottom – wire angle problem, settled into bottom.

7. Altimeter on bottom.

8. Transmissivity zero at bottom and frequently during upcast.

11/12. Transmissometer signal crashed during upcast. Cables cleaned after 12.

12. Couldn’t get close to bottom due to strong current.

13. Transmissometer trace looks good.

15. Remote pressure no good

17. Oxygen samples – duplicates taken for some bottles to test system.

20. Bottle #7 did not fire but console display showed that it had fired.

21. Poured hot water over top of rosette. Fired 12 bottles at 10m to test trip. No sampling was done.

22. Large wire angle

27. Upcast T/S look odd; not a rosette cast.

33. Niskin #7 misfired.

34. The carousel was soaked in hot water.

51. Bottle #7 did not fire.

56. Bottle #13 fired but was empty.

57. Fired 16 bottles at surface, all fired correctly.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2007-67 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Dates: Start: 27 November 2007 End: 1 November 2007 |  |  |  |  |  |  |
| Location: SoG/JdF |  |  |  |  |  |  |
| Vessel: Vector Party Chief: Chandler P. |  |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |  |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |  |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506         Cruise ID#:		2007-67

| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| --- | --- | --- | --- | --- | --- |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4484 | 19Mar05 | Factory |  |  |
| Conductivity | 1764 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2710 | 7Apr05 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 953DR | 22/04/07 |  |  |  |
| SBE 43 DO sensor | 1117 | 17/10/06 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| PAR | 4615 | 11/02/03 |  |  |  |
| Surface PAR | 16504 | 18/12/00 |  |  |  |
| Fluorometer | 2229 | ? | ? |  |  |
| Pressure Sensor | 69698 | 26/05/2006 | Factory |  |  |