## O2 analysis notes
Dissolved oxygen concentrations for the 2008 west coast leg of C3O (2008-02) were measured on board the Sir Wilfrid Laurier using an oxygen titrator on loan from the Scripps Insitution of Oceanography.
The SIO system is a basic Winkler titrator that differs in 3 significant ways from the titrators currently in use at IOS.
1. Instead of using a colorimeter probe inserted into the sample flask, the SIO system uses a UV light source and a photodiode detector
It capitalizes on the strong absorption of UV light at 365 nm by the tri-iodide ion.
UV light shines through the flask and as the titration proceeds (and tri-iodide is consumed), the quantity of light reaching the detector increases until the endpoint is reached.
2. The SIO system employs 2 thermisters to constantly monitor the temperature of both the KIO3 standard and the sodium thiosulfate solutions (and corrects mass of both delivered to 20C) so that the system need not be operated in a temperature controlled lab.
3. Aside from the dosimat used to dispense the sodium thiosulfate, the SIO system uses a second dosimat to dispense the KIO3 standard, minimizing the human factor and thereby improving system accuracy.
SAMPLING:
All samples were collected in ground glass stoppered Erlenmeyer flasks (nominal volume ~140 ml) supplied and calibrated by SIO.
Samples were collected at 62 stations in the North Pacific, Bering, Chukchi and Beaufort Seas along a transect from Victoria, BC to Cambridge Bay, Nu
Greater than 10% of samples were collected in duplicate (initially in triplicate) in order to evaluate the precision of the analysis.
Samples were collected immediately following the rosette cast, were reshaken ~30 minutes after pickling and water was placed on top of the stoppers to minimize oxygen exchange.
The ppt. was allowed to settle to the bottom of the flask in the CTD shack before samples were brought inside and stored in the lab at room temp.
A storage test was performed on duplicate samples taken at Stn CCL2. Duplicates stored in the dark at room temperature and run 4 days later were comparable to their counterparts run after 24 hours.
ISSUES DURING ANALYSIS:
Floaties
Since the SIO system uses optical detection any impurities or bubbles in the sample will possibly affect the final titre by absorbing/deflecting some of the UV light passing through the flask.
Impurities were commonly in the form of copepods, other zooplankton or unidentified particles termed "floaties". Whenever possible, samples were redrawn to remove these impurities.
Occasionally a mass of small bubbles was sucked down into the light path either as a result of excessive stirring or outgassing, these can interfere with the stability of the signal (termed "dancing bubbles")
Observation of the data indicates the magnitude of these latter two effects on the accuracy of the data is likely small but is noted where observed
Turbidity
A bigger problem for the analysis was turbidity either in the form of sediment in samples collected close to the bottom or highly productive samples full of large diatoms as observed at several stations in the Bering Sea.
Turbidity likely has a significant impact on the titre as evidenced by the poor replication at some of these stations.
In one extreme case a sample was so turbid it could not be titrated as the voltage on the detector did not increase above the critical value needed to reach endpoint.
Stalled burette
A computer glitch was encountered during the course of the cruise whereby the thiosulfate burette would occasionally stall after dispensing 1.00ml volume and the titration would not proceed further.
This occurred exclusively at or near one of the 3 UV shift changing voltages where the software directs the dosimat to change its rate of thiosulfate delivery as the endpoint is approached.
This problem was overcome by trial and error by clicking the STOP THIO button on the software, the titration would then jump to "endpoint" mode where only small additions of thiosulfate were made to a finite volume.
One of 3 situations occurred at this point
1. If the titration was already at a high enough voltage the titration resumed until the endpoint was reached
2. If the titration was still nowhere near the endpoint the sample would finish before the endpoint was reached prompting the user to either "FINISH SAMPLE" or OVER-TITRATE",
if FS was selected then the sample info was re-entered and sample was rerun and the two titres added together
3. If OT was selected, a 1 ml aliquot of KIO3 standard was added to the sample and it was re-titrated
Either of the 3 options gave good results as evidenced by the low Sp of replicates where one or more of the samples experienced a stalled burette.
For further details on reagent prep, sampling, analysis, troubleshooting and quality control see full 2008-02 oxygen sampling report
QA/QC flags (had to be adapted slightly from standard IOS format to fit the situations observed)
b | acceptable measurement | - possible problem with the sample or analysis but impact on result is likely small i.e. floaties/dancing bubbles/burette stall
c | questionable measurement | -no known problem with sample or analysis but result is questionable either because it doesn't agree with CTD profile or the reps are highly variable
d | questionable measurement | - known problem with sample or analysis but not severe enough to discard i.e. turbidity, bottle integrity issues, result is likely not accurate
e | bad measurement | -known problem with sample or analysis major enough to discard measurement (i.e. sample taken from wrong bottle, major bubble in sample), or sample lost due to operator error

## 2008-02 OXY
| no |  | no | yes | [14 empty] | yes | yes | no | [6 empty] | yes | [27 empty] | no | yes
| EVENT NUMBER |  | STATION | Pressure:Nominal | [14 empty] | Bottle_Number | Sample_Number | [7 empty] | Bottle_Number | [27 empty] | STATION | Pressure:Nominal
|  |  |  | db | [52 empty] | db
| R4 |  | C | R4 | [14 empty] | R4 | R4 | [7 empty] | R4 | [27 empty] | C | R4
| R4:F6.0 |  | C | R4:F7.0 | [14 empty] | R4:F6.0 | R4:F6.0 | [7 empty] | R4:F6.0 | [27 empty] | C | R4:F7.0
| -99 |  | -99 | -99 | [14 empty] | -99 | -99 | [7 empty] | -99 | [27 empty] | -99 | -99
2008-02 chemistry sheet |  |  |  |  | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW | RAW
[5 empty] | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD | CTD |  |  |  | IOS |  |  |  | Final
| Event # | Cast # | Station | Target Pressure | Sal PSU | ptemp degC | O2 ml/l | O2 % | O2 uM/Kg | Scan | Raw Pressure dbar | Temp degC | Conductivity mS/cm | O2 V | Fluo mg/l | Trans % | PAR-ctd | PAR-ref | Sample #'s | Bottle #'s | Flask ID | Oxy-1 | Oxy-2 | Oxy-3 | QA/QC | Oxy | comments | Nut Sample #'s | Depth | NO3-1 | SiO4-1 | PO4-1 | O18 | Ba | Chl-a (JG) | D.O bottle (JG) | Nut Samp # | Depth | NO3-1 | NO3-2 | NO3 diff dups | SiO4-1 | SiO4-2 | SiO4 diff dups | PO4-1 | PO4-2 | PO4 diff dups | Nutrient Notes | O18 | Ba | Phosphate | Silicate | Nitrate + Nitrite | Ammonia | Chl-A | Station | Raw Nominal Press | END
[30 empty] | um/l | um/l | um/l |  |  | ug/L |  |  |  | um/l | um/l |  | um/l | um/l |  | um/l | um/l | [8 empty] | mg/L
| 001 | 001 | NP1 |  | 32.0323 | 13.0089 | 6.8979 | 114.3089 | 300.812 | 18912 | 7.796 | 13.0099 | 37.8495 | 4.4604 | 4.936 | 50.2702 | 6.2631 | 1.5684e+02 |  | NONE | [35 empty] | Test 1
[5 empty] | 32.0331 | 13.0038 | 6.9004 | 114.3381 | 300.919 | 18944 | 7.912 | 13.0049 | 37.8474 | 4.4613 | 4.9359 | 50.3335 | 5.7688 | 1.5684e+02
[5 empty] | 32.0327 | 13.0039 | 6.9031 | 114.3837 | 301.04 | 18982 | 7.997 | 13.005 | 37.8467 | 4.4629 | 4.9358 | 50.2371 | 5.6942 | 1.5684e+02
[5 empty] | 32.0297 | 13.0084 | 6.8982 | 114.3115 | 300.827 | 19560 | 7.977 | 13.0095 | 37.8379 | 4.4604 | 4.9356 | 49.9517 | 5.2611 | 1.5684e+02
[5 empty] | 32.0277 | 13.0244 | 6.8834 | 114.103 | 300.183 | 19703 | 7.856 | 13.0254 | 37.8483 | 4.4534 | 4.9356 | 49.8559 | 5.864 | 1.5702e+02
[5 empty] | 32.0289 | 13.017 | 6.9065 | 114.468 | 301.187 | 19753 | 7.954 | 13.0181 | 37.8477 | 4.4657 | 4.9356 | 49.7716 | 5.2289 | 1.5741e+02
[5 empty] | 32.0303 | 13.0154 | 6.9036 | 114.4187 | 301.064 | 19798 | 8.105 | 13.0165 | 37.849 | 4.4639 | 4.9356 | 49.6461 | 5.157 | 1.5901e+02
[5 empty] | 32.0301 | 13.014 | 6.8969 | 114.3037 | 300.771 | 19842 | 7.866 | 13.0151 | 37.8495 | 4.4602 | 4.9351 | 49.5754 | 6.2248 | 1.5901e+02
[5 empty] | 32.0281 | 13.0222 | 6.8954 | 114.2961 | 300.705 | 19885 | 7.634 | 13.0232 | 37.8545 | 4.46 | 4.9356 | 49.5729 | 6.9611 | 1.5901e+02
[5 empty] | 32.0259 | 13.031 | 6.8837 | 114.1221 | 300.196 | 19931 | 7.917 | 13.032 | 37.8506 | 4.454 | 4.9356 | 49.6017 | 5.646 | 1.5928e+02
[5 empty] | 32.0262 | 13.0325 | 6.8717 | 113.9272 | 299.674 | 19981 | 7.981 | 13.0335 | 37.8548 | 4.4474 | 4.9355 | 49.4729 | 5.7856 | 1.6119e+02
[5 empty] | 32.0265 | 13.036 | 6.8961 | 114.3397 | 300.736 | 20036 | 7.798 | 13.037 | 37.8633 | 4.4614 | 4.9355 | 49.3757 | 6.3573 | 1.6119e+02
[5 empty] | 32.0267 | 13.0337 | 6.9056 | 114.493 | 301.153 | 20091 | 8.049 | 13.0347 | 37.8608 | 4.4664 | 4.9355 | 49.2329 | 5.4991 | 1.6119e+02
[5 empty] | 32.0299 | 13.0211 | 6.8983 | 114.3439 | 300.833 | 20138 | 8.055 | 13.0222 | 37.8633 | 4.4614 | 4.9355 | 49.174 | 5.7218 | 1.6201e+02
[5 empty] | 32.025 | 13.0389 | 6.8986 | 114.3866 | 300.845 | 20184 | 7.807 | 13.04 | 37.8672 | 4.463 | 4.9355 | 49.2224 | 6.3652 | 1.6337e+02
[5 empty] | 32.0268 | 13.0351 | 6.9024 | 114.4431 | 301.013 | 20227 | 7.726 | 13.0362 | 37.8679 | 4.4649 | 4.9354 | 49.149 | 6.5796 | 1.6337e+02
[5 empty] | 32.0272 | 13.0314 | 6.9125 | 114.602 | 301.453 | 20270 | 7.926 | 13.0325 | 37.8641 | 4.4702 | 4.9353 | 49.0202 | 5.9362 | 1.6337e+02
[5 empty] | 32.0285 | 13.0236 | 6.9216 | 114.7339 | 301.847 | 20315 | 7.914 | 13.0246 | 37.8602 | 4.4747 | 4.9353 | 48.8968 | 6.3526 | 1.6342e+02
[5 empty] | 32.0265 | 13.0344 | 6.909 | 114.5503 | 301.3 | 20371 | 7.818 | 13.0354 | 37.8649 | 4.4685 | 4.9354 | 48.9042 | 6.1627 | 1.6481e+02
[5 empty] | 32.0293 | 13.0235 | 6.9152 | 114.6294 | 301.57 | 20414 | 8.013 | 13.0245 | 37.8552 | 4.4711 | 4.9355 | 49.1308 | 5.4015 | 1.6555e+02
[5 empty] | 32.0292 | 13.0224 | 6.9088 | 114.5207 | 301.291 | 20452 | 8.024 | 13.0235 | 37.8599 | 4.4674 | 4.9355 | 49.0636 | 5.6678 | 1.6555e+02
[5 empty] | 32.0274 | 13.0287 | 6.897 | 114.3384 | 300.776 | 20487 | 7.858 | 13.0298 | 37.8663 | 4.4613 | 4.9356 | 48.9827 | 6.3408 | 1.6555e+02
[5 empty] | 32.0286 | 13.0267 | 6.9074 | 114.5062 | 301.228 | 20527 | 7.872 | 13.0278 | 37.8607 | 4.467 | 4.9354 | 49.0418 | 6.1742 | 1.6679e+02
[5 empty] | 32.0291 | 13.0244 | 6.9119 | 114.5754 | 301.423 | 20573 | 7.924 | 13.0254 | 37.8616 | 4.4693 | 4.9355 | 49.0987 | 6.158 | 1.6773e+02
| 002 | 002 | NP1 | bot-5 | 33.9151 | 6.4359 | 1.6284 | 23.6082 | 70.837 | 11033 | 136.604 | 6.4478 | 33.8618 | 1.3839 | 0.2888 | 86.1764 | 0.0911 | 298.42 | 1 | 1 | 1336 | 1.669 |  |  |  | 1.669
|  |  |  | 100 | 33.8216 | 6.5819 | 1.5433 | 22.4365 | 67.143 | 13861 | 99.845 | 6.5907 | 33.8899 | 1.3484 | 0.26349 | 87.1598 | 0.089567 | 333.63 | 2 | 2 | 1135 | 1.624 |  |  | d | 1.624 | small bubble in sample
|  |  |  | 75 | 33.3898 | 7.0882 | 3.4522 | 50.6389 | 150.247 | 15464 | 75.193 | 7.0951 | 33.9352 | 2.2994 | 0.31826 | 89.7311 | 0.089633 | 302.78 | 3 | 3 | 1094 | 3.575 |  |  | d | 3.575 | tiny bubble added during reagent dispensing
|  |  |  | 50 | 32.8088 | 7.493 | 4.2955 | 63.3614 | 187.043 | 17123 | 49.999 | 7.4976 | 33.7427 | 2.734 | 0.37582 | 89.4663 | 0.091532 | 272.28 | 4 | 4 | 1293 | 4.431 |  |  |  | 4.431
|  |  |  | 25 | 32.3245 | 9.5414 | 5.6658 | 87.2763 | 246.879 | 19060 | 25.087 | 9.5441 | 35.0584 | 3.5468 | 0.46999 | 88.355 | 0.22573 | 265.75 | 5 | 5 | 1654 | 5.935 |  |  |  | 5.935
|  |  |  | 10 | 32.0882 | 12.4088 | 6.5918 | 107.9025 | 287.419 | 20873 | 10.218 | 12.4102 | 37.3644 | 4.2442 | 3.4145 | 71.302 | 2.8633 | 254.01 | 6 | 6 | 1127, 1284, 1143 | 6.795 | 6.798 | 6.796 |  | 6.795
|  |  |  | 5 | 32.015 | 13.1287 | 6.932 | 115.1497 | 302.311 | 23078 | 5.992 | 13.1295 | 37.9427 | 4.4895 | 4.7647 | 45.4772 | 31.249 | 274.9 | 7 | 7 | 1089 | 7.204 |  |  |  | 7.204
| 003 | 003 | NP1 | 133 | 33.9157 | 6.4611 | 1.713 | 24.8492 | 74.517 | 15312 | 133.189 | 6.4727 | 33.8825 | 1.4257 | 0.073584 | 86.3301 | 0.080733 | 589.06
|  |  |  | 100 | 33.8919 | 6.5265 | 1.8294 | 26.5727 | 79.583 | 18253 | 100.217 | 6.5353 | 33.9033 | 1.4872 | 0.070943 | 87.6794 | 0.068546 | 593.96
|  |  |  | 80 | 33.4362 | 7.0701 | 3.2642 | 47.8761 | 142.059 | 20041 | 79.708 | 7.0774 | 33.9633 | 2.2055 | 0.10525 | 89.582 | 0.069628 | 609.91
|  |  |  | 80 | 33.4352 | 7.0708 | 3.2611 | 47.8315 | 141.926 | 20086 | 79.5 | 7.0781 | 33.9627 | 2.204 | 0.092109 | 89.5417 | 0.067835 | 610.95
|  |  |  | 75 | 33.3222 | 7.1302 | 3.4481 | 50.606 | 150.077 | 21526 | 72.654 | 7.1368 | 33.9099 | 2.2988 | 0.11525 | 89.6368 | 0.067715 | 541.63
|  |  |  | 75 | 33.3245 | 7.1289 | 3.4548 | 50.7033 | 150.368 | 21576 | 72.495 | 7.1355 | 33.9079 | 2.3022 | 0.11859 | 89.6554 | 0.06724 | 537.99
|  |  |  | 60 | 33.0742 | 7.23 | 3.9306 | 57.7265 | 171.112 | 23364 | 60.506 | 7.2356 | 33.7629 | 2.5415 | 0.13344 | 89.7933 | 0.069282 | 495.73
|  |  |  | 60 | 33.0749 | 7.2297 | 3.9287 | 57.6999 | 171.033 | 23407 | 60.484 | 7.2353 | 33.763 | 2.5406 | 0.14337 | 89.7792 | 0.066026 | 498.59
|  |  |  | 50 | 32.9086 | 7.3595 | 4.1368 | 60.871 | 180.115 | 24965 | 49.561 | 7.3641 | 33.7182 | 2.6504 | 0.15537 | 89.6248 | 0.069763 | 427.87
|  |  |  | 50 | 32.913 | 7.355 | 4.1346 | 60.8351 | 180.022 | 25024 | 49.526 | 7.3595 | 33.7184 | 2.6492 | 0.16277 | 89.6803 | 0.070969 | 426.94
|  |  |  | 25 | 32.264 | 10.109 | 5.9666 | 93.0384 | 260.022 | 27044 | 25.028 | 10.1118 | 35.5144 | 3.7404 | 0.65417 | 87.1833 | 0.18659 | 373.55
|  |  |  | 25 | 32.2632 | 10.1143 | 5.9568 | 92.8951 | 259.593 | 27090 | 25.21 | 10.1171 | 35.5167 | 3.7354 | 0.64962 | 87.2126 | 0.19261 | 377.34
|  |  |  | 20 | 32.1659 | 11.4569 | 6.4098 | 102.8633 | 279.417 | 28636 | 19.234 | 11.4592 | 36.6041 | 4.0719 | 1.402 | 83.8747 | 0.32642 | 346.34
|  |  |  | 20 | 32.1658 | 11.4578 | 6.42 | 103.0305 | 279.866 | 28679 | 19.271 | 11.4601 | 36.6046 | 4.0775 | 1.4874 | 83.8678 | 0.33936 | 345.12
|  |  |  | 5 | 32.0069 | 13.1837 | 6.9629 | 115.7906 | 303.665 | 30478 | 4.564 | 13.1843 | 37.9657 | 4.5118 | 4.8103 | 44.6373 | 102.87 | 482.96
|  |  |  | 5 | 32.0088 | 13.175 | 6.9643 | 115.7939 | 303.724 | 30518 | 4.682 | 13.1757 | 37.9663 | 4.5118 | 4.8089 | 44.5701 | 91.995 | 477.88
|  |  |  | 5 | 32.0095 | 13.173 | 6.9649 | 115.7998 | 303.751 | 30557 | 4.819 | 13.1737 | 37.9693 | 4.5119 | 4.8084 | 44.4698 | 89.365 | 488.56
|  |  |  | 5 | 32.0094 | 13.1731 | 6.9673 | 115.8394 | 303.855 | 30596 | 4.865 | 13.1737 | 37.9696 | 4.5132 | 4.8087 | 44.479 | 86.249 | 488.67
|  |  |  | 5 | 32.0108 | 13.1708 | 6.9718 | 115.9092 | 304.049 | 30641 | 4.792 | 13.1715 | 37.9665 | 4.5157 | 4.8081 | 44.5466 | 91.253 | 477.36
|  |  |  | 5 | 32.0093 | 13.1741 | 6.9701 | 115.8878 | 303.975 | 30800 | 4.65 | 13.1748 | 37.9689 | 4.515 | 4.8095 | 44.5597 | 99.198 | 488.29
|  |  |  | 0 | 32.0087 | 13.1819 | 7.0648 | 117.4799 | 308.105 | 31987 | 2.47 | 13.1822 | 37.9682 | 4.57 | 4.8512 | 44.3781 | 301.2 | 522.27
|  |  |  | 0 | 32.009 | 13.1789 | 7.057 | 117.3434 | 307.765 | 32032 | 2.505 | 13.1792 | 37.97 | 4.5654 | 4.8518 | 44.3242 | 299.09 | 518.25
| 004 | 004 | NP2 | 1000 | 34.3622 | 3.3273 | 0.21846 | 2.9505 | 9.497 | 34871 | 1002.011 | 3.3983 | 31.9293 | 0.6837 | 0.056266 | 90.6585 | 0.069708 | 8.7131 | 30 | 1 | 1746 | 0.202 |  |  |  | 0.202
|  |  |  | 800 | 34.2457 | 3.7816 | 0.27952 | 3.8139 | 12.153 | 42375 | 801.014 | 3.8397 | 32.1341 | 0.7125 | 0.055418 | 90.694 | 0.067656 | 7.6239 | 31 | 2 | 1145 | 0.277 |  |  |  | 0.277
|  |  |  | 600 | 34.1086 | 4.1189 | 0.59124 | 8.1241 | 25.709 | 48405 | 600.466 | 4.1629 | 32.2134 | 0.851 | 0.054596 | 90.7529 | 0.067047 | 6.5348 | 32 | 3 | 1748 | 0.597 |  |  |  | 0.597
|  |  |  | 400 | 33.9627 | 4.8074 | 1.3456 | 18.7779 | 58.52 | 54551 | 401.075 | 4.8382 | 32.5952 | 1.2014 | 0.052777 | 90.6533 | 0.068148 | 6.7526 | 33 | 4 | 1658 | 1.387 |  |  |  | 1.387
|  |  |  | 300 | 33.8841 | 5.4611 | 2.473 | 35.0366 | 107.569 | 58272 | 300.522 | 5.4855 | 33.0539 | 1.7405 | 0.054222 | 90.4709 | 0.066297 | 7.4061 | 34 | 5 | 1156 | 2.559 |  |  |  | 2.559
|  |  |  | 250 | 33.877 | 6.0643 | 2.8382 | 40.7875 | 123.463 | 61268 | 250.972 | 6.0857 | 33.5581 | 1.9368 | 0.052378 | 90.3765 | 0.068001 | 7.1794 | 35 | 6 | 663 | 2.972 |  |  |  | 2.972
|  |  |  | 200 | 33.8441 | 6.5312 | 3.2843 | 47.7068 | 142.882 | 64163 | 200.323 | 6.5489 | 33.919 | 2.1745 | 0.053973 | 90.3301 | 0.067116 | 7.1883 | 36 | 7 | 1376 | 3.445 |  |  |  | 3.445
|  |  |  | 175 | 33.789 | 6.7743 | 3.6377 | 53.1196 | 158.265 | 65905 | 175.352 | 6.7901 | 34.0737 | 2.3593 | 0.055468 | 90.3315 | 0.066458 | 7.1883 | 37 | 8 | 1137 | 3.785 |  |  |  | 3.785
|  |  |  | 150 | 33.672 | 6.9152 | 3.996 | 58.4976 | 173.876 | 67710 | 150.759 | 6.9289 | 34.077 | 2.544 | 0.057561 | 90.3593 | 0.066794 | 7.3661 | 38 | 9 | 1023 | 4.199 |  |  |  | 4.199
|  |  |  | 125 | 33.2766 | 6.8778 | 4.9794 | 72.6386 | 216.729 | 69408 | 126.001 | 6.8891 | 33.672 | 3.0232 | 0.072213 | 90.4364 | 0.067876 | 6.9705 | 39 | 10 | 1022, 1666, 1155 | 5.15 | 5.149 | 5.149 |  | 5.15
|  |  |  | 100 | 32.6955 | 6.584 | 6.1391 | 88.6007 | 267.315 | 71422 | 100.161 | 6.5926 | 32.8724 | 3.5675 | 0.23665 | 90.3234 | 0.070915 | 6.7526 | 40 | 11 | 1136 | 6.384 |  |  |  | 6.384
|  |  |  | 75 | 32.5411 | 7.0076 | 6.5251 | 95.0101 | 284.169 | 73364 | 75.115 | 7.0143 | 33.0851 | 3.7923 | 0.61499 | 89.8359 | 0.14155 | 6.5348 | 41 | 12 | 970 | 6.829 |  |  |  | 6.829
|  |  |  | 50 | 32.4879 | 8.7804 | 6.5883 | 99.8786 | 287.007 | 75205 | 51.059 | 8.7856 | 34.5648 | 3.9628 | 1.782 | 88.3976 | 0.51266 | 6.5348 | 42 | 13 | 1699 | 6.881 |  |  |  | 6.881
|  |  |  | 25 | 32.4185 | 10.601 | 6.6123 | 104.3279 | 288.148 | 77557 | 23.607 | 10.6037 | 36.2386 | 4.1214 | 1.5261 | 86.3674 | 4.1364 | 6.9705 | 43 | 14 | 1724 | 6.912 |  |  |  | 6.912
|  |  |  | 10 | 32.3763 | 11.8696 | 6.3353 | 102.7063 | 276.15 | 78839 | 10.333 | 11.8709 | 37.1868 | 4.0697 | 1.2226 | 85.8867 | 11.754 | 7.1883 | 44 | 15 | 1738 | 6.649 |  |  |  | 6.649
|  |  |  | 5 | 32.378 | 11.8812 | 6.3906 | 103.6264 | 278.557 | 80305 | 2.638 | 11.8815 | 37.1928 | 4.1045 | 1.2503 | 85.6899 | 27.841 | 6.7659 | 45 | 16 | 1722 | 6.652 |  |  |  | 6.652
| 005 | 005 | NP3 | 1000 | 34.3713 | 3.2292 | 0.24133 | 3.2516 | 10.49 | 33004 | 1001.731 | 3.2994 | 31.8502 | 0.6927 | 0.055169 | 90.9683 | 0.066725 | 116.12 | 46 | 1 | 1697 | 0.231 |  |  | b | 0.231 | dancing bubbles
|  |  |  | 800 | 34.2738 | 3.5929 | 0.2738 | 3.7191 | 11.904 | 40051 | 804.687 | 3.6501 | 31.9919 | 0.7095 | 0.055518 | 91.0081 | 0.074178 | 103.76 | 47 | 2 | 1743 | 0.268 |  |  |  | 0.268
|  |  |  | 600 | 34.1793 | 4.269 | 0.385 | 5.3124 | 16.741 | 46308 | 600.449 | 4.3138 | 32.4068 | 0.7627 | 0.054795 | 90.9861 | 0.065945 | 106.74 | 48 | 3 | 1739 | 0.396 |  |  |  | 0.396
|  |  |  | 400 | 34.0534 | 5.1663 | 0.84615 | 11.9188 | 36.799 | 52504 | 399.95 | 5.1982 | 32.992 | 0.9802 | 0.055219 | 90.897 | 0.070226 | 119.8 | 49 | 4 | 1725 | 0.87 |  |  |  | 0.87
|  |  |  | 300 | 33.98 | 5.6076 | 1.4086 | 20.0392 | 61.266 | 56231 | 301.889 | 5.6325 | 33.2701 | 1.2504 | 0.054272 | 90.7883 | 0.066198 | 124.16 | 50 | 5 | 1731 | 1.459 |  |  |  | 1.459
|  |  |  | 250 | 33.9416 | 5.9524 | 1.8956 | 27.1804 | 82.454 | 59080 | 250.349 | 5.9736 | 33.5154 | 1.4896 | 0.05193 | 90.7467 | 0.068115 | 124.16 | 51 | 6 | 1193 | 1.937 |  |  |  | 1.937
|  |  |  | 200 | 33.8859 | 6.4043 | 2.5365 | 36.7443 | 110.343 | 61677 | 199.867 | 6.4218 | 33.8431 | 1.8119 | 0.050908 | 90.6772 | 0.067076 | 117.63 | 52 | 7 | 1750 | 2.647 |  |  |  | 2.647
|  |  |  | 175 | 33.8512 | 6.6311 | 2.712 | 39.4858 | 117.985 | 63581 | 175.296 | 6.6467 | 34.001 | 1.9068 | 0.049986 | 90.6331 | 0.065282 | 123.09 | 53 | 8 | 1744 | 2.782 |  |  |  | 2.782
|  |  |  | 150 | 33.733 | 6.7911 | 3.2651 | 47.6773 | 142.061 | 65455 | 150.122 | 6.8046 | 34.0235 | 2.184 | 0.054247 | 90.5656 | 0.06581 | 140.92 | 54 | 9 | 1740 | 3.263 |  |  |  | 3.263
|  |  |  | 125 | 33.5209 | 6.9614 | 3.6282 | 53.1153 | 157.891 | 67291 | 124.455 | 6.9726 | 33.9689 | 2.3711 | 0.077695 | 90.6222 | 0.06632 | 139.41 | 55 | 10 | 1726, 1732, 1752 | 3.719 | 3.718 | 3.719 |  | 3.719
|  |  |  | 100 | 32.9189 | 6.6777 | 5.3719 | 77.8115 | 233.869 | 68880 | 99.758 | 6.6864 | 33.1571 | 3.2056 | 0.19877 | 90.5412 | 0.069691 | 137.67 | 56 | 11 | 1709 | 5.452 |  |  | d | 5.452 | ugly titration curve
|  |  |  | 75 | 32.5802 | 6.5339 | 6.4283 | 92.5906 | 279.93 | 70800 | 75.227 | 6.5403 | 32.7124 | 3.7119 | 0.4132 | 90.3492 | 0.16136 | 141.59 | 57 | 12 | 1745 | 6.662 |  |  |  | 6.662
|  |  |  | 50 | 32.4265 | 7.6574 | 6.689 | 98.7974 | 291.357 | 72596 | 50.037 | 7.6621 | 33.5327 | 3.9295 | 1.5177 | 89.0421 | 0.75596 | 146.16 | 58 | 13 | 1741 | 6.984 |  |  |  | 6.984
|  |  |  | 25 | 32.2833 | 9.6984 | 6.6619 | 102.9522 | 290.297 | 74560 | 24.857 | 9.7011 | 35.1697 | 4.0765 | 2.1111 | 86.8352 | 7.4402 | 152.48 | 59 | 14 | 1727 | 6.998 |  |  |  | 6.998
|  |  |  | 10 | 32.2405 | 12.0229 | 6.2678 | 101.8569 | 273.24 | 76054 | 9.378 | 12.0241 | 37.1893 | 4.0411 | 0.82617 | 87.8331 | 33.97 | 159.23 | 60 | 15 | 1733 | 6.572 |  |  |  | 6.572
|  |  |  | 5 | 32.2432 | 11.9439 | 6.2904 | 102.0536 | 274.222 | 78616 | 5.778 | 11.9447 | 37.1225 | 4.0496 | 0.84399 | 87.5094 | 44.955 | 161.19 | 61 | 16 | 1753 | 6.581 |  |  |  | 6.581
|  |  |  | Chl max | 32.3224 | 8.9506 | 6.7211 | 102.1704 | 292.834 | 80574 | 32.814 | 8.954 | 34.5449 | 4.0482 | 2.6604 | 87.1139 | 3.2857 | 152.48
| 006 | 006 | NP4 | 1000 | 34.3733 | 3.0429 | 0.25082 | 3.3637 | 10.903 | 37805 | 998.76 | 3.1114 | 31.6856 | 0.6961 | 0.092821 | 91.0157 | 0.068813 | 424.76 | 62 | 1 | 1135 | 0.24 |  |  |  | 0.24
|  |  |  | 800 | 34.2986 | 3.4439 | 0.25016 | 3.3859 | 10.876 | 43733 | 802.737 | 3.5 | 31.8798 | 0.6994 | 0.078941 | 91.0234 | 0.071455 | 411.65 | 63 | 2 | 1094 | 0.24 |  |  |  | 0.24
|  |  |  | 600 | 34.175 | 3.8126 | 0.46228 | 6.307 | 20.1 | 49589 | 601.213 | 3.8553 | 32.0001 | 0.794 | 0.075602 | 91.0445 | 0.06884 | 263.57 | 64 | 3 | 1293 | 0.466 |  |  |  | 0.466
|  |  |  | 400 | 34.0325 | 4.5119 | 0.85168 | 11.8059 | 37.038 | 55446 | 400.884 | 4.5418 | 32.3937 | 0.9767 | 0.077396 | 90.9966 | 0.065772 | 243.26 | 65 | 4 | 1654 | 0.864 |  |  |  | 0.864
|  |  |  | 300 | 33.934 | 4.9025 | 1.4799 | 20.6932 | 64.367 | 59789 | 300.706 | 4.9255 | 32.6029 | 1.2722 | 0.076275 | 90.8146 | 0.0661 | 224.32 | 66 | 5 | 1127 | 1.527 |  |  |  | 1.527
|  |  |  | 250 | 33.8992 | 5.1387 | 1.8888 | 26.553 | 82.154 | 62110 | 250.393 | 5.1582 | 32.7558 | 1.4695 | 0.078194 | 90.8084 | 0.065177 | 222.18 | 67 | 6 | 1284 | 1.941 |  |  |  | 1.941
|  |  |  | 200 | 33.8568 | 5.4823 | 2.5182 | 35.6817 | 109.538 | 64235 | 200.397 | 5.4984 | 32.9966 | 1.7775 | 0.084324 | 90.8103 | 0.065286 | 220 | 68 | 7 | 1143 | 2.614 |  |  |  | 2.614
|  |  |  | 175 | 33.816 | 5.6339 | 3.323 | 47.2419 | 144.553 | 65801 | 176.079 | 5.6482 | 33.0833 | 2.1652 | 0.099749 | 90.7988 | 0.067776 | 231.3 | 69 | 8 | 1089 | 3.404 |  |  |  | 3.404
|  |  |  | 150 | 33.7607 | 5.8995 | 3.6235 | 51.8197 | 157.636 | 67243 | 150.808 | 5.9121 | 33.2494 | 2.3229 | 0.11475 | 90.8184 | 0.070527 | 236.81 | 70 | 9 | 1509 | 3.737 |  |  |  | 3.737
|  |  |  | 125 | 33.5535 | 6.0892 | 4.1224 | 59.1365 | 179.37 | 68777 | 126.048 | 6.0998 | 33.2261 | 2.573 | 0.11286 | 90.8578 | 0.066383 | 249.65 | 71 | 10 | 879 | 4.269 |  |  |  | 4.269
|  |  |  | 100 | 32.9079 | 6.1553 | 5.6052 | 80.1909 | 244.013 | 70849 | 100.443 | 6.1637 | 32.6922 | 3.2861 | 0.25193 | 90.7294 | 0.11141 | 263.61 | 72 | 11 | 1746 | 5.935 |  |  |  | 5.935
|  |  |  | 75 | 32.546 | 6.3401 | 6.7298 | 96.4688 | 293.057 | 72405 | 76.153 | 6.3465 | 32.5147 | 3.8424 | 0.93399 | 89.7316 | 0.33199 | 281.4 | 73 | 12 | 1145 | 6.985 |  |  |  | 6.985
|  |  |  | 50 | 32.5188 | 7.2146 | 6.8485 | 100.1818 | 298.268 | 73978 | 50.675 | 7.2192 | 33.2395 | 3.9769 | 1.8237 | 88.6202 | 1.9088 | 306.56 | 74 | 13 | 1748, 1658, 1156 | 7.115 | 7.112 | 7.111 |  | 7.115
|  |  |  | 25 | 32.4935 | 9.41 | 6.6083 | 101.6056 | 287.902 | 75830 | 25.817 | 9.4128 | 35.1228 | 4.0312 | 1.1081 | 88.5422 | 9.3173 | 302.78 | 75 | 14 | 663 | 6.877 |  |  |  | 6.877
|  |  |  | 10 | 32.4628 | 10.9479 | 6.341 | 100.8304 | 276.332 | 79250 | 7.097 | 10.9487 | 36.4463 | 4.01 | 0.70048 | 88.0998 | 41.13 | 341.99 | 76 | 15 | 1376 | 6.654 |  |  |  | 6.654
| 007 | 007 | NP5 |  | 32.5597 | 7.5327 | 6.9489 | 102.4255 | 302.643 | 11479 | 30.317 | 7.5355 | 33.5439 | 4.0614 | 1.4963 | 86.6159 | 1.8349 | 116.67
[5 empty] | 32.5604 | 7.5281 | 6.9417 | 102.3094 | 302.33 | 11525 | 30.073 | 7.5309 | 33.5418 | 4.0576 | 1.4445 | 86.626 | 1.9219 | 117.63
[5 empty] | 32.5599 | 7.5315 | 6.9435 | 102.3435 | 302.409 | 11573 | 30.029 | 7.5343 | 33.5403 | 4.0588 | 1.4379 | 86.6911 | 1.9724 | 117.63
[5 empty] | 32.5601 | 7.5406 | 6.9516 | 102.4844 | 302.761 | 11642 | 30.528 | 7.5434 | 33.546 | 4.0633 | 1.4708 | 86.705 | 1.8409 | 117.63
[5 empty] | 32.5601 | 7.5405 | 6.9379 | 102.2818 | 302.163 | 11724 | 30.172 | 7.5433 | 33.5469 | 4.0566 | 1.4753 | 86.5122 | 1.9604 | 119.3
[5 empty] | 32.5604 | 7.5425 | 6.9306 | 102.179 | 301.845 | 11764 | 30.382 | 7.5453 | 33.5474 | 4.053 | 1.4585 | 86.454 | 1.8944 | 119.8
[5 empty] | 32.5605 | 7.5417 | 6.9369 | 102.2708 | 302.122 | 11802 | 30.238 | 7.5445 | 33.5472 | 4.0562 | 1.4653 | 86.5924 | 1.8977 | 119.8
[5 empty] | 32.5609 | 7.5396 | 6.9219 | 102.0451 | 301.469 | 11838 | 29.8 | 7.5424 | 33.5455 | 4.0488 | 1.5702 | 86.6301 | 2.0212 | 119.8
[5 empty] | 32.5607 | 7.5397 | 6.93 | 102.164 | 301.82 | 11874 | 29.772 | 7.5425 | 33.5466 | 4.0528 | 1.3937 | 86.6215 | 2.0813 | 121.62
[5 empty] | 32.5608 | 7.5424 | 6.9487 | 102.4471 | 302.637 | 11909 | 30.305 | 7.5452 | 33.5453 | 4.0621 | 1.4021 | 86.555 | 2.0038 | 121.98
[5 empty] | 32.5611 | 7.5395 | 6.9528 | 102.5007 | 302.815 | 11945 | 30.655 | 7.5423 | 33.5415 | 4.0638 | 1.5887 | 86.4173 | 1.9148 | 121.98
[5 empty] | 32.5611 | 7.5382 | 6.9334 | 102.2107 | 301.967 | 11983 | 30.46 | 7.541 | 33.5357 | 4.0541 | 1.3637 | 86.5146 | 1.9152 | 121.98
[5 empty] | 32.5613 | 7.5366 | 6.9355 | 102.2389 | 302.061 | 12027 | 30.148 | 7.5394 | 33.5392 | 4.0552 | 1.449 | 86.52 | 1.9974 | 121.98
[5 empty] | 32.5575 | 7.5307 | 6.951 | 102.4511 | 302.737 | 12084 | 30.067 | 7.5335 | 33.5358 | 4.0624 | 1.4754 | 86.6522 | 2.0619 | 123.34
[5 empty] | 32.5636 | 7.5073 | 6.9513 | 102.4031 | 302.745 | 12125 | 30.249 | 7.5101 | 33.5271 | 4.0608 | 1.4705 | 86.553 | 2.0294 | 124.1
[5 empty] | 32.5643 | 7.4757 | 6.9498 | 102.3067 | 302.678 | 12162 | 30.415 | 7.4785 | 33.5091 | 4.0575 | 1.5157 | 86.7036 | 1.9682 | 124.16
[5 empty] | 32.5632 | 7.4953 | 6.9296 | 102.0552 | 301.801 | 12197 | 30.305 | 7.4981 | 33.4835 | 4.049 | 1.4067 | 86.6296 | 1.9932 | 124.16
[5 empty] | 32.5629 | 7.4977 | 6.9189 | 101.903 | 301.335 | 12230 | 29.95 | 7.5004 | 33.5058 | 4.044 | 1.4726 | 86.2809 | 2.0834 | 124.16
[5 empty] | 32.5638 | 7.4964 | 6.9293 | 102.0535 | 301.787 | 12262 | 29.849 | 7.4992 | 33.5055 | 4.0491 | 1.4147 | 86.3845 | 2.1686 | 124.16
[5 empty] | 32.565 | 7.4815 | 6.9481 | 102.2955 | 302.603 | 12297 | 30.288 | 7.4843 | 33.5042 | 4.0571 | 1.9962 | 86.6367 | 2.059 | 124.16
[5 empty] | 32.5641 | 7.4845 | 6.9535 | 102.3827 | 302.842 | 12335 | 30.665 | 7.4873 | 33.4872 | 4.0599 | 1.5654 | 86.6302 | 1.9252 | 124.16
[5 empty] | 32.5629 | 7.493 | 6.9571 | 102.4546 | 302.998 | 12372 | 30.24 | 7.4958 | 33.4912 | 4.0626 | 1.4265 | 86.5277 | 1.9964 | 124.16
[5 empty] | 32.5634 | 7.4987 | 6.9597 | 102.5067 | 303.112 | 12415 | 29.756 | 7.5015 | 33.5061 | 4.0645 | 1.4211 | 86.516 | 2.1921 | 124.16
| 008 | 008 | NP5 | 250 | 34.3726 | 2.9899 | 0.26885 | 3.6007 | 11.687 | 29402 | 999.538 | 3.058 | 31.6388 | 0.7031 | 0.10077 | 91.0287 | 0.066202 | 213.47
|  |  |  | 250 | 33.8823 | 4.9952 | 1.7166 | 24.0461 | 74.664 | 49235 | 250.011 | 5.0144 | 32.6139 | 1.3871 | 0.11163 | 90.8534 | 0.069505 | 278.38
|  |  |  | 110 | 33.8823 | 4.9951 | 1.7211 | 24.1089 | 74.859 | 49287 | 250.229 | 5.0143 | 32.6139 | 1.3892 | 0.11143 | 90.8347 | 0.066759 | 276.64
|  |  |  | 110 | 32.9418 | 5.4072 | 5.561 | 78.169 | 242.061 | 53914 | 109.561 | 5.4157 | 32.0809 | 3.2165 | 0.066428 | 90.7564 | 0.092615 | 309.17
|  |  |  | 100 | 32.9423 | 5.4083 | 5.5589 | 78.1418 | 241.97 | 53962 | 109.587 | 5.4168 | 32.086 | 3.2155 | 0.06718 | 90.79 | 0.086941 | 309.31
|  |  |  | 100 | 32.7284 | 5.1856 | 6.3736 | 88.988 | 277.473 | 55395 | 99.395 | 5.1931 | 31.694 | 3.5837 | 0.061985 | 90.7743 | 0.11659 | 326.74
|  |  |  | 43 | 32.7254 | 5.18 | 6.3666 | 88.8764 | 277.168 | 55443 | 99.531 | 5.1874 | 31.6881 | 3.5799 | 0.061915 | 90.7695 | 0.11948 | 326.74
|  |  |  | 43 | 32.5681 | 7.26 | 6.9463 | 101.7506 | 302.517 | 58268 | 42.96 | 7.2639 | 33.3108 | 4.0334 | 1.4235 | 87.1217 | 2.096 | 337.83
|  |  |  | 40 | 32.5694 | 7.2523 | 6.9399 | 101.6395 | 302.238 | 58314 | 43.28 | 7.2562 | 33.3103 | 4.0295 | 1.8233 | 87.1495 | 2.071 | 338.72
|  |  |  | 40 | 32.561 | 7.5759 | 6.9904 | 103.1435 | 304.454 | 59987 | 38.268 | 7.5795 | 33.5815 | 4.0819 | 1.5043 | 86.34 | 3.4468 | 354.83
|  |  |  | 13 | 32.5606 | 7.5766 | 6.9892 | 103.1272 | 304.402 | 60035 | 37.909 | 7.5801 | 33.5819 | 4.0815 | 1.618 | 86.388 | 3.5702 | 355.06
|  |  |  | 13 | 32.4822 | 10.3133 | 6.4653 | 101.4056 | 281.715 | 61820 | 12.934 | 10.3147 | 35.9027 | 4.0283 | 0.7858 | 87.1179 | 33.185 | 363.82
|  |  |  | 5 | 32.4912 | 10.2241 | 6.4502 | 100.9763 | 281.051 | 61865 | 13.598 | 10.2256 | 35.7963 | 4.0137 | 0.83354 | 87.0625 | 31.506 | 363.77
|  |  |  | 5 | 32.4871 | 10.4565 | 6.4409 | 101.342 | 280.658 | 64008 | 4.17 | 10.4569 | 36.0242 | 4.0298 | 0.65953 | 87.2007 | 67.495 | 370.71
|  |  |  | 5 | 32.4893 | 10.4501 | 6.449 | 101.457 | 281.011 | 64045 | 4.394 | 10.4506 | 36.0267 | 4.0337 | 0.64743 | 87.189 | 64.886 | 370.31
|  |  |  | 5 | 32.489 | 10.436 | 6.4473 | 101.3984 | 280.935 | 64081 | 5.437 | 10.4366 | 36.013 | 4.0312 | 0.58496 | 87.1932 | 59.426 | 370.31
|  |  |  | 5 | 32.4891 | 10.4261 | 6.4428 | 101.3066 | 280.741 | 64118 | 5.146 | 10.4267 | 36.0138 | 4.0283 | 0.62378 | 87.1968 | 59.236 | 370.42
|  |  |  | 5 | 32.4884 | 10.4363 | 6.4351 | 101.2074 | 280.406 | 64157 | 4.488 | 10.4368 | 36.0134 | 4.0252 | 0.63765 | 87.1803 | 65.892 | 370.31
|  |  |  | 5 | 32.4886 | 10.4365 | 6.4463 | 101.3842 | 280.894 | 64200 | 4.756 | 10.437 | 36.0149 | 4.031 | 0.61317 | 87.191 | 62.669 | 371.44
|  |  |  | 5 | 32.4892 | 10.4203 | 6.4519 | 101.4368 | 281.137 | 64273 | 4.672 | 10.4208 | 36.006 | 4.0329 | 0.69258 | 87.1723 | 63.73 | 372.48
[5 empty] | 32.489 | 10.4218 | 6.4628 | 101.6106 | 281.61 | 64359 | 4.948 | 10.4223 | 36.0026 | 4.0387 | 0.58503 | 87.1991 | 63.28 | 372.48
| 009 | 009 | NP5 | 3000 | 34.6515 | 1.3805 | 2.3535 | 30.4317 | 102.272 | 78891 | 3000.848 | 1.5994 | 31.4005 | 1.2883 | 0.080337 | 91.2853 | 0.055301 | 5.1509 |  | 1
|  |  |  | 3000 | 34.6516 | 1.3803 | 2.3543 | 30.441 | 102.304 | 78944 | 2999.917 | 1.5992 | 31.4001 | 1.2886 | 0.080262 | 91.2863 | 0.055301 | 5.1603 |  | 2
|  |  |  | 3000 | 34.6514 | 1.3805 | 2.3545 | 30.4435 | 102.312 | 79000 | 2999.992 | 1.5994 | 31.4001 | 1.2886 | 0.080337 | 91.282 | 0.055301 | 5.1434 | 100 | 3 | 1135 | 2.416 |  |  |  | 2.416
|  |  |  | 2500 | 34.628 | 1.5354 | 1.8708 | 24.2547 | 81.296 | 92363 | 2501.405 | 1.7103 | 31.2881 | 1.1862 | 0.078443 | 91.293 | 0.055301 | 5.8378 | 101 | 4 | 1094 | 1.919 |  |  |  | 1.919
|  |  |  | 2000 | 34.5834 | 1.7995 | 1.2329 | 16.0716 | 53.578 | 106185 | 2000.773 | 1.9349 | 31.2495 | 1.0143 | 0.070544 | 91.2288 | 0.055301 | 6.1725 | 102 | 5 | 1293, 1654, 1127 | 1.261 | 1.258 | 1.262 |  | 1.261
|  |  |  | 1500 | 34.5027 | 2.2336 | 0.57234 | 7.5332 | 24.875 | 119922 | 1503.075 | 2.3338 | 31.3276 | 0.8056 | 0.067255 | 91.2058 | 0.055301 | 6.9305 | 103 | 6 | 1284 | 0.579 |  |  |  | 0.579
|  |  |  | 1250 | 34.4453 | 2.5496 | 0.36012 | 4.7742 | 15.653 | 127404 | 1251.725 | 2.6334 | 31.4371 | 0.7333 | 0.066956 | 91.1661 | 0.055301 | 9.1709 | 104 | 7 | 1143 | 0.357 |  |  |  | 0.357
|  |  |  | 1000 | 34.3713 | 2.998 | 0.26085 | 3.4943 | 11.339 | 134600 | 1000.971 | 3.0662 | 31.6467 | 0.6999 | 0.072438 | 91.1254 | 0.055301 | 7.7635 | 105 | 8 | 1089 | 0.254 |  |  |  | 0.254
|  |  |  | 800 | 34.2912 | 3.3918 | 0.26441 | 3.5739 | 11.495 | 140929 | 802.212 | 3.4475 | 31.8281 | 0.7051 | 0.076873 | 91.101 | 0.055301 | 9.2803 | 106 | 9 | 1509 | 0.263 |  |  |  | 0.263
|  |  |  | 600 | 34.1774 | 3.8973 | 0.38808 | 5.3059 | 16.874 | 146789 | 601.039 | 3.9404 | 32.0783 | 0.7625 | 0.08243 | 91.1182 | 0.055301 | 9.1429 | 107 | 10 | 879 | 0.4 |  |  |  | 0.4
|  |  |  | 400 | 34.0213 | 4.3472 | 0.7819 | 10.7941 | 34.002 | 153155 | 402.56 | 4.3767 | 32.2401 | 0.944 | 0.085146 | 91.0502 | 0.055301 | 9.5648 | 108 | 11 | 1746 | 0.809 |  |  |  | 0.809
|  |  |  | 300 | 33.9294 | 4.7046 | 1.2521 | 17.4226 | 54.455 | 156987 | 299.618 | 4.7271 | 32.423 | 1.1655 | 0.082081 | 90.9205 | 0.055301 | 1.0085 | 109 | 12 | 1145 | 1.303 |  |  |  | 1.303
|  |  |  | 250 | 33.8829 | 5.0113 | 1.7251 | 24.1747 | 75.034 | 159494 | 250.117 | 5.0306 | 32.6289 | 1.3913 | 0.074581 | 90.8822 | 0.055301 | 9.1242 | 110 | 13 | 1748 | 1.796 |  |  |  | 1.796
|  |  |  | 200 | 33.8392 | 5.3553 | 2.3998 | 33.8968 | 104.388 | 161777 | 199.585 | 5.3711 | 32.8681 | 1.7186 | 0.078044 | 90.8458 | 0.055301 | 9.6057 | 111 | 14 | 1658 | 2.525 |  |  |  | 2.525
|  |  |  | 175 | 33.8205 | 5.714 | 2.5008 | 35.6215 | 108.786 | 163769 | 174.513 | 5.7283 | 33.1563 | 1.7794 | 0.07934 | 90.8462 | 0.055301 | 1.0421 | 112 | 15 | 1156 | 2.591 |  |  |  | 2.591
|  |  |  | 150 | 33.7726 | 5.954 | 2.8227 | 40.4231 | 122.798 | 165549 | 150.374 | 5.9666 | 33.3143 | 1.9431 | 0.070245 | 90.7945 | 0.055301 | 1.0545 | 113 | 16 | 663 | 2.927 |  |  |  | 2.927
|  |  |  | 125 | 33.4755 | 5.9028 | 3.6423 | 51.992 | 158.486 | 167265 | 124.685 | 5.9131 | 32.9902 | 2.3348 | 0.072787 | 90.818 | 0.055301 | 1.1073 | 114 | 17 | 1376 | 3.739 |  |  |  | 3.739
|  |  |  | 100 | 32.7041 | 5.231 | 6.4823 | 90.5901 | 282.21 | 168827 | 100.477 | 5.2385 | 31.7055 | 3.637 | 0.074381 | 90.7754 | 0.055301 | 1.0908 | 115 | 18 | 1137 | 6.707 |  |  |  | 6.707
|  |  |  | 75 | 32.6166 | 5.7439 | 6.7967 | 96.1013 | 295.934 | 170610 | 75.754 | 5.7499 | 32.0683 | 3.8316 | 0.05639 | 90.5269 | 0.055301 | 1.0905 | 116 | 19 | 1023 | 7.104 |  |  |  | 7.104
|  |  |  | 50 | 32.5928 | 6.6246 | 6.8786 | 99.2916 | 299.54 | 172508 | 50.278 | 6.6289 | 32.7887 | 3.9485 | 0.074132 | 88.9463 | 0.055301 | 1.1082 | 117 | 20 | 1022 | 7.16 |  |  |  | 7.16
|  |  |  | 25 | 32.5377 | 8.558 | 6.8627 | 103.542 | 298.936 | 174555 | 25.781 | 8.5606 | 34.4079 | 4.0988 | 0.08395 | 85.2005 | 0.055301 | 1.1474 | 118 | 21 | 1666 | 7.181 |  |  |  | 7.181
|  |  |  | 10 | 32.4805 | 10.4318 | 6.3814 | 100.3494 | 278.067 | 176010 | 10.601 | 10.433 | 36.0067 | 3.9933 | 0.18784 | 86.8611 | 0.055301 | 1.0939 | 119 | 22 | 1155 | 6.705 |  |  |  | 6.705
|  |  |  | 5 | 32.4824 | 10.4244 | 6.5062 | 102.294 | 283.502 | 177147 | 4.794 | 10.425 | 35.997 | 4.0619 | 0.14769 | 86.7409 | 0.055301 | 1.1688 | 120 | 23 | 1136 | 6.667 |  |  |  | 6.667
| 010 | 010 | NP6 |  | 32.6354 | 5.8303 | 6.8889 | 97.6132 | 299.947 | 7505 | 49.674 | 5.8343 | 32.14 | 3.8939 | 0.94174 | 88.6177 | 9.1227 | 787.15
[5 empty] | 32.6353 | 5.8327 | 6.8689 | 97.3358 | 299.078 | 7544 | 49.261 | 5.8367 | 32.14 | 3.8847 | 1.018 | 88.6259 | 8.6057 | 785.97
[5 empty] | 32.6356 | 5.8315 | 6.8709 | 97.362 | 299.167 | 7581 | 49.284 | 5.8354 | 32.1409 | 3.8856 | 0.95597 | 88.6232 | 8.1948 | 784.63
[5 empty] | 32.6358 | 5.8296 | 6.883 | 97.5284 | 299.69 | 7614 | 49.68 | 5.8336 | 32.141 | 3.891 | 0.95978 | 88.6453 | 8.1644 | 784.18
[5 empty] | 32.6359 | 5.8284 | 6.8917 | 97.649 | 300.069 | 7648 | 49.872 | 5.8324 | 32.1386 | 3.895 | 0.97153 | 88.659 | 8.1685 | 782.9
[5 empty] | 32.6356 | 5.8297 | 6.8862 | 97.5744 | 299.832 | 7681 | 49.597 | 5.8337 | 32.14 | 3.8927 | 0.92334 | 88.6615 | 7.9909 | 782
[5 empty] | 32.6356 | 5.8301 | 6.8925 | 97.6644 | 300.106 | 7711 | 49.252 | 5.834 | 32.1404 | 3.8958 | 1.0525 | 88.6745 | 7.2398 | 782.16
[5 empty] | 32.6357 | 5.8304 | 6.8933 | 97.677 | 300.142 | 7739 | 49.137 | 5.8343 | 32.1398 | 3.8963 | 1.0203 | 88.6571 | 6.0774 | 782.45
[5 empty] | 32.6358 | 5.8301 | 6.8953 | 97.7042 | 300.227 | 7768 | 49.211 | 5.834 | 32.1402 | 3.8972 | 0.97621 | 88.6345 | 5.7937 | 782
[5 empty] | 32.6358 | 5.8297 | 6.9044 | 97.8322 | 300.624 | 7797 | 49.38 | 5.8336 | 32.1406 | 3.9014 | 1.0095 | 88.6264 | 5.6878 | 781.85
[5 empty] | 32.6358 | 5.8289 | 6.9082 | 97.8845 | 300.79 | 7827 | 49.591 | 5.8329 | 32.1401 | 3.9031 | 1.0368 | 88.6599 | 5.5872 | 780.87
[5 empty] | 32.6361 | 5.8278 | 6.9001 | 97.7663 | 300.434 | 7857 | 49.744 | 5.8317 | 32.1388 | 3.8991 | 1.0376 | 88.7038 | 8.3857 | 779.82
[5 empty] | 32.6361 | 5.8264 | 6.899 | 97.7476 | 300.386 | 7886 | 49.761 | 5.8304 | 32.1363 | 3.8984 | 1.0303 | 88.6723 | 9.5031 | 779.17
[5 empty] | 32.6359 | 5.8288 | 6.899 | 97.7529 | 300.386 | 7915 | 49.608 | 5.8328 | 32.1391 | 3.8987 | 1.0045 | 88.6644 | 9.5975 | 778.17
[5 empty] | 32.6357 | 5.8296 | 6.8981 | 97.7429 | 300.35 | 7945 | 49.375 | 5.8335 | 32.14 | 3.8984 | 1.0481 | 88.6645 | 9.6614 | 777.64
[5 empty] | 32.6358 | 5.8303 | 6.8932 | 97.6751 | 300.137 | 7975 | 49.276 | 5.8342 | 32.14 | 3.8962 | 1.0253 | 88.6418 | 9.6971 | 776.71
[5 empty] | 32.6359 | 5.8293 | 6.8983 | 97.7447 | 300.357 | 8004 | 49.335 | 5.8333 | 32.1397 | 3.8985 | 1.0134 | 88.6615 | 9.4548 | 777.15
[5 empty] | 32.6359 | 5.8292 | 6.8966 | 97.7204 | 300.283 | 8035 | 49.422 | 5.8332 | 32.1395 | 3.8976 | 1.1004 | 88.4971 | 9.3057 | 776.55
[5 empty] | 32.6359 | 5.8291 | 6.8916 | 97.6489 | 300.064 | 8067 | 49.453 | 5.833 | 32.1395 | 3.8952 | 1.0648 | 88.6432 | 9.3345 | 775.46
[5 empty] | 32.6359 | 5.829 | 6.8986 | 97.7476 | 300.368 | 8099 | 49.56 | 5.833 | 32.1382 | 3.8985 | 1.0009 | 88.6569 | 8.3058 | 775.53
[5 empty] | 32.636 | 5.8274 | 6.9133 | 97.9523 | 301.008 | 8134 | 49.732 | 5.8314 | 32.1383 | 3.9054 | 1.0414 | 88.7066 | 6.9655 | 775.46
[5 empty] | 32.6359 | 5.8284 | 6.9084 | 97.8858 | 300.797 | 8174 | 49.485 | 5.8323 | 32.1389 | 3.9032 | 0.97504 | 88.6723 | 6.0471 | 773.84
[5 empty] | 32.6357 | 5.831 | 6.8889 | 97.6149 | 299.947 | 8220 | 48.949 | 5.8349 | 32.1405 | 3.8943 | 0.98575 | 88.6349 | 5.7393 | 773.51
| 011 | 011 | NP6 | 1000 | 34.3659 | 2.934 | 0.3096 | 4.1404 | 13.458 | 30340 | 999.792 | 3.0017 | 31.584 | 0.7192 | 0.054746 | 91.0282 | 0.074197 | 1178.8 | 120 | 1 | 1351 | 0.295 |  |  |  | 0.295
|  |  |  | 800 | 34.2881 | 3.3579 | 0.29616 | 3.9996 | 12.875 | 36288 | 801.456 | 3.4134 | 31.795 | 0.7182 | 0.057088 | 91.022 | 0.071745 | 1271.2 | 121 | 2 | 1697 | 0.288 |  |  |  | 0.288
|  |  |  | 600 | 34.1936 | 3.9247 | 0.35556 | 4.8651 | 15.46 | 42029 | 600.272 | 3.9678 | 32.1141 | 0.7487 | 0.056241 | 91.0062 | 0.071472 | 1388.9 | 122 | 3 | 1743 | 0.352 |  |  |  | 0.352
|  |  |  | 400 | 34.0201 | 4.0906 | 0.83625 | 11.4717 | 36.365 | 48065 | 400.315 | 4.1191 | 32.0122 | 0.966 | 0.055269 | 90.9693 | 0.073524 | 1510.8 | 123 | 4 | 1739 | 0.849 |  |  |  | 0.849
|  |  |  | 300 | 33.9165 | 4.4242 | 1.4114 | 19.5037 | 61.382 | 51468 | 300.553 | 4.4461 | 32.1654 | 1.2335 | 0.055468 | 90.7941 | 0.070245 | 1391.3 | 124 | 5 | 1725 | 1.446 |  |  |  | 1.446
|  |  |  | 250 | 33.8504 | 4.5805 | 2.1294 | 29.5226 | 92.615 | 53844 | 250.567 | 4.5989 | 32.222 | 1.5675 | 0.053051 | 90.6834 | 0.069185 | 1313.5 | 125 | 6 | 1731 | 2.182 |  |  |  | 2.182
|  |  |  | 200 | 33.8116 | 4.9684 | 2.4953 | 34.9108 | 108.536 | 55963 | 201.385 | 4.9837 | 32.5049 | 1.7522 | 0.052329 | 90.6451 | 0.068188 | 1349.3 | 126 | 7 | 1193 | 2.588 |  |  |  | 2.588
|  |  |  | 175 | 33.7986 | 5.3389 | 2.7082 | 38.2247 | 117.803 | 57428 | 176.31 | 5.3528 | 32.8042 | 1.8659 | 0.052852 | 90.6945 | 0.064808 | 1455.1 | 127 | 8 | 1750 | 2.809 |  |  |  | 2.809
|  |  |  | 150 | 33.7172 | 5.7387 | 2.9901 | 42.5859 | 130.083 | 58918 | 150.674 | 5.751 | 33.0722 | 2.0153 | 0.058683 | 90.671 | 0.10487 | 1533.7 | 128 | 9 | 1744 | 3.117 |  |  | d | 3.117 | leaky Niskin
|  |  |  | 125 | 33.3027 | 5.1143 | 4.6454 | 64.9958 | 202.145 | 60454 | 124.281 | 5.1238 | 32.1474 | 2.7709 | 0.062321 | 90.6092 | 0.22308 | 1318.5 | 129 | 10 | 1740 | 4.844 |  |  |  | 4.844
|  |  |  | 100 | 32.7681 | 4.9127 | 6.7663 | 93.8725 | 294.549 | 61970 | 99.439 | 4.9199 | 31.496 | 3.7483 | 0.11164 | 90.4785 | 0.557 | 1260.1 | 130 | 11 | 1726 | 7.022 |  |  |  | 7.022
|  |  |  | 75 | 32.687 | 5.2335 | 6.8099 | 95.159 | 296.476 | 63524 | 75.794 | 5.2392 | 31.693 | 3.801 | 0.22479 | 90.3172 | 1.4745 | 1225.1 | 131 | 12 | 1732 | 7.042 |  |  |  | 7.042
|  |  |  | 50 | 32.6374 | 5.74 | 6.8864 | 97.3691 | 299.836 | 65105 | 49.372 | 5.7439 | 32.0678 | 3.886 | 0.87177 | 88.9243 | 6.8355 | 1255.6 | 132 | 13 | 1752 | 7.143 |  |  |  | 7.143
|  |  |  | 25 | 32.5673 | 8.6128 | 6.7135 | 101.436 | 292.433 | 66640 | 25.158 | 8.6153 | 34.4869 | 4.0277 | 0.66906 | 88.6188 | 37.373 | 1223.4 | 133 | 14 | 1709, 1745, 1741 | 6.998 | * | 6.997 | *e | 6.998 | operator error, lost sample
|  |  |  | 10 | 32.518 | 9.496 | 6.458 | 99.498 | 281.354 | 68072 | 10.379 | 9.4972 | 35.2131 | 3.9668 | 0.37659 | 88.4551 | 87.988 | 1146.1 | 134 | 15 | 1727 | 6.712 |  |  |  | 6.712
|  |  |  | 5 | 32.5177 | 9.5221 | 6.4592 | 99.5732 | 281.408 | 69547 | 5.884 | 9.5227 | 35.2326 | 3.9713 | 0.23493 | 88.2903 | 144.89 | 1133.3 | 135 | 16 | 1733 | 6.711 |  |  |  | 6.711
| 012 | 012 | NP6 | 1000 | 34.3683 | 2.9152 | 0.31287 | 4.1823 | 13.6 | 28975 | 999.946 | 2.9827 | 31.5715 | 0.7204 | 0.056316 | 91.0253 | 0.066972 | 1985.8
|  |  |  | 300 | 33.917 | 4.506 | 1.2747 | 17.6501 | 55.438 | 47750 | 300.807 | 4.5282 | 32.2388 | 1.1729 | 0.055344 | 90.7495 | 0.067261 | 2018.1
|  |  |  | 140 | 33.712 | 5.7823 | 2.9764 | 42.4327 | 129.489 | 53188 | 140.418 | 5.7938 | 33.1038 | 2.0121 | 0.060552 | 90.6585 | 0.23314 | 2643.2
|  |  |  | 140 | 33.7128 | 5.7817 | 2.9729 | 42.382 | 129.335 | 53908 | 140.554 | 5.7932 | 33.1032 | 2.0104 | 0.059057 | 90.672 | 0.24926 | 2794.8
|  |  |  | 75 | 32.6817 | 5.1728 | 6.7597 | 94.3159 | 294.291 | 56763 | 74.452 | 5.1783 | 31.633 | 3.7733 | 0.16167 | 90.3789 | 3.9219 | 2328.3
|  |  |  | 75 | 32.6824 | 5.1668 | 6.7553 | 94.2411 | 294.098 | 57524 | 75.458 | 5.1724 | 31.6263 | 3.7704 | 0.15866 | 90.3818 | 4.2305 | 2707.4
|  |  |  | 71 | 32.6774 | 5.3567 | 6.8464 | 95.9457 | 298.07 | 58964 | 69.753 | 5.362 | 31.7839 | 3.8298 | 0.38153 | 89.9566 | 5.0479 | 2631.3
|  |  |  | 71 | 32.6774 | 5.3613 | 6.8604 | 96.1534 | 298.681 | 59679 | 69.67 | 5.3666 | 31.7879 | 3.8369 | 0.42989 | 89.9447 | 4.8317 | 2412.5
|  |  |  | 45 | 32.6505 | 6.1033 | 6.9421 | 99.0171 | 302.27 | 61476 | 45.431 | 6.107 | 32.3926 | 3.9426 | 0.9396 | 88.7313 | 19.577 | 1991
|  |  |  | 45 | 32.651 | 6.0972 | 6.9419 | 99.0006 | 302.262 | 62203 | 45.277 | 6.1009 | 32.3935 | 3.9421 | 0.90332 | 88.717 | 22.626 | 1493.2
|  |  |  | 35 | 32.6385 | 7.1351 | 6.9998 | 102.2812 | 304.824 | 64013 | 35.142 | 7.1383 | 33.2694 | 4.0552 | 0.58683 | 88.7773 | 39.949 | 2190.9
|  |  |  | 35 | 32.6391 | 7.1257 | 6.9803 | 101.975 | 303.976 | 64779 | 35.855 | 7.1289 | 33.1946 | 4.0446 | 0.59582 | 88.649 | 40.638 | 2307.3
|  |  |  | 7 | 32.5194 | 9.6032 | 6.4637 | 99.823 | 281.605 | 66861 | 6.869 | 9.6039 | 35.2996 | 3.9792 | 0.12511 | 88.3665 | 441.21 | 2803
|  |  |  | 7 | 32.5194 | 9.6246 | 6.4677 | 99.9333 | 281.782 | 67614 | 7.143 | 9.6254 | 35.3268 | 3.9827 | 0.12479 | 88.3296 | 319.66 | 2162
|  |  |  | 7 | 32.52 | 9.6143 | 6.4597 | 99.7864 | 281.43 | 68306 | 7.368 | 9.6151 | 35.3207 | 3.9777 | 0.12546 | 88.3071 | 263.62 | 1748.6
|  |  |  | 7 | 32.5189 | 9.6444 | 6.4621 | 99.8897 | 281.537 | 68561 | 7.188 | 9.6452 | 35.3402 | 3.9812 | 0.12888 | 88.3047 | 265.95 | 1659.8
|  |  |  | 5 | 32.5202 | 9.6393 | 6.4691 | 99.9873 | 281.842 | 69995 | 6.379 | 9.64 | 35.3406 | 3.9848 | 0.11931 | 88.2712 | 347.47 | 2540.1
|  |  |  | 5 | 32.5204 | 9.6331 | 6.4708 | 100.0006 | 281.918 | 70180 | 6.148 | 9.6338 | 35.3318 | 3.9854 | 0.12265 | 88.2693 | 338.27 | 2311.5
|  |  |  | 3 | 32.5195 | 9.6198 | 6.5538 | 101.2518 | 285.533 | 71702 | 2.776 | 9.6201 | 35.3181 | 4.0294 | 0.1148 | 88.2381 | 542.05 | 1859.4
|  |  |  | 3 | 32.519 | 9.6259 | 6.5383 | 101.0249 | 284.856 | 71987 | 2.591 | 9.6262 | 35.319 | 4.0218 | 0.11889 | 88.2233 | 565.9 | 1838.2
|  |  |  | 0 | 32.5204 | 9.611 | 6.5678 | 101.4483 | 286.141 | 73934 | 2.843 | 9.6113 | 35.3091 | 4.0361 | 0.11963 | 88.2156 | 470.82 | 2101.4
|  |  |  | 0 | 32.5201 | 9.6093 | 6.5637 | 101.382 | 285.965 | 74158 | 2.726 | 9.6096 | 35.3084 | 4.0339 | 0.11978 | 88.2128 | 489.24 | 2161.5
| 013 | 013 | NP7 | 1000 | 34.3656 | 2.9236 | 0.29275 | 3.9141 | 12.726 | 34365 | 1000.68 | 2.9912 | 31.5756 | 0.7124 | 0.055144 | 91.0589 | 0.067958 | 370.31 | 158 | 1 | 1697 | 0.279 |  |  |  | 0.279
|  |  |  | 800 | 34.2924 | 3.2568 | 0.36342 | 4.8956 | 15.799 | 40423 | 802.644 | 3.3117 | 31.7096 | 0.7456 | 0.054671 | 91.055 | 0.068072 | 413.43 | 159 | 2 | 1743 | 0.354 |  |  |  | 0.354
|  |  |  | 600 | 34.173 | 3.6523 | 0.4708 | 6.3976 | 20.47 | 46347 | 600.444 | 3.6942 | 31.8586 | 0.7969 | 0.054771 | 91.0454 | 0.066454 | 451.52 | 160 | 3 | 1739 | 0.464 |  |  |  | 0.464
|  |  |  | 400 | 33.9952 | 4.1006 | 0.90296 | 12.3878 | 39.267 | 52440 | 399.394 | 4.129 | 31.9993 | 0.9956 | 0.055443 | 90.9727 | 0.066763 | 551.1 | 161 | 4 | 17251731 | 0.919 | 0.922 |  |  | 0.919
|  |  |  | 300 | 33.8908 | 4.3072 | 1.5692 | 21.6187 | 68.247 | 56394 | 299.051 | 4.3287 | 32.0404 | 1.3029 | 0.054098 | 90.8458 | 0.068558 | 579.42 | 162 | 5 | 1193 | 1.613 |  |  |  | 1.613
|  |  |  | 250 | 33.8452 | 4.4111 | 2.071 | 28.593 | 90.073 | 58581 | 250.704 | 4.4292 | 32.0692 | 1.537 | 0.057786 | 90.7878 | 0.070178 | 616.63 | 163 | 6 | 1750 | 2.139 |  |  |  | 2.139
|  |  |  | 200 | 33.7786 | 4.6427 | 2.7944 | 38.7799 | 121.548 | 60842 | 200.6 | 4.6575 | 32.1869 | 1.8808 | 0.049936 | 90.6787 | 0.069308 | 651.3 | 164 | 7 | 1744 | 2.906 |  |  |  | 2.906
|  |  |  | 175 | 33.7486 | 4.8264 | 3.3954 | 47.3197 | 147.696 | 62544 | 175.419 | 4.8395 | 32.3109 | 2.1688 | 0.054621 | 90.6547 | 0.069288 | 675.57 | 165 | 8 | 1740 | 3.533 |  |  |  | 3.533
|  |  |  | 150 | 33.5947 | 4.9109 | 4.0277 | 56.1864 | 175.22 | 64381 | 150.436 | 4.9222 | 32.2395 | 2.4698 | 0.052129 | 90.7069 | 0.071007 | 717.54 | 166 | 9 | 1726 | 4.201 |  |  |  | 4.201
|  |  |  | 125 | 33.0982 | 4.9103 | 5.5102 | 76.6119 | 239.808 | 65953 | 125.058 | 4.9195 | 31.8007 | 3.1598 | 0.058882 | 90.7122 | 0.099551 | 665.17 | 167 | 10 | 1732 | 5.576 |  |  |  | 5.576
|  |  |  | 100 | 32.7453 | 4.8475 | 6.6546 | 92.1637 | 289.692 | 67586 | 100.501 | 4.8548 | 31.4256 | 3.6906 | 0.08542 | 90.7079 | 0.17203 | 593.16 | 168 | 11 | 1752 | 6.958 |  |  |  | 6.958
|  |  |  | 75 | 32.6888 | 5.2817 | 6.8492 | 95.8209 | 298.19 | 69439 | 75.641 | 5.2874 | 31.7295 | 3.8233 | 0.24388 | 90.3186 | 0.39702 | 529.14 | 169 | 12 | 1709 | 7.146 |  |  |  | 7.146
|  |  |  | 50 | 32.6631 | 5.9273 | 6.9781 | 99.124 | 303.828 | 71530 | 49.596 | 5.9313 | 32.2524 | 3.9448 | 1.0859 | 88.5518 | 1.7514 | 541.32 | 170 | 13 | 1745, 1741 | 7.25 | 7.245 |  |  | 7.25
|  |  |  | 25 | 32.6014 | 7.9264 | 6.8942 | 102.5717 | 300.268 | 73324 | 26.058 | 7.9289 | 33.9232 | 4.0674 | 1.4478 | 86.9712 | 12.073 | 577.73 | 171 | 14 | 1727 | 7.151 |  |  |  | 7.151
|  |  |  | 10 | 32.5281 | 9.4534 | 6.5324 | 100.5552 | 284.592 | 74662 | 10.345 | 9.4545 | 35.1835 | 4.0027 | 0.88316 | 87.2532 | 37.859 | 592.09 | 172 | 15 | 1733 | 6.767 |  |  |  | 6.767
|  |  |  | 5 | 32.5287 | 9.6225 | 6.4867 | 100.227 | 282.606 | 76148 | 6.092 | 9.6231 | 35.3373 | 3.9932 | 0.73564 | 87.4582 | 57.062 | 563.73 | 173 | 16 | 1753 | 6.742 |  |  |  | 6.742
| 014 | 014 | NP8 |  | 32.6843 | 6.5224 | 7.1432 | 102.9213 | 311.034 | 12556 | 37.135 | 6.5256 | 32.7761 | 4.0774 | 1.1038 | 88.2543 | 1.4417 | 115.45
[5 empty] | 32.685 | 6.5149 | 7.1516 | 103.0252 | 311.402 | 12589 | 37.356 | 6.518 | 32.7672 | 4.0809 | 1.0988 | 88.357 | 1.4015 | 115.45
[5 empty] | 32.6844 | 6.5117 | 7.147 | 102.9509 | 311.201 | 12617 | 37.345 | 6.5149 | 32.7684 | 4.0784 | 1.2268 | 88.381 | 1.4156 | 115.45
[5 empty] | 32.6842 | 6.5188 | 7.1323 | 102.7553 | 310.56 | 12645 | 37.141 | 6.5219 | 32.7737 | 4.0718 | 1.0961 | 88.4057 | 1.4338 | 115.45
[5 empty] | 32.6844 | 6.5271 | 7.1283 | 102.7179 | 310.386 | 12676 | 36.93 | 6.5302 | 32.7799 | 4.0706 | 1.1399 | 88.3344 | 1.4655 | 115.45
[5 empty] | 32.6844 | 6.529 | 7.1209 | 102.6163 | 310.065 | 12705 | 36.878 | 6.5322 | 32.7793 | 4.0673 | 1.2471 | 88.3174 | 1.4626 | 115.45
[5 empty] | 32.6845 | 6.5314 | 7.1279 | 102.7227 | 310.369 | 12735 | 36.908 | 6.5345 | 32.7735 | 4.0708 | 1.152 | 88.3618 | 1.4572 | 115.45
[5 empty] | 32.6855 | 6.5286 | 7.1329 | 102.7883 | 310.585 | 12765 | 36.867 | 6.5318 | 32.7797 | 4.0731 | 1.1459 | 88.3224 | 1.4628 | 115.45
[5 empty] | 32.6855 | 6.5313 | 7.1354 | 102.8311 | 310.694 | 12795 | 36.824 | 6.5345 | 32.785 | 4.0745 | 1.1226 | 88.3133 | 1.4659 | 115.45
[5 empty] | 32.6855 | 6.5297 | 7.1429 | 102.9352 | 311.021 | 12825 | 36.955 | 6.5329 | 32.7844 | 4.078 | 1.0788 | 88.3073 | 1.4693 | 115.45
[5 empty] | 32.6856 | 6.5284 | 7.144 | 102.9485 | 311.07 | 12854 | 37.184 | 6.5316 | 32.7813 | 4.0783 | 1.0972 | 88.3476 | 1.4425 | 115.45
[5 empty] | 32.6858 | 6.5261 | 7.134 | 102.7984 | 310.633 | 12881 | 37.277 | 6.5292 | 32.7803 | 4.0732 | 1.1262 | 88.3334 | 1.4198 | 115.45
[5 empty] | 32.6855 | 6.5267 | 7.1325 | 102.7788 | 310.57 | 12910 | 37.156 | 6.5299 | 32.7817 | 4.0726 | 1.0762 | 88.3149 | 1.4245 | 115.45
[5 empty] | 32.6851 | 6.5328 | 7.1285 | 102.7356 | 310.395 | 12938 | 36.957 | 6.536 | 32.7852 | 4.0712 | 1.1274 | 88.2629 | 1.4513 | 115.45
[5 empty] | 32.6849 | 6.5344 | 7.1313 | 102.779 | 310.516 | 12965 | 36.901 | 6.5375 | 32.7859 | 4.0727 | 1.1589 | 88.4237 | 1.4615 | 115.45
[5 empty] | 32.6855 | 6.5334 | 7.1353 | 102.8346 | 310.69 | 12994 | 36.983 | 6.5366 | 32.7837 | 4.0746 | 1.2085 | 88.3133 | 1.4664 | 115.45
[5 empty] | 32.6857 | 6.533 | 7.1302 | 102.76 | 310.467 | 13025 | 37.064 | 6.5361 | 32.7838 | 4.072 | 1.1636 | 88.3055 | 1.457 | 115.45
[5 empty] | 32.6856 | 6.5328 | 7.1326 | 102.7946 | 310.573 | 13057 | 37.048 | 6.536 | 32.7862 | 4.0732 | 1.1239 | 88.3436 | 1.4512 | 115.45
[5 empty] | 32.6856 | 6.5342 | 7.1408 | 102.9165 | 310.932 | 13089 | 36.959 | 6.5373 | 32.7888 | 4.0773 | 1.1251 | 88.3309 | 1.465 | 115.45
[5 empty] | 32.6853 | 6.5371 | 7.1409 | 102.9243 | 310.935 | 13123 | 36.853 | 6.5402 | 32.7907 | 4.0776 | 1.1341 | 88.2918 | 1.4852 | 115.45
[5 empty] | 32.6852 | 6.5374 | 7.1253 | 102.7005 | 310.257 | 13158 | 36.903 | 6.5405 | 32.7908 | 4.0701 | 1.2117 | 88.3185 | 1.4761 | 115.45
[5 empty] | 32.6858 | 6.5342 | 7.1208 | 102.6282 | 310.059 | 13198 | 37.232 | 6.5374 | 32.7867 | 4.0675 | 1.2989 | 88.3182 | 1.4307 | 115.45
[5 empty] | 32.6852 | 6.5347 | 7.1336 | 102.8135 | 310.617 | 13279 | 37.058 | 6.5379 | 32.788 | 4.0738 | 1.0767 | 88.2429 | 1.4569 | 115.45
| 015 | 015 | NP8 | 1000 | 34.3766 | 2.8358 | 0.3691 | 4.9243 | 16.045 | 25375 | 1001.274 | 2.9029 | 31.5079 | 0.7425 | 0.053998 | 91.0574 | 0.068724 | 575.59 | 174 | 1 | 1135 | 0.354 |  |  |  | 0.354
|  |  |  | 800 | 34.3768 | 2.8352 | 0.36948 | 4.9293 | 16.061 | 25448 | 1002.472 | 2.9023 | 31.5081 | 0.7427 | 0.055867 | 91.0541 | 0.069315 | 577.24 | 175 | 2 | 1094 | 0.328 |  |  |  | 0.328
|  |  |  | 600 | 34.3765 | 2.8359 | 0.36881 | 4.9204 | 16.032 | 25525 | 1001.688 | 2.903 | 31.5081 | 0.7424 | 0.054696 | 91.0397 | 0.075647 | 579.29 | 176 | 3 | 1293 | 0.411 |  |  |  | 0.411
|  |  |  | 400 | 34.3137 | 3.1548 | 0.3371 | 4.5301 | 14.654 | 31194 | 802.414 | 3.2091 | 31.6376 | 0.7344 | 0.054795 | 91.0421 | 0.068211 | 655.79 | 177 | 4 | 1654 | 0.58 |  |  |  | 0.58
|  |  |  | 300 | 34.2185 | 3.5376 | 0.41626 | 5.642 | 18.098 | 37199 | 599.444 | 3.5788 | 31.7933 | 0.7732 | 0.055194 | 91.0373 | 0.066465 | 623.02 | 178 | 5 | 1127 | 0.878 |  |  |  | 0.878
|  |  |  | 250 | 34.0662 | 3.8618 | 0.56165 | 7.6637 | 24.423 | 43196 | 400.531 | 3.8897 | 31.8503 | 0.8432 | 0.056066 | 90.9846 | 0.066591 | 809.12 | 179 | 6 | 1284 | 1.095 |  |  |  | 1.095
|  |  |  | 200 | 33.9721 | 4.1018 | 0.86028 | 11.7987 | 37.411 | 46906 | 300.79 | 4.123 | 31.9305 | 0.9817 | 0.055443 | 90.8802 | 0.06897 | 948.61 | 180 | 7 | 1143 | 1.541 |  |  |  | 1.541
|  |  |  | 175 | 33.9274 | 4.184 | 1.0688 | 14.6826 | 46.482 | 49324 | 250.746 | 4.2017 | 31.939 | 1.0792 | 0.057661 | 90.8506 | 0.066568 | 860.42 | 181 | 8 | 1089 | 1.851 |  |  |  | 1.851
|  |  |  | 150 | 33.8695 | 4.2817 | 1.5233 | 20.9668 | 66.252 | 51692 | 198.823 | 4.2958 | 31.9503 | 1.2909 | 0.054197 | 90.7625 | 0.066831 | 864.77 | 182 | 9 | 1509 | 2.361 |  |  |  | 2.361
|  |  |  | 125 | 33.844 | 4.3449 | 1.7902 | 24.6732 | 77.862 | 53347 | 175.185 | 4.3573 | 31.9703 | 1.4164 | 0.056316 | 90.704 | 0.066975 | 893.31 | 183 | 10 | 879, 1746, 1145 | 3.013 | 3.014 | 3.011 |  | 3.013
|  |  |  | 100 | 33.7942 | 4.4577 | 2.2913 | 31.6543 | 99.659 | 55111 | 149.785 | 4.4684 | 32.0149 | 1.652 | 0.057337 | 90.6686 | 0.069466 | 941.1 | 184 | 11 | 1748 | 4.19 |  |  |  | 4.19
|  |  |  | 75 | 33.7058 | 4.6073 | 2.9317 | 40.6244 | 127.525 | 56852 | 124.082 | 4.6163 | 32.0559 | 1.9559 | 0.057935 | 90.7357 | 0.13643 | 941.05 | 185 | 12 | 1658 | 6.372 |  |  |  | 6.372
|  |  |  | 50 | 33.4156 | 4.5601 | 4.2413 | 58.5887 | 184.531 | 58310 | 100.444 | 4.5673 | 31.7578 | 2.5636 | 0.075303 | 90.729 | 0.34486 | 960.35 | 186 | 13 | 1156 | 7.202 |  |  |  | 7.202
|  |  |  | 25 | 32.9122 | 4.6015 | 6.1888 | 85.2908 | 269.373 | 59907 | 75.568 | 4.6068 | 31.3438 | 3.47 | 0.11131 | 90.6787 | 1.024 | 994 | 187 | 14 | 663 | 7.308 |  |  |  | 7.308
|  |  |  | 10 | 32.7085 | 5.2176 | 6.9238 | 96.7225 | 301.429 | 61500 | 49.439 | 5.2213 | 31.682 | 3.8653 | 0.98475 | 89.4151 | 3.9026 | 948.3 | 188 | 15 | 1376 | 6.891 |  |  |  | 6.891
|  |  |  | 5 | 32.674 | 6.0437 | 7.0565 | 100.5205 | 307.244 | 62901 | 37.176 | 6.0467 | 32.3568 | 3.9974 | 1.1571 | 88.5293 | 9.964 | 953.99 | 189 | 16 | 1137 | 6.852 |  |  | b | 6.852 | floaties
|  |  |  | 1000 | 32.6606 | 7.2771 | 7.0434 | 103.2721 | 306.726 | 64329 | 24.85 | 7.2793 | 33.3939 | 4.0933 | 0.86866 | 88.3363 | 26.129 | 934.57
|  |  |  | 1000 | 32.6276 | 9.0835 | 6.6082 | 100.947 | 287.856 | 65786 | 10.04 | 9.0845 | 34.9539 | 4.0171 | 0.44888 | 88.3957 | 72.542 | 910.48
|  |  |  | 37 | 32.6267 | 9.1222 | 6.5978 | 100.8727 | 287.402 | 67227 | 3.953 | 9.1226 | 34.9795 | 4.0172 | 0.29765 | 88.3684 | 136.89 | 882.69
| 016 | 016 | NP8
| 017 | 017 | NP8 | 400 | 34.0498 | 3.9024 | 0.61409 | 8.3867 | 26.703 | 963 | 401.125 | 3.9304 | 31.8713 | 0.8665 | 0.056216 | 90.9578 | 0.072757 | 413.34
|  |  |  | 400 | 34.0497 | 3.9021 | 0.61334 | 8.3764 | 26.671 | 1012 | 401.322 | 3.9301 | 31.8713 | 0.8661 | 0.056395 | 90.9924 | 0.071019 | 411.96
|  |  |  | 120 | 33.5627 | 4.6113 | 3.5833 | 49.6118 | 155.888 | 9087 | 119.035 | 4.6199 | 31.9355 | 2.2579 | 0.08085 | 90.7851 | 0.10547 | 402.98
|  |  |  | 120 | 33.5638 | 4.6118 | 3.5837 | 49.618 | 155.905 | 9123 | 119.123 | 4.6204 | 31.9344 | 2.2581 | 0.076016 | 90.7279 | 0.10788 | 402.79
|  |  |  | 120 | 33.5658 | 4.6125 | 3.5822 | 49.5988 | 155.84 | 9162 | 119.044 | 4.6211 | 31.9339 | 2.2575 | 0.070876 | 90.7266 | 0.10752 | 400.8
|  |  |  | 120 | 33.5638 | 4.6117 | 3.5819 | 49.5933 | 155.828 | 9207 | 118.928 | 4.6203 | 31.9335 | 2.2573 | 0.070971 | 90.7626 | 0.10552 | 400.57
|  |  |  | 100 | 33.2564 | 4.5183 | 4.8497 | 66.8541 | 211.026 | 10869 | 99.652 | 4.5253 | 31.5827 | 2.8415 | 0.094796 | 90.7489 | 0.19281 | 402.56
|  |  |  | 100 | 33.2594 | 4.518 | 4.8665 | 67.0865 | 211.756 | 10917 | 99.661 | 4.5251 | 31.5912 | 2.8493 | 0.084936 | 90.7617 | 0.19105 | 402.98
|  |  |  | 79 | 32.9441 | 4.5752 | 6.0481 | 83.3153 | 263.238 | 12524 | 79.021 | 4.5807 | 31.3495 | 3.4022 | 0.1074 | 90.7902 | 0.48314 | 461.92
|  |  |  | 79 | 32.9465 | 4.5737 | 6.0771 | 83.7136 | 264.501 | 12607 | 79.512 | 4.5793 | 31.3505 | 3.4154 | 0.099673 | 90.7888 | 0.47587 | 460.37
|  |  |  | 60 | 32.7291 | 4.9723 | 6.8323 | 94.8942 | 297.433 | 14166 | 59.969 | 4.9767 | 31.4934 | 3.7995 | 0.55877 | 90.2008 | 1.2378 | 486.51
|  |  |  | 60 | 32.7264 | 4.9911 | 6.8032 | 94.5317 | 296.169 | 14233 | 59.408 | 4.9955 | 31.5072 | 3.7875 | 0.56186 | 90.1548 | 1.2575 | 489.4
|  |  |  | 40 | 32.6801 | 5.7653 | 7.0615 | 99.9317 | 307.45 | 15724 | 40.498 | 5.7685 | 32.1189 | 3.9766 | 1.1246 | 88.8127 | 5.1146 | 582.09
|  |  |  | 40 | 32.6811 | 5.7667 | 7.0689 | 100.0402 | 307.772 | 15777 | 40.675 | 5.77 | 32.1255 | 3.9802 | 1.1594 | 88.8587 | 5.2003 | 591.33
|  |  |  | 31 | 32.674 | 6.629 | 7.1199 | 102.8347 | 310.025 | 17164 | 30.833 | 6.6317 | 32.8551 | 4.0772 | 1.0785 | 88.5599 | 14.629 | 792.13
|  |  |  | 31 | 32.6748 | 6.627 | 7.119 | 102.8182 | 309.988 | 17216 | 30.931 | 6.6297 | 32.8587 | 4.0766 | 1.1132 | 88.5154 | 14.739 | 800.36
|  |  |  | 18 | 32.637 | 9.0095 | 6.617 | 100.9213 | 288.233 | 18511 | 17.85 | 9.0114 | 34.9014 | 4.0128 | 0.65364 | 88.4468 | 45.574 | 1007.8
|  |  |  | 18 | 32.6376 | 9.0088 | 6.6076 | 100.7763 | 287.822 | 18559 | 17.683 | 9.0107 | 34.9002 | 4.0079 | 0.7559 | 88.435 | 46.232 | 1013.1
|  |  |  | 10 | 32.6392 | 9.0135 | 6.6129 | 100.8675 | 288.055 | 20331 | 8.394 | 9.0144 | 34.9016 | 4.0153 | 0.53512 | 88.2922 | 98.882 | 1104.4
|  |  |  | 10 | 32.6399 | 9.0129 | 6.6115 | 100.845 | 287.994 | 20431 | 8.366 | 9.0138 | 34.9016 | 4.0145 | 0.51329 | 88.2266 | 95.951 | 1107.3
|  |  |  | 5 | 32.6382 | 9.0125 | 6.6469 | 101.3822 | 289.535 | 22101 | 6.373 | 9.0132 | 34.8972 | 4.0337 | 0.34482 | 88.2252 | 109.01 | 1028.4
|  |  |  | 5 | 32.6365 | 9.0131 | 6.6577 | 101.5465 | 290.004 | 22200 | 6.348 | 9.0137 | 34.8975 | 4.0393 | 0.32733 | 88.2626 | 107.83 | 1022.7
|  |  |  | 0 | 32.639 | 9.0138 | 6.7155 | 102.4314 | 292.524 | 23083 | 3.326 | 9.0141 | 34.8979 | 4.0707 | 0.3511 | 88.1989 | 143.44 | 956.88
|  |  |  | 0 | 32.6389 | 9.0133 | 6.7177 | 102.4641 | 292.621 | 23177 | 3.356 | 9.0136 | 34.8977 | 4.0718 | 0.3289 | 88.23 | 140.98 | 956.08
| 018 | 018 | NP9 |  | 32.8487 | 7.3582 | 7.5188 | 110.5832 | 327.386 | 8808 | 16.348 | 7.3597 | 33.6686 | 4.3449 | 3.1613 | 82.6008 | 2.0365 | 63.17
[5 empty] | 32.8475 | 7.3515 | 7.5218 | 110.6079 | 327.513 | 8847 | 16.373 | 7.353 | 33.6623 | 4.3458 | 3.0927 | 82.6091 | 2.0133 | 63.17
[5 empty] | 32.8484 | 7.3821 | 7.5235 | 110.7121 | 327.588 | 8887 | 16.309 | 7.3836 | 33.6655 | 4.3493 | 3.15 | 82.6704 | 2.0623 | 63.17
[5 empty] | 32.8504 | 7.3553 | 7.5269 | 110.6959 | 327.737 | 8922 | 16.383 | 7.3569 | 33.6432 | 4.3488 | 3.1583 | 82.5801 | 2.0689 | 63.17
[5 empty] | 32.8491 | 7.3404 | 7.5262 | 110.6455 | 327.704 | 8957 | 16.487 | 7.3419 | 33.6323 | 4.3471 | 3.2644 | 82.5806 | 2.0145 | 63.17
[5 empty] | 32.8474 | 7.3509 | 7.5295 | 110.7208 | 327.851 | 8992 | 16.373 | 7.3524 | 33.6414 | 4.3496 | 3.025 | 82.6369 | 1.9735 | 63.17
[5 empty] | 32.8486 | 7.4242 | 7.5281 | 110.8892 | 327.794 | 9031 | 15.98 | 7.4257 | 33.7066 | 4.3553 | 2.7457 | 82.8037 | 2.0941 | 63.17
[5 empty] | 32.8536 | 7.4678 | 7.5222 | 110.9168 | 327.535 | 9066 | 15.765 | 7.4692 | 33.7129 | 4.3562 | 3.0025 | 82.7017 | 2.2547 | 63.17
[5 empty] | 32.8513 | 7.4229 | 7.5245 | 110.835 | 327.636 | 9094 | 15.989 | 7.4244 | 33.6915 | 4.3535 | 2.9206 | 82.7092 | 2.2685 | 63.17
[5 empty] | 32.854 | 7.3511 | 7.5318 | 110.7588 | 327.947 | 9124 | 16.476 | 7.3526 | 33.6617 | 4.3508 | 3.0738 | 82.628 | 2.1033 | 63.17
[5 empty] | 32.8489 | 7.3598 | 7.5282 | 110.7252 | 327.794 | 9157 | 16.667 | 7.3613 | 33.6037 | 4.3496 | 2.948 | 82.6555 | 1.9481 | 63.17
[5 empty] | 32.8482 | 7.3948 | 7.511 | 110.5604 | 327.044 | 9188 | 16.312 | 7.3963 | 33.6731 | 4.3441 | 2.7649 | 82.69 | 2.0059 | 63.17
[5 empty] | 32.8514 | 7.4307 | 7.5266 | 110.8847 | 327.725 | 9217 | 15.943 | 7.4322 | 33.6933 | 4.3552 | 2.9394 | 82.7122 | 2.158 | 63.17
[5 empty] | 32.8528 | 7.3776 | 7.5521 | 111.1246 | 328.832 | 9245 | 15.968 | 7.3791 | 33.6609 | 4.3635 | 2.907 | 82.6547 | 2.2009 | 63.17
[5 empty] | 32.8529 | 7.3485 | 7.5465 | 110.9686 | 328.59 | 9274 | 16.305 | 7.35 | 33.6444 | 4.3581 | 2.9859 | 82.6696 | 2.1144 | 63.17
[5 empty] | 32.8493 | 7.3303 | 7.5358 | 110.7612 | 328.123 | 9304 | 16.519 | 7.3318 | 33.6142 | 4.3509 | 3.1203 | 82.6453 | 2.0285 | 63.17
[5 empty] | 32.8499 | 7.3384 | 7.5334 | 110.7468 | 328.017 | 9333 | 16.343 | 7.3399 | 33.6165 | 4.3505 | 2.8766 | 82.6388 | 2.0547 | 63.17
[5 empty] | 32.85 | 7.3643 | 7.5352 | 110.8408 | 328.1 | 9362 | 16.023 | 7.3657 | 33.6561 | 4.3538 | 3.0147 | 82.6347 | 2.1413 | 63.17
[5 empty] | 32.8518 | 7.382 | 7.5525 | 111.1412 | 328.85 | 9393 | 15.948 | 7.3835 | 33.6591 | 4.364 | 3.039 | 82.6297 | 2.1614 | 63.17
[5 empty] | 32.8519 | 7.3487 | 7.5683 | 111.2879 | 329.537 | 9424 | 16.214 | 7.3501 | 33.6469 | 4.3689 | 2.9419 | 82.5758 | 2.0654 | 63.17
[5 empty] | 32.8503 | 7.3164 | 7.5624 | 111.117 | 329.279 | 9455 | 16.447 | 7.3179 | 33.6402 | 4.3631 | 3.1775 | 82.5818 | 1.9888 | 63.17
[5 empty] | 32.8496 | 7.3257 | 7.5426 | 110.8493 | 328.417 | 9495 | 16.387 | 7.3272 | 33.6192 | 4.354 | 3.0177 | 82.5349 | 2.0322 | 63.17
[5 empty] | 32.8509 | 7.3442 | 7.5503 | 111.012 | 328.756 | 9575 | 16.312 | 7.3457 | 33.6321 | 4.3596 | 3.0632 | 82.5489 | 2.044 | 63.17
| 019 | 019 | NP9 | 1000 | 34.4053 | 2.6825 | 0.40753 | 5.4169 | 17.714 | 29552 | 1000.088 | 2.7482 | 31.3964 | 0.7572 | 0.054322 | 91.0708 | 0.072837 | 169.9 | 214 | 1 | 1697 | 0.396 |  |  |  | 0.396
|  |  |  | 1000 | 34.4055 | 2.6821 | 0.40749 | 5.4164 | 17.712 | 29597 | 1000.828 | 2.7479 | 31.3965 | 0.7572 | 0.05416 | 91.0673 | 0.073474 | 169.9 | 215 | 2
|  |  |  | 1000 | 34.4054 | 2.6821 | 0.40765 | 5.4184 | 17.719 | 29642 | 1001.181 | 2.7479 | 31.3969 | 0.7572 | 0.055098 | 91.072 | 0.07233 | 169.9 | 216 | 3
|  |  |  | 800 | 34.3416 | 2.9939 | 0.33634 | 4.5025 | 14.621 | 35314 | 800.779 | 3.0471 | 31.5182 | 0.7336 | 0.055344 | 91.078 | 0.070406 | 176.88 | 217 | 4 | 1743 | 0.328 |  |  |  | 0.328
|  |  |  | 600 | 34.2585 | 3.3518 | 0.30597 | 4.1291 | 13.302 | 41135 | 600.913 | 3.3923 | 31.6643 | 0.7256 | 0.057312 | 91.0541 | 0.07164 | 172.08 | 218 | 5 | 1739, 1753 | 0.299 | 0.299 |  |  | 0.299
|  |  |  | 400 | 34.139 | 3.7384 | 0.33537 | 4.5643 | 14.582 | 47168 | 400.574 | 3.7659 | 31.803 | 0.7432 | 0.05801 | 91.0095 | 0.069819 | 204.76 | 219 | 6 | 1725 | 0.332 |  |  |  | 0.332
|  |  |  | 300 | 34.0618 | 3.9108 | 0.36502 | 4.9857 | 15.873 | 50762 | 302.495 | 3.9317 | 31.8396 | 0.7589 | 0.057237 | 90.9549 | 0.073362 | 211.29 | 220 | 7 | 1731 | 0.36 |  |  |  | 0.36
|  |  |  | 250 | 34.0258 | 4.0227 | 0.39534 | 5.413 | 17.192 | 53310 | 252.709 | 4.0402 | 31.8821 | 0.7741 | 0.056839 | 90.8936 | 0.069651 | 213.11 | 221 | 8 | 1193 | 0.396 |  |  |  | 0.396
|  |  |  | 200 | 33.964 | 4.0518 | 0.51253 | 7.019 | 22.289 | 55946 | 200.313 | 4.0657 | 31.8286 | 0.8286 | 0.056864 | 90.9008 | 0.069815 | 211.29 | 222 | 9 | 1750 | 0.514 |  |  |  | 0.514
|  |  |  | 175 | 33.9355 | 4.0866 | 0.64364 | 8.8201 | 27.991 | 57854 | 174.833 | 4.0987 | 31.8218 | 0.8893 | 0.066382 | 90.8831 | 0.069735 | 211.29 | 223 | 10 | 1744 | 0.65 |  |  |  | 0.65
|  |  |  | 150 | 33.89 | 4.1523 | 0.92146 | 12.6433 | 40.075 | 60242 | 150.925 | 4.1628 | 31.8276 | 1.0178 | 0.056889 | 90.8597 | 0.069667 | 209.11 | 224 | 11 | 1740 | 0.94 |  |  |  | 0.94
|  |  |  | 125 | 33.8382 | 4.1663 | 1.2437 | 17.0644 | 54.093 | 61869 | 124.659 | 4.1749 | 31.784 | 1.1673 | 0.058733 | 90.7921 | 0.069635 | 211.78 | 225 | 12 | 1726 | 1.287 |  |  |  | 1.287
|  |  |  | 100 | 33.6016 | 4.0736 | 2.6112 | 35.6874 | 113.588 | 63375 | 99.402 | 4.0803 | 31.4875 | 1.7951 | 0.088561 | 90.7603 | 0.069566 | 213.47 | 226 | 13
|  |  |  | 100 | 33.6035 | 4.0756 | 2.6087 | 35.6553 | 113.478 | 63423 | 99.349 | 4.0824 | 31.4901 | 1.794 | 0.086284 | 90.7431 | 0.07278 | 213.47 | 227 | 14 | 1732, 1752 | 2.586 | 2.58 |  |  | 2.586
|  |  |  | 75 | 32.9438 | 3.7471 | 6.6391 | 89.6111 | 288.94 | 65267 | 73.554 | 3.7518 | 30.6401 | 3.6182 | 0.3507 | 90.4761 | 0.096675 | 213.47 | 228 | 15
|  |  |  | 75 | 32.9483 | 3.7376 | 6.6425 | 89.6382 | 289.085 | 65365 | 74.072 | 3.7423 | 30.638 | 3.6189 | 0.32185 | 90.4546 | 0.089254 | 213.29 | 229 | 16 | 1709 | 6.89 |  |  |  | 6.89
|  |  |  | 50 | 32.8806 | 4.0708 | 7.0759 | 96.2297 | 307.972 | 66922 | 49.518 | 4.0741 | 30.8504 | 3.8513 | 0.87862 | 89.4208 | 0.26331 | 213.51 | 230 | 17
|  |  |  | 50 | 32.881 | 4.0709 | 7.0836 | 96.3353 | 308.309 | 66984 | 49.598 | 4.0742 | 30.8509 | 3.8548 | 0.86746 | 89.439 | 0.26862 | 213.87 | 231 | 18 | 1745 | 7.363 |  |  |  | 7.363
|  |  |  | 25 | 32.8572 | 5.9723 | 7.5591 | 107.6236 | 329.078 | 68728 | 24.76 | 5.9743 | 32.4544 | 4.244 | 2.8221 | 85.2506 | 2.3665 | 217.83 | 232 | 19
|  |  |  | 25 | 32.8568 | 5.9873 | 7.5526 | 107.5693 | 328.796 | 68772 | 24.846 | 5.9893 | 32.455 | 4.2421 | 3.0165 | 85.2931 | 2.3615 | 217.83 | 233 | 20 | 1741 | 7.795 |  |  | b | 7.795 | floaties
|  |  |  | 10 | 32.8491 | 7.8657 | 7.4082 | 110.2369 | 322.589 | 70241 | 9.434 | 7.8666 | 34.1545 | 4.3354 | 2.7455 | 82.806 | 15.361 | 222.98 | 234 | 21
|  |  |  | 10 | 32.8538 | 7.905 | 7.4044 | 110.2835 | 322.425 | 70298 | 9.691 | 7.9059 | 34.1682 | 4.3367 | 2.7398 | 82.7423 | 14.414 | 224 | 235 | 22 | 1727 | 7.694 |  |  | b | 7.694 | floaties
|  |  |  | 5 | 32.8249 | 9.8373 | 6.6357 | 103.214 | 289.045 | 71847 | 3.295 | 9.8376 | 35.8012 | 4.0952 | 1.8113 | 83.6894 | 38.297 | 230.9 | 236 | 23
|  |  |  | 5 | 32.8251 | 9.831 | 6.6419 | 103.2954 | 289.312 | 71898 | 3.473 | 9.8314 | 35.7927 | 4.0979 | 1.7792 | 83.6411 | 36.869 | 230.9 | 237 | 24 | 1733 | 6.89 |  |  | b | 6.89 | floaties
| 020 | 020 | NP9 | 300 | 34.0684 | 3.9066 | 0.38259 | 5.2254 | 16.636 | 14952 | 300.865 | 3.9274 | 31.839 | 0.7668 | 0.056922 | 90.9478 | 0.075005 | 498.56
|  |  |  | 300 | 34.0685 | 3.9062 | 0.38267 | 5.2263 | 16.639 | 14997 | 301.089 | 3.9269 | 31.8393 | 0.7668 | 0.05843 | 90.9587 | 0.075184 | 498.59
|  |  |  | 120 | 33.811 | 4.1804 | 1.4234 | 19.5325 | 61.909 | 20319 | 119.976 | 4.1887 | 31.7697 | 1.2503 | 0.068127 | 90.807 | 0.077434 | 509.71
|  |  |  | 120 | 33.8111 | 4.1814 | 1.4244 | 19.5473 | 61.954 | 20369 | 119.916 | 4.1897 | 31.7706 | 1.2508 | 0.068476 | 90.7893 | 0.074875 | 509.8
|  |  |  | 80 | 33.0551 | 3.7193 | 6.0361 | 81.4766 | 262.671 | 22444 | 80.611 | 3.7244 | 30.7132 | 3.3413 | 0.23127 | 90.5886 | 0.13967 | 518.43
|  |  |  | 80 | 33.0559 | 3.7191 | 6.0338 | 81.4459 | 262.572 | 22496 | 80.653 | 3.7242 | 30.7132 | 3.3402 | 0.19843 | 90.6255 | 0.1334 | 518.43
|  |  |  | 70 | 32.9335 | 3.7882 | 6.8044 | 91.9286 | 296.136 | 23559 | 69.303 | 3.7927 | 30.6645 | 3.6981 | 0.39613 | 90.419 | 0.20921 | 520.61
|  |  |  | 70 | 32.9336 | 3.7883 | 6.8134 | 92.0508 | 296.529 | 23605 | 69.36 | 3.7927 | 30.6644 | 3.7022 | 0.39398 | 90.4469 | 0.21026 | 521.42
|  |  |  | 36 | 32.8611 | 5.575 | 7.4242 | 104.7105 | 323.19 | 25642 | 35.649 | 5.5778 | 32.1195 | 4.141 | 2.4052 | 86.5111 | 2.3114 | 524.96
|  |  |  | 36 | 32.861 | 5.577 | 7.4297 | 104.7936 | 323.431 | 25704 | 35.784 | 5.5798 | 32.1188 | 4.1438 | 2.3896 | 86.5695 | 2.2978 | 524.96
|  |  |  | 15 | 32.8327 | 7.066 | 7.6534 | 111.788 | 333.235 | 27461 | 14.891 | 7.0674 | 33.3922 | 4.3873 | 2.9767 | 82.6197 | 22.136 | 529.32
|  |  |  | 15 | 32.8406 | 7.0422 | 7.653 | 111.7256 | 333.214 | 27504 | 15.221 | 7.0435 | 33.3361 | 4.3851 | 2.8561 | 82.6336 | 22.423 | 529.32
|  |  |  | 11 | 32.8231 | 9.4137 | 6.8536 | 105.6034 | 298.513 | 28812 | 11.652 | 9.4149 | 35.3824 | 4.1733 | 2.1497 | 83.0937 | 33.045 | 529.5
|  |  |  | 11 | 32.8284 | 9.3719 | 6.8298 | 105.1427 | 297.474 | 28854 | 11.714 | 9.3731 | 35.3932 | 4.1577 | 2.1776 | 83.0994 | 32.442 | 529.37
|  |  |  | 5 | 32.8258 | 9.8477 | 6.6683 | 103.7459 | 290.464 | 31046 | 6.951 | 9.8485 | 35.8247 | 4.1115 | 1.4428 | 83.5319 | 55.631 | 529.32
|  |  |  | 5 | 32.8256 | 9.8465 | 6.666 | 103.7079 | 290.365 | 31106 | 7.345 | 9.8473 | 35.8233 | 4.11 | 1.4971 | 83.5352 | 52.202 | 528.52
|  |  |  | 5 | 32.8255 | 9.8469 | 6.6633 | 103.6659 | 290.245 | 31213 | 7.228 | 9.8477 | 35.8231 | 4.1086 | 1.4462 | 83.5343 | 52.417 | 529.01
|  |  |  | 3 | 32.8258 | 9.8408 | 6.6596 | 103.5949 | 290.084 | 31548 | 7.13 | 9.8416 | 35.819 | 4.1063 | 1.5536 | 83.4955 | 53.072 | 527.14
|  |  |  | 3 | 32.8247 | 9.8433 | 6.743 | 104.8969 | 293.718 | 32948 | 4.918 | 9.8438 | 35.8185 | 4.1515 | 1.4694 | 83.4111 | 70.292 | 520.61
|  |  |  | 3 | 32.8252 | 9.841 | 6.7386 | 104.8229 | 293.525 | 32990 | 4.992 | 9.8415 | 35.8181 | 4.1489 | 1.5931 | 83.4494 | 67.831 | 520.47
|  |  |  | 0 | 32.8257 | 9.8543 | 6.7514 | 105.0529 | 294.084 | 34452 | 2.38 | 9.8546 | 35.8301 | 4.1579 | 1.1939 | 83.2192 | 135.12 | 516.25
|  |  |  | 0 | 32.8248 | 9.8667 | 6.7487 | 105.0399 | 293.968 | 34509 | 2.849 | 9.867 | 35.8339 | 4.1573 | 1.2337 | 83.1565 | 100.04 | 516.25
| 021 | 021 | NP10 | 1000 | 34.3996 | 2.7024 | 0.40603 | 5.3995 | 17.649 | 31271 | 1000.63 | 2.7683 | 31.4089 | 0.7567 | 0.055394 | 91.0665 | 0.068474 | 1285.5 | 260 | 1 | 1697 | 0.392 |  |  |  | 0.392
|  |  |  | 800 | 34.3407 | 2.9919 | 0.34976 | 4.6819 | 15.205 | 37230 | 802.395 | 3.0452 | 31.5183 | 0.7391 | 0.055867 | 91.0325 | 0.068424 | 1292.5 | 261 | 2 | 1743 | 0.336 |  |  |  | 0.336
|  |  |  | 600 | 34.2657 | 3.3578 | 0.3043 | 4.1073 | 13.229 | 43135 | 600.296 | 3.3983 | 31.6754 | 0.725 | 0.05649 | 91.0157 | 0.069742 | 1246.6 | 262 | 3 | 1739 | 0.296 |  |  |  | 0.296
|  |  |  | 400 | 34.1449 | 3.7181 | 0.36217 | 4.9267 | 15.747 | 49227 | 399.844 | 3.7455 | 31.7898 | 0.7549 | 0.056739 | 91.0138 | 0.067689 | 1172.2 | 263 | 4 | 1725 | 0.357 |  |  |  | 0.357
|  |  |  | 300 | 34.0556 | 3.9008 | 0.42524 | 5.8066 | 18.491 | 52795 | 299.691 | 3.9214 | 31.8243 | 0.7858 | 0.059405 | 90.9439 | 0.068927 | 1151.5 | 264 | 5 | 1731 | 0.42 |  |  |  | 0.42
|  |  |  | 250 | 33.9968 | 3.9122 | 0.49875 | 6.809 | 21.689 | 54977 | 249.949 | 3.9294 | 31.7593 | 0.8201 | 0.056889 | 90.9435 | 0.067572 | 1128.7 | 265 | 6 | 1193 | 0.496 |  |  |  | 0.496
|  |  |  | 200 | 33.9419 | 4.0898 | 0.63912 | 8.7596 | 27.794 | 57043 | 199.717 | 4.1036 | 31.8435 | 0.8862 | 0.058085 | 90.8989 | 0.067561 | 1100.7 | 266 | 7 | 1750 | 0.638 |  |  |  | 0.638
|  |  |  | 175 | 33.9068 | 4.0585 | 0.81238 | 11.1226 | 35.33 | 58515 | 174.636 | 4.0705 | 31.776 | 0.9658 | 0.056166 | 90.873 | 0.06779 | 1079.3 | 267 | 8 | 1744, 1740 | 0.814 | 0.812 |  |  | 0.814
|  |  |  | 150 | 33.8406 | 4.0518 | 1.2199 | 16.6918 | 53.057 | 60006 | 148.818 | 4.062 | 31.6996 | 1.153 | 0.063118 | 90.8209 | 0.070485 | 1060.6 | 268 | 9 | 1726 | 1.226 |  |  | b | 1.226 | copepod in sample
|  |  |  | 125 | 33.7531 | 4.1703 | 1.7801 | 24.4126 | 77.428 | 61520 | 124.512 | 4.1789 | 31.7165 | 1.4134 | 0.079988 | 90.7912 | 0.13686 | 1034.9 | 269 | 10 | 1732 | 1.772 |  |  |  | 1.772
|  |  |  | 100 | 33.548 | 3.9579 | 2.8494 | 38.818 | 123.953 | 62955 | 100.686 | 3.9646 | 31.3423 | 1.9001 | 0.11163 | 90.7323 | 0.32297 | 1020.7 | 270 | 11 | 1752 | 2.972 |  |  |  | 2.972
|  |  |  | 75 | 33.1939 | 3.7935 | 5.0017 | 67.7003 | 217.636 | 64770 | 76.454 | 3.7984 | 30.8931 | 2.8783 | 0.19578 | 90.7002 | 1.0023 | 1010.9 | 271 | 12 | 1709 | 5.246 |  |  |  | 5.246
|  |  |  | 50 | 32.9441 | 3.9225 | 6.9291 | 93.9285 | 301.563 | 66208 | 50.92 | 3.9258 | 30.7801 | 3.7731 | 0.82158 | 90.0553 | 3.4749 | 1006.7 | 272 | 13 | 1745 | 7.197 |  |  |  | 7.197
|  |  |  | 25 | 32.7619 | 4.4695 | 7.339 | 100.7074 | 319.468 | 67781 | 25.057 | 4.4712 | 31.0593 | 4.013 | 0.93033 | 88.4014 | 20.677 | 1009.9 | 273 | 14 | 1741 | 7.534 |  |  |  | 7.534
|  |  |  | 10 | 32.7166 | 9.8815 | 6.5359 | 101.6924 | 284.722 | 69125 | 10.283 | 9.8826 | 35.7616 | 4.0403 | 0.62739 | 87.8939 | 69.185 | 1006.3 | 274 | 15 | 1727, 1733 | 6.958 | 6.961 |  |  | 6.958
|  |  |  | 5 | 32.7091 | 10.4553 | 6.4478 | 101.5901 | 280.911 | 70086 | 2.854 | 10.4557 | 36.259 | 4.0389 | 0.36894 | 87.8819 | 159.46 | 1007 | 275 | 16 | 1753 | 6.72 |  |  |  | 6.72
| 022 | 022 | NP11 | 1000 | 34.4115 | 2.6753 | 0.37024 | 4.9206 | 16.093 | 30808 | 999.034 | 2.7409 | 31.3942 | 0.7425 | 0.05507 | 91.0526 | 0.074077 | 317.14 | 276 | 1 | 1135 | 0.366 |  |  |  | 0.366
|  |  |  | 800 | 34.3479 | 2.9921 | 0.31583 | 4.2279 | 13.729 | 36923 | 802.074 | 3.0454 | 31.5229 | 0.7252 | 0.055219 | 91.0445 | 0.071882 | 392.36 | 277 | 2 | 1094 | 0.306 |  |  |  | 0.306
|  |  |  | 600 | 34.2657 | 3.3558 | 0.2951 | 3.9829 | 12.829 | 42904 | 600.516 | 3.3964 | 31.674 | 0.721 | 0.056664 | 91.0234 | 0.070169 | 480.37 | 278 | 3 | 1293, 1654 | 0.289 | 0.295 |  |  | 0.289
|  |  |  | 400 | 34.1535 | 3.7532 | 0.31973 | 4.3536 | 13.902 | 48940 | 400.593 | 3.7807 | 31.8285 | 0.7364 | 0.056664 | 91.0004 | 0.071384 | 565.91 | 279 | 4 | 1127 | 0.316 |  |  |  | 0.316
|  |  |  | 300 | 34.0742 | 3.9523 | 0.36014 | 4.9245 | 15.66 | 52289 | 300.721 | 3.9732 | 31.8862 | 0.7569 | 0.05639 | 90.9564 | 0.070012 | 587.24 | 280 | 5 | 1284 | 0.363 |  |  |  | 0.363
|  |  |  | 250 | 34.0324 | 4.0429 | 0.40183 | 5.5048 | 17.474 | 54927 | 252.878 | 4.0605 | 31.9058 | 0.7771 | 0.056864 | 90.9353 | 0.07139 | 637.39 | 281 | 6 | 1143 | 0.404 |  |  |  | 0.404
|  |  |  | 200 | 33.9616 | 4.1486 | 0.55678 | 7.6431 | 24.213 | 57102 | 200.655 | 4.1626 | 31.912 | 0.8492 | 0.059306 | 90.8683 | 0.069407 | 686.37 | 282 | 7 | 1089 | 0.559 |  |  |  | 0.559
|  |  |  | 175 | 33.9338 | 4.186 | 0.6504 | 8.9344 | 28.285 | 58792 | 175.301 | 4.1983 | 31.9082 | 0.8931 | 0.062221 | 90.8434 | 0.068537 | 689.66 | 283 | 8 | 1509 | 0.666 |  |  |  | 0.666
|  |  |  | 150 | 33.8913 | 4.2296 | 0.86199 | 11.8498 | 37.489 | 60798 | 150.269 | 4.2402 | 31.8975 | 0.9913 | 0.066956 | 90.8046 | 0.069565 | 690.69 | 284 | 9 | 879 | 0.893 |  |  |  | 0.893
|  |  |  | 125 | 33.7294 | 4.1379 | 1.7608 | 24.125 | 76.589 | 62573 | 125.706 | 4.1466 | 31.667 | 1.4036 | 0.082579 | 90.6619 | 0.066759 | 722.38 | 285 | 10 | 1746 | 1.866 |  |  |  | 1.866
|  |  |  | 100 | 33.3987 | 3.84 | 3.8081 | 51.6762 | 165.674 | 64212 | 100.374 | 3.8465 | 31.1154 | 2.3324 | 0.15519 | 90.6729 | 0.086484 | 740.61 | 286 | 11 | 1145 | 3.959 |  |  |  | 3.959
|  |  |  | 75 | 32.9967 | 3.7064 | 6.2573 | 84.4022 | 272.31 | 65953 | 75.498 | 3.7112 | 30.6523 | 3.4419 | 0.67531 | 90.0064 | 0.17579 | 764.57 | 287 | 12 | 1748 | 6.524 |  |  |  | 6.524
|  |  |  | 50 | 32.8424 | 4.4541 | 7.1134 | 97.6311 | 309.629 | 67545 | 50.243 | 4.4575 | 31.1465 | 3.8974 | 1.4364 | 88.7768 | 0.96782 | 785.38 | 288 | 13 | 1658 | 7.373 |  |  |  | 7.373
|  |  |  | 25 | 32.7526 | 7.0217 | 7.1947 | 104.927 | 313.279 | 69476 | 25.034 | 7.0239 | 33.3862 | 4.1499 | 1.524 | 88.5795 | 6.424 | 809.7 | 289 | 14 | 1156, 663 | 7.429 | 7.426 |  |  | 7.429
|  |  |  | 10 | 32.6826 | 9.3947 | 6.71 | 103.2544 | 292.289 | 71261 | 9.909 | 9.3958 | 35.2474 | 4.0945 | 0.83128 | 88.5714 | 21.844 | 831.25 | 290 | 15 | 1376 | 6.963 |  |  |  | 6.963
|  |  |  | 5 | 32.6892 | 10.0082 | 6.569 | 102.4742 | 286.175 | 72719 | 3.182 | 10.0085 | 35.8487 | 4.0698 | 0.97398 | 88.5192 | 50.78 | 851.7 | 291 | 16 | 1137 | 6.852 |  |  |  | 6.852
| 023 | 023 | NP12 | 1000 | 34.4503 | 2.472 | 0.46111 | 6.0983 | 20.042 | 33730 | 1000.838 | 2.5362 | 31.2482 | 0.7776 | 0.055892 | 91.0761 | 0.072893 | 1123.3 | 292 | 1 | 1697 | 0.47 |  |  |  | 0.47
|  |  |  | 800 | 34.3822 | 2.8109 | 0.35447 | 4.7246 | 15.408 | 39642 | 801.249 | 2.863 | 31.3916 | 0.7404 | 0.055394 | 91.0459 | 0.07018 | 1051.1 | 293 | 2 | 1743 | 0.351 |  |  |  | 0.351
|  |  |  | 600 | 34.2639 | 3.3525 | 0.3064 | 4.135 | 13.32 | 45442 | 600.242 | 3.393 | 31.6695 | 0.7258 | 0.055992 | 90.9951 | 0.069197 | 1058.8 | 294 | 3 | 1739 | 0.303 |  |  |  | 0.303
|  |  |  | 400 | 34.1302 | 3.8356 | 0.37546 | 5.1219 | 16.325 | 51296 | 400.164 | 3.8634 | 31.8814 | 0.7612 | 0.05659 | 90.9664 | 0.069615 | 1038.4 | 295 | 4 | 1725 | 0.372 |  |  |  | 0.372
|  |  |  | 300 | 34.0556 | 4.0334 | 0.46839 | 6.4167 | 20.368 | 54738 | 300.617 | 4.0545 | 31.9414 | 0.8058 | 0.057237 | 90.9085 | 0.069427 | 1021 | 296 | 5 | 1731 | 0.475 |  |  |  | 0.475
|  |  |  | 250 | 33.9951 | 4.0976 | 0.49705 | 6.8168 | 21.615 | 56964 | 252.498 | 4.1152 | 31.922 | 0.8203 | 0.056764 | 90.9234 | 0.069038 | 984.98 | 297 | 6 | 1193 | 0.5 |  |  | b | 0.5 | dancing bubbles
|  |  |  | 200 | 33.9309 | 4.2373 | 0.7497 | 10.3117 | 32.604 | 59247 | 199.351 | 4.2513 | 31.9629 | 0.9377 | 0.055468 | 90.9123 | 0.068659 | 955.28 | 298 | 7 | 1750 | 0.752 |  |  |  | 0.752
|  |  |  | 175 | 33.8953 | 4.2702 | 0.92917 | 12.787 | 40.41 | 60949 | 174.752 | 4.2825 | 31.9495 | 1.0212 | 0.055518 | 90.8898 | 0.068726 | 946.26 | 299 | 8 | 1744 | 0.941 |  |  |  | 0.941
|  |  |  | 150 | 33.8176 | 4.3949 | 1.4067 | 19.4066 | 61.182 | 62766 | 150.059 | 4.4056 | 31.9802 | 1.2434 | 0.062794 | 90.8621 | 0.070653 | 936.16 | 300 | 9 | 1740 | 1.401 |  |  |  | 1.401
|  |  |  | 125 | 33.5543 | 4.3242 | 2.9331 | 40.3227 | 127.596 | 64501 | 124.408 | 4.3329 | 31.6804 | 1.946 | 0.075852 | 90.8367 | 0.11131 | 928.3 | 301 | 10 | 1726, 1732 | 3.01 | 3.015 |  |  | 3.01
|  |  |  | 100 | 33.2434 | 3.9581 | 4.6215 | 62.8327 | 201.088 | 65992 | 99.815 | 3.9647 | 31.0893 | 2.7072 | 0.1308 | 90.7059 | 0.26541 | 936.03 | 302 | 11 | 1752 | 4.753 |  |  |  | 4.753
|  |  |  | 75 | 33.2409 | 3.9554 | 4.6307 | 62.9525 | 201.489 | 66205 | 99.877 | 3.962 | 31.0853 | 2.7112 | 0.12098 | 90.7141 | 0.264 | 934.83 | 303 | 12 | 1709 | 6.08 |  |  |  | 6.08
|  |  |  | 50 | 33.2415 | 3.9561 | 4.6146 | 62.7349 | 200.788 | 66312 | 99.835 | 3.9627 | 31.075 | 2.704 | 0.11941 | 90.6983 | 0.25614 | 936.61 | 304 | 13 | 1745 | 7.195 |  |  |  | 7.195
|  |  |  | 25 | 32.9849 | 3.9966 | 5.8921 | 80.043 | 256.427 | 67981 | 74.895 | 4.0016 | 30.888 | 3.2945 | 0.4514 | 90.2003 | 0.85471 | 925.58 | 305 | 14 | 1741 | 7.174 |  |  |  | 7.174
|  |  |  | 10 | 32.7231 | 4.8236 | 6.9177 | 95.7295 | 301.149 | 69435 | 50.337 | 4.8273 | 31.3589 | 3.8323 | 1.8493 | 88.3502 | 4.6234 | 935.45 | 306 | 15 | 1727, 1733 | 6.746 | 6.747 |  |  | 6.746
|  |  |  | 5 | 32.7233 | 4.8243 | 6.9251 | 95.8326 | 301.469 | 69594 | 50.384 | 4.8279 | 31.3579 | 3.8357 | 1.8466 | 88.2109 | 4.6704 | 934.92 | 307 | 16 | 1753 | 6.73 |  |  |  | 6.73
|  |  |  | 100 | 32.7244 | 4.8203 | 6.9327 | 95.9301 | 301.802 | 69683 | 50.371 | 4.8239 | 31.3545 | 3.839 | 1.8747 | 88.2745 | 4.718 | 932.83
|  |  |  | 100 | 32.6608 | 6.7698 | 6.9309 | 100.427 | 301.805 | 71545 | 25.543 | 6.772 | 32.958 | 3.9978 | 0.84055 | 88.942 | 26.132 | 918.74
|  |  |  | 50 | 32.6617 | 6.769 | 6.9258 | 100.3524 | 301.585 | 71702 | 25.795 | 6.7712 | 32.9576 | 3.9952 | 0.87149 | 89.0004 | 26.404 | 915.14
|  |  |  | 50 | 32.6617 | 6.7587 | 6.9244 | 100.3075 | 301.522 | 71776 | 25.739 | 6.761 | 32.9498 | 3.9937 | 0.83698 | 88.9612 | 25.887 | 917.58
|  |  |  | 25 | 32.6004 | 9.9917 | 6.4861 | 101.0896 | 282.584 | 73588 | 9.74 | 9.9928 | 35.7534 | 4.0198 | 0.44355 | 88.6351 | 64.498 | 900.33
|  |  |  | 25 | 32.6018 | 10.0944 | 6.5416 | 102.185 | 285.006 | 75113 | 2.39 | 10.0947 | 35.8065 | 4.0601 | 0.38713 | 88.4613 | 143.95 | 891.76
|  |  |  | 5 | 32.6026 | 10.0776 | 6.5506 | 102.2872 | 285.395 | 75329 | 2.905 | 10.0779 | 35.81 | 4.0634 | 0.40223 | 88.4536 | 139.95 | 893.27
|  |  |  | 5 | 32.6028 | 10.0751 | 6.5366 | 102.063 | 284.785 | 75389 | 2.383 | 10.0753 | 35.8096 | 4.056 | 0.38641 | 88.4402 | 151.95 | 897.49
| 024 | 024 | NP12 | 101 | 33.3965 | 4.1636 | 3.8464 | 52.6133 | 167.345 | 9405 | 102.632 | 4.1706 | 31.3906 | 2.3629 | 0.083651 | 90.7309 | 0.19152 | 500.56
|  |  |  | 101 | 33.3957 | 4.1623 | 3.8502 | 52.6642 | 167.513 | 9473 | 102.387 | 4.1693 | 31.3886 | 2.3647 | 0.08846 | 90.7405 | 0.18774 | 498.06
|  |  |  | 50 | 32.7456 | 4.6031 | 6.8814 | 94.7305 | 299.556 | 12428 | 49.76 | 4.6066 | 31.1872 | 3.7992 | 1.8494 | 88.3294 | 2.2174 | 389.91
|  |  |  | 50 | 32.7456 | 4.6045 | 6.8868 | 94.808 | 299.791 | 12459 | 49.659 | 4.608 | 31.1886 | 3.8019 | 1.9113 | 88.3564 | 2.2382 | 389.14
|  |  |  | 50 | 32.7461 | 4.6035 | 6.8983 | 94.964 | 300.29 | 12490 | 49.748 | 4.607 | 31.1888 | 3.8071 | 1.9123 | 88.3321 | 2.2213 | 387.73
|  |  |  | 50 | 32.7455 | 4.6025 | 6.9028 | 95.0235 | 300.487 | 12523 | 49.871 | 4.606 | 31.188 | 3.8091 | 1.9467 | 88.1914 | 2.1688 | 385.91
|  |  |  | 50 | 32.7458 | 4.6035 | 6.8824 | 94.7454 | 299.6 | 12556 | 49.78 | 4.607 | 31.1889 | 3.7997 | 1.8459 | 88.3094 | 2.1704 | 385.55
|  |  |  | 50 | 32.7458 | 4.6054 | 6.8642 | 94.4986 | 298.806 | 12588 | 49.578 | 4.6089 | 31.1899 | 3.7915 | 1.8493 | 88.3194 | 2.214 | 384.06
|  |  |  | 50 | 32.746 | 4.6049 | 6.8812 | 94.7318 | 299.546 | 12626 | 49.578 | 4.6084 | 31.1897 | 3.7993 | 1.985 | 88.3333 | 2.2225 | 382.84
|  |  |  | 50 | 32.7461 | 4.6042 | 6.8841 | 94.7703 | 299.673 | 12667 | 49.776 | 4.6077 | 31.1889 | 3.8006 | 1.9553 | 88.3236 | 2.1834 | 381.2
|  |  |  | 50 | 32.7461 | 4.6041 | 6.8743 | 94.6348 | 299.245 | 12707 | 49.796 | 4.6076 | 31.1893 | 3.796 | 1.9213 | 88.3362 | 2.1475 | 379.13
|  |  |  | 50 | 32.746 | 4.605 | 6.8796 | 94.7109 | 299.479 | 12751 | 49.689 | 4.6085 | 31.1908 | 3.7986 | 1.9131 | 88.3402 | 2.1657 | 378.84
|  |  |  | 50 | 32.7462 | 4.6054 | 6.8961 | 94.9386 | 300.196 | 12819 | 49.632 | 4.6089 | 31.1892 | 3.8063 | 2.0334 | 88.2511 | 2.176 | 376.31
|  |  |  | 20 | 32.6216 | 8.5399 | 6.6532 | 100.3921 | 289.79 | 14568 | 19.586 | 8.5418 | 34.4778 | 3.9951 | 0.80133 | 88.9641 | 15.492 | 333.21
|  |  |  | 20 | 32.6199 | 8.5445 | 6.6662 | 100.5987 | 290.36 | 14611 | 19.627 | 8.5465 | 34.4764 | 4.0021 | 0.72093 | 88.9434 | 15.41 | 331.37
|  |  |  | 0 | 32.6008 | 9.968 | 6.5972 | 102.7644 | 287.42 | 16228 | 2.389 | 9.9683 | 35.7081 | 4.0801 | 0.57283 | 88.2572 | 74.641 | 322.38
|  |  |  | 0 | 32.6007 | 9.9685 | 6.5974 | 102.7691 | 287.43 | 16273 | 2.46 | 9.9688 | 35.7093 | 4.0802 | 0.52897 | 88.2786 | 77.734 | 322.38
| 025 | 025 | NP13 |  | 34.3277 | 3.0197 | 0.3752 | 5.0274 | 16.311 | 29237 | 1002.974 | 3.0882 | 31.6296 | 0.7455 | 0.05654 | 91.0995 | 0.075342 | 34.852
[5 empty] | 34.2499 | 3.4416 | 0.31788 | 4.3007 | 13.82 | 35392 | 800.355 | 3.4975 | 31.8358 | 0.7274 | 0.056565 | 90.9674 | 0.072443 | 41.387
[5 empty] | 34.1436 | 3.7999 | 0.35789 | 4.8802 | 15.561 | 41464 | 601.015 | 3.8424 | 31.9626 | 0.7492 | 0.057462 | 91.02 | 0.069022 | 47.922
[5 empty] | 34.0065 | 4.3475 | 0.75793 | 10.4622 | 32.961 | 47253 | 400.089 | 4.3768 | 32.2264 | 0.9334 | 0.056639 | 90.6672 | 0.067911 | 50.1
[5 empty] | 33.9191 | 4.6565 | 1.2388 | 17.2171 | 53.88 | 50501 | 301.3 | 4.679 | 32.3734 | 1.1586 | 0.058409 | 90.6049 | 0.067506 | 52.278
[5 empty] | 33.8415 | 4.8535 | 1.7958 | 25.0621 | 78.109 | 52829 | 250.887 | 4.8725 | 32.454 | 1.4205 | 0.056664 | 90.7352 | 0.06917 | 54.457
[5 empty] | 33.7476 | 5.218 | 2.2514 | 31.6757 | 97.935 | 55180 | 200.45 | 5.2337 | 32.6676 | 1.645 | 0.059879 | 90.4737 | 0.065458 | 56.813
[5 empty] | 33.6165 | 5.4744 | 2.8967 | 40.9707 | 126.026 | 56829 | 175.648 | 5.4884 | 32.7657 | 1.9572 | 0.064065 | 90.4579 | 0.065213 | 57.613
[5 empty] | 33.3394 | 5.3073 | 3.8992 | 54.825 | 169.671 | 58777 | 149.574 | 5.3189 | 32.3623 | 2.4241 | 0.07468 | 90.4924 | 0.065334 | 60.991
[5 empty] | 33.0286 | 4.9653 | 4.6419 | 64.5955 | 202.03 | 60333 | 125.608 | 4.9746 | 31.781 | 2.7574 | 0.10154 | 90.3042 | 0.066164 | 63.17
[5 empty] | 32.8078 | 4.846 | 5.3148 | 73.6356 | 231.355 | 62243 | 101.085 | 4.8534 | 31.4759 | 3.0683 | 0.37211 | 89.9509 | 0.065839 | 65.348
[5 empty] | 32.8075 | 4.8454 | 5.307 | 73.5263 | 231.016 | 62501 | 101.141 | 4.8528 | 31.4746 | 3.0646 | 0.26543 | 89.9863 | 0.0663 | 65.348
[5 empty] | 32.807 | 4.8459 | 5.3167 | 73.6611 | 231.437 | 62714 | 101.333 | 4.8533 | 31.4751 | 3.0691 | 0.2597 | 89.97 | 0.069492 | 65.348
[5 empty] | 32.5585 | 4.6793 | 6.1998 | 85.4042 | 269.926 | 65379 | 74.87 | 4.6846 | 31.0981 | 3.4739 | 0.56403 | 89.8508 | 0.066646 | 67.526
[5 empty] | 32.3043 | 4.6033 | 6.5914 | 90.4748 | 287.03 | 67079 | 50.568 | 4.6068 | 30.8322 | 3.655 | 0.77152 | 88.6365 | 0.11284 | 69.705
[5 empty] | 32.3103 | 4.8126 | 6.5826 | 90.8189 | 286.649 | 67339 | 50.138 | 4.8161 | 30.9904 | 3.6663 | 0.89976 | 88.6174 | 0.11315 | 69.749
[5 empty] | 32.3069 | 4.7375 | 6.5737 | 90.5285 | 286.26 | 67478 | 50.464 | 4.7411 | 30.9048 | 3.6566 | 1.1069 | 88.6739 | 0.11082 | 70.816
[5 empty] | 32.1563 | 7.7187 | 6.8367 | 100.9409 | 297.857 | 69824 | 25.679 | 7.7211 | 33.3274 | 4.0129 | 1.0995 | 88.9736 | 0.70705 | 74.061
[5 empty] | 32.2242 | 6.5188 | 6.9108 | 99.2657 | 301.023 | 71217 | 31.696 | 6.5215 | 32.3561 | 3.9563 | 1.352 | 88.7375 | 0.4762 | 76.239
[5 empty] | 32.2168 | 6.5733 | 6.9153 | 99.4524 | 301.22 | 71338 | 31.337 | 6.576 | 32.3939 | 3.9626 | 1.3024 | 88.7165 | 0.49037 | 76.239
[5 empty] | 32.1194 | 8.5788 | 6.9774 | 105.0346 | 304.03 | 73382 | 5.32 | 8.5793 | 34.0201 | 4.1591 | 1.295 | 87.2489 | 6.713 | 78.418
[5 empty] | 32.12 | 8.5799 | 7.0203 | 105.6839 | 305.9 | 74183 | 5.673 | 8.5805 | 34.0215 | 4.181 | 1.7633 | 87.1254 | 6.7552 | 78.462
[5 empty] | 32.12 | 8.5802 | 7.0065 | 105.4777 | 305.301 | 74405 | 5.682 | 8.5808 | 34.0219 | 4.174 | 1.2981 | 87.1575 | 6.6737 | 79.485
[5 empty] | 32.1201 | 8.5805 | 7.0076 | 105.4951 | 305.349 | 74484 | 5.615 | 8.5811 | 34.0225 | 4.1746 | 1.3239 | 87.0981 | 6.6147 | 80.018
| 026 | 026 | NP13 | 1000 | 34.3316 | 3.0014 | 0.37493 | 5.0216 | 16.299 | 28407 | 1000.658 | 3.0696 | 31.616 | 0.7454 | 0.056465 | 91.0957 | 0.063713 | 111.09 | 333 | 1 | 1135, 1094 | 0.364 | 0.364 |  |  | 0.364
|  |  |  | 800 | 34.255 | 3.4202 | 0.31495 | 4.259 | 13.692 | 33977 | 801.551 | 3.476 | 31.8222 | 0.7261 | 0.056689 | 90.9545 | 0.065052 | 124.16 | 334 | 2 | 1293 | 0.307 |  |  |  | 0.307
|  |  |  | 600 | 34.1522 | 3.7706 | 0.34724 | 4.7318 | 15.098 | 39638 | 601.191 | 3.813 | 31.9438 | 0.7445 | 0.057213 | 91.0258 | 0.063338 | 141.59 | 335 | 3 | 1654 | 0.34 |  |  |  | 0.34
|  |  |  | 400 | 34.0123 | 4.3207 | 0.7212 | 9.9491 | 31.363 | 45311 | 401.829 | 4.3501 | 32.2078 | 0.9168 | 0.05634 | 90.9042 | 0.064671 | 161.19 | 336 | 4 | 1127 | 0.732 |  |  |  | 0.732
|  |  |  | 300 | 33.9134 | 4.6792 | 1.2901 | 17.9395 | 56.112 | 48563 | 300.399 | 4.7017 | 32.3879 | 1.1823 | 0.05806 | 90.1711 | 0.065354 | 178.62 | 337 | 5 | 1284 | 1.332 |  |  |  | 1.332
|  |  |  | 250 | 33.8375 | 4.8394 | 1.8107 | 25.2611 | 78.758 | 50580 | 251.136 | 4.8584 | 32.4386 | 1.4271 | 0.055319 | 90.7926 | 0.064798 | 193.87 | 338 | 6 | 1143 | 1.878 |  |  |  | 1.878
|  |  |  | 200 | 33.7594 | 5.2032 | 2.2046 | 31.0098 | 95.902 | 52649 | 200.684 | 5.2188 | 32.6649 | 1.6229 | 0.060826 | 90.5408 | 0.062935 | 210.71 | 339 | 7 | 1089 | 2.299 |  |  |  | 2.299
|  |  |  | 175 | 33.6651 | 5.5219 | 2.7676 | 39.2017 | 120.405 | 54098 | 175.173 | 5.5359 | 32.8501 | 1.8984 | 0.058608 | 90.5877 | 0.062763 | 223.96 | 340 | 8 | 1509 | 2.889 |  |  |  | 2.889
|  |  |  | 150 | 33.4311 | 5.3655 | 3.5385 | 49.854 | 153.968 | 55552 | 150.366 | 5.3773 | 32.4951 | 2.2581 | 0.07782 | 90.5355 | 0.062942 | 235.25 | 341 | 9 | 879 | 3.686 |  |  |  | 3.686
|  |  |  | 125 | 33.0528 | 5.0181 | 4.587 | 63.9241 | 199.639 | 56971 | 126.489 | 5.0276 | 31.845 | 2.7346 | 0.10595 | 90.3598 | 0.062207 | 248.01 | 342 | 10 | 1746 | 4.856 |  |  |  | 4.856
|  |  |  | 100 | 32.7849 | 4.8191 | 5.3841 | 74.5354 | 234.374 | 58458 | 100.663 | 4.8264 | 31.4322 | 3.0987 | 0.16516 | 90.0208 | 0.062547 | 254.86 | 343 | 11 | 1145 | 5.609 |  |  |  | 5.609
|  |  |  | 75 | 32.7851 | 4.8193 | 5.3863 | 74.5661 | 234.47 | 58630 | 100.833 | 4.8266 | 31.4326 | 3.0996 | 0.1591 | 89.9892 | 0.066214 | 254.86 | 344 | 12 | 1748 | 6.547 |  |  |  | 6.547
|  |  |  | 50 | 32.7849 | 4.8196 | 5.3814 | 74.4982 | 234.255 | 58743 | 100.577 | 4.8268 | 31.4324 | 3.0975 | 0.16469 | 90.0567 | 0.062239 | 256.64 | 345 | 13 | 1658 | 6.892 |  |  |  | 6.892
|  |  |  | 25 | 32.5373 | 4.9089 | 3.1002 | 42.9393 | 134.98 | 60251 | 75.952 | 4.9144 | 31.322 | 2.0425 | 0.45312 | 89.6526 | 0.08598 | 265.75 | 346 | 14 | 1156 | 7.167 |  |  |  | 7.167
|  |  |  | 10 | 32.2561 | 4.2623 | 2.7778 | 37.7993 | 120.963 | 61708 | 50.798 | 4.2657 | 31.2357 | 1.8742 | 0.8862 | 88.7045 | 0.33886 | 276.64 | 347 | 15 | 663 | 7.203 |  |  |  | 7.203
|  |  |  | 5 | 32.2554 | 4.2685 | 2.735 | 37.2226 | 119.1 | 61928 | 50.526 | 4.2718 | 31.2028 | 1.8547 | 1.2046 | 88.5111 | 0.351 | 278.82 | 348 | 16 | 1376, 1137 | 7.227 | 7.206* |  | *c | 7.227 | bad rep, maybe drawn from Bot 15?
|  |  |  | 100 | 32.2556 | 4.266 | 2.7246 | 37.0782 | 118.646 | 61995 | 50.7 | 4.2694 | 31.1934 | 1.8498 | 0.89131 | 88.7294 | 0.35606 | 278.64
|  |  |  | 100 | 32.1955 | 5.0802 | 2.5197 | 34.9628 | 109.737 | 63390 | 30.406 | 5.0824 | 31.1607 | 1.7808 | 0.93426 | 88.6662 | 1.9717 | 287.53
|  |  |  | 50 | 32.1926 | 5.0313 | 2.5148 | 34.8531 | 109.524 | 63486 | 30.639 | 5.0336 | 31.1738 | 1.7772 | 0.98754 | 88.6968 | 1.9683 | 287.57
|  |  |  | 50 | 32.1932 | 6.6697 | 2.3592 | 34.0007 | 102.768 | 64512 | 25.119 | 6.6719 | 31.479 | 1.7478 | 1.6736 | 88.5982 | 3.3595 | 294.07
|  |  |  | 30 | 32.1267 | 8.6242 | 2.5138 | 37.8834 | 109.537 | 65759 | 10.666 | 8.6253 | 32.391 | 1.8801 | 1.2817 | 87.6018 | 15.342 | 300.6
|  |  |  | 30 | 32.1286 | 8.6611 | 2.6468 | 39.9208 | 115.332 | 67173 | 5.241 | 8.6616 | 33.3988 | 1.9501 | 1.2497 | 87.4266 | 33.333 | 307.14
|  |  |  | 5 | 32.1284 | 8.659 | 2.6327 | 39.7056 | 114.716 | 67323 | 5.397 | 8.6596 | 33.4777 | 1.9428 | 1.1507 | 87.3638 | 31.574 | 307.18
|  |  |  | 5 | 32.1288 | 8.6539 | 2.6254 | 39.5912 | 114.398 | 67398 | 4.913 | 8.6544 | 33.5155 | 1.939 | 1.1509 | 87.3241 | 34.621 | 309.31
| 027 | 027 | NP13 | 80 | 32.4867 | 4.8701 | 6.4021 | 88.5612 | 278.756 | 7030 | 80.709 | 4.8759 | 31.205 | 3.5775 | 0.43595 | 89.5472 | 0.088921 | 420.41
|  |  |  | 80 | 32.4872 | 4.8691 | 6.4112 | 88.6851 | 279.152 | 7120 | 80.796 | 4.8749 | 31.2039 | 3.5817 | 0.5248 | 89.5817 | 0.086858 | 418.36
|  |  |  | 30 | 32.1755 | 6.7645 | 7.0099 | 101.2401 | 305.359 | 9356 | 30.537 | 6.7672 | 32.5213 | 4.0231 | 1.9137 | 88.0904 | 3.0908 | 422.58
|  |  |  | 30 | 32.1744 | 6.7651 | 6.9925 | 100.9893 | 304.601 | 9401 | 30.504 | 6.7677 | 32.5194 | 4.0146 | 1.7702 | 87.9966 | 3.0797 | 422.58
|  |  |  | 30 | 32.1737 | 6.7814 | 6.9882 | 100.9647 | 304.413 | 9449 | 30.328 | 6.784 | 32.5594 | 4.0138 | 1.7894 | 88.002 | 3.1382 | 422.58
|  |  |  | 30 | 32.1748 | 6.7872 | 7.0165 | 101.3881 | 305.646 | 9534 | 30.376 | 6.7898 | 32.5436 | 4.0281 | 1.7517 | 87.9887 | 3.1486 | 422.58
|  |  |  | 30 | 32.1738 | 6.7611 | 7.0044 | 101.1509 | 305.118 | 9611 | 30.547 | 6.7637 | 32.492 | 4.02 | 1.8663 | 88.0619 | 3.1255 | 422.58
|  |  |  | 30 | 32.1746 | 6.7614 | 6.9864 | 100.8925 | 304.335 | 9680 | 30.388 | 6.764 | 32.5241 | 4.0114 | 1.7094 | 87.9983 | 3.118 | 423.11
|  |  |  | 30 | 32.1761 | 6.7837 | 6.9737 | 100.7635 | 303.785 | 9750 | 30.306 | 6.7863 | 32.5385 | 4.007 | 1.7167 | 87.9609 | 3.1924 | 422.62
|  |  |  | 30 | 32.1733 | 6.716 | 6.985 | 100.7635 | 304.272 | 9818 | 30.607 | 6.7186 | 32.4926 | 4.007 | 1.7417 | 87.9556 | 3.1187 | 423.29
|  |  |  | 30 | 32.174 | 6.7863 | 6.9695 | 100.7063 | 303.599 | 9874 | 30.099 | 6.7889 | 32.5465 | 4.0051 | 1.6628 | 87.9935 | 3.252 | 424.76
|  |  |  | 30 | 32.1739 | 6.7761 | 6.9964 | 101.0718 | 304.773 | 9961 | 30.449 | 6.7787 | 32.5209 | 4.0174 | 1.8587 | 88.0691 | 3.2122 | 422.58
|  |  |  | 30 | 32.1733 | 6.7236 | 6.9671 | 100.5241 | 303.495 | 10079 | 30.448 | 6.7262 | 32.485 | 3.999 | 1.6955 | 87.9269 | 3.0977 | 424.49
|  |  |  | 11 | 32.111 | 8.6577 | 7.0413 | 106.1833 | 306.822 | 11661 | 11.613 | 8.6589 | 34.0838 | 4.1948 | 1.5202 | 86.6686 | 22.959 | 424.76
|  |  |  | 11 | 32.1107 | 8.6553 | 7.0379 | 106.1252 | 306.672 | 11740 | 11.628 | 8.6565 | 34.0836 | 4.1928 | 1.4705 | 86.6777 | 23.359 | 425.87
|  |  |  | 0 | 32.1117 | 8.6581 | 7.1489 | 107.8051 | 311.511 | 13051 | 3.334 | 8.6584 | 34.0798 | 4.2539 | 1.281 | 86.4239 | 73.803 | 429.48
|  |  |  | 0 | 32.1116 | 8.6586 | 7.154 | 107.8837 | 311.734 | 13208 | 3.683 | 8.6589 | 34.0801 | 4.2564 | 1.2758 | 86.4665 | 67.926 | 431.3
| 028 | 028 | NP14 | 1000 | 34.3382 | 2.9599 | 0.39469 | 5.2809 | 17.157 | 25892 | 999.474 | 3.0277 | 31.5839 | 0.7532 | 0.055942 | 91.0565 | 0.071023 | 1704.5 | 374 | 1 | 1135 | 0.39 |  |  |  | 0.39
|  |  |  | 800 | 34.25 | 3.4567 | 0.31185 | 4.2208 | 13.558 | 32720 | 801.578 | 3.5127 | 31.85 | 0.725 | 0.057786 | 90.9277 | 0.071188 | 1377.1 | 375 | 2 | 1094 | 0.303 |  |  |  | 0.303
|  |  |  | 600 | 34.1571 | 3.7935 | 0.3732 | 5.0885 | 16.227 | 38323 | 597.123 | 3.8358 | 31.966 | 0.7558 | 0.056914 | 90.9578 | 0.072945 | 1735.3 | 376 | 3 | 1293 | 0.377 |  |  |  | 0.377
|  |  |  | 400 | 34.0342 | 4.2158 | 0.65737 | 9.0465 | 28.586 | 43829 | 397.758 | 4.2445 | 32.133 | 0.8879 | 0.056963 | 90.7179 | 0.068557 | 1787.7 | 377 | 4 | 1654 | 0.666 |  |  |  | 0.666
|  |  |  | 300 | 33.9419 | 4.4963 | 1.0257 | 14.2021 | 44.61 | 47946 | 301.097 | 4.5185 | 32.2494 | 1.0602 | 0.069747 | 90.8175 | 0.070389 | 1684.5 | 378 | 5 | 1127 | 1.06 |  |  |  | 1.06
|  |  |  | 250 | 33.8806 | 4.7562 | 1.4568 | 20.2883 | 63.362 | 50763 | 244.909 | 4.7746 | 32.3987 | 1.264 | 0.071092 | 90.6959 | 0.066591 | 1645.3 | 379 | 6 | 1284, 1143 | 1.5 | 1.502 |  |  | 1.5
|  |  |  | 200 | 33.8137 | 5.045 | 1.8979 | 26.6031 | 82.554 | 52984 | 200.063 | 5.0603 | 32.5726 | 1.4771 | 0.094092 | 90.5374 | 0.068214 | 1758.9 | 380 | 7 | 1089 | 1.966 |  |  |  | 1.966
|  |  |  | 175 | 33.7406 | 5.2059 | 2.2388 | 31.4861 | 97.387 | 54627 | 175.435 | 5.2196 | 32.6379 | 1.6422 | 0.067803 | 90.4723 | 0.067151 | 1812.7 | 381 | 8 | 1509 | 2.315 |  |  |  | 2.315
|  |  |  | 150 | 33.6253 | 5.261 | 2.6472 | 37.2498 | 115.165 | 56084 | 149.785 | 5.2726 | 32.5744 | 1.838 | 0.087937 | 90.3928 | 0.067699 | 1897.7 | 382 | 9 | 879 | 2.723 |  |  |  | 2.723
|  |  |  | 125 | 33.4723 | 5.327 | 3.1939 | 44.9672 | 138.968 | 57678 | 125.184 | 5.3367 | 32.4844 | 2.1003 | 0.11846 | 90.3957 | 0.10044 | 1959.2 | 383 | 10 | 1746 | 3.351 |  |  |  | 3.351
|  |  |  | 100 | 33.0422 | 4.9883 | 4.5842 | 63.8312 | 199.518 | 59076 | 100.281 | 4.9958 | 31.7987 | 2.7391 | 0.14448 | 90.253 | 0.29206 | 2084.1 | 384 | 11 | 1145 | 4.759 |  |  |  | 4.759
|  |  |  | 75 | 33.0437 | 4.9942 | 4.5821 | 63.8116 | 199.426 | 59223 | 100.234 | 5.0016 | 31.8057 | 2.7384 | 0.11864 | 90.2343 | 0.29066 | 2098.6 | 385 | 12 | 1748 | 5.815 |  |  |  | 5.815
|  |  |  | 50 | 33.0504 | 5.0116 | 4.5732 | 63.718 | 199.04 | 59271 | 100.525 | 5.0191 | 31.8252 | 2.7352 | 0.1267 | 90.3701 | 0.27486 | 2090.3 | 386 | 13 | 1658, 1156 | 6.434 | 6.431 |  |  | 6.434
|  |  |  | 25 | 32.7789 | 4.7541 | 5.6206 | 77.6794 | 244.668 | 60949 | 75.555 | 4.7595 | 31.3606 | 3.2132 | 0.17752 | 90.0993 | 1.0112 | 2025.3 | 387 | 14 | 663 | 6.928 |  |  |  | 6.928
|  |  |  | 10 | 32.6011 | 5.0154 | 6.1794 | 85.8413 | 269.036 | 62682 | 50.518 | 5.0191 | 31.4148 | 3.4975 | 0.34019 | 89.8417 | 3.4525 | 1856.2 | 388 | 15 | 1376 | 7.101 |  |  |  | 7.101
|  |  |  | 5 | 32.6019 | 5.0137 | 6.1863 | 85.9345 | 269.338 | 62809 | 50.692 | 5.0174 | 31.4148 | 3.5006 | 0.3476 | 89.9194 | 3.4623 | 1853.9 | 389 | 16 | 1137 | 6.955 |  |  |  | 6.955
|  |  |  | 100 | 32.6018 | 5.0137 | 6.1778 | 85.8165 | 268.968 | 62855 | 50.632 | 5.0175 | 31.4143 | 3.4966 | 0.33293 | 89.8691 | 3.4487 | 1853.1
|  |  |  | 100 | 32.4244 | 6.0066 | 6.6667 | 94.7277 | 290.328 | 64715 | 24.976 | 6.0086 | 32.0936 | 3.8067 | 0.82617 | 88.3296 | 15.035 | 1559.5
|  |  |  | 50 | 32.4255 | 6.0072 | 6.6719 | 94.8024 | 290.55 | 64885 | 25.243 | 6.0092 | 32.0937 | 3.8091 | 0.83359 | 88.5925 | 15.177 | 1560.2
|  |  |  | 50 | 32.4259 | 6.0072 | 6.6752 | 94.8504 | 290.697 | 64937 | 25.374 | 6.0093 | 32.0962 | 3.8107 | 0.84451 | 88.5417 | 14.601 | 1564.9
|  |  |  | 25 | 32.2013 | 9.3518 | 6.8166 | 104.4742 | 297.043 | 66590 | 10.9 | 9.3529 | 34.7776 | 4.1355 | 0.62316 | 87.1201 | 64.511 | 1555.1
|  |  |  | 25 | 32.1783 | 9.8191 | 6.6843 | 103.501 | 291.301 | 67779 | 5.786 | 9.8198 | 35.1623 | 4.1038 | 0.40527 | 87.0852 | 117.98 | 1489.5
|  |  |  | 5 | 32.1785 | 9.8181 | 6.6825 | 103.4711 | 291.223 | 67904 | 5.647 | 9.8188 | 35.1632 | 4.1029 | 0.405 | 87.0895 | 118.49 | 1468
|  |  |  | 5 | 32.1781 | 9.8148 | 6.6847 | 103.4979 | 291.32 | 67953 | 5.901 | 9.8155 | 35.1602 | 4.1037 | 0.41232 | 87.0972 | 116.59 | 1463.6
| 029 | 029 | NP14 | 100 | 33.2041 | 5.0674 | 4.0883 | 57.0962 | 177.914 | 8555 | 100.347 | 5.0749 | 32.0104 | 2.5128 | 0.12499 | 90.299 | 0.11212 | 437.83
|  |  |  | 100 | 33.2073 | 5.069 | 4.0842 | 57.043 | 177.737 | 8604 | 100.61 | 5.0765 | 32.0163 | 2.511 | 0.20818 | 90.3051 | 0.10225 | 438.74
|  |  |  | 100 | 33.2322 | 5.0845 | 4.0717 | 56.8991 | 177.19 | 8653 | 101.175 | 5.0921 | 32.0259 | 2.506 | 0.1297 | 90.2946 | 0.097381 | 439.88
|  |  |  | 56 | 32.712 | 4.9527 | 5.9397 | 82.4486 | 258.58 | 10858 | 57.991 | 4.9569 | 31.4607 | 3.3801 | 0.3197 | 90.1108 | 0.71629 | 464.99
|  |  |  | 25 | 32.347 | 6.0544 | 6.8042 | 96.7418 | 296.334 | 13717 | 24.794 | 6.0564 | 32.0115 | 3.875 | 0.89184 | 88.0969 | 7.043 | 527.62
|  |  |  | 25 | 32.331 | 6.0061 | 6.8076 | 96.6686 | 296.482 | 13753 | 24.776 | 6.0081 | 31.9961 | 3.8726 | 0.94103 | 88.2087 | 7.1261 | 530.2
|  |  |  | 25 | 32.3401 | 6.0426 | 6.789 | 96.4943 | 295.673 | 13790 | 24.905 | 6.0446 | 32.0527 | 3.8665 | 0.92192 | 88.14 | 7.0465 | 531.2
|  |  |  | 25 | 32.3448 | 6.0713 | 6.7896 | 96.5719 | 295.7 | 13828 | 24.893 | 6.0733 | 32.0318 | 3.8691 | 0.89591 | 88.0752 | 7.0675 | 530.03
|  |  |  | 25 | 32.3411 | 6.0354 | 6.7923 | 96.5263 | 295.819 | 13867 | 24.777 | 6.0374 | 32.0046 | 3.8677 | 0.90686 | 88.119 | 7.1635 | 531.16
|  |  |  | 25 | 32.3331 | 6.0206 | 6.7923 | 96.4869 | 295.819 | 13905 | 24.804 | 6.0226 | 32.009 | 3.8664 | 0.92519 | 88.163 | 7.1451 | 533.85
|  |  |  | 25 | 32.3345 | 6.0252 | 6.7903 | 96.4693 | 295.73 | 13948 | 24.879 | 6.0272 | 32.0146 | 3.8657 | 0.89995 | 88.1241 | 7.124 | 535.4
|  |  |  | 25 | 32.3274 | 5.9709 | 6.8108 | 96.6316 | 296.624 | 14019 | 24.751 | 5.9729 | 31.9632 | 3.8714 | 0.92565 | 88.0695 | 7.3178 | 534.25
|  |  |  | 25 | 32.3244 | 5.9854 | 6.7928 | 96.4075 | 295.841 | 14060 | 24.869 | 5.9874 | 31.9927 | 3.8638 | 0.91176 | 88.2207 | 7.1936 | 536.2
|  |  |  | 25 | 32.3382 | 6.0256 | 6.7786 | 96.3069 | 295.221 | 14099 | 24.97 | 6.0276 | 32.0153 | 3.8602 | 0.88306 | 88.1731 | 7.1948 | 538.03
|  |  |  | 25 | 32.3281 | 5.9714 | 6.7973 | 96.4411 | 296.034 | 14140 | 24.842 | 5.9734 | 31.9618 | 3.8649 | 0.88054 | 88.1589 | 7.3349 | 538.03
|  |  |  | 21 | 32.3208 | 5.9462 | 6.7931 | 96.3192 | 295.853 | 14186 | 24.826 | 5.9482 | 31.9726 | 3.8609 | 1.1802 | 88.0909 | 7.4139 | 538.3
|  |  |  | 21 | 32.2495 | 6.3892 | 6.8138 | 97.5871 | 296.786 | 15449 | 21.014 | 6.3909 | 32.2827 | 3.9045 | 1.0969 | 87.281 | 11.148 | 565.64
|  |  |  | 21 | 32.2513 | 6.4691 | 6.8124 | 97.7528 | 296.728 | 15508 | 20.77 | 6.4709 | 32.3333 | 3.91 | 1.0862 | 87.3734 | 11.566 | 563.85
|  |  |  | 0 | 32.2549 | 6.5966 | 6.811 | 98.0286 | 296.67 | 15556 | 20.658 | 6.5983 | 32.3359 | 3.9192 | 1.129 | 87.3411 | 11.73 | 566.08
|  |  |  | 0 | 32.1757 | 9.95 | 6.8155 | 105.8377 | 297.027 | 16851 | 4.556 | 9.9505 | 35.2809 | 4.1833 | 0.74023 | 87.2646 | 69.578 | 589.11
|  |  |  | 0 | 32.1766 | 9.9477 | 6.8156 | 105.835 | 297.033 | 16892 | 4.456 | 9.9482 | 35.2845 | 4.1833 | 0.71798 | 87.3841 | 71.105 | 591.12
|  |  |  | ??????? | 32.1769 | 9.9459 | 6.8148 | 105.8177 | 296.996 | 16938 | 4.637 | 9.9464 | 35.2725 | 4.1826 | 0.83728 | 87.3768 | 68.012 | 590.22
| 030 | 030 | UN-8 | 1000 | 34.3431 | 2.9557 | 0.3596 | 4.8116 | 15.634 | 32672 | 1001.033 | 3.0236 | 31.5854 | 0.7391 | 0.055344 | 91.1234 | 0.069651 | 19.56 | 419 | 1 | 1135 | 0.351 |  |  |  | 0.351
|  |  |  | 800 | 34.2467 | 3.3834 | 0.3319 | 4.4845 | 14.432 | 37888 | 801.314 | 3.439 | 31.7825 | 0.733 | 0.057537 | 90.9138 | 0.068149 | 19.604 | 420 | 2 | 1094 | 0.325 |  |  |  | 0.325
|  |  |  | 600 | 34.1461 | 3.6913 | 0.3485 | 4.7404 | 15.156 | 44283 | 601.259 | 3.7334 | 31.8695 | 0.7448 | 0.057188 | 91.0823 | 0.067167 | 19.56 | 421 | 3 | 1293, 1654 | 0.348 | 0.351 |  |  | 0.348
|  |  |  | 400 | 34.0204 | 4.1901 | 0.6562 | 9.0246 | 28.538 | 49756 | 399.585 | 4.2188 | 32.0993 | 0.8871 | 0.058184 | 90.603 | 0.066198 | 19.604 | 422 | 4 | 1127 | 0.665 |  |  |  | 0.665
|  |  |  | 300 | 33.9642 | 4.393 | 0.8926 | 12.33 | 38.821 | 53109 | 302.111 | 4.4149 | 32.1806 | 0.9989 | 0.058782 | 90.2875 | 0.066102 | 19.604 | 423 | 5 | 1284 | 0.919 |  |  |  | 0.919
|  |  |  | 250 | 33.9252 | 4.5233 | 1.1139 | 15.4302 | 48.446 | 55671 | 251.377 | 4.5417 | 32.2359 | 1.1037 | 0.060876 | 90.4144 | 0.064916 | 19.604 | 424 | 6 | 1143 | 1.149 |  |  |  | 1.149
|  |  |  | 200 | 33.8844 | 4.7305 | 1.4172 | 19.7243 | 61.642 | 57782 | 201.044 | 4.7455 | 32.3573 | 1.2493 | 0.063791 | 90.457 | 0.065734 | 19.604 | 425 | 7 | 1089 | 1.465 |  |  |  | 1.465
|  |  |  | 175 | 33.8642 | 4.864 | 1.6158 | 22.5574 | 70.283 | 59582 | 176.049 | 4.8773 | 32.4444 | 1.3456 | 0.07493 | 90.6394 | 0.064297 | 19.604 | 426 | 8 | 1509 | 1.673 |  |  |  | 1.673
|  |  |  | 150 | 33.8081 | 5.2178 | 2.1382 | 30.0929 | 93.009 | 61381 | 150.679 | 5.2295 | 32.6948 | 1.5993 | 0.078493 | 90.6748 | 0.064006 | 19.604 | 427 | 9 | 879 | 2.224 |  |  |  | 2.224
|  |  |  | 125 | 33.6133 | 5.287 | 2.9624 | 41.7069 | 128.882 | 63179 | 126.277 | 5.2968 | 32.5718 | 1.9911 | 0.18014 | 90.7352 | 0.06329 | 19.604 | 428 | 10 | 1746 | 3.107 |  |  |  | 3.107
|  |  |  | 100 | 33.2543 | 4.9359 | 4.3286 | 60.2808 | 188.364 | 64988 | 100.83 | 4.9434 | 31.9433 | 2.6198 | 0.17002 | 90.5953 | 0.063316 | 19.56 | 429 | 11 | 1145 | 4.463 |  |  |  | 4.463
|  |  |  | 75 | 32.8625 | 4.7344 | 6.092 | 84.2023 | 265.176 | 66667 | 75.528 | 4.7398 | 31.4155 | 3.4331 | 0.32728 | 90.4043 | 0.06329 | 19.604 | 430 | 12 | 1748 | 6.314 |  |  |  | 6.314
|  |  |  | 50 | 32.7039 | 5.0541 | 6.5633 | 91.3233 | 285.734 | 68363 | 50.942 | 5.0578 | 31.5355 | 3.6825 | 0.48546 | 89.9418 | 0.062755 | 19.604 | 431 | 13 | 1658 | 6.871 |  |  |  | 6.871
|  |  |  | 25 | 32.4118 | 7.8606 | 6.7123 | 99.5932 | 292.387 | 70399 | 24.839 | 7.863 | 33.6818 | 3.9673 | 1.6439 | 88.0069 | 0.064378 | 19.604 | 432 | 14 | 1156, 663 | 7.026 | 7.017 |  |  | 7.026
|  |  |  | 10 | 32.1996 | 9.0014 | 6.7718 | 102.9738 | 295.076 | 71947 | 10.97 | 9.0025 | 34.4694 | 4.0855 | 1.5279 | 85.4514 | 0.068193 | 19.604 | 433 | 15 | 1376 | 7.064 |  |  |  | 7.064
|  |  |  | 5 | 32.2009 | 8.9999 | 6.7811 | 103.111 | 295.48 | 73275 | 5.833 | 9.0004 | 34.4639 | 4.0926 | 1.6371 | 85.2906 | 0.07226 | 19.604 | 434 | 16 | 1137 | 7.079 |  |  |  | 7.079
| 031 | 031 | UN-7 | 1000 | 34.3652 | 2.8682 | 0.38449 | 5.1335 | 16.714 | 26860 | 999.776 | 2.9353 | 31.5263 | 0.7488 | 0.057088 | 90.4148 | 0.072939 | 19.604 | 435 | 1 | 1023 | 0.382 |  |  |  | 0.382
|  |  |  | 800 | 34.2623 | 3.3293 | 0.32576 | 4.3954 | 14.162 | 32318 | 801.164 | 3.3845 | 31.7477 | 0.7303 | 0.057287 | 90.9545 | 0.069871 | 19.604 | 436 | 2 | 1022 | 0.324 |  |  |  | 0.324
|  |  |  | 600 | 34.1105 | 3.8275 | 0.43131 | 5.8841 | 18.754 | 37837 | 601.209 | 3.8702 | 31.9589 | 0.7807 | 0.058259 | 90.7385 | 0.070534 | 19.604 | 437 | 3 | 1666 | 0.434 |  |  |  | 0.434
|  |  |  | 400 | 33.9898 | 4.2018 | 0.71663 | 9.8558 | 31.165 | 43400 | 400.782 | 4.2307 | 32.0836 | 0.9138 | 0.06115 | 90.6494 | 0.071067 | 19.604 | 438 | 4 | 1155, 1136 | 0.735 | 0.736 |  |  | 0.735
|  |  |  | 300 | 33.9125 | 4.5367 | 1.1667 | 16.1664 | 50.742 | 46653 | 300.479 | 4.5588 | 32.2644 | 1.1244 | 0.062196 | 90.6102 | 0.066153 | 19.604 | 439 | 5 | 970 | 1.178 |  |  |  | 1.178
|  |  |  | 250 | 33.7663 | 4.8694 | 1.6539 | 23.0794 | 71.942 | 49191 | 250.565 | 4.8884 | 32.4757 | 1.3553 | 0.064065 | 90.5245 | 0.065301 | 19.782 | 440 | 6 | 1699 | 1.691 |  |  |  | 1.691
|  |  |  | 200 | 33.6569 | 5.1237 | 2.2054 | 30.9398 | 95.941 | 51447 | 200.379 | 5.1392 | 32.5736 | 1.6207 | 0.064414 | 90.3483 | 0.064349 | 21.071 | 441 | 7 | 1724 | 2.242 |  |  |  | 2.242
|  |  |  | 175 | 33.3893 | 5.1804 | 3.1833 | 44.64 | 138.515 | 53242 | 173.994 | 5.1938 | 32.3722 | 2.0797 | 0.078443 | 90.3636 | 0.06416 | 21.783 | 442 | 8 | 1738 | 3.323 |  |  |  | 3.323
|  |  |  | 150 | 33.1088 | 5.0597 | 4.0797 | 56.935 | 177.553 | 55040 | 151.018 | 5.071 | 32.0067 | 2.4944 | 0.07169 | 89.6674 | 0.064604 | 21.783 | 443 | 9 | 1735 | 4.263 |  |  |  | 4.263
|  |  |  | 125 | 32.8626 | 4.7831 | 4.8171 | 66.6644 | 209.678 | 57129 | 125.502 | 4.7922 | 31.5399 | 2.8269 | 0.0915 | 89.5712 | 0.063654 | 21.783 | 444 | 10 | 1152 | 5.027 |  |  |  | 5.027
|  |  |  | 100 | 32.7126 | 4.655 | 5.3314 | 73.4765 | 232.091 | 58887 | 99.353 | 4.662 | 31.2857 | 3.0638 | 0.095039 | 89.474 | 0.063452 | 21.783 | 445 | 11 | 1169 | 5.566 |  |  |  | 5.566
|  |  |  | 75 | 32.4965 | 4.4289 | 5.8978 | 80.7149 | 256.782 | 60686 | 74.764 | 4.434 | 30.895 | 3.3164 | 0.11784 | 89.2025 | 0.061947 | 21.783 | 446 | 12 | 1704 | 6.147 |  |  |  | 6.147
|  |  |  | 50 | 32.2291 | 4.4757 | 6.2311 | 85.2198 | 271.351 | 62362 | 50.951 | 4.4792 | 30.6849 | 3.4774 | 0.21856 | 89.3433 | 0.063186 | 21.827 | 447 | 13 | 1698, 1706 | 6.497 | 6.497 |  |  | 6.497
|  |  |  | 25 | 32.0185 | 5.1561 | 6.7434 | 93.6328 | 293.731 | 64038 | 25.649 | 5.158 | 31.072 | 3.7712 | 0.53171 | 88.5585 | 0.072823 | 23.872 | 448 | 14 | 1351 | 6.991 |  |  |  | 6.991
|  |  |  | 10 | 31.9328 | 7.2746 | 7.2287 | 105.4801 | 314.968 | 65861 | 10.715 | 7.2755 | 32.8102 | 4.1749 | 1.3147 | 83.8627 | 0.22436 | 23.961 | 449 | 15 | 1730 | 7.495 |  |  |  | 7.495
|  |  |  | 5 | 31.9321 | 7.831 | 7.2593 | 107.2978 | 316.327 | 67165 | 4.95 | 7.8315 | 33.2539 | 4.238 | 1.6109 | 84.195 | 0.50244 | 23.961 | 450 | 16 | 1743 | 7.515 |  |  |  | 7.515
|  |  |  | ????? | 31.9348 | 7.8339 | 7.2628 | 107.359 | 316.479 | 69566 | 6.418 | 7.8345 | 33.2581 | 4.2393 | 1.4323 | 84.3353 | 0.47966 | 23.961
|  |  |  | ????? | 31.9355 | 7.8332 | 7.2533 | 107.2165 | 316.063 | 69643 | 6.418 | 7.8338 | 33.2574 | 4.2345 | 1.5409 | 84.353 | 0.49435 | 23.961
| 032 | 032 | UN-6 | bot | 32.2633 | 4.4444 | 6.3038 | 86.1735 | 274.512 | 7007 | 86.918 | 4.4503 | 30.6557 | 3.4956 | 0.17779 | 89.5257 | 0.078183 | 115.14 | 453 | 1 | 1739 | 6.6 |  |  |  | 6.6
|  |  |  | 75 | 32.21 | 4.3652 | 6.418 | 87.5309 | 279.491 | 8314 | 75.439 | 4.3703 | 30.5393 | 3.546 | 0.20829 | 89.2733 | 0.075181 | 123.09 | 454 | 2 | 1725 | 6.643 |  |  |  | 6.643
|  |  |  | 50 | 32.0938 | 5.0404 | 6.6977 | 92.7874 | 291.717 | 9891 | 50.414 | 5.0441 | 30.9992 | 3.7323 | 0.36289 | 88.8812 | 0.09483 | 104.83 | 455 | 3 | 1731 | 6.935 |  |  |  | 6.935
|  |  |  | 40 | 32.1571 | 6.6848 | 6.8933 | 99.3599 | 300.281 | 11331 | 40.33 | 6.6883 | 32.4429 | 3.9551 | 0.79779 | 88.1912 | 0.1767 | 102.34 | 456 | 4 | 1193 | 7.144 |  |  |  | 7.144
|  |  |  | 30 | 32.0612 | 7.5769 | 7.2113 | 106.0594 | 314.193 | 12651 | 30.292 | 7.5797 | 33.1135 | 4.1843 | 1.0808 | 86.0126 | 0.43467 | 102.38 | 457 | 5 | 1750 | 7.51 |  |  |  | 7.51
|  |  |  | 20 | 31.9752 | 7.8423 | 7.1676 | 106.0026 | 312.32 | 14011 | 20.216 | 7.8442 | 33.3068 | 4.1866 | 0.97216 | 85.9125 | 1.3528 | 102.38 | 458 | 6 | 1744, 1740 | 7.463 | 7.459 |  |  | 7.463
|  |  |  | 10 | 31.9757 | 7.8508 | 7.1844 | 106.2698 | 313.053 | 15265 | 9.838 | 7.8517 | 33.3087 | 4.2006 | 1.0424 | 85.9911 | 4.8067 | 103.58 | 459 | 7 | 1726 | 7.452 |  |  |  | 7.452
|  |  |  | 5 | 31.976 | 7.8488 | 7.3111 | 108.1386 | 318.575 | 16463 | 4.72 | 7.8493 | 33.3048 | 4.2666 | 0.99225 | 85.7814 | 9.8752 | 100.2 | 460 | 8 | 1732 | 7.455 |  |  |  | 7.455
|  |  |  | surface | 31.9758 | 7.851 | 7.2909 | 107.844 | 317.693 | 17676 | 2.927 | 7.8513 | 33.306 | 4.2574 | 1.0133 | 85.665 | 13.879 | 98.022 | 461 | 9 | 1752 | 7.452 |  |  |  | 7.452
| 033 | 033 | UN-5 | bot | 32.2829 | 4.6489 | 6.2507 | 85.885 | 272.199 | 7539 | 75.357 | 4.6542 | 30.8878 | 3.49 | 0.71989 | 89.384 | 0.098353 | 499 | 462 | 1 | 1709 | 6.498 |  |  |  | 6.498
|  |  |  | 50 | 32.1801 | 4.5838 | 6.4678 | 88.6632 | 281.674 | 9567 | 51.402 | 4.5873 | 30.7283 | 3.5934 | 0.32028 | 88.8874 | 0.25872 | 427.34 | 463 | 2 | 1745 | 6.685 |  |  |  | 6.685
|  |  |  | 40 | 32.0974 | 4.9867 | 6.6643 | 92.2047 | 290.26 | 11257 | 39.843 | 4.9895 | 30.9905 | 3.7171 | 0.52443 | 87.9633 | 0.69756 | 468.33 | 464 | 3 | 1741 | 6.89 |  |  |  | 6.89
|  |  |  | 30 | 32.0885 | 5.7728 | 6.7855 | 95.6698 | 295.57 | 12810 | 29.974 | 5.7751 | 31.6429 | 3.837 | 1.5194 | 87.5716 | 2.0455 | 538.03 | 465 | 4 | 1727 | 7.017 |  |  |  | 7.017
|  |  |  | chl max | 32.062 | 7.7125 | 7.2443 | 106.8762 | 315.637 | 14529 | 15.42 | 7.714 | 33.274 | 4.2189 | 1.8858 | 85.246 | 9.9842 | 532.92 | 466 | 5 | 1733, 1753 | 7.495 | 7.498 |  |  | 7.495
|  |  |  | 10 | 32.082 | 8.1267 | 7.2984 | 108.7167 | 318.008 | 15816 | 10.759 | 8.1277 | 33.6412 | 4.2825 | 2.212 | 84.1572 | 20.054 | 529.32 | 467 | 6 | 1700 | 7.526 |  |  |  | 7.526
|  |  |  | 5 | 32.0831 | 8.1541 | 7.3707 | 109.8615 | 321.156 | 17370 | 5.854 | 8.1547 | 33.6639 | 4.3237 | 1.6117 | 83.3054 | 48.741 | 544.57 | 468 | 7 | 1696 | 7.514 |  |  |  | 7.514
|  |  |  | surface | 32.0839 | 8.1575 | 7.3497 | 109.5576 | 320.244 | 18830 | 3.07 | 8.1578 | 33.6659 | 4.3148 | 1.5635 | 83.2288 | 100.04 | 577.24 | 469 | 8 | 1742 | 7.513 |  |  |  | 7.513
| 034 | 034 | UN-4 | bot | 32.4352 | 4.4781 | 5.9452 | 81.4327 | 258.859 | 5958 | 104.611 | 4.4853 | 30.8415 | 3.3295 | 0.21206 | 88.1381 | 0.090203 | 956.48 | 470 | 1 | 1336 | 6.157 |  |  |  | 6.157
|  |  |  | 75 | 32.4 | 4.4364 | 6.0814 | 83.1897 | 264.796 | 9090 | 75.417 | 4.4416 | 30.76 | 3.3996 | 0.252 | 88.6629 | 0.079693 | 888.73 | 471 | 2 | 1135 | 6.267 |  |  |  | 6.267
|  |  |  | 50 | 32.279 | 4.4957 | 6.3917 | 87.4884 | 278.336 | 10785 | 50.742 | 4.4992 | 30.6961 | 3.5542 | 0.35945 | 88.896 | 0.27144 | 861.26 | 472 | 3 | 1094 | 6.62 |  |  |  | 6.62
|  |  |  | 40 | 32.1766 | 4.7531 | 6.6012 | 90.8628 | 287.489 | 12315 | 41.055 | 4.756 | 30.8348 | 3.6717 | 0.42015 | 88.6815 | 0.52637 | 791.96 | 473 | 4 | 1293, 1654 | 6.765 | 6.764 |  |  | 6.765
|  |  |  | 30 | 32.1567 | 6.1464 | 6.7726 | 96.384 | 295.003 | 13775 | 30.522 | 6.1488 | 31.9782 | 3.8601 | 0.95734 | 87.6928 | 1.316 | 725.05 | 474 | 5 | 1127 | 7.005 |  |  |  | 7.005
|  |  |  | chl max | 32.1083 | 8.2228 | 7.4589 | 111.3716 | 324.999 | 15810 | 15.089 | 8.2242 | 33.7083 | 4.37 | 3.0378 | 82.0701 | 14.763 | 726.56 | 475 | 6 | 1284 | 7.741 |  |  |  | 7.741
|  |  |  | 10 | 32.1101 | 8.2465 | 7.5405 | 112.6517 | 328.557 | 17235 | 9.888 | 8.2474 | 33.7257 | 4.4161 | 1.7619 | 82.9573 | 35.278 | 747.15 | 476 | 7 | 1143 | 7.73 |  |  |  | 7.73
|  |  |  | 5 | 32.1102 | 8.2648 | 7.5859 | 113.3761 | 330.535 | 18681 | 5.133 | 8.2653 | 33.7402 | 4.443 | 2.0035 | 82.7174 | 83.577 | 734.12 | 477 | 8 | 1089 | 7.729 |  |  |  | 7.729
|  |  |  | surface | 32.1103 | 8.262 | 7.5651 | 113.0576 | 329.629 | 20202 | 2.851 | 8.2622 | 33.7381 | 4.4334 | 1.6477 | 82.9405 | 134.2 | 727.58 | 478 | 9 | 1509 | 7.716 |  |  |  | 7.716
|  |  |  | surface | 32.1102 | 8.2669 | 7.565 | 113.068 | 329.623 | 20417 | 2.885 | 8.2672 | 33.7401 | 4.4337 | 1.6338 | 82.5345 | 132.83 | 727.23
|  |  |  | surface | 32.1103 | 8.2672 | 7.5659 | 113.0829 | 329.664 | 20514 | 2.848 | 8.2675 | 33.7401 | 4.4343 | 1.6336 | 83.1005 | 132.93 | 727.5
| 035 | 035 | UN-3 | bot-5 | 32.2967 | 4.4991 | 6.2899 | 86.1187 | 273.9 | 9746 | 94.982 | 4.5057 | 30.7346 | 3.4905 | 0.22975 | 88.9435 | 0.092165 | 827.74 | 480 | 1 | 879 | 6.524 |  |  |  | 6.524
|  |  |  | 75 | 32.294 | 4.5002 | 6.3041 | 86.3107 | 274.518 | 11479 | 74.584 | 4.5054 | 30.7239 | 3.5049 | 0.19947 | 89.1575 | 0.12263 | 833.79 | 481 | 2 | 1746 | 6.526 |  |  |  | 6.526
|  |  |  | 50 | 32.2242 | 5.0071 | 6.5474 | 90.7108 | 285.143 | 13595 | 52.764 | 5.0109 | 31.085 | 3.6612 | 0.33593 | 89.0042 | 0.45658 | 866.1 | 482 | 3 | 1145, 1748 | 6.784 | 6.790* |  | *b | 6.784 | copepod
|  |  |  | 40 | 32.1284 | 6.1168 | 6.8477 | 97.3686 | 298.28 | 14956 | 40.482 | 6.1201 | 31.933 | 3.8891 | 0.58112 | 88.6389 | 1.2941 | 898.96 | 483 | 4 | 1658 | 7.096 |  |  |  | 7.096
|  |  |  | 30 | 32.1126 | 6.4189 | 6.9809 | 99.9635 | 304.099 | 16408 | 30.686 | 6.4214 | 32.1671 | 3.9806 | 0.70095 | 88.3119 | 3.0643 | 926.07 | 484 | 5 | 1156 | 7.234 |  |  |  | 7.234
|  |  |  | 20 | 32.0747 | 6.8082 | 7.1041 | 102.6366 | 309.489 | 17726 | 20.244 | 6.81 | 32.4691 | 4.0751 | 0.99504 | 87.4476 | 8.6633 | 953.68 | 485 | 6 | 663 | 7.322 |  |  |  | 7.322
|  |  |  | 10 | 31.9216 | 8.1088 | 7.4019 | 110.0994 | 322.556 | 19220 | 9.839 | 8.1098 | 33.4428 | 4.3299 | 3.1298 | 78.9022 | 37.181 | 967.64 | 486 | 7 | 1376 | 7.622 |  |  | b | 7.622 | big floatie, anthropod?
|  |  |  | 5 | 31.9039 | 8.2434 | 7.5222 | 112.2204 | 327.811 | 20615 | 5.403 | 8.2439 | 33.5239 | 4.4037 | 3.0857 | 77.7516 | 101.44 | 1004.8 | 487 | 8 | 1137 | 7.643 |  |  |  | 7.643
|  |  |  | surface | 31.9017 | 8.2825 | 7.5088 | 112.118 | 327.228 | 21783 | 3.285 | 8.2828 | 33.5465 | 4.4012 | 2.2188 | 77.3035 | 154.35 | 1026.5 | 488 | 9 | 1023 | 7.667 |  |  |  | 7.667
| 036 | 036 | UN-2 | bot-5 | 32.0803 | 4.027 | 6.5608 | 88.661 | 285.729 | 7558 | 83.529 | 4.0324 | 30.1482 | 3.5815 | 0.13959 | 88.9358 | 0.098452 | 1199.2 | 489 | 1 | 1022 | 6.822 |  |  |  | 6.822
|  |  |  | 50 | 32.0565 | 4.0724 | 6.6517 | 89.9708 | 289.693 | 9599 | 50.363 | 4.0756 | 30.15 | 3.6392 | 0.17769 | 89.0301 | 0.46518 | 1285.5 | 490 | 2 | 1666 | 6.962 |  |  |  | 6.962
|  |  |  | 40 | 31.8934 | 5.5148 | 7.0846 | 99.1456 | 308.635 | 11021 | 39.307 | 5.5178 | 31.256 | 3.9511 | 0.75742 | 88.9765 | 1.1044 | 1347.4 | 491 | 3 | 1155, 1136 | 7.328 | 7.333 |  |  | 7.328
|  |  |  | 30 | 31.8759 | 6.3436 | 7.251 | 103.4858 | 315.917 | 12374 | 29.502 | 6.346 | 31.8956 | 4.1006 | 1.8737 | 86.918 | 3.2193 | 1382.9 | 492 | 4 | 970 | 7.51 |  |  |  | 7.51
|  |  |  | 20 | 31.8114 | 7.1154 | 7.5063 | 109.0418 | 327.089 | 13780 | 18.866 | 7.1171 | 32.5233 | 4.2921 | 3.6124 | 83.1732 | 14.04 | 1358 | 493 | 5 | 1699 | 7.789 |  |  |  | 7.789
|  |  |  | 10 | 31.7881 | 8.0428 | 7.7475 | 114.9662 | 337.647 | 15193 | 8.523 | 8.0436 | 33.2464 | 4.4958 | 1.3011 | 80.6954 | 78.091 | 1323.6 | 494 | 6 | 1724 | 7.878 |  |  |  | 7.878
|  |  |  | 5 | 31.7877 | 8.1093 | 7.7266 | 114.8301 | 336.741 | 16304 | 4.989 | 8.1098 | 33.2937 | 4.4929 | 0.75206 | 80.6408 | 173.85 | 1374.5 | 495 | 7 | 1738 | 7.864 |  |  |  | 7.864
|  |  |  | surface | 31.7862 | 8.1541 | 7.6888 | 114.384 | 335.096 | 17070 | 2.634 | 8.1543 | 33.3373 | 4.4788 | 0.55779 | 80.8322 | 312.21 | 1664.4 | 496 | 8 | 1735 | 7.858 |  |  |  | 7.858
|  |  |  | surface | 31.7856 | 8.1623 | 7.6871 | 114.3794 | 335.022 | 17106 | 2.745 | 8.1625 | 33.3376 | 4.4786 | 0.56994 | 80.6723 | 352.65 | 1593.7
|  |  |  | surface | 31.7865 | 8.1584 | 7.6873 | 114.373 | 335.03 | 17149 | 2.674 | 8.1586 | 33.3418 | 4.4784 | 0.52327 | 80.7459 | 357.64 | 1506.9
| 037 | 037 | UN-1 | bot-5 | 31.9252 | 4.7279 | 6.8176 | 93.6317 | 296.971 | 7842 | 62.908 | 4.7323 | 30.5898 | 3.7562 | 0.39952 | 88.862 | 1.0248 | 1518.8 | 499 | 1 | 1728 | 7.061 |  |  |  | 7.061
|  |  |  | 50 | 31.919 | 4.7463 | 6.8244 | 93.7612 | 297.269 | 10099 | 50.254 | 4.7498 | 30.5941 | 3.7659 | 0.3809 | 88.7495 | 2.0809 | 1253.2 | 500 | 2 | 1734 | 7.063 |  |  |  | 7.063
|  |  |  | 40 | 31.9164 | 4.7733 | 6.8367 | 93.9889 | 297.806 | 11446 | 40.303 | 4.7761 | 30.6095 | 3.7778 | 0.3351 | 88.7395 | 4.9206 | 1315.4 | 501 | 3 | 1751 | 7.073 |  |  |  | 7.073
|  |  |  | 30 | 31.9178 | 4.7631 | 6.8384 | 93.9881 | 297.879 | 12888 | 30.501 | 4.7652 | 30.5981 | 3.782 | 0.39082 | 88.8046 | 11.48 | 1318.6 | 502 | 4 | 1423 | 7.067 |  |  |  | 7.067
|  |  |  | 20 | 31.8951 | 5.0622 | 6.8605 | 94.9646 | 298.858 | 14222 | 20.494 | 5.0636 | 30.8262 | 3.8188 | 0.39528 | 88.8132 | 28.999 | 1369.2 | 503 | 5 | 1737 | 7.115 |  |  |  | 7.115
|  |  |  | 10 | 31.8893 | 5.2425 | 6.9146 | 96.1265 | 301.222 | 15670 | 10.038 | 5.2433 | 30.966 | 3.8624 | 0.23436 | 88.7874 | 95.371 | 1435 | 504 | 6 | 1723, 1729 | 7.149 | 7.145 |  |  | 7.149
|  |  |  | 5 | 31.8892 | 5.4388 | 7.0409 | 98.3458 | 306.73 | 17010 | 5.455 | 5.4392 | 31.1182 | 3.9395 | 0.17034 | 88.9205 | 161.91 | 1309.5 | 505 | 7 | 1698 | 7.149 |  |  |  | 7.149
|  |  |  | surface | 31.8909 | 5.5052 | 7.0298 | 98.3475 | 306.247 | 18290 | 2.65 | 5.5054 | 31.1772 | 3.9407 | 0.16122 | 88.9506 | 175.73 | 1412.8 | 506 | 8 | 1706 | 7.148 |  |  |  | 7.148
| 038 | 038 | BCL-1 | 1000 | 34.3114 | 2.8235 | 0.54834 | 7.3102 | 23.837 | 24360 | 1002.851 | 2.8904 | 31.4438 | 0.8135 | 0.058633 | 90.4024 | 0.07679 | 1446.3 | 507 | 1 | 1351, 1742 | 0.543 | 0.543 |  |  | 0.543
|  |  |  | 800 | 34.2007 | 3.1604 | 0.61589 | 8.2714 | 26.776 | 31213 | 802.468 | 3.2146 | 31.5489 | 0.8488 | 0.059704 | 90.7198 | 0.069752 | 1660.5 | 508 | 2 | 1743 | 0.62 |  |  |  | 0.62
|  |  |  | 600 | 34.0142 | 3.5348 | 1.0547 | 14.2753 | 45.863 | 38504 | 604.655 | 3.5762 | 31.6226 | 1.0441 | 0.057786 | 90.6499 | 0.068344 | 1912.5 | 509 | 3 | 1739 | 1.09 |  |  |  | 1.09
|  |  |  | 400 | 33.6931 | 3.6969 | 2.6973 | 36.5632 | 117.319 | 44161 | 403.498 | 3.7242 | 31.3889 | 1.7753 | 0.057836 | 90.8486 | 0.067193 | 2067.1 | 510 | 4 | 1725 | 2.8 |  |  |  | 2.8
|  |  |  | 300 | 33.4543 | 3.7263 | 4.0716 | 55.1356 | 177.129 | 47875 | 300.927 | 3.7463 | 31.166 | 2.3993 | 0.068949 | 90.6648 | 0.071424 | 2098.6 | 511 | 5 | 1731 | 4.233 |  |  |  | 4.233
|  |  |  | 250 | 33.3893 | 4.0358 | 4.1099 | 56.0525 | 178.808 | 50302 | 249.927 | 4.0529 | 31.3525 | 2.4415 | 0.099025 | 90.4867 | 0.065945 | 2074.6 | 512 | 6 | 1193 | 4.289 |  |  |  | 4.289
|  |  |  | 200 | 33.3212 | 4.1781 | 4.2946 | 58.7467 | 186.859 | 53006 | 200.123 | 4.192 | 31.393 | 2.5431 | 0.12761 | 90.4532 | 0.066254 | 1818.4 | 513 | 7 | 1750 | 4.477 |  |  |  | 4.477
|  |  |  | 175 | 33.2173 | 4.0822 | 4.7359 | 64.5822 | 206.073 | 55143 | 171.127 | 4.0938 | 31.2073 | 2.745 | 0.089507 | 90.5374 | 0.066517 | 1618.2 | 514 | 8 | 1744 | 4.95 |  |  |  | 4.95
|  |  |  | 150 | 33.1199 | 4.1602 | 4.9726 | 67.8934 | 216.389 | 56870 | 149.439 | 4.1704 | 31.1828 | 2.8617 | 0.10909 | 90.468 | 0.067605 | 1564.9 | 515 | 9 | 1740 | 5.116 |  |  |  | 5.116
|  |  |  | 125 | 32.9963 | 4.3723 | 5.145 | 70.5537 | 223.92 | 59170 | 121.187 | 4.3807 | 31.2433 | 2.9592 | 0.16003 | 90.2851 | 0.068198 | 1579.2 | 516 | 10 | 1726, 1732 | 5.349 | 5.342 |  |  | 5.349
|  |  |  | 100 | 32.8491 | 4.4275 | 5.5517 | 76.1561 | 241.646 | 61150 | 100.239 | 4.4345 | 31.1532 | 3.1539 | 0.17221 | 90.1395 | 0.069535 | 1606.1 | 517 | 11 | 1752 | 5.764 |  |  |  | 5.764
|  |  |  | 75 | 32.7133 | 4.5972 | 5.9499 | 81.8808 | 259.011 | 63048 | 74.091 | 4.6024 | 31.168 | 3.3556 | 0.1976 | 90.0725 | 0.18527 | 1675.7 | 518 | 12 | 1709 | 6.183 |  |  |  | 6.183
|  |  |  | 50 | 32.5883 | 5.4182 | 6.4843 | 90.9496 | 282.328 | 64817 | 49.361 | 5.422 | 31.7476 | 3.6698 | 1.511 | 87.2628 | 1.1849 | 1715.6 | 519 | 13 | 1745 | 6.721 |  |  |  | 6.721
|  |  |  | 25 | 32.576 | 5.5133 | 6.6106 | 92.9217 | 287.834 | 66831 | 23.867 | 5.5151 | 31.8062 | 3.7471 | 1.771 | 86.513 | 17.709 | 1699.2 | 520 | 14 | 1741 | 6.854 |  |  |  | 6.854
|  |  |  | 10 | 32.5373 | 5.6194 | 6.7359 | 94.8977 | 293.303 | 68899 | 9.698 | 5.6201 | 31.8433 | 3.82 | 0.87571 | 86.7624 | 99.686 | 1639.3 | 521 | 15 | 1727 | 6.963 |  |  |  | 6.963
|  |  |  | 5 | 32.5367 | 5.6424 | 6.7317 | 94.8886 | 293.118 | 70006 | 3.83 | 5.6427 | 31.8705 | 3.8222 | 0.48018 | 86.4675 | 243.2 | 1632.2 | 522 | 16 | 1733 | 6.967 |  |  |  | 6.967
| 039 | 039 | BCL-2 |  | 32.8552 | 5.7074 | 6.5434 | 92.5743 | 284.851 | 14242 | 23.642 | 5.7093 | 32.2183 | 3.7349 | 1.1208 | 87.6023 | 0.42936 | 37.031
[5 empty] | 32.8576 | 5.6725 | 6.5357 | 92.3897 | 284.514 | 14306 | 23.695 | 5.6744 | 32.1913 | 3.7287 | 0.95978 | 87.5061 | 0.43218 | 37.031
[5 empty] | 32.8504 | 5.7357 | 6.5164 | 92.2519 | 283.678 | 14354 | 23.52 | 5.7375 | 32.2731 | 3.724 | 1.0493 | 87.4922 | 0.4541 | 37.031
[5 empty] | 32.8577 | 5.7013 | 6.5492 | 92.6449 | 285.104 | 14397 | 23.744 | 5.7032 | 32.199 | 3.7373 | 1.0236 | 87.5528 | 0.43896 | 37.031
[5 empty] | 32.8596 | 5.6609 | 6.5679 | 92.8211 | 285.917 | 14438 | 23.803 | 5.6628 | 32.1611 | 3.7434 | 0.97112 | 87.4992 | 0.43213 | 37.031
[5 empty] | 32.8523 | 5.7307 | 6.5509 | 92.7313 | 285.182 | 14481 | 23.504 | 5.7326 | 32.2713 | 3.7403 | 1.0107 | 87.4937 | 0.45381 | 37.031
[5 empty] | 32.8567 | 5.7202 | 6.5644 | 92.9012 | 285.767 | 14525 | 23.634 | 5.722 | 32.2359 | 3.746 | 1.0444 | 87.5369 | 0.4431 | 37.031
[5 empty] | 32.8585 | 5.6885 | 6.5833 | 93.0992 | 286.587 | 14568 | 23.722 | 5.6904 | 32.1873 | 3.7528 | 0.98785 | 87.4385 | 0.43272 | 37.031
[5 empty] | 32.8562 | 5.7012 | 6.5534 | 92.7035 | 285.289 | 14609 | 23.587 | 5.7031 | 32.2243 | 3.7394 | 1.0307 | 87.4962 | 0.44852 | 37.031
[5 empty] | 32.8627 | 5.6629 | 6.5423 | 92.4652 | 284.801 | 14651 | 23.752 | 5.6647 | 32.1716 | 3.7313 | 1.0714 | 87.5762 | 0.44847 | 37.031
[5 empty] | 32.8571 | 5.6697 | 6.5511 | 92.6019 | 285.187 | 14694 | 23.679 | 5.6716 | 32.1798 | 3.736 | 0.94828 | 87.4871 | 0.44532 | 37.031
[5 empty] | 32.8497 | 5.7432 | 6.5363 | 92.5499 | 284.545 | 14735 | 23.356 | 5.7451 | 32.2853 | 3.7341 | 1.0689 | 87.5549 | 0.45509 | 37.031
[5 empty] | 32.8582 | 5.722 | 6.5468 | 92.6572 | 285 | 14777 | 23.48 | 5.7239 | 32.2508 | 3.7378 | 1.0359 | 87.5599 | 0.45795 | 37.031
[5 empty] | 32.8542 | 5.7224 | 6.5645 | 92.9063 | 285.773 | 14820 | 23.619 | 5.7242 | 32.2105 | 3.7462 | 1.0546 | 87.4765 | 0.44648 | 37.031
[5 empty] | 32.8547 | 5.719 | 6.551 | 92.7078 | 285.184 | 14865 | 23.526 | 5.7209 | 32.2307 | 3.7395 | 1.1117 | 87.5127 | 0.45243 | 37.031
[5 empty] | 32.8507 | 5.7284 | 6.5357 | 92.5098 | 284.52 | 14913 | 23.535 | 5.7302 | 32.2242 | 3.7328 | 1.0663 | 87.5137 | 0.44621 | 37.031
[5 empty] | 32.8525 | 5.7315 | 6.544 | 92.635 | 284.88 | 14962 | 23.553 | 5.7334 | 32.2221 | 3.737 | 1.1307 | 87.508 | 0.45171 | 37.031
[5 empty] | 32.8612 | 5.6818 | 6.5522 | 92.6467 | 285.235 | 15012 | 23.742 | 5.6836 | 32.1647 | 3.7374 | 1.0623 | 87.507 | 0.42136 | 37.031
[5 empty] | 32.8628 | 5.6444 | 6.5563 | 92.6222 | 285.41 | 15059 | 23.657 | 5.6462 | 32.1749 | 3.7367 | 0.98394 | 87.4864 | 0.44367 | 37.031
[5 empty] | 32.8479 | 5.7524 | 6.5417 | 92.6452 | 284.78 | 15106 | 23.381 | 5.7542 | 32.2761 | 3.7373 | 1.1018 | 87.5146 | 0.4598 | 37.031
[5 empty] | 32.8544 | 5.7269 | 6.5588 | 92.8357 | 285.524 | 15156 | 23.519 | 5.7288 | 32.2366 | 3.7438 | 1.0338 | 87.4965 | 0.45369 | 37.031
| 040 | 040 | BCL-2 | 1000 | 34.3507 | 2.7516 | 0.43783 | 5.8278 | 19.032 | 25452 | 1001.22 | 2.8179 | 31.4124 | 0.7694 | 0.055942 | 91.1282 | 0.063743 | 78.418 | 523 | 1 | 1135 | 0.432 |  |  |  | 0.432
|  |  |  | 800 | 34.2778 | 3.0021 | 0.45224 | 6.0526 | 19.66 | 32364 | 800.528 | 3.0552 | 31.4731 | 0.781 | 0.056864 | 91.0785 | 0.064025 | 84.685 | 524 | 2 | 1094 | 0.447 |  |  |  | 0.447
|  |  |  | 800 | 34.1633 | 3.3072 | 0.59207 | 7.976 | 25.742 | 38219 | 601.998 | 3.3475 | 31.5462 | 0.8465 | 0.056266 | 91.0311 | 0.065563 | 88.02 | 525 | 3 | 1293 | 0.601 |  |  |  | 0.601
|  |  |  | 400 | 33.9568 | 3.5985 | 1.2645 | 17.1289 | 54.987 | 45172 | 400.125 | 3.6254 | 31.5272 | 1.1488 | 0.056191 | 91.0033 | 0.065965 | 87.131 | 526 | 4 | 1654, 1127 | 1.304 | 1.315 |  | c | 1.304 | reps a bit off
|  |  |  | 300 | 33.7745 | 3.6725 | 2.2367 | 30.3131 | 97.282 | 48732 | 300.737 | 3.6925 | 31.3887 | 1.5874 | 0.057836 | 90.9564 | 0.061401 | 89.576 | 527 | 5 | 1284 | 2.32 |  |  |  | 2.32
|  |  |  | 250 | 33.6804 | 3.7025 | 2.7371 | 37.0958 | 119.054 | 51180 | 251.534 | 3.7192 | 31.3104 | 1.8173 | 0.057537 | 90.9061 | 0.063115 | 91.487 | 528 | 6 | 1143 | 2.868 |  |  |  | 2.868
|  |  |  | 200 | 33.5689 | 3.7149 | 3.2927 | 44.6017 | 143.231 | 53612 | 200.418 | 3.7281 | 31.2045 | 2.0747 | 0.053101 | 90.7878 | 0.061867 | 93.665 | 529 | 7 | 1089 | 3.398 |  |  |  | 3.398
|  |  |  | 175 | 33.4948 | 3.6239 | 3.8065 | 51.4183 | 165.591 | 55659 | 175.575 | 3.6352 | 31.0489 | 2.3065 | 0.05639 | 90.7079 | 0.064957 | 95.844 | 530 | 8 | 1509 | 3.962 |  |  |  | 3.962
|  |  |  | 150 | 33.4206 | 3.4742 | 4.3726 | 58.8135 | 190.224 | 57483 | 150.574 | 3.4837 | 30.8452 | 2.5593 | 0.056066 | 90.6585 | 0.061275 | 103.54 | 531 | 9 | 879 | 4.531 |  |  |  | 4.531
|  |  |  | 125 | 33.3209 | 3.2415 | 5.205 | 69.5542 | 226.448 | 59441 | 122.893 | 3.249 | 30.5513 | 2.927 | 0.056664 | 90.6729 | 0.063004 | 115.45 | 532 | 10 | 1746 | 5.439 |  |  |  | 5.439
|  |  |  | 100 | 33.269 | 3.2447 | 5.4271 | 72.5005 | 236.121 | 61057 | 100.342 | 3.2508 | 30.4988 | 3.0332 | 0.066358 | 90.6504 | 0.066233 | 119.53 | 533 | 11 | 1145 | 5.633 |  |  |  | 5.633
|  |  |  | 75 | 33.2273 | 3.4695 | 5.4432 | 73.1025 | 236.836 | 62611 | 75.272 | 3.4742 | 30.6427 | 3.0613 | 0.076973 | 90.6336 | 0.12876 | 119.8 | 534 | 12 | 1748, 1658 | 5.665 | 5.677 |  | c | 5.665 | reps a bit off
|  |  |  | 50 | 33.1434 | 3.6863 | 5.6267 | 75.9294 | 244.838 | 64337 | 52.638 | 3.6896 | 30.7454 | 3.1641 | 0.10872 | 90.3646 | 0.30958 | 121.98 | 535 | 13 | 1156 | 5.852 |  |  |  | 5.852
|  |  |  | 25 | 32.9487 | 4.6341 | 6.2325 | 85.9745 | 271.268 | 66132 | 25.296 | 4.6359 | 31.3803 | 3.5127 | 0.68254 | 88.5728 | 1.5749 | 121.98 | 536 | 14 | 663 | 6.521 |  |  | d | 6.521 | large air bubble in sample, profile reasonable
|  |  |  | 10 | 32.7052 | 7.5686 | 6.6814 | 98.6529 | 290.961 | 67660 | 11.264 | 7.5697 | 33.6964 | 3.9423 | 0.75652 | 87.9542 | 6.3654 | 128.52 | 537 | 15 | 1376 | 6.954 |  |  |  | 6.954
|  |  |  | 5 | 32.7048 | 7.5832 | 6.7014 | 98.9805 | 291.834 | 69177 | 5.036 | 7.5837 | 33.7034 | 3.9562 | 0.76898 | 87.6803 | 14.965 | 134.07 | 538 | 16 | 1137 | 6.931 |  |  |  | 6.931
| 041 | 041 | BCL-2 | 600 | 34.1445 | 3.3488 | 0.63874 | 8.6126 | 27.772 | 22824 | 600.799 | 3.3892 | 31.5661 | 0.8665 | 0.056963 | 91.0469 | 0.071483 | 522.56
|  |  |  | 600 | 34.1447 | 3.3478 | 0.63945 | 8.622 | 27.802 | 22984 | 600.687 | 3.3882 | 31.566 | 0.8668 | 0.057262 | 91.0507 | 0.065588 | 525.58
|  |  |  | 250 | 33.6698 | 3.7343 | 2.7478 | 37.2665 | 119.518 | 33072 | 250.4 | 3.751 | 31.3284 | 1.8232 | 0.05492 | 90.8065 | 0.068197 | 361.59
|  |  |  | 250 | 33.6689 | 3.7351 | 2.7526 | 37.333 | 119.729 | 36095 | 251.286 | 3.7519 | 31.3295 | 1.8252 | 0.05644 | 90.8539 | 0.064481 | 344.17
|  |  |  | 100 | 33.2347 | 3.4641 | 5.4449 | 73.1211 | 236.906 | 41046 | 99.747 | 3.4703 | 30.6583 | 3.0538 | 0.074905 | 90.649 | 0.081832 | 306.12
|  |  |  | 100 | 33.2359 | 3.4648 | 5.4415 | 73.0773 | 236.758 | 41149 | 99.96 | 3.471 | 30.6582 | 3.0523 | 0.079166 | 90.6648 | 0.092865 | 306.52
|  |  |  | 100 | 33.2366 | 3.4625 | 5.4236 | 72.8336 | 235.98 | 41226 | 99.918 | 3.4688 | 30.6579 | 3.0441 | 0.073709 | 90.6624 | 0.082201 | 307.14
|  |  |  | 100 | 33.2362 | 3.4657 | 5.4217 | 72.814 | 235.899 | 41355 | 99.867 | 3.4719 | 30.658 | 3.0435 | 0.080636 | 90.6255 | 0.087035 | 307.36
|  |  |  | 75 | 33.1693 | 3.7116 | 5.5459 | 74.9019 | 241.318 | 42843 | 75.435 | 3.7165 | 30.8007 | 3.1215 | 0.10122 | 90.4407 | 0.21445 | 337.36
|  |  |  | 75 | 33.167 | 3.6982 | 5.5449 | 74.8628 | 241.277 | 42949 | 74.918 | 3.703 | 30.794 | 3.1204 | 0.11046 | 90.4589 | 0.21022 | 340.03
|  |  |  | 53 | 33.1026 | 4.0911 | 5.7691 | 78.6131 | 251.052 | 44535 | 53.299 | 4.0947 | 31.0568 | 3.2539 | 0.20132 | 90.1328 | 0.71832 | 353.73
|  |  |  | 53 | 33.1024 | 4.0903 | 5.772 | 78.6518 | 251.181 | 44735 | 53.335 | 4.0939 | 31.0579 | 3.2552 | 0.20926 | 90.1012 | 0.71407 | 351.32
|  |  |  | 25 | 32.7415 | 7.2756 | 6.6546 | 97.618 | 289.775 | 46414 | 25.679 | 7.2779 | 33.4752 | 3.9014 | 0.90496 | 88.1773 | 4.5845 | 331.1
|  |  |  | 25 | 32.7497 | 7.1193 | 6.6697 | 97.4908 | 290.425 | 46484 | 25.605 | 7.1216 | 33.4001 | 3.8975 | 0.86559 | 88.1237 | 5.3705 | 329.5
|  |  |  | 15 | 32.711 | 7.7202 | 6.7074 | 99.3889 | 292.1 | 47954 | 15.074 | 7.7216 | 33.8358 | 3.9652 | 0.80574 | 88.0241 | 14.057 | 326.78
|  |  |  | 15 | 32.7111 | 7.7208 | 6.7112 | 99.4467 | 292.266 | 48036 | 14.834 | 7.7222 | 33.8358 | 3.9672 | 0.69764 | 88.0064 | 14.022 | 326.74
|  |  |  | 5 | 32.7116 | 7.7215 | 6.7093 | 99.4182 | 292.183 | 49553 | 4.554 | 7.7219 | 33.8312 | 3.9709 | 0.68224 | 87.9149 | 45.749 | 335.05
|  |  |  | 5 | 32.7117 | 7.7215 | 6.7118 | 99.4553 | 292.292 | 49614 | 4.75 | 7.7219 | 33.8312 | 3.9721 | 0.67197 | 87.8853 | 43.312 | 335.45
|  |  |  | 5 | 32.7115 | 7.7212 | 6.7103 | 99.4315 | 292.225 | 49700 | 4.403 | 7.7216 | 33.8302 | 3.9715 | 0.85839 | 87.9149 | 31.475 | 335.67
|  |  |  | 5 | 32.7116 | 7.7215 | 6.7076 | 99.3918 | 292.106 | 49854 | 4.283 | 7.7219 | 33.8299 | 3.9702 | 0.6552 | 87.8891 | 49.018 | 338.16
|  |  |  | 5 | 32.7116 | 7.7224 | 6.7164 | 99.5242 | 292.489 | 50056 | 4.522 | 7.7228 | 33.8307 | 3.9746 | 0.64995 | 87.7641 | 33.9 | 340.79
|  |  |  | 5 | 32.7116 | 7.7222 | 6.7143 | 99.4936 | 292.4 | 50197 | 4.539 | 7.7226 | 33.8307 | 3.9735 | 0.67295 | 87.879 | 35.377 | 342.17
|  |  |  | 5 | 32.7118 | 7.7219 | 6.7164 | 99.5239 | 292.49 | 50340 | 4.552 | 7.7224 | 33.8309 | 3.9746 | 0.65194 | 87.8695 | 42.449 | 344.17
|  |  |  | 5 | 32.7118 | 7.7216 | 6.7136 | 99.4821 | 292.37 | 50471 | 4.394 | 7.7221 | 33.8303 | 3.9732 | 0.64421 | 87.8584 | 45.234 | 346.34
| 042 | 042 | BS-0 | 3000 | 34.6533 | 1.3794 | 1.9678 | 25.4431 | 85.508 | 62910 | 3000.632 | 1.5984 | 31.4008 | 1.1748 | 0.050634 | 91.2399 | 0.055301 | 146.21 | 563 | 1 | 1135 | 2.013 |  |  |  | 2.013
|  |  |  | 2500 | 34.6281 | 1.5194 | 1.5956 | 20.6787 | 69.34 | 75131 | 2499.156 | 1.6938 | 31.2728 | 1.0993 | 0.052553 | 91.2092 | 0.055301 | 119.8 | 564 | 2 | 1094 | 1.641 |  |  |  | 1.641
|  |  |  | 2000 | 34.589 | 1.717 | 1.211 | 15.7538 | 52.629 | 86911 | 1997.619 | 1.8508 | 31.1797 | 1.0062 | 0.054148 | 91.1416 | 0.055301 | 135.05 | 565 | 3 | 1293 | 1.245 |  |  |  | 1.245
|  |  |  | 1500 | 34.5103 | 2.0826 | 0.75419 | 9.8886 | 32.778 | 96646 | 1499.97 | 2.1808 | 31.1993 | 0.8713 | 0.054272 | 90.9755 | 0.055301 | 119.8 | 566 | 4 | 1654 | 0.77 |  |  |  | 0.77
|  |  |  | 1250 | 34.4453 | 2.3725 | 0.52353 | 6.9092 | 22.755 | 101848 | 1250.496 | 2.4544 | 31.2799 | 0.7949 | 0.055319 | 91.1833 | 0.055301 | 127.99 | 567 | 5 | 1127 | 0.528 |  |  |  | 0.528
|  |  |  | 1000 | 34.363 | 2.702 | 0.42049 | 5.5904 | 18.278 | 107599 | 999.648 | 2.7678 | 31.3784 | 0.7624 | 0.056664 | 91.1503 | 0.055301 | 113.05 | 568 | 6 | 1284, 1143 | 0.424 | 0.422 |  |  | 0.424
|  |  |  | 800 | 34.2819 | 2.9904 | 0.40983 | 5.4835 | 17.816 | 113055 | 801.053 | 3.0435 | 31.4664 | 0.7636 | 0.056465 | 91.1144 | 0.055301 | 98.022 | 569 | 7 | 1089 | 0.412 |  |  |  | 0.412
|  |  |  | 600 | 34.1467 | 3.3513 | 0.57708 | 7.7817 | 25.091 | 117731 | 601.135 | 3.3917 | 31.5709 | 0.8403 | 0.057686 | 91.0397 | 0.055301 | 87.531 | 570 | 8 | 1509 | 0.592 |  |  |  | 0.592
|  |  |  | 400 | 33.9087 | 3.706 | 1.3371 | 18.1548 | 58.146 | 122757 | 400.411 | 3.7332 | 31.581 | 1.1818 | 0.057412 | 90.8276 | 0.055301 | 102.38 | 571 | 9 | 879 | 1.387 |  |  |  | 1.387
|  |  |  | 300 | 33.7344 | 3.851 | 2.216 | 30.1575 | 96.386 | 125786 | 299.207 | 3.8713 | 31.5095 | 1.5824 | 0.05487 | 90.6609 | 0.055301 | 130.7 | 572 | 10 | 1746 | 2.304 |  |  |  | 2.304
|  |  |  | 250 | 33.6459 | 3.8887 | 2.6584 | 36.1872 | 115.636 | 128136 | 249.55 | 3.9056 | 31.4423 | 1.7876 | 0.055817 | 90.5877 | 0.055301 | 143.68 | 573 | 11 | 1145 | 2.767 |  |  |  | 2.767
|  |  |  | 200 | 33.5063 | 3.7307 | 3.6399 | 49.3038 | 158.342 | 131147 | 200.02 | 3.7439 | 31.1629 | 2.2306 | 0.053226 | 90.3339 | 0.055301 | 148.12 | 574 | 12 | 1748 | 3.774 |  |  |  | 3.774
|  |  |  | 175 | 33.3867 | 3.2795 | 4.8183 | 64.4819 | 209.615 | 133201 | 175.132 | 3.2903 | 30.6629 | 2.7417 | 0.058135 | 90.422 | 0.15234 | 156.84 | 575 | 13 | 1658 | 5.016 |  |  |  | 5.016
|  |  |  | 150 | 33.3174 | 2.9299 | 5.6932 | 75.4868 | 247.682 | 135196 | 150.689 | 2.9388 | 30.2933 | 3.1169 | 0.054148 | 90.4867 | 0.076608 | 168.93 | 576 | 14 | 1156 | 5.886 |  |  |  | 5.886
|  |  |  | 125 | 33.2523 | 2.7122 | 6.2518 | 82.4 | 271.993 | 136781 | 125.063 | 2.7193 | 30.0436 | 3.3577 | 0.06399 | 90.5005 | 0.06658 | 165.06 | 577 | 15 | 663 | 6.499 |  |  |  | 6.499
|  |  |  | 100 | 33.2081 | 2.6066 | 6.4439 | 84.6776 | 280.36 | 139111 | 100.081 | 2.6122 | 29.9044 | 3.4439 | 0.080611 | 90.4335 | 0.058402 | 163.37 | 578 | 16 | 1376, 1137 | 6.714 | 6.694 |  | c | 6.714 | bad reps
|  |  |  | 75 | 33.1674 | 2.6042 | 6.6922 | 87.908 | 291.172 | 140693 | 75.178 | 2.6084 | 29.8577 | 3.5625 | 0.17483 | 90.2152 | 0.055301 | 152.48 | 579 | 17 | 1023 | 6.922 |  |  |  | 6.922
|  |  |  | 50 | 33.1183 | 3.5636 | 7.0756 | 95.175 | 307.889 | 142324 | 50.328 | 3.5667 | 30.6222 | 3.8163 | 0.90294 | 88.5479 | 0.055301 | 143.77 | 580 | 18 | 1022 | 7.318 |  |  |  | 7.318
|  |  |  | 25 | 32.9342 | 7.3994 | 6.7191 | 98.9717 | 292.545 | 144177 | 25.341 | 7.4018 | 33.7692 | 3.9471 | 1.0721 | 87.6712 | 0.083773 | 139.77 | 581 | 19 | 1666 | 7.109 |  |  |  | 7.109
|  |  |  | 10 | 32.9369 | 7.4634 | 6.7183 | 99.1057 | 292.514 | 145683 | 10.517 | 7.4644 | 33.8136 | 3.9583 | 0.98948 | 87.6233 | 0.055301 | 137.23 | 582 | 20 | 1155 | 7.055 |  |  |  | 7.055
|  |  |  | 5 | 32.9336 | 7.4602 | 6.7384 | 99.391 | 293.387 | 146868 | 4.867 | 7.4607 | 33.8116 | 3.9705 | 0.95796 | 87.563 | 0.27557 | 135.05 | 583 | 21 | 1136 | 7.057 |  |  |  | 7.057
| 043 | 043 | BS-1 | 1000 | 34.3413 | 2.7906 | 0.42951 | 5.7223 | 18.671 | 35275 | 999.834 | 2.8571 | 31.4377 | 0.7663 | 0.056091 | 91.1306 | 0.055301 | 19.604 | 584 | 1 | 1351 | 0.446 |  |  |  | 0.446
|  |  |  | 150 | 33.2446 | 3.3544 | 5.2545 | 70.3821 | 228.619 | 56650 | 150.803 | 3.3637 | 30.5964 | 2.9457 | 0.057337 | 90.4728 | 0.055301 | 19.604 | 585 | 2 | 1742 | 5.467 |  |  |  | 5.467
|  |  |  | 100 | 33.1396 | 3.5246 | 5.3265 | 71.5942 | 231.774 | 59161 | 100.829 | 3.5309 | 30.6298 | 3.002 | 0.061474 | 90.5374 | 0.055301 | 19.604 | 586 | 3 | 1743 | 5.527 |  |  |  | 5.527
|  |  |  | 80 | 33.0601 | 3.4418 | 5.3655 | 71.9288 | 233.482 | 60951 | 81.538 | 3.4468 | 30.4919 | 3.0197 | 0.066482 | 90.1539 | 0.055301 | 19.604 | 587 | 4 | 1739, 1725 | 5.536 | 5.543 |  |  | 5.536
|  |  |  | 60 | 33.0014 | 3.7818 | 5.4381 | 73.49 | 236.66 | 62564 | 60.52 | 3.7857 | 30.7121 | 3.0788 | 0.1359 | 90.2281 | 0.055301 | 19.604 | 588 | 5 | 1731 | 5.681 |  |  |  | 5.681
|  |  |  | 50 | 32.8887 | 3.7243 | 5.5443 | 74.761 | 241.301 | 63983 | 50.035 | 3.7275 | 30.5619 | 3.1254 | 0.17991 | 89.6928 | 0.055301 | 19.604 | 589 | 6 | 1193 | 5.826 |  |  |  | 5.826
|  |  |  | 40 | 32.8489 | 4.1267 | 5.851 | 79.664 | 254.67 | 65197 | 40.379 | 4.1294 | 30.8689 | 3.2941 | 0.33784 | 89.5089 | 0.055301 | 19.604 | 590 | 7 | 1750 | 6.111 |  |  |  | 6.111
|  |  |  | 30 | 32.6246 | 5.0095 | 6.4143 | 89.1036 | 279.262 | 66412 | 30.14 | 5.0117 | 31.4218 | 3.6161 | 0.69425 | 88.4522 | 0.055301 | 19.604 | 591 | 8 | 1744 | 6.697 |  |  |  | 6.697
|  |  |  | 20 | 32.2558 | 6.4166 | 7.0002 | 100.3264 | 304.906 | 68219 | 20.391 | 6.4183 | 32.2952 | 3.9976 | 1.8091 | 85.3552 | 0.055301 | 19.604 | 592 | 9 | 1740 | 7.26 |  |  |  | 7.26
|  |  |  | 15 | 32.1601 | 6.6973 | 7.0592 | 101.7775 | 307.507 | 69624 | 14.148 | 6.6985 | 32.441 | 4.049 | 2.6262 | 82.6944 | 0.055301 | 19.604 | 593 | 10 | 1726 | 7.326 |  |  |  | 7.326
|  |  |  | 10 | 32.1593 | 6.6982 | 7.0609 | 101.8027 | 307.581 | 70846 | 10.375 | 6.6991 | 32.4386 | 4.0517 | 2.7497 | 82.6858 | 0.055301 | 19.604 | 594 | 11 | 1732 | 7.328 |  |  | b | 7.328 | copepods/floaties
|  |  |  | 5 | 32.1587 | 6.6986 | 7.0825 | 102.1138 | 308.522 | 72251 | 4.983 | 6.6991 | 32.4359 | 4.0647 | 2.6009 | 82.7452 | 0.055301 | 19.604 | 595 | 12 | 1752 | 7.325 |  |  |  | 7.325
|  |  |  | 2 | 32.1588 | 6.6998 | 7.0787 | 102.0615 | 308.357 | 74520 | 3.422 | 6.7001 | 32.4357 | 4.0637 | 2.6484 | 82.7236 | 0.055428 | 19.604 | 596 | 13 | 1709, 1745 | 7.326 | 7.328 |  |  | 7.326
|  |  |  | 2 | 32.1589 | 6.6998 | 7.0935 | 102.2756 | 309.004 | 74621 | 3.629 | 6.7001 | 32.4359 | 4.0709 | 2.8649 | 82.7337 | 0.055301 | 19.604
|  |  |  | 2 | 32.1589 | 6.6999 | 7.0786 | 102.0612 | 308.355 | 74726 | 3.398 | 6.7002 | 32.4354 | 4.0637 | 2.8263 | 82.7322 | 0.055301 | 19.604
| 044 | 044
| 045 | 045 |  | 3 | 32.287 | 6.6626 | 6.9867 | 100.7316 | 304.319 | 45087 | 3.915 | 6.6629 | 32.5221 | 4.0184 | 1.7511 | 85.0239 | 0.055301 | 19.604
|  |  |  | 3 | 32.2874 | 6.6659 | 6.9925 | 100.8234 | 304.572 | 45184 | 3.527 | 6.6662 | 32.525 | 4.0217 | 1.7416 | 85.0545 | 0.49274 | 19.604
| 046 | 046
| 047 | 047
| 048 | 048 | BS-5 | bot-5 | 33.0915 | 3.3105 | 5.4373 | 72.6729 | 236.598 | 12279 | 126.223 | 3.3182 | 30.4176 | 3.0303 | 0.05948 | 89.5716 | 0.076923 | 637.16 | 601 | 1 | 1741, 1727 | 5.633 | 5.634 |  |  | 5.633 | **resorted target pressure to match
|  |  |  | bot-5 | 33.0919 | 3.3096 | 5.4329 | 72.6126 | 236.406 | 12514 | 126.452 | 3.3173 | 30.4172 | 3.0282 | 0.057138 | 89.5094 | 0.076169 | 632.63 | [8 empty] | CTD data raw pressure (col. O)
|  |  |  | 120 | 33.0869 | 3.3101 | 5.4429 | 72.7442 | 236.843 | 14081 | 119.023 | 3.3174 | 30.4106 | 3.0351 | 0.057512 | 89.7464 | 0.073044 | 629.56 | 602 | 2 | 1733 | 5.641 |  |  |  | 5.641
|  |  |  | 100 | 33.0263 | 3.3861 | 5.5033 | 73.6602 | 239.486 | 15731 | 101.136 | 3.3923 | 30.4145 | 3.0716 | 0.061673 | 90.1448 | 0.073282 | 654.86 | 603 | 3 | 1753 | 5.707 |  |  |  | 5.707
|  |  |  | 90 | 32.9876 | 3.4589 | 5.5867 | 74.8916 | 243.123 | 17122 | 90.936 | 3.4645 | 30.441 | 3.1163 | 0.06865 | 90.1965 | 0.074937 | 642.46
|  |  |  | 75 | 32.8787 | 3.3348 | 5.6665 | 75.6698 | 246.615 | 18909 | 76.017 | 3.3393 | 30.2379 | 3.1479 | 0.078493 | 90.1783 | 0.10614 | 594.09 | 604 | 4 | 1700 | 5.894 |  |  |  | 5.894
|  |  |  | 50 | 32.6336 | 3.3941 | 6.1548 | 82.1738 | 267.916 | 20783 | 50.433 | 3.3971 | 30.0702 | 3.3767 | 0.20792 | 89.6439 | 0.33581 | 577.06 | 605 | 5 | 1696 | 6.349 |  |  |  | 6.349
|  |  |  | 40 | 32.4911 | 3.6757 | 6.4476 | 86.6062 | 280.702 | 22468 | 39.504 | 3.6781 | 30.1769 | 3.5305 | 0.44836 | 88.8922 | 0.65543 | 535.85 | 606 | 6 | 1706 | 6.684 |  |  |  | 6.684
|  |  |  | 40 | 32.4943 | 3.6634 | 6.448 | 86.5867 | 280.717 | 22601 | 39.783 | 3.6659 | 30.1708 | 3.5297 | 0.49737 | 88.9904 | 0.69458 | 535.85
|  |  |  | 40 | 32.4943 | 3.6603 | 6.4471 | 86.5685 | 280.68 | 22738 | 39.631 | 3.6627 | 30.1707 | 3.5292 | 0.45513 | 88.9655 | 0.682 | 535.72
|  |  |  | 35 | 32.4612 | 3.9906 | 6.5525 | 88.6857 | 285.282 | 24636 | 33.492 | 3.9927 | 30.428 | 3.6027 | 0.75488 | 88.479 | 0.98591 | 544.61
|  |  |  | 35 | 32.4617 | 3.9942 | 6.5502 | 88.6635 | 285.184 | 24709 | 33.342 | 3.9964 | 30.4324 | 3.602 | 0.7469 | 88.4057 | 1.0039 | 546.7
|  |  |  | 30 | 32.4698 | 4.0911 | 6.5629 | 89.053 | 285.739 | 26594 | 30.387 | 4.0931 | 30.4981 | 3.6162 | 0.72223 | 88.3842 | 0.7452 | 582.89 | 607 | 7 | 1728 | 6.713 |  |  |  | 6.713
|  |  |  | 20 | 32.2362 | 6.9449 | 6.9985 | 101.5409 | 304.854 | 27927 | 20.843 | 6.9467 | 32.7274 | 4.0373 | 2.0177 | 84.5508 | 1.5998 | 596.71 | 608 | 8 | 1734, 1751 | 7.242 | 7.26 |  | c | 7.242 | bad reps
|  |  |  | 20 | 32.237 | 6.944 | 7.014 | 101.7646 | 305.53 | 28068 | 20.862 | 6.9458 | 32.7365 | 4.0449 | 2.1225 | 84.5465 | 1.4118 | 598.93
|  |  |  | 20 | 32.2378 | 6.9505 | 7.0265 | 101.9618 | 306.074 | 28132 | 20.736 | 6.9523 | 32.7357 | 4.0516 | 2.1557 | 84.579 | 1.5648 | 600.76
|  |  |  | 15 | 32.1718 | 6.9817 | 6.9976 | 101.5725 | 304.834 | 29943 | 13.634 | 6.9829 | 32.6939 | 4.0416 | 1.8979 | 83.7579 | 2.1204 | 664.24 | 609 | 9 | 1423 | 7.225 |  |  |  | 7.225
|  |  |  | 10 | 32.1709 | 6.9727 | 7.0073 | 101.6903 | 305.256 | 31387 | 10.383 | 6.9736 | 32.6882 | 4.0472 | 1.8547 | 83.7454 | 1.656 | 681.8 | 610 | 10 | 1737 | 7.227 |  |  |  | 7.227
|  |  |  | 5 | 32.1719 | 7.0421 | 7.1053 | 103.2792 | 309.528 | 33716 | 4.451 | 7.0425 | 32.7699 | 4.1038 | 1.3847 | 83.7627 | 1.7281 | 778.49 | 611 | 11 | 1723 | 7.238 |  |  |  | 7.238
|  |  |  | 5 | 32.1703 | 7.0366 | 7.1049 | 103.2584 | 309.509 | 33869 | 4.462 | 7.037 | 32.7523 | 4.1031 | 1.4508 | 83.7938 | 1.6961 | 789.86
|  |  |  | 5 | 32.1718 | 7.0244 | 7.1013 | 103.1784 | 309.354 | 33968 | 4.316 | 7.0247 | 32.7401 | 4.1004 | 1.4327 | 83.7354 | 1.7357 | 796
|  |  |  | 2 | 32.1701 | 7.0036 | 7.1044 | 103.1709 | 309.485 | 35489 | 3.196 | 7.0039 | 32.7043 | 4.1007 | 1.1997 | 83.6741 | 1.856 | 933.19 | 612 | 12 | 1729 | 7.235 |  |  |  | 7.235
| 049 | 049 | BS-5 |  | 32.4995 | 3.7146 | 6.3917 | 85.9451 | 278.269 | 9054 | 50.273 | 3.7178 | 30.2317 | 3.5038 | 0.52062 | 88.9837 | 1.4188 | 646.59
[5 empty] | 32.5006 | 3.7108 | 6.3932 | 85.9569 | 278.331 | 9104 | 50.483 | 3.714 | 30.2237 | 3.5041 | 0.50248 | 88.9717 | 1.4222 | 648.01
[5 empty] | 32.2162 | 7.002 | 6.9435 | 100.8678 | 302.468 | 10861 | 32.891 | 7.0049 | 32.7661 | 4.0088 | 2.0346 | 84.0792 | 1.4494 | 686.37
[5 empty] | 32.2158 | 6.9983 | 6.9535 | 101.0037 | 302.902 | 10926 | 32.775 | 7.0012 | 32.7633 | 4.0134 | 2.2617 | 84.0801 | 1.479 | 688.91
[5 empty] | 32.1922 | 6.9908 | 7.0048 | 101.7111 | 305.14 | 12906 | 15.443 | 6.9921 | 32.7167 | 4.0455 | 1.9606 | 83.8273 | 1.1692 | 709.49
[5 empty] | 32.1915 | 6.9878 | 7.0033 | 101.683 | 305.078 | 13061 | 15.393 | 6.9892 | 32.72 | 4.0446 | 2.0119 | 83.8173 | 1.1806 | 710.11
[5 empty] | 32.1669 | 6.9399 | 7.0115 | 101.6703 | 305.439 | 14460 | 9.53 | 6.9407 | 32.6515 | 4.047 | 1.8271 | 83.8948 | 1.1659 | 721.41
[5 empty] | 32.1687 | 6.9419 | 7.0143 | 101.7169 | 305.56 | 14583 | 9.488 | 6.9427 | 32.6535 | 4.0485 | 1.8471 | 83.8249 | 1.128 | 719.41
[5 empty] | 32.172 | 6.9518 | 7.127 | 103.3765 | 310.47 | 16617 | 4.847 | 6.9522 | 32.6641 | 4.107 | 1.2658 | 83.7751 | 0.90632 | 731.1
[5 empty] | 32.1719 | 6.9518 | 7.1229 | 103.316 | 310.289 | 16670 | 4.296 | 6.9522 | 32.6635 | 4.1053 | 1.3395 | 83.8024 | 0.87154 | 731.72
[5 empty] | 32.1719 | 6.9533 | 7.1236 | 103.33 | 310.32 | 16726 | 4.769 | 6.9537 | 32.6642 | 4.1055 | 1.2812 | 83.8216 | 0.92314 | 730.97
[5 empty] | 32.172 | 6.9552 | 7.125 | 103.3558 | 310.383 | 16789 | 4.611 | 6.9556 | 32.6661 | 4.1064 | 1.1572 | 83.788 | 0.94059 | 729.94
| 050 | 050 | BCL-3 | bot-5 | 32.7818 | 2.0127 | 5.9218 | 76.4278 | 257.715 | 6935 | 109.478 | 2.0181 | 29.0656 | 3.1642 | 0.31883 | 86.5972 | 0.14275 | 828.32
|  |  |  | bot-5 | 32.7815 | 2.0125 | 5.9217 | 76.427 | 257.715 | 7013 | 109.618 | 2.0179 | 29.0644 | 3.1641 | 0.31499 | 86.5249 | 0.13546 | 830.9
|  |  |  | bot-5 | 32.7809 | 2.0122 | 5.9125 | 76.3072 | 257.313 | 7364 | 109.724 | 2.0177 | 29.0632 | 3.1601 | 0.29287 | 86.4129 | 0.12528 | 835.79 | 638 | 3 | 1135 | 6.092 |  |  |  | 6.092
|  |  |  | 100 | 32.7824 | 2.0144 | 5.9018 | 76.1732 | 256.847 | 9363 | 100.365 | 2.0194 | 29.0622 | 3.1588 | 0.29267 | 86.5407 | 0.086654 | 857.35 | 639 | 4 | 1094 | 6.093 |  |  |  | 6.093
|  |  |  | 80 | 32.7733 | 2.0033 | 5.9019 | 76.1461 | 256.853 | 12356 | 80.957 | 2.0073 | 29.037 | 3.1646 | 0.28975 | 87.0789 | 0.17548 | 867.44 | 640 | 5 | 1293 | 6.089 |  |  |  | 6.089
|  |  |  | 60 | 32.7358 | 2.0619 | 6.1462 | 79.3947 | 267.492 | 14925 | 60.496 | 2.0649 | 29.0461 | 3.2815 | 0.28843 | 88.7534 | 0.58454 | 870.02 | 641 | 6 | 1654, 1127 | 6.38 | 6.38 |  |  | 6.38
|  |  |  | 50 | 32.6274 | 2.2195 | 6.5621 | 85.0463 | 285.62 | 16964 | 49.035 | 2.222 | 29.075 | 3.4768 | 0.2887 | 89.4984 | 1.2481 | 867.08 | 642 | 7 | 1284 | 6.792 |  |  |  | 6.792
|  |  |  | 40 | 32.6215 | 3.6406 | 7.3243 | 98.3818 | 318.835 | 18709 | 40.815 | 3.6432 | 30.2672 | 3.929 | 1.6377 | 87.0967 | 2.4696 | 860.82 | 643 | 8 | 1143 | 7.583 |  |  |  | 7.583
|  |  |  | 30 | 32.5188 | 4.3709 | 7.74 | 105.7833 | 336.98 | 19961 | 30.494 | 4.373 | 30.7767 | 4.1829 | 4.0799 | 82.9276 | 8.0718 | 850.63 | 644 | 9 | 1089 | 8.009 |  |  |  | 8.009
|  |  |  | 20 | 32.2586 | 7.2515 | 7.0322 | 102.7775 | 306.332 | 21106 | 19.727 | 7.2532 | 33.0105 | 4.079 | 1.1341 | 86.58 | 23.89 | 836.81 | 645 | 10 | 1509 | 7.273 |  |  |  | 7.273
|  |  |  | 15 | 32.2635 | 7.3034 | 7.0032 | 102.4794 | 305.069 | 22381 | 14.971 | 7.3047 | 33.0595 | 4.071 | 0.98535 | 86.4785 | 35.873 | 821.61 | 646 | 11 | 879 | 7.235 |  |  |  | 7.235
|  |  |  | 10 | 32.2694 | 7.436 | 6.9711 | 102.3263 | 303.674 | 23515 | 10.23 | 7.4369 | 33.1776 | 4.0677 | 0.66335 | 86.9937 | 53.459 | 810.14 | 647 | 12 | 1746 | 7.211 |  |  |  | 7.211
|  |  |  | 5 | 32.2656 | 7.4213 | 7.0561 | 103.5347 | 307.375 | 24889 | 5.123 | 7.4218 | 33.1541 | 4.1111 | 0.6496 | 86.8247 | 92.403 | 795.03 | 648 | 13 | 1145 | 7.208 |  |  |  | 7.208
| 051 | 051 | BCL-4 | 70 | 32.5863 | 0.9191 | 5.858 | 73.391 | 254.962 | 9961 | 70.591 | 0.922 | 27.9886 | 3.077 | 0.088385 | 84.3181 | 0.078594 | 19.604 | 649 | 1 | 1748 | 6.062 |  |  |  | 6.062
|  |  |  | 60 | 32.5843 | 0.9122 | 5.8731 | 73.5649 | 255.619 | 11683 | 59.944 | 0.9146 | 27.9777 | 3.0865 | 0.090553 | 84.738 | 0.072227 | 19.604 | 650 | 2 | 1658 | 6.055 |  |  |  | 6.055
|  |  |  | 50 | 32.4318 | 1.918 | 7.3048 | 93.8191 | 317.992 | 13868 | 40.676 | 1.9199 | 28.6325 | 3.7783 | 1.6021 | 86.6758 | 0.068975 | 19.604 | 651 | 3 | 1156 | 7.647 |  |  |  | 7.647
|  |  |  | 40 | 32.4101 | 2.7452 | 7.9014 | 103.6298 | 343.989 | 15279 | 35.013 | 2.7471 | 29.3356 | 4.1117 | 2.4506 | 85.6071 | 0.065911 | 19.604 | 652 | 4 | 663 | 8.184 |  |  |  | 8.184
|  |  |  | 30 | 32.361 | 3.1045 | 8.7359 | 115.5777 | 380.342 | 16412 | 29.682 | 3.1062 | 29.6049 | 4.5189 | 2.077 | 85.7239 | 0.066104 | 19.604 | 653 | 5 | 1376, 1137 | 8.754 | 8.723 |  | b | 8.754 | 1-copepod/big floaties, 2-dancing bubbles, bad reps
|  |  |  | 25 | 32.248 | 3.6327 | 8.8135 | 118.0653 | 383.773 | 17586 | 24.981 | 3.6343 | 29.8989 | 4.6044 | 1.7577 | 86.0284 | 0.066951 | 19.604 | 654 | 6 | 1023 | 9.121 |  |  |  | 9.121
|  |  |  | 20 | 31.8707 | 3.5104 | 8.2192 | 109.4925 | 357.996 | 18684 | 19.536 | 3.5116 | 29.594 | 4.3164 | 1.0096 | 86.31 | 0.066886 | 19.604 | 655 | 7 | 1022 | 8.017 |  |  |  | 8.017
|  |  |  | 15 | 31.294 | 5.8307 | 7.1799 | 100.8419 | 312.943 | 20057 | 15.415 | 5.8319 | 30.9314 | 4.0188 | 0.47671 | 87.8565 | 0.070265 | 19.604 | 656 | 8 | 1666 | 7.443 |  |  | b | 7.443 | copepod
|  |  |  | 10 | 31.2948 | 5.8318 | 7.2291 | 101.535 | 315.087 | 21450 | 9.849 | 5.8325 | 30.9286 | 4.0449 | 0.45037 | 87.7296 | 0.066298 | 19.604 | 657 | 9 | 1155 | 7.432 |  |  |  | 7.432
|  |  |  | 5 | 31.2952 | 5.8349 | 7.3217 | 102.8425 | 319.123 | 22834 | 4.522 | 5.8352 | 30.9283 | 4.0918 | 0.48596 | 87.7809 | 0.069121 | 19.604 | 658 | 10 | 1136 | 7.431 |  |  |  | 7.431
| 052 | 052 | BCL-5 | bot-5 | 32.1607 | -1.4663 | 7.0578 | 82.6988 | 307.254 | 8074 | 82.575 | -1.4647 | 25.7583 | 3.3918 | 0.13513 | 82.7763 | 0.08669 | 940.21
|  |  |  | bot-5 | 32.1606 | -1.4674 | 7.0582 | 82.701 | 307.272 | 8146 | 82.521 | -1.4659 | 25.7578 | 3.3919 | 0.1379 | 82.8725 | 0.085917 | 939.63
|  |  |  | bot-5 | 32.1606 | -1.4675 | 7.0667 | 82.8006 | 307.643 | 8211 | 82.393 | -1.466 | 25.7578 | 3.3954 | 0.13042 | 82.9046 | 0.085401 | 940.43 | 661 | 3 | 1742 | 7.295 |  |  |  | 7.295
|  |  |  | 70 | 32.1471 | -1.3069 | 6.9679 | 81.9949 | 303.345 | 10549 | 69.373 | -1.3055 | 25.8699 | 3.3727 | 0.11562 | 82.997 | 0.089794 | 970.93 | 662 | 4 | 1743 | 7.188 |  |  |  | 7.188
|  |  |  | 60 | 32.136 | -1.218 | 6.9262 | 81.6966 | 301.535 | 12325 | 59.621 | -1.2168 | 25.9272 | 3.3661 | 0.092198 | 82.4847 | 0.19956 | 1005 | 663 | 5 | 1739 | 7.148 |  |  |  | 7.148
|  |  |  | 50 | 32.0983 | -1.1837 | 7.0158 | 82.8086 | 305.444 | 13760 | 48.157 | -1.1827 | 25.9218 | 3.4081 | 0.085171 | 84.4081 | 0.52397 | 949.94 | 664 | 6 | 1725 | 7.235 |  |  |  | 7.235
|  |  |  | 40 | 32.005 | -1.3457 | 7.2298 | 84.9014 | 314.78 | 15175 | 40.18 | -1.3449 | 25.7171 | 3.4826 | 0.23055 | 87.9595 | 0.94226 | 874.06 | 665 | 7 | 1731 | 7.493 |  |  |  | 7.493
|  |  |  | 30 | 31.8329 | -1.4965 | 7.7869 | 90.9579 | 339.084 | 16789 | 29.979 | -1.496 | 25.4743 | 3.6931 | 4.9411 | 78.1093 | 4.9154 | 806.72 | 666 | 8 | 1193, 1750 | 8.47 | 8.609 |  | d | 8.47 | very turbid, bad reps
|  |  |  | 25 | 31.7524 | -1.4187 | 9.2313 | 108.0001 | 402.007 | 18256 | 24.708 | -1.4183 | 25.4735 | 4.2758 | 0.59131 | 88.3104 | 9.3201 | 771.11 | 667 | 9 | 1744 | 9.4 |  |  |  | 9.4
|  |  |  | 20 | 31.6392 | -0.918 | 9.131 | 108.2131 | 397.678 | 19528 | 19.726 | -0.9176 | 25.753 | 4.2842 | 1.1445 | 86.965 | 17.522 | 781.15 | 668 | 10 | 1740 | 9.344 |  |  |  | 9.344
|  |  |  | 15 | 31.1121 | 4.5053 | 7.8196 | 106.2281 | 340.822 | 20728 | 13.028 | 4.5061 | 29.6834 | 4.2061 | 0.4059 | 88.4306 | 34.455 | 792.04 | 669 | 11 | 1726 | 8.175 |  |  |  | 8.175
|  |  |  | 10 | 31.0343 | 4.4424 | 7.8893 | 106.9543 | 343.88 | 21753 | 9.015 | 4.443 | 29.6422 | 4.2329 | 0.31086 | 88.1204 | 53.101 | 793.47 | 670 | 12 | 1732 | 8.055 |  |  |  | 8.055
|  |  |  | 5 | 31.0406 | 4.7797 | 7.8749 | 107.6488 | 343.264 | 23095 | 5.337 | 4.78 | 29.8295 | 4.2575 | 0.26354 | 88.2381 | 77.884 | 794.45 | 671 | 13 | 1752 | 8.006 |  |  |  | 8.006
| 053 | 053 | SLIP-1 | bot-5 | 32.4016 | -1.3837 | 5.7363 | 67.478 | 249.675 | 4060 | 73.868 | -1.3823 | 25.9956 | 2.8798 | 0.79861 | 79.1407 | 0.094644 | 355.06
|  |  |  | bot-5 | 32.4013 | -1.3845 | 5.746 | 67.5914 | 250.1 | 4130 | 73.556 | -1.3831 | 25.9942 | 2.8838 | 0.82684 | 78.7318 | 0.093607 | 355.06
|  |  |  | bot-5 | 32.401 | -1.3841 | 5.7496 | 67.6338 | 250.255 | 4208 | 73.838 | -1.3826 | 25.9942 | 2.8851 | 1.0441 | 78.6298 | 0.092879 | 355.06 | 674 | 3 | 1742 | 5.995 |  |  |  | 5.995
|  |  |  | 70 | 32.4003 | -1.3871 | 5.7155 | 67.227 | 248.771 | 6356 | 71.046 | -1.3857 | 25.9914 | 2.8722 | 0.76709 | 79.6104 | 0.078183 | 333.71 | 675 | 4 | 1743 | 5.953 |  |  |  | 5.953
|  |  |  | 60 | 32.3831 | -1.3623 | 6.0667 | 71.3978 | 264.063 | 8220 | 59.888 | -1.3611 | 25.9918 | 3.017 | 0.8614 | 85.0655 | 0.13635 | 348.52 | 676 | 5 | 1739, 1725 | 6.336 | 6.325 |  | d | 6.336 | turbid, reps a bit off
|  |  |  | 50 | 32.2894 | -1.4723 | 6.6154 | 77.5695 | 287.964 | 12382 | 49.646 | -1.4714 | 25.8348 | 3.2301 | 2.4089 | 86.2037 | 0.38297 | 371.2 | 677 | 6 | 1731 | 6.832 |  |  |  | 6.832
|  |  |  | 40 | 32.1464 | -1.6402 | 6.9338 | 80.8482 | 301.858 | 13680 | 37.786 | -1.6396 | 25.5922 | 3.3462 | 2.4176 | 85.2848 | 1.703 | 359.41 | 678 | 7 | 1193 | 7.224 |  |  |  | 7.224
|  |  |  | 30 | 32.0987 | -1.5643 | 7.9158 | 92.4612 | 344.625 | 14812 | 29.935 | -1.5638 | 25.6151 | 3.7444 | 1.1163 | 85.9614 | 4.5396 | 387.73 | 679 | 8 | 1750 | 8.402 |  |  |  | 8.402
|  |  |  | 25 | 32.0893 | -1.3614 | 9.891 | 116.1701 | 430.621 | 17395 | 25.371 | -1.3609 | 25.7671 | 4.5537 | 0.95116 | 86.0466 | 7.8634 | 370.04 | 680 | 9 | 1744 | 10.087 |  |  |  | 10.087
|  |  |  | 20 | 31.7705 | 2.8824 | 8.3476 | 109.3885 | 363.597 | 19015 | 19.33 | 2.8835 | 29.41 | 4.3146 | 0.37196 | 88.9592 | 11.428 | 322.42 | 681 | 10 | 1740, 1726 | 8.778 | 8.788 |  | b | 8.778 | stalled burette 1-added titres, 2-finished titrating
|  |  |  | 15 | 31.6789 | 5.3222 | 7.4093 | 103.0596 | 322.825 | 20007 | 14.723 | 5.3233 | 30.8438 | 4.0957 | 0.38848 | 89.1828 | 15.498 | 324.56 | 682 | 11 | 1732 | 7.732 |  |  |  | 7.732
|  |  |  | 10 | 31.6807 | 5.4703 | 7.3676 | 102.8465 | 321.014 | 21058 | 9.818 | 5.4711 | 30.9735 | 4.0904 | 0.2358 | 89.009 | 22.412 | 322.51 | 683 | 12 | 1752 | 7.622 |  |  |  | 7.622
|  |  |  | 5 | 31.6652 | 5.9243 | 7.4189 | 104.6841 | 323.272 | 22372 | 4.542 | 5.9247 | 31.35 | 4.1542 | 0.20999 | 89.394 | 38.693 | 333.27 | 684 | 13 | 1709 | 7.574 |  |  |  | 7.574
| 054 | 054 | SLIP-2 | bot-5 | 32.4381 | -1.3954 | 5.5863 | 65.7097 | 243.14 | 9561 | 76.59 | -1.3939 | 26.0175 | 2.8192 | 0.4034 | 80.3415 | 0.078836 | 43.521
|  |  |  | bot-5 | 32.438 | -1.3956 | 5.5791 | 65.6248 | 242.828 | 9653 | 76.242 | -1.3941 | 26.0179 | 2.8164 | 0.30473 | 80.2084 | 0.078246 | 43.565
|  |  |  | bot-5 | 32.4374 | -1.394 | 5.5655 | 65.4667 | 242.234 | 9739 | 76.422 | -1.3926 | 26.0177 | 2.811 | 0.36984 | 80.2592 | 0.077745 | 43.565 | 687 | 3 | 1745 | 5.853 |  |  |  | 5.853
|  |  |  | 70 | 32.4293 | -1.3796 | 5.6099 | 66.0119 | 244.171 | 11921 | 69.602 | -1.3782 | 26.0185 | 2.8315 | 0.32847 | 81.0464 | 0.071736 | 41.387 | 688 | 4 | 1741 | 5.842 |  |  |  | 5.842
|  |  |  | 60 | 32.4277 | -1.373 | 5.6371 | 66.3423 | 245.353 | 13199 | 60.786 | -1.3718 | 26.0186 | 2.8453 | 0.28776 | 82.1826 | 0.070601 | 39.209 | 689 | 5 | 1727 | 5.86 |  |  |  | 5.86
|  |  |  | 50 | 32.2779 | -1.553 | 6.8395 | 80.0131 | 297.72 | 14448 | 49.946 | -1.5521 | 25.7605 | 3.3132 | 0.5268 | 88.5192 | 0.084496 | 39.209 | 690 | 6 | 1733 | 6.914 |  |  |  | 6.914
|  |  |  | 40 | 32.216 | -1.6354 | 7.0738 | 82.5315 | 307.937 | 15621 | 39.93 | -1.6347 | 25.6484 | 3.4026 | 0.87616 | 88.3626 | 0.14179 | 38.587 | 691 | 7 | 1753 | 7.368 |  |  |  | 7.368
|  |  |  | 35 | 32.1565 | -1.5789 | 7.2502 | 84.6863 | 315.631 | 16794 | 34.561 | -1.5784 | 25.6449 | 3.4779 | 0.77484 | 86.9947 | 0.20588 | 37.031 | 692 | 8 | 1700 | 7.501 |  |  |  | 7.501
|  |  |  | 30 | 32.1114 | -1.5012 | 8.1934 | 95.8776 | 356.705 | 17860 | 29.2 | -1.5007 | 25.6748 | 3.8609 | 0.8784 | 87.2829 | 0.32602 | 37.031 | 693 | 9 | 1696 | 8.714 |  |  |  | 8.714
|  |  |  | 25 | 32.0193 | -0.3522 | 9.4456 | 113.9639 | 411.263 | 18984 | 24.541 | -0.3515 | 26.4954 | 4.4762 | 0.97563 | 85.7818 | 0.52326 | 37.031 | 694 | 10 | 1706, 1728 | 9.711 | 9.351* |  | *e | 9.711 | bad rep, likely drawn from Bot 11, used as 20m rep
|  |  |  | 20 | 31.9327 | 0.8329 | 9.0985 | 113.2205 | 396.2 | 20112 | 19.553 | 0.8336 | 27.4241 | 4.4503 | 1.2042 | 86.7045 | 0.86268 | 37.031 | 695 | 11 | 1734 | 9.369 |  |  |  | 9.369
|  |  |  | 15 | 31.6181 | 4.8657 | 7.5527 | 103.857 | 329.074 | 21324 | 14.306 | 4.8667 | 30.4193 | 4.1241 | 0.39575 | 88.8031 | 1.409 | 37.031 | 696 | 12 | 1751 | 8.08 |  |  |  | 8.08
|  |  |  | 10 | 31.6181 | 5.3371 | 7.5148 | 104.5216 | 327.437 | 22630 | 9.522 | 5.3378 | 30.7695 | 4.1477 | 0.2378 | 89.1603 | 2.1358 | 39.209 | 697 | 13 | 1423 | 7.748 |  |  |  | 7.748
|  |  |  | 5 | 31.6268 | 5.5022 | 7.5724 | 105.7488 | 329.955 | 23901 | 4.895 | 5.5025 | 30.9489 | 4.1913 | 0.28474 | 89.224 | 3.1576 | 37.12 | 698 | 14 | 1737 | 7.729 |  |  |  | 7.729
| 055 | 055 | SLIP-3 | bot-5 | 32.5416 | -1.6396 | 5.654 | 66.1076 | 246.064 | 7332 | 67.656 | -1.6384 | 25.8924 | 2.8357 | 0.38389 | 79.1368 | 0.068458 | 19.604
|  |  |  | bot-5 | 32.5413 | -1.642 | 5.6565 | 66.1328 | 246.174 | 7418 | 67.61 | -1.6408 | 25.8925 | 2.8366 | 0.38526 | 78.9554 | 0.068278 | 19.604
|  |  |  | bot-5 | 32.5406 | -1.642 | 5.6499 | 66.0559 | 245.89 | 7494 | 67.561 | -1.6409 | 25.8916 | 2.834 | 0.46607 | 79.1368 | 0.068114 | 19.604 | 699 | 3 | 1336 | 5.87 |  |  |  | 5.87
|  |  |  | 60 | 32.5295 | -1.6436 | 5.6578 | 66.1389 | 246.232 | 9067 | 58.179 | -1.6426 | 25.8757 | 2.8396 | 0.34784 | 81.8402 | 0.065439 | 19.604 | 700 | 4 | 1135, 1094 | 5.839 | 5.831 |  |  | 5.839
|  |  |  | 50 | 32.5222 | -1.643 | 5.6611 | 66.1755 | 246.378 | 10292 | 50.762 | -1.6421 | 25.868 | 2.8431 | 0.41686 | 83.7904 | 0.064767 | 19.604 | 701 | 5 | 1293 | 5.864 |  |  |  | 5.864
|  |  |  | 40 | 32.319 | -1.4725 | 7.6125 | 89.2789 | 331.361 | 11546 | 40.042 | -1.4718 | 25.8485 | 3.6317 | 1.531 | 87.5501 | 0.065338 | 19.604 | 702 | 6 | 1654 | 7.158 |  |  |  | 7.158
|  |  |  | 35 | 32.2445 | -1.5101 | 7.5516 | 88.4271 | 328.729 | 12743 | 33.837 | -1.5095 | 25.7671 | 3.6054 | 1.6644 | 86.3047 | 0.064715 | 19.604 | 703 | 7 | 1127 | 7.802 |  |  |  | 7.802
|  |  |  | 30 | 32.222 | -1.3121 | 8.8702 | 104.4164 | 386.136 | 13971 | 29.744 | -1.3115 | 25.907 | 4.151 | 2.3979 | 82.8433 | 0.063826 | 19.604 | 704 | 8 | 1284 | 8.771 |  |  |  | 8.771
|  |  |  | 25 | 32.1785 | -1.2116 | 9.046 | 106.748 | 393.808 | 15198 | 24.489 | -1.2111 | 25.9497 | 4.2327 | 2.2671 | 82.6001 | 0.06419 | 19.604 | 705 | 9 | 1143 | 9.65 |  |  |  | 9.65
|  |  |  | 20 | 32.105 | -0.7863 | 9.6243 | 114.8352 | 419.01 | 17257 | 20.92 | -0.7858 | 26.2266 | 4.509 | 1.504 | 84.2525 | 0.06453 | 19.604 | 706 | 10 | 1089, 1509 | 9.721 | 7.925* |  | *e | 9.721 | bad rep, likely drawn from Bot 11, used as 15m rep
|  |  |  | 15 | 31.8675 | 5.6692 | 7.3621 | 103.3894 | 320.736 | 18740 | 15.525 | 5.6704 | 31.3131 | 4.1056 | 0.3824 | 89.3677 | 0.06401 | 19.604 | 707 | 11 | 879 | 7.929 |  |  |  | 7.929
|  |  |  | 10 | 31.8737 | 5.7337 | 7.3539 | 103.4365 | 320.378 | 20457 | 10.261 | 5.7345 | 31.3654 | 4.1096 | 0.27722 | 89.1134 | 0.065941 | 19.604 | 708 | 12 | 1746 | 7.534 |  |  |  | 7.534
|  |  |  | 5 | 31.8588 | 5.9054 | 7.3411 | 103.6701 | 319.83 | 21802 | 5.777 | 5.9059 | 31.4933 | 4.1192 | 0.28579 | 89.0675 | 0.066385 | 19.604 | 709 | 13 | 1145 | 7.498 |  |  |  | 7.498
| 056 | 056 | BRS-5 | bot-5 | 32.2474 | -1.6377 | 5.9582 | 69.527 | 259.367 | 8227 | 61.099 | -1.6367 | 25.6779 | 2.9536 | 1.27 | 80.6279 | 0.084702 | 241.39
|  |  |  | bot-5 | 32.247 | -1.6365 | 5.9604 | 69.5546 | 259.462 | 8299 | 60.901 | -1.6355 | 25.6778 | 2.9546 | 1.2428 | 80.5781 | 0.083152 | 241.79
|  |  |  | 50 | 32.1797 | -1.6437 | 6.3072 | 73.5521 | 274.571 | 10139 | 50.389 | -1.6428 | 25.619 | 3.0938 | 1.2429 | 88.3847 | 0.073904 | 247.03
|  |  |  | 40 | 32.1691 | -1.6427 | 6.3762 | 74.3532 | 277.578 | 11794 | 40.142 | -1.642 | 25.6079 | 3.1244 | 1.9903 | 87.4165 | 0.08127 | 257.04
|  |  |  | 25 | 32.1499 | -1.5811 | 7.6372 | 89.1968 | 332.481 | 13424 | 24.985 | -1.5807 | 25.6404 | 3.6353 | 4.9356 | 58.6431 | 1.4757 | 263.3
|  |  |  | 10 | 31.8957 | 3.8053 | 7.9904 | 107.2452 | 348.033 | 15135 | 10.34 | 3.806 | 29.8753 | 4.2438 | 0.54464 | 88.0337 | 16.251 | 265.75
|  |  |  | 5 | 31.8624 | 4.856 | 7.6935 | 105.9373 | 335.145 | 16473 | 5.747 | 4.8564 | 30.7071 | 4.1989 | 0.41001 | 86.1778 | 27.089 | 263.57
| 057 | 057 | SLIP-4 | bot-5 | 32.7545 | -1.6923 | 5.5491 | 64.8823 | 241.457 | 5786 | 65.918 | -1.6911 | 26.003 | 2.7948 | 0.27226 | 73.8201 | 0.12752 | 553.28 | 710 | 1 | 1748 | 5.741 |  |  |  | 5.741 | **resorted target pressure to match
|  |  |  | 60 | 32.7389 | -1.688 | 5.5753 | 65.1899 | 242.603 | 7277 | 60.654 | -1.6869 | 25.9929 | 2.8068 | 0.282 | 79.5338 | 0.18925 | 577.28 | 711 | 2 | 1658 | 5.752 |  |  |  | 5.752 | CTD data raw pressure (col. O)
|  |  |  | 50 | 32.6699 | -1.6525 | 5.5507 | 64.9346 | 241.547 | 8997 | 50.224 | -1.6516 | 25.9671 | 2.8012 | 0.47467 | 82.8793 | 0.5344 | 603.96 | 712 | 3 | 1156, 663 | 5.749 | 5.759 |  |  | 5.749
|  |  |  | 40 | 32.4629 | -1.4517 | 6.7883 | 79.7368 | 295.45 | 10869 | 39.102 | -1.451 | 25.9724 | 3.3076 | 1.0306 | 87.5932 | 1.5949 | 642.41 | 713 | 4 | 1376 | 6.995 |  |  |  | 6.995
|  |  |  | 35 | 32.431 | -1.4488 | 7.3584 | 86.4217 | 320.273 | 12862 | 35.153 | -1.4482 | 25.9498 | 3.5364 | 0.83778 | 88.1668 | 2.5982 | 690.07 | 714 | 5 | 1137 | 7.782 |  |  | d | 7.782 | turbid
|  |  |  | 30 | 32.3878 | -1.3553 | 7.6434 | 89.9719 | 332.688 | 14531 | 30.193 | -1.3547 | 25.9916 | 3.6591 | 0.94901 | 87.5472 | 4.5336 | 736.78 | 715 | 6 | 1023 | 7.878 |  |  |  | 7.878
|  |  |  | 25 | 32.3053 | -1.3803 | 7.8001 | 91.7013 | 339.532 | 15598 | 24.288 | -1.3798 | 25.9076 | 3.7205 | 1.838 | 81.5036 | 9.8695 | 771.33 | 716 | 7 | 1022 | 8.117 |  |  |  | 8.117
|  |  |  | 20 | 32.2417 | -1.1603 | 8.8038 | 104.0796 | 383.241 | 17587 | 20.257 | -1.1598 | 26.0345 | 4.1437 | 1.4324 | 81.4615 | 18.126 | 831.12 | 717 | 8 | 1666 | 9.057 |  |  |  | 9.057
|  |  |  | 20 | 32.2377 | -1.1294 | 9.2459 | 109.3959 | 402.49 | 18058 | 20.069 | -1.129 | 26.0594 | 4.3249 | 1.5578 | 80.4675 | 19.633 | 843.84
|  |  |  | 20 | 32.2426 | -1.1477 | 9.1602 | 108.3316 | 398.758 | 18345 | 20.322 | -1.1472 | 26.0455 | 4.2886 | 1.4657 | 80.1706 | 18.763 | 851.48
|  |  |  | 20 | 32.2403 | -1.1297 | 9.2352 | 109.27 | 402.023 | 18426 | 20.123 | -1.1293 | 26.0582 | 4.3206 | 1.5969 | 80.1256 | 20 | 853.41
|  |  |  | 20 | 32.2438 | -1.1447 | 9.2025 | 108.8418 | 400.6 | 18467 | 20.253 | -1.1443 | 26.0445 | 4.306 | 1.6343 | 80.1466 | 19.542 | 854.24
|  |  |  | 20 | 32.2454 | -1.1552 | 9.1583 | 108.2881 | 398.672 | 18508 | 20.371 | -1.1548 | 26.0435 | 4.2871 | 1.6114 | 80.0997 | 18.844 | 855.78
|  |  |  | 20 | 32.2411 | -1.1369 | 9.1602 | 108.3616 | 398.757 | 18554 | 20.159 | -1.1365 | 26.0553 | 4.2897 | 1.5397 | 79.618 | 19.582 | 856.11
|  |  |  | 20 | 32.2413 | -1.1325 | 9.2122 | 108.9904 | 401.022 | 18594 | 20.062 | -1.1321 | 26.055 | 4.3111 | 1.6088 | 79.8873 | 20.337 | 857.95
|  |  |  | 20 | 32.2433 | -1.1438 | 9.2022 | 108.8401 | 400.585 | 18637 | 20.297 | -1.1434 | 26.045 | 4.3059 | 1.6139 | 80.2111 | 19.726 | 859.24
|  |  |  | 20 | 32.244 | -1.1461 | 9.3306 | 110.3518 | 406.173 | 18988 | 20.189 | -1.1457 | 26.043 | 4.3575 | 1.6422 | 80.2922 | 20.112 | 865.26
|  |  |  | 20 | 32.2447 | -1.1486 | 9.3481 | 110.5518 | 406.934 | 19056 | 20.229 | -1.1481 | 26.0441 | 4.3643 | 1.6749 | 80.2644 | 20.023 | 866.99
|  |  |  | 20 | 32.2445 | -1.1529 | 9.351 | 110.5736 | 407.063 | 19106 | 20.245 | -1.1524 | 26.0436 | 4.3651 | 1.6682 | 80.2544 | 19.831 | 868.69
|  |  |  | 20 | 32.2446 | -1.1452 | 9.3681 | 110.7986 | 407.805 | 19170 | 20.147 | -1.1447 | 26.0478 | 4.3727 | 1.7687 | 80.3214 | 20.134 | 870.73
|  |  |  | 20 | 32.2438 | -1.1436 | 9.3435 | 110.5124 | 406.736 | 19274 | 20.161 | -1.1431 | 26.0482 | 4.363 | 1.7366 | 80.3712 | 20.29 | 873.88
|  |  |  | 15 | 32.1905 | -0.4882 | 9.141 | 110.0152 | 397.949 | 20784 | 15.425 | -0.4878 | 26.449 | 4.3467 | 1.3912 | 80.2807 | 40.493 | 914.16 | 718 | 9 | 1155 | 9.357 |  |  |  | 9.357
|  |  |  | 10 | 31.9463 | 6.0208 | 7.2896 | 103.2876 | 317.573 | 21872 | 10.404 | 6.0217 | 31.6749 | 4.1037 | 0.23481 | 88.6638 | 62.221 | 947.15 | 719 | 10 | 1136 | 7.645 |  |  |  | 7.645
|  |  |  | 5 | 31.9638 | 6.0455 | 7.4011 | 104.9388 | 322.424 | 22982 | 4.553 | 6.0458 | 31.7054 | 4.1625 | 0.1975 | 88.4397 | 116.7 | 973.06 | 720 | 11 | 970 | 7.596 |  |  |  | 7.596
| 058 | 58 | SLIP-4 | 65 | 32.7553 | -1.6914 | 5.5707 | 65.1372 | 242.399 | 5737 | 65.212 | -1.6903 | 26.003 | 2.8036 | 0.32057 | 72.6623 | 0.21463 | 1577.5
|  |  |  | 59 | 32.7539 | -1.6922 | 5.5665 | 65.0865 | 242.218 | 7466 | 60.207 | -1.6911 | 25.9999 | 2.8034 | 0.28636 | 73.6645 | 0.45957 | 1591.9
|  |  |  | 59 | 32.7539 | -1.692 | 5.5686 | 65.1105 | 242.306 | 7626 | 60.256 | -1.691 | 26.0002 | 2.8042 | 0.31168 | 73.8546 | 0.45905 | 1590
|  |  |  | 40 | 32.6076 | -1.5892 | 5.941 | 69.5916 | 258.545 | 10225 | 39.204 | -1.5885 | 25.9674 | 2.9627 | 0.85054 | 86.8118 | 5.1886 | 1673.5
|  |  |  | 30 | 32.452 | -1.4108 | 7.0944 | 83.419 | 308.775 | 11540 | 29.548 | -1.4103 | 25.9928 | 3.4364 | 0.72231 | 87.9935 | 14.62 | 1749.2
|  |  |  | 30 | 32.4527 | -1.4086 | 7.0677 | 83.1107 | 307.613 | 11685 | 29.651 | -1.4081 | 25.9927 | 3.4258 | 0.6343 | 88.1017 | 14.678 | 1760.5
|  |  |  | 28 | 32.4402 | -1.4481 | 7.2958 | 85.6934 | 317.546 | 13254 | 28.599 | -1.4475 | 25.954 | 3.5143 | 0.72792 | 88.0945 | 17.149 | 1756.4
|  |  |  | 28 | 32.4406 | -1.4465 | 7.2746 | 85.4478 | 316.621 | 13301 | 28.78 | -1.446 | 25.9587 | 3.5058 | 0.69012 | 88.1075 | 16.236 | 1752.3
|  |  |  | 28 | 32.4397 | -1.4494 | 7.2774 | 85.4731 | 316.743 | 13347 | 28.402 | -1.4489 | 25.9519 | 3.5068 | 0.7163 | 88.1403 | 17.026 | 1753.5
|  |  |  | 28 | 32.4393 | -1.4495 | 7.29 | 85.6216 | 317.295 | 13397 | 28.345 | -1.449 | 25.951 | 3.5119 | 0.67656 | 88.0964 | 17.542 | 1750.2
|  |  |  | 15 | 32.0175 | 3.648 | 8.1356 | 108.8564 | 354.319 | 15275 | 15.483 | 3.6489 | 29.8456 | 4.2964 | 0.21656 | 87.7542 | 76.276 | 1669
|  |  |  | 15 | 31.9943 | 3.9502 | 8.1081 | 109.2886 | 353.136 | 15319 | 15.197 | 3.9512 | 30.0096 | 4.3105 | 0.21368 | 87.8584 | 79.446 | 1667.2
|  |  |  | 15 | 32.0068 | 3.9568 | 8.1278 | 109.5818 | 353.993 | 15363 | 15.234 | 3.9578 | 30.0271 | 4.3204 | 0.21079 | 87.7253 | 79.816 | 1667.1
|  |  |  | 15 | 32.014 | 3.8313 | 8.1221 | 109.1697 | 353.737 | 15411 | 15.471 | 3.8323 | 29.9133 | 4.3066 | 0.20716 | 87.8597 | 77.711 | 1667.2
|  |  |  | 5 | 31.9566 | 6.1229 | 7.3326 | 104.1548 | 319.446 | 17262 | 5.385 | 6.1233 | 31.769 | 4.1353 | 0.096161 | 88.5843 | 205.37 | 1691.3
|  |  |  | 5 | 31.9567 | 6.1266 | 7.3345 | 104.1911 | 319.529 | 17315 | 5.234 | 6.127 | 31.7672 | 4.1366 | 0.094541 | 88.5537 | 209.41 | 1692.3
|  |  |  | 4 | 31.957 | 6.1474 | 7.3657 | 104.6855 | 320.887 | 18438 | 4.736 | 6.1477 | 31.7856 | 4.1536 | 0.089769 | 88.3847 | 225.86 | 1724.2
|  |  |  | 4 | 31.9569 | 6.1535 | 7.3549 | 104.5473 | 320.417 | 18482 | 4.562 | 6.1539 | 31.7883 | 4.1489 | 0.084815 | 88.4839 | 231.41 | 1723.6
|  |  |  | 4 | 31.9572 | 6.1512 | 7.3574 | 104.5772 | 320.526 | 18527 | 4.471 | 6.1515 | 31.7891 | 4.15 | 0.08631 | 88.4927 | 237.44 | 1724.7
|  |  |  | 4 | 31.9568 | 6.1548 | 7.3655 | 104.7006 | 320.878 | 18587 | 4.772 | 6.1551 | 31.7884 | 4.1541 | 0.089008 | 88.4488 | 221.93 | 1715.9
|  |  |  | 3 | 31.9576 | 6.1478 | 7.3892 | 105.0209 | 321.911 | 19581 | 3.311 | 6.1481 | 31.7862 | 4.1657 | 0.08604 | 88.3686 | 272.55 | 1701
|  |  |  | 3 | 31.9579 | 6.1425 | 7.3916 | 105.0426 | 322.017 | 19628 | 3.499 | 6.1428 | 31.7838 | 4.1664 | 0.086487 | 88.4282 | 267.46 | 1703.4
| 059 | 59 | BCL-6 | bottom | 32.7158 | -1.3002 | 7.2043 | 85.1245 | 313.495 | 3972 | 49.406 | -1.2992 | 26.2823 | 3.4865 | 2.2698 | 79.6831 | 0.2041 | 248.32 | 743 | 1 | 1742 | 7.464 |  |  |  | 7.464
|  |  |  | bottom | 32.7153 | -1.2988 | 7.2023 | 85.1042 | 313.409 | 4018 | 49.759 | -1.2978 | 26.28 | 3.4856 | 2.6321 | 79.3587 | 0.19874 | 248.32 |  | 2
|  |  |  | bottom | 32.7152 | -1.2996 | 7.2051 | 85.1349 | 313.529 | 4071 | 49.742 | -1.2985 | 26.2794 | 3.4867 | 2.0212 | 79.63 | 0.19702 | 248.32 |  | 3
|  |  |  | 40 | 32.7075 | -1.2944 | 7.1932 | 85.0013 | 313.013 | 5706 | 40.57 | -1.2935 | 26.2758 | 3.4857 | 2.3042 | 80.1701 | 0.67014 | 237.34 | 744 | 4 | 1743 | 7.435 |  |  |  | 7.435
|  |  |  | 35 | 32.7043 | -1.2848 | 7.1997 | 85.0986 | 313.298 | 6916 | 35.591 | -1.2841 | 26.2778 | 3.4909 | 2.288 | 80.15 | 1.4126 | 228.59 | 745 | 5 | 1739 | 7.448 |  |  |  | 7.448
|  |  |  | 30 | 32.7019 | -1.2808 | 7.237 | 85.5479 | 314.924 | 8468 | 31.099 | -1.2801 | 26.2788 | 3.508 | 2.0903 | 80.3186 | 2.6097 | 214.76 | 746 | 6 | 1725, 1731 | 7.476 | 7.48 |  |  | 7.476
|  |  |  | 25 | 32.3033 | -0.2347 | 9.5331 | 115.6081 | 414.983 | 9531 | 25.089 | -0.234 | 26.815 | 4.5316 | 0.52939 | 88.2314 | 4.7237 | 211.29 | 747 | 7 | 1193 | 9.676 |  |  |  | 9.676
|  |  |  | 20 | 32.2784 | -0.0651 | 9.7663 | 118.9537 | 425.145 | 11269 | 21.006 | -0.0645 | 26.9381 | 4.6472 | 0.58272 | 88.7897 | 6.4304 | 204.76 | 748 | 8 | 1750 | 10.188 |  |  |  | 10.188
|  |  |  | 15 | 32.0516 | 5.2662 | 7.8247 | 108.9576 | 340.826 | 12847 | 15.198 | 5.2673 | 30.8183 | 4.2958 | 0.23007 | 89.1939 | 10.197 | 200.4 | 749 | 9 | 1744 | 8.155 |  |  |  | 8.155
|  |  |  | 10 | 32.0491 | 6.2763 | 7.2478 | 103.3889 | 315.735 | 14206 | 9.556 | 6.2771 | 31.9819 | 4.1069 | 0.30024 | 88.796 | 15.955 | 199.73 | 750 | 10 | 1740 | 7.548 |  |  |  | 7.548
|  |  |  | 5 | 32.0539 | 6.2797 | 7.4067 | 105.6658 | 322.654 | 15874 | 4.906 | 6.2801 | 31.9841 | 4.1865 | 0.27147 | 88.5301 | 27.295 | 198.06 | 751 | 11 | 1726 | 7.535 |  |  |  | 7.535
|  |  |  | 5 | 32.0536 | 6.28 | 7.4056 | 105.6513 | 322.609 | 15914 | 4.684 | 6.2804 | 31.9843 | 4.1861 | 0.29182 | 88.5097 | 28.678 | 198.22
|  |  |  | 5 | 32.0538 | 6.2796 | 7.4078 | 105.6819 | 322.704 | 15957 | 4.745 | 6.28 | 31.9846 | 4.1871 | 0.29103 | 88.5763 | 28.203 | 198.22
| 060 | 060 | UTBS-5 | bottom | 32.6444 | 1.2546 | 6.8546 | 86.6631 | 298.327 | 11872 | 41.909 | 1.2564 | 28.2968 | 3.5364 | 0.4954 | 80.7682 | 0.072206 | 2.1783e+01
|  |  |  | bottom | 32.6447 | 1.2536 | 6.8608 | 86.7396 | 298.599 | 12044 | 41.929 | 1.2553 | 28.2966 | 3.539 | 0.48912 | 80.7686 | 0.071052 | 2.1783e+01
|  |  |  | bottom | 32.6448 | 1.2538 | 6.8536 | 86.6498 | 298.288 | 12135 | 41.907 | 1.2555 | 28.296 | 3.536 | 0.53487 | 80.6786 | 0.071011 | 2.1783e+01 | 754 | 3 | 1732 | 7.108 |  |  |  | 7.108
|  |  |  | 35 | 32.6448 | 1.2517 | 6.8543 | 86.6529 | 298.316 | 13981 | 35.094 | 1.2532 | 28.2917 | 3.5388 | 0.5281 | 81.099 | 0.069679 | 2.1783e+01 | 755 | 4 | 1752 | 7.106 |  |  |  | 7.106
|  |  |  | 30 | 32.6448 | 1.2519 | 6.8593 | 86.7157 | 298.533 | 15238 | 30.246 | 1.2532 | 28.2895 | 3.5428 | 0.46216 | 81.2949 | 0.068478 | 2.1783e+01 | 756 | 5 | 1709 | 7.107 |  |  |  | 7.107
|  |  |  | 25 | 32.6448 | 1.2512 | 6.8722 | 86.8774 | 299.097 | 16540 | 24.66 | 1.2523 | 28.2866 | 3.5506 | 0.55072 | 81.3561 | 0.075496 | 2.1783e+01 | 757 | 6 | 1745, 1741 | 7.109 | 7.105 |  |  | 7.109
|  |  |  | 20 | 32.6448 | 1.2528 | 6.8779 | 86.9518 | 299.342 | 17932 | 20.507 | 1.2537 | 28.2858 | 3.5548 | 0.57457 | 81.3997 | 0.11173 | 2.1783e+01 | 758 | 7 | 1727 | 7.111 |  |  |  | 7.111
|  |  |  | 15 | 32.6447 | 1.2492 | 6.8908 | 87.107 | 299.906 | 19501 | 15.788 | 1.2499 | 28.2809 | 3.562 | 0.45977 | 81.6123 | 0.18102 | 2.1827e+01 | 759 | 8 | 1733 | 7.11 |  |  |  | 7.11
|  |  |  | 10 | 32.6424 | 2.0484 | 6.916 | 89.2448 | 301.018 | 21094 | 10.991 | 2.049 | 29.1148 | 3.635 | 0.45685 | 85.586 | 0.3362 | 2.2761e+01 | 760 | 9 | 1753 | 7.126 |  |  |  | 7.126
|  |  |  | 5 | 32.5797 | 3.7876 | 6.9687 | 93.9157 | 303.369 | 22550 | 5.334 | 3.7879 | 30.2341 | 3.7926 | 0.37878 | 88.2204 | 0.63347 | 2.3872e+01 | 761 | 10 | 1700 | 7.239 |  |  |  | 7.239
| 061 | 061 | UTBS-4 | bottom | 32.8927 | 1.5857 | 6.1315 | 78.3221 | 266.814 | 7938 | 44.529 | 1.5877 | 28.7653 | 3.2519 | 0.53423 | 72.0446 | 0.22374 | 1.6337e+02
|  |  |  | bottom | 32.8932 | 1.5849 | 6.1214 | 78.1914 | 266.373 | 8006 | 44.676 | 1.5869 | 28.7643 | 3.2474 | 0.50146 | 71.9743 | 0.21728 | 1.6337e+02
|  |  |  | bottom | 32.8928 | 1.5863 | 6.1211 | 78.1905 | 266.361 | 8064 | 44.585 | 1.5884 | 28.7663 | 3.2474 | 0.5189 | 72.3204 | 0.22113 | 1.6337e+02
|  |  |  | 35 | 32.8179 | 2.0805 | 6.51 | 84.1779 | 283.308 | 9918 | 34.847 | 2.0822 | 29.1148 | 3.4532 | 0.63806 | 82.4311 | 0.63805 | 1.5901e+02
|  |  |  | 30 | 32.8146 | 2.07 | 6.5219 | 84.307 | 283.827 | 11362 | 30.64 | 2.0715 | 29.1015 | 3.4592 | 0.66549 | 82.7078 | 1.0019 | 1.5675e+02
|  |  |  | 25 | 32.7975 | 2.1326 | 6.6649 | 86.2832 | 290.056 | 12577 | 24.999 | 2.1338 | 29.1372 | 3.5284 | 0.94186 | 85.3945 | 1.7682 | 1.5466e+02
|  |  |  | 20 | 32.7899 | 2.2879 | 6.6922 | 86.9765 | 291.25 | 14162 | 19.921 | 2.289 | 29.2593 | 3.5537 | 1.1081 | 85.8671 | 2.9121 | 1.5030e+02
|  |  |  | 15 | 32.773 | 2.5897 | 6.8303 | 89.4441 | 297.271 | 15681 | 14.514 | 2.5904 | 29.4468 | 3.6392 | 0.86068 | 86.9406 | 5.0628 | 1.4377e+02
|  |  |  | 10 | 32.7643 | 2.6745 | 6.9323 | 90.969 | 301.713 | 17142 | 9.428 | 2.675 | 29.5297 | 3.693 | 0.79225 | 86.6557 | 9.1195 | 1.3941e+02
|  |  |  | 5 | 32.6396 | 3.7937 | 7.1555 | 96.4874 | 311.49 | 18553 | 4.961 | 3.794 | 30.3886 | 3.8803 | 0.48529 | 87.2925 | 17.945 | 1.3505e+02
| 062 | 062 | UTBS-1 | b-5 | 32.7805 | 1.6856 | 6.3253 | 80.9446 | 275.273 | 6470 | 41.177 | 1.6875 | 28.7596 | 3.3418 | 0.67202 | 73.0411 | 0.2753 | 3.8773e+02
|  |  |  | b-5 | 32.7803 | 1.6857 | 6.3409 | 81.1443 | 275.951 | 6537 | 41.167 | 1.6876 | 28.7595 | 3.3486 | 0.76447 | 73.1181 | 0.27666 | 3.8773e+02
|  |  |  | b-5 | 32.7805 | 1.685 | 6.3452 | 81.197 | 276.135 | 6610 | 41.15 | 1.6869 | 28.7598 | 3.3504 | 0.71369 | 73.4078 | 0.27889 | 3.8773e+02 | 770 | 3 | 1696 | 6.66 |  |  |  | 6.66
|  |  |  | 40 | 32.7803 | 1.6858 | 6.3519 | 81.2843 | 276.427 | 7635 | 40.402 | 1.6877 | 28.7602 | 3.3536 | 0.68286 | 73.3973 | 0.29565 | 3.9427e+02 | 771 | 4 | 1706 | 6.636 |  |  |  | 6.636
|  |  |  | 35 | 32.7803 | 1.688 | 6.3595 | 81.3865 | 276.761 | 9357 | 35.591 | 1.6897 | 28.7594 | 3.3589 | 0.77404 | 73.6362 | 0.73448 | 4.1169e+02 | 772 | 5 | 1728 | 6.619 |  |  |  | 6.619
|  |  |  | 30 | 32.7815 | 1.7411 | 6.3857 | 81.834 | 277.902 | 10976 | 29.585 | 1.7424 | 28.8071 | 3.3763 | 0.77025 | 75.9527 | 2.4484 | 4.1605e+02 | 773 | 6 | 1734 | 6.605 |  |  |  | 6.605
|  |  |  | 25 | 32.771 | 2.0123 | 6.5114 | 84.022 | 283.379 | 12763 | 25.483 | 2.0135 | 29.0186 | 3.4516 | 1.2157 | 83.1742 | 4.1288 | 4.2014e+02 | 774 | 7 | 1751 | 6.817 |  |  |  | 6.817
|  |  |  | 20 | 32.7186 | 2.5275 | 7.0611 | 92.2875 | 307.326 | 13871 | 20.502 | 2.5286 | 29.401 | 3.7335 | 1.6796 | 84.6729 | 7.1785 | 4.2041e+02 | 775 | 8 | 1423 | 7.096 |  |  |  | 7.096
|  |  |  | 15 | 32.7135 | 2.6693 | 7.1703 | 94.0481 | 312.082 | 14984 | 15.336 | 2.6702 | 29.5071 | 3.7953 | 1.1075 | 86.1802 | 13.339 | 4.1899e+02 | 776 | 9 | 1737 | 7.39 |  |  |  | 7.39
|  |  |  | 10 | 32.7126 | 2.6911 | 7.1672 | 94.0589 | 311.949 | 15573 | 15.322 | 2.6919 | 29.5287 | 3.7956 | 1.2514 | 85.971 | 13.806 | 4.2249e+02 | 777 | 10 | 1723 | 7.397 |  |  |  | 7.397
|  |  |  | 5 | 32.7102 | 2.6989 | 7.1528 | 93.8866 | 311.322 | 15766 | 15.33 | 2.6997 | 29.5321 | 3.7897 | 1.2675 | 85.7205 | 13.867 | 4.2258e+02 | 778 | 11 | 1729, 1738 | 7.415 | 7.417 |  |  | 7.415
|  |  |  | 15 | 32.7121 | 2.6944 | 7.1509 | 93.8528 | 311.24 | 15905 | 15.29 | 2.6952 | 29.5303 | 3.7886 | 1.0448 | 85.843 | 13.981 | 4.2249e+02
|  |  |  | 15 | 32.7136 | 2.6879 | 7.1632 | 94.0006 | 311.777 | 15949 | 15.302 | 2.6888 | 29.53 | 3.7937 | 1.3958 | 85.8435 | 13.892 | 4.2163e+02
|  |  |  | 15 | 32.7143 | 2.6801 | 7.1386 | 93.6586 | 310.703 | 15988 | 15.348 | 2.681 | 29.5282 | 3.782 | 0.95072 | 85.9267 | 13.911 | 4.2198e+02
|  |  |  | 15 | 32.7135 | 2.6804 | 7.129 | 93.5328 | 310.286 | 16024 | 15.331 | 2.6812 | 29.5249 | 3.7778 | 1.3876 | 85.739 | 13.839 | 4.2149e+02
|  |  |  | 15 | 32.7142 | 2.676 | 7.1305 | 93.5425 | 310.35 | 16059 | 15.291 | 2.6769 | 29.5234 | 3.7781 | 1.2948 | 85.9306 | 13.986 | 4.2143e+02
|  |  | Caro chlM | 15 | 32.7144 | 2.6753 | 7.1472 | 93.7605 | 311.079 | 16098 | 15.308 | 2.6761 | 29.5206 | 3.7855 | 1.1497 | 86.098 | 14.007 | 4.2166e+02
|  |  |  | 15 | 32.7141 | 2.6743 | 7.1652 | 93.9945 | 311.864 | 16145 | 15.365 | 2.6751 | 29.5182 | 3.7935 | 1.3276 | 86.015 | 13.927 | 4.2208e+02
|  |  |  | 15 | 32.7136 | 2.6626 | 7.1583 | 93.8751 | 311.561 | 16406 | 15.357 | 2.6634 | 29.5088 | 3.7894 | 1.0593 | 86.1591 | 14.142 | 4.2476e+02
|  |  |  | 15 | 32.714 | 2.6621 | 7.1414 | 93.6522 | 310.824 | 16524 | 15.347 | 2.6629 | 29.507 | 3.7818 | 1.1468 | 86.1725 | 14.154 | 4.2476e+02
|  |  |  | 15 | 32.7148 | 2.6609 | 7.1335 | 93.5474 | 310.484 | 16707 | 15.347 | 2.6617 | 29.5066 | 3.7783 | 1.0067 | 86.2563 | 13.98 | 4.2476e+02
|  |  |  | 15 | 32.7145 | 2.6592 | 7.1428 | 93.6645 | 310.886 | 16841 | 15.345 | 2.66 | 29.5064 | 3.7823 | 1.0181 | 86.2501 | 13.884 | 4.2476e+02
|  |  |  | 15 | 32.7064 | 2.721 | 7.2058 | 94.6333 | 313.634 | 18054 | 10.594 | 2.7215 | 29.5482 | 3.8171 | 1.492 | 85.6224 | 25.564 | 4.3130e+02
|  |  |  | 15 | 32.583 | 3.2716 | 7.3078 | 97.2306 | 318.117 | 19528 | 5.51 | 3.2719 | 29.9169 | 3.9066 | 0.58294 | 86.0955 | 53.921 | 4.3721e+02
| 063 | 063 | UTBS-1 | bot-5 | 32.7803 | 1.7157 | 6.3566 | 81.4075 | 276.632 | 3694 | 43.906 | 1.7177 | 28.7865 | 3.3565 | 0.69577 | 72.9405 | 0.19198 | 4.2041e+02
|  |  |  | bot-5 | 32.7799 | 1.7162 | 6.3452 | 81.2623 | 276.136 | 3744 | 43.935 | 1.7182 | 28.7863 | 3.3515 | 0.85044 | 73.0324 | 0.18843 | 4.2041e+02
|  |  |  | 40 | 32.7791 | 1.7176 | 6.3069 | 80.7752 | 274.474 | 5374 | 40.965 | 1.7195 | 28.7846 | 3.3361 | 0.64613 | 73.1679 | 0.30788 | 4.1605e+02
|  |  |  | 40 | 32.779 | 1.7171 | 6.3062 | 80.7649 | 274.442 | 5452 | 40.973 | 1.719 | 28.7849 | 3.3358 | 0.58982 | 73.2225 | 0.30442 | 4.1605e+02
|  |  |  | 32 | 32.7796 | 1.7276 | 6.3513 | 81.363 | 276.402 | 6926 | 32.045 | 1.7291 | 28.7894 | 3.3594 | 0.65035 | 73.889 | 1.5854 | 4.1298e+02
|  |  |  | 32 | 32.7799 | 1.7302 | 6.364 | 81.532 | 276.957 | 7123 | 32.055 | 1.7317 | 28.7903 | 3.3651 | 0.63908 | 73.9192 | 1.5779 | 4.1298e+02
|  |  |  | 30 | 32.7808 | 1.7517 | 6.3421 | 81.2968 | 276.004 | 8416 | 29.834 | 1.7531 | 28.8107 | 3.3579 | 0.75381 | 75.2422 | 2.3467 | 4.1169e+02
|  |  |  | 30 | 32.7806 | 1.7584 | 6.3489 | 81.3974 | 276.298 | 8478 | 29.873 | 1.7598 | 28.8105 | 3.3613 | 0.71892 | 75.2302 | 2.3094 | 4.1169e+02
|  |  |  | 20 | 32.7154 | 2.5827 | 7.0219 | 91.9015 | 305.62 | 10431 | 20.378 | 2.5838 | 29.4409 | 3.7203 | 1.2292 | 85.8671 | 7.4776 | 4.1387e+02
|  |  |  | 20 | 32.7134 | 2.5943 | 7.0319 | 92.0593 | 306.06 | 10470 | 20.302 | 2.5954 | 29.4467 | 3.7257 | 1.1092 | 85.8162 | 7.5894 | 4.1387e+02
|  |  |  | 20 | 32.7141 | 2.5974 | 7.0306 | 92.0495 | 306.002 | 10513 | 20.336 | 2.5985 | 29.4429 | 3.7254 | 1.2554 | 85.731 | 7.5247 | 4.1387e+02
|  |  |  | 20 | 32.7132 | 2.6119 | 7.0477 | 92.3062 | 306.745 | 10730 | 20.271 | 2.613 | 29.4655 | 3.734 | 1.0483 | 85.9571 | 7.559 | 4.1605e+02
|  |  |  | 15 | 32.5471 | 3.4494 | 7.1158 | 95.0767 | 309.773 | 12359 | 15.157 | 3.4503 | 30.0255 | 3.8286 | 0.59469 | 85.6957 | 13.397 | 4.2252e+02
|  |  |  | 15 | 32.546 | 3.4526 | 7.1263 | 95.2237 | 310.23 | 12397 | 15.079 | 3.4535 | 30.0255 | 3.8336 | 0.57477 | 85.8066 | 13.505 | 4.2247e+02
|  |  |  | 15 | 32.5467 | 3.4524 | 7.1404 | 95.412 | 310.844 | 12434 | 15.099 | 3.4533 | 30.0267 | 3.84 | 0.60255 | 85.7135 | 13.473 | 4.2217e+02
|  |  |  | 15 | 32.5463 | 3.4543 | 7.1377 | 95.3798 | 310.725 | 12477 | 15.189 | 3.4552 | 30.0269 | 3.8389 | 0.58598 | 85.7775 | 13.356 | 4.2163e+02
|  |  |  | 6 | 32.543 | 3.4713 | 7.3032 | 97.6294 | 317.931 | 14549 | 6.397 | 3.4717 | 30.0379 | 3.9193 | 0.76801 | 85.5552 | 42.761 | 4.2892e+02
|  |  |  | 6 | 32.5434 | 3.4703 | 7.3015 | 97.605 | 317.859 | 14595 | 6.328 | 3.4707 | 30.038 | 3.9185 | 0.64225 | 85.5106 | 44.054 | 4.2912e+02
|  |  |  | 5 | 32.5442 | 3.468 | 7.3106 | 97.7219 | 318.256 | 15903 | 5.552 | 3.4683 | 30.0357 | 3.9228 | 0.68783 | 85.5266 | 49.838 | 4.2518e+02
|  |  |  | 5 | 32.544 | 3.4683 | 7.3113 | 97.7315 | 318.286 | 15938 | 5.523 | 3.4686 | 30.0358 | 3.9232 | 0.67016 | 85.5126 | 50.022 | 4.2482e+02
|  |  |  | 5 | 32.5442 | 3.4679 | 7.3115 | 97.7326 | 318.292 | 15980 | 5.537 | 3.4683 | 30.0359 | 3.9232 | 0.59867 | 85.5502 | 49.721 | 4.2476e+02
|  |  |  | 5 | 32.5441 | 3.4688 | 7.3098 | 97.713 | 318.221 | 16094 | 5.551 | 3.4692 | 30.0362 | 3.9225 | 0.6405 | 85.5582 | 48.762 | 4.2476e+02
|  |  |  | 3 | 32.5432 | 3.4703 | 7.3106 | 97.7252 | 318.252 | 17423 | 3.414 | 3.4705 | 30.0353 | 3.9239 | 0.62465 | 85.4768 | 75.716 | 4.1623e+02
|  |  |  | 3 | 32.5436 | 3.4694 | 7.3135 | 97.7621 | 318.378 | 17482 | 3.608 | 3.4696 | 30.0353 | 3.925 | 0.66293 | 85.5262 | 72.749 | 4.1605e+02
| 064 | 064 | BRS-1 | bot-5 | 32.4504 | 1.5993 | 7.9985 | 101.9011 | 348.176 | 5280 | 44.023 | 1.6013 | 28.4283 | 4.0516 | 3.9523 | 72.8658 | 0.081579 | 2.9482e+02 | 803 | 1 | 1742 | 8.272 |  |  | d | 8.272 | entire water column very turbid!
|  |  |  | 35 | 32.4479 | 1.6268 | 8.0074 | 102.0847 | 348.565 | 6755 | 35.511 | 1.6284 | 28.4456 | 4.0618 | 3.2195 | 73.7215 | 0.076717 | 2.8713e+02 | 804 | 2 | 1743 | 8.275 |  |  | d | 8.275
|  |  |  | 30 | 32.4462 | 1.6514 | 8.027 | 102.397 | 349.417 | 7853 | 31.212 | 1.6528 | 28.4596 | 4.0743 | 3.7551 | 74.6738 | 0.075446 | 2.7882e+02 | 805 | 3 | 1739 | 8.269 |  |  | d | 8.269
|  |  |  | 25 | 32.4431 | 1.7457 | 7.9947 | 102.2305 | 348.014 | 8931 | 24.569 | 1.7468 | 28.5295 | 4.0716 | 3.5644 | 79.1924 | 0.18912 | 2.6615e+02 | 806 | 4 | 1725, 1731 | 8.219 | 8.221 |  | d | 8.219 | but did not seem to affect reps.
|  |  |  | 20 | 32.3594 | 2.7487 | 8.3017 | 108.8496 | 361.429 | 10031 | 19.565 | 2.7498 | 29.2819 | 4.2965 | 4.8301 | 77.5395 | 0.46661 | 2.5992e+02 | 807 | 5 | 1193 | 8.543 |  |  | d | 8.543
|  |  |  | 15 | 32.3364 | 3.1096 | 8.4639 | 111.9723 | 368.507 | 11089 | 14.515 | 3.1104 | 29.5692 | 4.4043 | 4.9229 | 75.9115 | 1.6471 | 2.5268e+02 | 808 | 6 | 1750 | 8.745 |  |  | d | 8.745
|  |  |  | 10 | 32.3372 | 3.1197 | 8.458 | 111.9233 | 368.252 | 12265 | 9.866 | 3.1203 | 29.576 | 4.405 | 4.9314 | 75.213 | 5.1523 | 2.4179e+02 | 809 | 7 | 1744 | 8.78 |  |  | d | 8.78
|  |  |  | 5 | 32.3375 | 3.1156 | 8.6061 | 113.8703 | 374.698 | 13353 | 5.851 | 3.116 | 29.5701 | 4.4733 | 4.9187 | 75.6515 | 14.781 | 2.3329e+02 | 810 | 8 | 1740 | 8.776 |  |  | d | 8.776
| 065 | 065 | BRS-2 | bot-5 | 32.3775 | 2.1375 | 8.0814 | 104.3417 | 351.817 | 4980 | 45.954 | 2.1397 | 28.8132 | 4.1322 | 1.8752 | 75.9896 | 0.076103 | 7.1883e+01 | 811 | 1 | 1726 | 8.412 |  |  |  | 8.412
|  |  |  | 40 | 32.371 | 2.1759 | 8.1012 | 104.6957 | 352.683 | 6520 | 40.317 | 2.1779 | 28.833 | 4.1467 | 2.4812 | 76.2975 | 0.073697 | 6.9749e+01 | 812 | 2 | 1732 | 8.425 |  |  |  | 8.425
|  |  |  | 35 | 32.3735 | 2.1618 | 8.1031 | 104.6823 | 352.762 | 8286 | 33.416 | 2.1634 | 28.8244 | 4.1497 | 1.8501 | 75.3193 | 0.081234 | 6.9705e+01 | 813 | 3 | 1752 | 8.44 |  |  |  | 8.44
|  |  |  | 30 | 32.3677 | 2.2118 | 8.1276 | 105.1293 | 353.834 | 9617 | 30.314 | 2.2133 | 28.8582 | 4.1662 | 2.2627 | 76.2922 | 0.10014 | 6.9661e+01 | 814 | 4 | 1709 | 8.451 |  |  | b | 8.451 | dancing bubbles
|  |  |  | 25 | 32.3672 | 2.225 | 8.1248 | 105.1267 | 353.71 | 11389 | 25.259 | 2.2263 | 28.8675 | 4.1685 | 2.3305 | 77.0473 | 0.21173 | 6.7526e+01 | 815 | 5 | 1745, 1741 | 8.458 | 8.452 |  |  | 8.458
|  |  |  | 20 | 32.3655 | 2.2442 | 8.169 | 105.7489 | 355.636 | 12609 | 20.365 | 2.2452 | 28.8788 | 4.192 | 2.6139 | 77.0693 | 0.69346 | 6.7526e+01 | 816 | 6 | 1727 | 8.459 |  |  |  | 8.459
|  |  |  | 15 | 32.3036 | 3.0176 | 8.1711 | 107.8266 | 355.768 | 14055 | 15.667 | 3.0185 | 29.4768 | 4.263 | 1.9343 | 82.095 | 1.4407 | 6.7526e+01 | 817 | 7 | 1733 | 8.459 |  |  |  | 8.459
|  |  |  | 10 | 32.1107 | 4.5449 | 7.8335 | 107.2252 | 341.165 | 15565 | 11.297 | 4.5456 | 30.5709 | 4.2408 | 0.75395 | 86.3143 | 2.7528 | 6.5348e+01 | 818 | 8 | 1753 | 8.173 |  |  |  | 8.173
|  |  |  | 5 | 32.1275 | 4.4694 | 7.7743 | 106.2296 | 338.581 | 16853 | 6.115 | 4.4698 | 30.5272 | 4.2096 | 0.90245 | 86.3813 | 5.7113 | 6.5348e+01 | 819 | 9 | 1700 | 8.172 |  |  |  | 8.172
| 066 | 066 | BRS-3 | bot-5 | 32.0682 | 3.8324 | 7.5777 | 101.8972 | 330.013 | 8050 | 50.387 | 3.8356 | 29.962 | 4.0431 | 1.0904 | 75.7914 | 0.084962 | 4.3565e+01 | 820 | 1 | 1696 | 7.938 |  |  | b | 7.938 | copepod
|  |  |  | 40 | 32.083 | 3.8836 | 7.6767 | 103.3681 | 334.322 | 9745 | 41.736 | 3.8862 | 30.0136 | 4.0969 | 0.91675 | 75.9259 | 0.079076 | 4.5744e+01 | 821 | 2 | 1706 | 7.965 |  |  |  | 7.965
|  |  |  | 35 | 32.0945 | 3.9079 | 7.7266 | 104.1103 | 336.495 | 11372 | 35.632 | 3.9101 | 30.0388 | 4.1248 | 1.0152 | 76.5838 | 0.077812 | 4.5744e+01 | 822 | 3 | 1728 | 8.001 |  |  |  | 8.001
|  |  |  | 30 | 32.0811 | 3.8854 | 7.7226 | 103.9878 | 336.321 | 13455 | 30.326 | 3.8873 | 30.0035 | 4.1233 | 1.2061 | 78.2147 | 0.077897 | 4.5744e+01 | 823 | 4 | 1734 | 8.02 |  |  |  | 8.02
|  |  |  | 25 | 32.0736 | 3.8674 | 7.7032 | 103.6753 | 335.481 | 14681 | 24.681 | 3.869 | 29.9823 | 4.1154 | 0.97077 | 78.5015 | 0.092423 | 4.7922e+01 | 824 | 5 | 1751, 1423 | 7.988 | 7.977 |  | c | 7.988 | reps a bit off
|  |  |  | 20 | 32.0234 | 3.8745 | 7.7235 | 103.9312 | 336.378 | 15899 | 20.863 | 3.8758 | 29.9469 | 4.1259 | 1.1081 | 80.1826 | 0.18928 | 4.7922e+01 | 825 | 6 | 1737 | 7.981 |  |  |  | 7.981
|  |  |  | 15 | 32.0263 | 3.7914 | 7.6548 | 102.795 | 333.38 | 17251 | 15.764 | 3.7924 | 29.8802 | 4.0899 | 1.0417 | 81.0679 | 0.42609 | 4.7922e+01 | 826 | 7 | 1723 | 7.949 |  |  |  | 7.949
|  |  |  | 10 | 31.9688 | 3.9509 | 7.6455 | 103.0375 | 332.998 | 18694 | 10.341 | 3.9515 | 29.8705 | 4.1003 | 1.0816 | 81.6956 | 1.5269 | 4.7922e+01 | 827 | 8 | 1729 | 7.944 |  |  |  | 7.944
|  |  |  | 5 | 31.9158 | 4.2212 | 7.7514 | 105.1257 | 337.632 | 20060 | 4.822 | 4.2215 | 30.124 | 4.1733 | 0.98819 | 82.6719 | 4.0267 | 4.7922e+01 | 828 | 9 | 1738 | 7.889 |  |  |  | 7.889
| 067 | 067 | BRS-4 | bot-10 | 31.7962 | 3.1303 | 7.6643 | 101.0869 | 333.837 | 8529 | 44.961 | 3.1329 | 29.1506 | 4.0199 | 0.89838 | 77.2187 | 0.07662 | 2.3961e+01 | 829 | 1 | 1336 | 7.942 |  |  |  | 7.942
|  |  |  | 40 | 31.7976 | 3.1411 | 7.6592 | 101.047 | 333.614 | 10106 | 41.122 | 3.1434 | 29.1581 | 4.0203 | 0.86133 | 76.4555 | 0.073254 | 2.3961e+01 | 830 | 2 | 1135 | 7.944 |  |  |  | 7.944
|  |  |  | 35 | 31.7948 | 3.13 | 7.6543 | 100.9517 | 333.401 | 11400 | 35.858 | 3.132 | 29.1526 | 4.0195 | 0.93484 | 77.5055 | 0.072969 | 2.3961e+01 | 831 | 3 | 1094 | 7.951 |  |  |  | 7.951
|  |  |  | 30 | 31.7917 | 3.1132 | 7.7037 | 101.5574 | 335.553 | 12668 | 30.556 | 3.1149 | 29.1259 | 4.0426 | 0.89225 | 78.0978 | 0.072045 | 2.3961e+01 | 832 | 4 | 1293, 1654 | 7.952 | 7.951 |  |  | 7.952
|  |  |  | 25 | 31.7938 | 3.1204 | 7.7248 | 101.8542 | 336.469 | 13972 | 26.261 | 3.1219 | 29.1302 | 4.0547 | 0.86711 | 77.9111 | 0.072012 | 2.3961e+01 | 833 | 5 | 1376 | 7.975 |  |  |  | 7.975
|  |  |  | 20 | 31.7865 | 3.112 | 7.7095 | 101.626 | 335.807 | 15425 | 20.378 | 3.1132 | 29.1169 | 4.0497 | 0.97774 | 78.2127 | 0.071444 | 2.3961e+01 | 834 | 6 | 1284 | 7.965 |  |  | d | 7.965 | pulled flask just prior to end of titration, profile reasonable
|  |  |  | 15 | 31.6009 | 3.4485 | 7.5969 | 100.8642 | 330.96 | 16926 | 13.842 | 3.4493 | 29.2524 | 4.026 | 1.0064 | 79.2814 | 0.07285 | 2.3961e+01 | 835 | 7 | 1143 | 7.866 |  |  |  | 7.866
|  |  |  | 10 | 31.3088 | 3.8659 | 7.584 | 101.5478 | 330.486 | 18449 | 11.035 | 3.8666 | 29.3316 | 4.0495 | 0.97991 | 78.4368 | 0.093531 | 2.3694e+01 | 836 | 8 | 1089 | 7.757 |  |  |  | 7.757
|  |  |  | 5 | 28.3067 | 8.2535 | 6.6596 | 97.0998 | 291.017 | 20698 | 4.275 | 8.2539 | 30.0553 | 3.8911 | 1.2521 | 74.6896 | 0.52241 | 2.2405e+01 | 837 | 9 | 1509 | 6.791 |  |  | b | 6.791 | lots of small copepod floaties
| 068 | 068 | BRS-5 | bot-5 | 30.9706 | 3.9882 | 7.404 | 99.219 | 322.728 | 11575 | 37.439 | 3.9906 | 29.1491 | 3.958 | 1.0873 | 69.6716 | 0.073718 | 2.1383e+01 | 838 | 1 | 879 | 7.646 |  |  |  | 7.646
|  |  |  | 35 | 30.8849 | 4.1391 | 7.3852 | 99.2813 | 321.936 | 13216 | 35.209 | 4.1413 | 29.1915 | 3.9608 | 1.0237 | 70.3683 | 0.072443 | 2.1516e+01 | 839 | 2 | 1746 | 7.629 |  |  |  | 7.629
|  |  |  | 30 | 30.8765 | 4.1667 | 7.3696 | 99.1323 | 321.257 | 14472 | 30.203 | 4.1686 | 29.2145 | 3.958 | 1.0152 | 70.3832 | 0.07286 | 2.1160e+01 | 840 | 3 | 1145 | 7.582 |  |  |  | 7.582
|  |  |  | 25 | 30.7072 | 4.4125 | 7.3404 | 99.2271 | 320.033 | 15946 | 25.531 | 4.4142 | 29.2579 | 3.9627 | 0.9674 | 71.2675 | 0.072306 | 2.1561e+01 | 841 | 4 | 1748, 1658 | 7.567 | 7.57 |  |  | 7.567
|  |  |  | 20 | 30.4323 | 4.8368 | 7.2185 | 98.42 | 314.801 | 17204 | 20.29 | 4.8381 | 29.3548 | 3.9367 | 0.92544 | 72.0839 | 0.071194 | 2.1427e+01 | 842 | 5 | 1156 | 7.42 |  |  |  | 7.42
|  |  |  | 15 | 29.8573 | 5.7335 | 7.0322 | 97.6147 | 306.842 | 18453 | 14.973 | 5.7346 | 29.5421 | 3.9097 | 0.86003 | 71.3192 | 0.070434 | 2.0671e+01 | 843 | 6 | 663 | 7.27 |  |  |  | 7.27
|  |  |  | 10 | 28.2796 | 8.4506 | 6.5612 | 96.0829 | 286.732 | 19834 | 10.863 | 8.4516 | 30.2254 | 3.8533 | 1.156 | 72.8083 | 0.070229 | 2.1116e+01 | 844 | 7 | 1137 | 6.936 |  |  |  | 6.936
|  |  |  | 5 | 27.4461 | 9.9754 | 6.3418 | 95.6211 | 277.383 | 21232 | 5.724 | 9.976 | 30.572 | 3.8364 | 1.4131 | 72.3693 | 0.1044 | 2.0716e+01 | 845 | 8 | 1023 | 6.503 |  |  |  | 6.503
| 069 | 069 | UTN-1 | bot-x | 30.057 | 6.1939 | 6.8983 | 96.9455 | 300.969 | 11618 | 28.092 | 6.1961 | 30.1156 | 3.8801 | 0.60188 | 76.1131 | 0.10847 | 4.7526e+02
|  |  |  | bot-x | 30.0655 | 6.1837 | 6.926 | 97.3153 | 302.171 | 11702 | 27.961 | 6.1858 | 30.1148 | 3.8927 | 0.58192 | 76.082 | 0.11287 | 4.7704e+02
|  |  |  | bot-x | 30.085 | 6.1626 | 6.9369 | 97.4322 | 302.642 | 11758 | 28.129 | 6.1648 | 30.116 | 3.8966 | 0.5968 | 75.9106 | 0.10883 | 4.7869e+02 | 846 | 3 | 1022 | 7.161 |  |  |  | 7.161
|  |  |  | 25 | 29.959 | 6.3055 | 6.9212 | 97.4623 | 301.992 | 13337 | 25.651 | 6.3075 | 30.1149 | 3.8984 | 0.59597 | 77.144 | 0.16133 | 4.8571e+02 | 847 | 4 | 1666 | 7.162 |  |  |  | 7.162
|  |  |  | 20 | 29.851 | 6.449 | 6.8771 | 97.1025 | 300.098 | 14898 | 20.686 | 6.4506 | 30.129 | 3.8881 | 0.65727 | 79.1675 | 0.30746 | 4.3783e+02 | 848 | 5 | 1155 | 7.139 |  |  |  | 7.139
|  |  |  | 15 | 29.4639 | 6.9701 | 6.8059 | 97.0406 | 297.097 | 16349 | 15.39 | 6.9714 | 30.1877 | 3.8871 | 0.62867 | 78.4287 | 1.0196 | 4.3957e+02 | 849 | 6 | 1136 | 7.027 |  |  |  | 7.027
|  |  |  | 10 | 29.337 | 7.121 | 6.7914 | 97.0977 | 296.501 | 18104 | 10.669 | 7.1219 | 30.1826 | 3.8908 | 0.82111 | 73.1258 | 4.4091 | 4.2543e+02 | 850 | 7 | 970, 1169 | 6.952 | 6.952 |  |  | 6.952
|  |  |  | 5 | 29.0773 | 7.5436 | 6.7637 | 97.4962 | 295.368 | 20078 | 5.951 | 7.5441 | 30.2737 | 3.9055 | 0.68516 | 74.5761 | 21.521 | 3.9084e+02 | 851 | 8 | 1724 | 6.906 |  |  |  | 6.906
| 070 | 070 | UTN-2 | bot-5 | 32.3087 | 2.8166 | 8.1475 | 106.9779 | 354.73 | 3839 | 40.669 | 2.8189 | 29.3203 | 4.2223 | 2.336 | 72.6154 | 0.18545 | 1.3578e+03
|  |  |  | bot-5 | 32.3075 | 2.8158 | 8.1663 | 107.2223 | 355.551 | 3942 | 40.7 | 2.818 | 29.3163 | 4.2306 | 2.1096 | 73.2024 | 0.18127 | 1.3562e+03
|  |  |  | bot-5 | 32.3075 | 2.815 | 8.1571 | 107.0989 | 355.15 | 4005 | 40.561 | 2.8172 | 29.3176 | 4.2265 | 2.3011 | 73.3149 | 0.18716 | 1.3553e+03
|  |  |  | 35 | 32.3039 | 2.8318 | 8.1778 | 107.4136 | 356.055 | 5216 | 34.249 | 2.8337 | 29.3188 | 4.2402 | 1.3694 | 76.2022 | 0.65204 | 1.3915e+03
|  |  |  | 30 | 32.28 | 3.0299 | 8.1233 | 107.2132 | 353.691 | 6700 | 29.929 | 3.0316 | 29.4598 | 4.2351 | 1.3873 | 78.3904 | 1.4934 | 1.3496e+03
|  |  |  | 25 | 32.2779 | 3.0715 | 8.1029 | 107.0535 | 352.804 | 7929 | 24.255 | 3.0729 | 29.4904 | 4.2323 | 1.7677 | 76.5934 | 4.6419 | 1.4387e+03
|  |  |  | 20 | 32.2122 | 3.527 | 7.9978 | 106.8297 | 348.261 | 9189 | 19.7 | 3.5282 | 29.7415 | 4.2258 | 1.3662 | 81.8086 | 11.926 | 1.4522e+03
|  |  |  | 15 | 32.1329 | 4.1173 | 7.7263 | 104.6688 | 336.476 | 10241 | 14.333 | 4.1182 | 30.2377 | 4.1535 | 0.83078 | 83.2546 | 28.948 | 1.4292e+03
|  |  |  | 10 | 32.1216 | 4.1794 | 7.6975 | 104.4302 | 335.227 | 11463 | 10.452 | 4.1801 | 30.2758 | 4.147 | 0.76851 | 82.9468 | 58.798 | 1.5313e+03
|  |  |  | 5 | 31.8434 | 4.9441 | 7.4845 | 103.2669 | 326.047 | 12316 | 5.248 | 4.9445 | 30.73 | 4.1081 | 0.32212 | 84.0643 | 170.45 | 1.6786e+03
| 071 | 071 | UTN-3 | bot-5 | 32.4146 | 1.8147 | 8.2216 | 105.3028 | 357.903 | 3252 | 49.337 | 1.8171 | 28.5776 | 4.1639 | 2.6171 | 69.3582 | 0.10231 | 2.0478e+03
|  |  |  | bot-5 | 32.4144 | 1.8144 | 8.2083 | 105.1308 | 357.321 | 3295 | 49.306 | 1.8167 | 28.5776 | 4.1581 | 1.9276 | 68.7046 | 0.1021 | 2.0601e+03
|  |  |  | bot-5 | 32.4144 | 1.8144 | 8.2089 | 105.1386 | 357.348 | 3341 | 49.449 | 1.8167 | 28.5775 | 4.1583 | 2.2112 | 69.1053 | 0.10211 | 2.0551e+03
|  |  |  | 40 | 32.4132 | 1.8125 | 8.2448 | 105.5918 | 358.913 | 5406 | 40.431 | 1.8144 | 28.5709 | 4.178 | 2.263 | 74.6666 | 0.20732 | 2.2030e+03
|  |  |  | 35 | 32.4131 | 1.8122 | 8.2663 | 105.8648 | 359.846 | 6719 | 35.437 | 1.8139 | 28.5685 | 4.1897 | 1.6845 | 73.7808 | 0.60186 | 1.9829e+03
|  |  |  | 30 | 32.4128 | 1.8122 | 8.3015 | 106.3145 | 361.379 | 7939 | 29.489 | 1.8135 | 28.5656 | 4.2078 | 1.5955 | 76.6676 | 2.1694 | 1.1543e+03
|  |  |  | 25 | 32.4085 | 1.8709 | 8.3751 | 107.4158 | 364.586 | 9355 | 24.916 | 1.8721 | 28.6056 | 4.2474 | 0.77332 | 83.27 | 4.8497 | 1.1298e+03
|  |  |  | 20 | 32.3762 | 2.7597 | 8.4977 | 111.4627 | 369.957 | 10819 | 20.733 | 2.7608 | 29.3151 | 4.3847 | 2.0908 | 81.8555 | 11.93 | 1.0321e+03
|  |  |  | 15 | 32.3788 | 2.8329 | 8.5209 | 111.9753 | 370.968 | 12021 | 15.715 | 2.8337 | 29.377 | 4.4046 | 1.1812 | 82.6236 | 31.134 | 9.7217e+02
|  |  |  | 10 | 32.3797 | 2.9664 | 8.5274 | 112.4391 | 371.257 | 13426 | 10.441 | 2.967 | 29.493 | 4.4227 | 0.57943 | 82.7107 | 100.62 | 8.3063e+02
|  |  |  | 5 | 32.3794 | 3.1284 | 8.6205 | 114.1301 | 375.316 | 15051 | 5.26 | 3.1287 | 29.6209 | 4.4825 | 0.25773 | 82.8826 | 335.07 | 2.0045e+03
| 072 | 072 | UTN-4 | bot-5 | 32.4484 | 1.5932 | 8.1154 | 103.3729 | 353.263 | 5170 | 49.373 | 1.5955 | 28.4214 | 4.099 | 4.3832 | 57.0807 | 0.099403 | 1.9953e+03 | 869 | 1 | 1730 | 8.516 |  |  |  | 8.516
|  |  |  | 40 | 32.4467 | 1.5754 | 8.1471 | 103.7271 | 354.645 | 6125 | 40.01 | 1.5772 | 28.4024 | 4.1155 | 2.1737 | 71.7799 | 0.19071 | 2.0309e+03 | 870 | 2 | 1742, 1743 | 8.525 | 8.532 |  |  | 8.525
|  |  | jackie | 40 | 32.4464 | 1.5751 | 8.1536 | 103.8083 | 354.926 | 6247 | 40.039 | 1.5769 | 28.4026 | 4.1183 | 2.4972 | 71.619 | 0.18955 | 2.0517e+03
|  |  | jackie | 40 | 32.4464 | 1.5755 | 8.1641 | 103.9437 | 355.385 | 6805 | 40.005 | 1.5773 | 28.4021 | 4.1229 | 2.7976 | 72.0169 | 0.20022 | 1.9974e+03
|  |  |  | 35 | 32.4463 | 1.5726 | 8.1924 | 104.2955 | 356.618 | 8069 | 33.714 | 1.5741 | 28.3972 | 4.1378 | 2.628 | 73.8277 | 0.74076 | 2.0846e+03 | 871 | 3 | 1739 | 8.53 |  |  |  | 8.53
|  |  |  | 30 | 32.4462 | 1.5824 | 8.2275 | 104.7684 | 358.146 | 9249 | 29.388 | 1.5837 | 28.4 | 4.1559 | 1.904 | 75.845 | 1.4724 | 2.0907e+03 | 872 | 4 | 1725 | 8.533 |  |  | b | 8.533 | burette stalled, finished titrating
|  |  |  | 25 | 32.4428 | 1.9329 | 8.1011 | 104.0917 | 352.651 | 10300 | 25.598 | 1.9341 | 28.671 | 4.1339 | 2.1266 | 82.5345 | 2.4417 | 2.0759e+03 | 873 | 5 | 1731 | 8.495 |  |  |  | 8.495
|  |  |  | 23 | 32.4486 | 1.9672 | 8.0969 | 104.1333 | 352.468 | 11829 | 23.323 | 1.9683 | 28.7195 | 4.1363 | 2.707 | 84.6446 | 3.38 | 2.0408e+03
|  |  |  | 23 | 32.4484 | 1.9674 | 8.1155 | 104.3732 | 353.278 | 11877 | 23.334 | 1.9685 | 28.7192 | 4.1444 | 1.8542 | 84.634 | 3.3484 | 2.0383e+03
|  |  |  | 23 | 32.4485 | 1.9678 | 8.1315 | 104.5803 | 353.975 | 11916 | 23.259 | 1.9689 | 28.7191 | 4.1515 | 1.6914 | 84.3555 | 3.4181 | 2.0496e+03
|  |  |  | 23 | 32.4485 | 1.9672 | 8.1348 | 104.6201 | 354.116 | 11954 | 23.307 | 1.9683 | 28.7191 | 4.1528 | 1.9586 | 84.6317 | 3.3999 | 2.0525e+03
|  |  |  | 23 | 32.4484 | 1.9672 | 8.1244 | 104.4871 | 353.666 | 11994 | 23.34 | 1.9683 | 28.719 | 4.1483 | 2.654 | 84.2779 | 3.38 | 2.0497e+03
|  |  |  | 23 | 32.4485 | 1.9676 | 8.1159 | 104.3786 | 353.295 | 12039 | 23.279 | 1.9687 | 28.7191 | 4.1447 | 1.8078 | 84.9098 | 3.4051 | 2.0539e+03
|  |  |  | 23 | 32.4485 | 1.9675 | 8.1174 | 104.3983 | 353.362 | 12080 | 23.28 | 1.9687 | 28.7191 | 4.1453 | 1.071 | 84.2205 | 3.4333 | 2.0605e+03
|  |  |  | 23 | 32.4485 | 1.9675 | 8.1097 | 104.2992 | 353.027 | 12122 | 23.269 | 1.9686 | 28.719 | 4.142 | 1.4511 | 85.0076 | 3.4526 | 2.0617e+03
|  |  |  | 23 | 32.4394 | 1.9676 | 8.112 | 104.3219 | 353.127 | 12177 | 23.244 | 1.9687 | 28.7189 | 4.1428 | 2.3475 | 84.8698 | 3.4353 | 2.0659e+03
|  |  |  | 23 | 32.4485 | 1.9678 | 8.1021 | 104.2019 | 352.695 | 12211 | 23.268 | 1.9689 | 28.7192 | 4.1387 | 2.5621 | 84.4509 | 3.4638 | 2.0667e+03
|  |  |  | 23 | 32.4483 | 1.9679 | 8.0947 | 104.1071 | 352.374 | 12251 | 23.279 | 1.969 | 28.7194 | 4.1354 | 2.809 | 84.5711 | 3.4414 | 2.0668e+03
|  |  |  | 23 | 32.4486 | 1.9678 | 8.0969 | 104.1347 | 352.467 | 12346 | 23.274 | 1.9689 | 28.7196 | 4.1364 | 1.9673 | 84.4009 | 3.4472 | 2.0597e+03
|  |  |  | 23 | 32.4487 | 1.9678 | 8.1039 | 104.2254 | 352.774 | 12393 | 23.289 | 1.9689 | 28.7194 | 4.1395 | 1.3898 | 84.7378 | 3.438 | 2.0613e+03
|  |  |  | 20 | 32.447 | 1.9758 | 8.1776 | 105.1925 | 355.98 | 14091 | 20.447 | 1.9768 | 28.7236 | 4.1737 | 1.9957 | 84.7835 | 5.3746 | 2.0589e+03 | 874 | 6 | 1193 | 8.442 |  |  |  | 8.442
|  |  |  | 15 | 32.3893 | 2.9535 | 8.3569 | 110.1626 | 363.83 | 15410 | 14.976 | 2.9543 | 29.4816 | 4.3429 | 1.0259 | 86.1476 | 11.177 | 2.0559e+03 | 875 | 7 | 1750, 1744 | 8.595 | 8.606 |  | c | 8.595 | reps a bit off
|  |  |  | 10 | 32.3831 | 3.1583 | 8.3268 | 110.328 | 362.529 | 16628 | 9.725 | 3.1589 | 29.6479 | 4.3506 | 0.42262 | 86.2515 | 18.12 | 2.0568e+03 | 876 | 8 | 1740 | 8.64 |  |  |  | 8.64
|  |  |  | 5 | 32.3741 | 3.5668 | 8.3805 | 112.1715 | 364.881 | 17692 | 5.327 | 3.5671 | 29.998 | 4.4146 | 0.3445 | 87.0703 | 31.736 | 2.0595e+03 | 877 | 9 | 1726 | 8.587 |  |  |  | 8.587
| 073 | 073 | UTN-4 | 41 | 32.446 | 1.5748 | 8.185 | 104.2073 | 356.294 | 8255 | 42.564 | 1.5767 | 28.4033 | 4.1306 | 3.1634 | 71.1196 | 0.083205 | 1.8951e+03
|  |  |  | 41 | 32.4461 | 1.5749 | 8.1643 | 103.9449 | 355.395 | 8367 | 42.554 | 1.5768 | 28.4037 | 4.1217 | 2.8985 | 70.7719 | 0.07993 | 1.8962e+03
|  |  |  | 40 | 32.4462 | 1.5726 | 8.1476 | 103.7248 | 354.664 | 9717 | 40.84 | 1.5744 | 28.401 | 4.1151 | 3.5691 | 71.6637 | 0.11284 | 1.8220e+03
|  |  |  | 40 | 32.4461 | 1.5723 | 8.1447 | 103.6876 | 354.54 | 9765 | 40.865 | 1.5741 | 28.4019 | 4.1138 | 2.7543 | 71.1392 | 0.1106 | 1.8901e+03
|  |  |  | 24 | 32.4465 | 1.6062 | 8.2042 | 104.5345 | 357.129 | 11599 | 24.808 | 1.6073 | 28.4216 | 4.1501 | 1.0198 | 80.3962 | 2.1269 | 1.8828e+03
|  |  |  | 24 | 32.4466 | 1.6067 | 8.1987 | 104.4668 | 356.892 | 11641 | 24.849 | 1.6078 | 28.422 | 4.1478 | 0.92202 | 80.5172 | 2.1012 | 1.8841e+03
|  |  |  | 24 | 32.4466 | 1.6071 | 8.1935 | 104.4011 | 356.664 | 11684 | 24.887 | 1.6082 | 28.423 | 4.1456 | 1.0047 | 80.6344 | 2.0939 | 1.8835e+03
|  |  |  | 24 | 32.4461 | 1.6133 | 8.1944 | 104.4287 | 356.703 | 11734 | 24.843 | 1.6144 | 28.4266 | 4.1465 | 1.3341 | 80.907 | 2.1136 | 1.8820e+03
|  |  |  | 11 | 32.3903 | 3.0167 | 8.337 | 110.0756 | 362.965 | 13874 | 11.818 | 3.0173 | 29.5349 | 4.3414 | 1.1328 | 85.8395 | 11.901 | 1.8733e+03
|  |  |  | 11 | 32.3904 | 3.0169 | 8.341 | 110.129 | 363.139 | 13920 | 11.793 | 3.0176 | 29.5353 | 4.3432 | 1.1454 | 85.4793 | 12.015 | 1.8740e+03
|  |  |  | 5 | 32.382 | 3.5891 | 8.3451 | 111.7655 | 363.338 | 15077 | 5.281 | 3.5894 | 29.974 | 4.4007 | 0.3623 | 86.5754 | 29.345 | 1.8711e+03
|  |  |  | 5 | 32.3831 | 3.5669 | 8.3475 | 111.7372 | 363.443 | 15115 | 5.281 | 3.5672 | 29.9597 | 4.3998 | 0.27443 | 86.6761 | 29.44 | 1.8702e+03
|  |  |  | 5 | 32.3798 | 3.5678 | 8.3429 | 111.6754 | 363.243 | 15155 | 5.258 | 3.5681 | 29.9769 | 4.3977 | 0.20817 | 86.7332 | 29.258 | 1.8709e+03
|  |  |  | 5 | 32.382 | 3.5751 | 8.3504 | 111.797 | 363.567 | 15201 | 5.244 | 3.5754 | 29.9994 | 4.4019 | 0.27017 | 86.4369 | 29.072 | 1.8733e+03
|  |  |  | 2 | 32.3829 | 3.5642 | 8.3951 | 112.3654 | 365.513 | 17394 | 2.534 | 3.5643 | 29.9876 | 4.4227 | 0.2683 | 86.6351 | 46.88 | 1.8628e+03
|  |  |  | 2 | 32.3826 | 3.5843 | 8.3922 | 112.3835 | 365.39 | 17505 | 2.468 | 3.5844 | 29.991 | 4.4233 | 0.22733 | 86.5786 | 47.34 | 1.8624e+03
| 074 | 074 | UTN-5 | bot-5 | 32.6056 | 1.8758 | 6.9567 | 89.3571 | 302.796 | 7795 | 48.091 | 1.8781 | 28.7811 | 3.624 | 0.91595 | 58.4454 | 0.079669 | 8.0947e+02
|  |  |  | bot-5 | 32.6087 | 1.8852 | 6.9505 | 89.3009 | 302.525 | 7880 | 48.014 | 1.8875 | 28.7892 | 3.6221 | 1.8631 | 57.6673 | 0.078981 | 9.0629e+02
|  |  |  | bot-5 | 32.6093 | 1.8862 | 6.9172 | 88.8755 | 301.075 | 8223 | 48.035 | 1.8885 | 28.7923 | 3.6077 | 1.2055 | 49.7539 | 0.077798 | 6.9585e+02
|  |  |  | 40 | 32.6049 | 1.8726 | 6.948 | 89.2358 | 302.414 | 9703 | 40.791 | 1.8745 | 28.774 | 3.6228 | 0.86474 | 60.4099 | 0.076413 | 6.2316e+02
|  |  |  | 35 | 32.5996 | 1.8535 | 6.9767 | 89.5577 | 303.668 | 10994 | 35.74 | 1.8552 | 28.7538 | 3.6359 | 1.7204 | 64.9089 | 0.075507 | 6.1689e+02
|  |  |  | 30 | 32.5979 | 1.8462 | 6.9839 | 89.6312 | 303.98 | 13312 | 30.565 | 1.8476 | 28.7425 | 3.6405 | 0.93207 | 66.5953 | 0.084253 | 1.1070e+03
|  |  |  | 25 | 32.5961 | 1.8407 | 6.9977 | 89.7938 | 304.58 | 15655 | 25.558 | 1.8419 | 28.7341 | 3.6481 | 1.3842 | 66.7844 | 0.16746 | 9.7645e+02
|  |  |  | 20 | 32.5768 | 1.7797 | 7.0813 | 90.7116 | 308.222 | 17313 | 20.547 | 1.7807 | 28.664 | 3.6815 | 1.0111 | 71.099 | 0.42218 | 4.0471e+02
|  |  |  | 15 | 32.5181 | 2.3227 | 7.7406 | 100.5072 | 336.95 | 19281 | 13.892 | 2.3234 | 29.0595 | 4.0165 | 4.2438 | 78.184 | 1.886 | 1.1713e+03
|  |  |  | 10 | 32.5185 | 2.3856 | 8.1399 | 105.8606 | 354.331 | 20764 | 10.687 | 2.3862 | 29.1157 | 4.2001 | 4.1942 | 78.0935 | 4.247 | 7.5946e+02
|  |  |  | 5 | 32.5158 | 2.6904 | 8.2661 | 108.3345 | 359.834 | 22732 | 5.816 | 2.6907 | 29.3776 | 4.286 | 3.2945 | 78.0322 | 15.839 | 4.7971e+02
| 075 | 075 | UTN-6 | bot-5 | 32.5527 | 1.631 | 6.9685 | 88.9133 | 303.316 | 8142 | 46.507 | 1.6331 | 28.5366 | 3.6101 | 1.9117 | 68.5924 | 0.0783 | 5.2278e+01
|  |  |  | bot-5 | 32.5526 | 1.631 | 6.966 | 88.8817 | 303.209 | 8226 | 46.568 | 1.6331 | 28.5365 | 3.609 | 1.4399 | 68.9783 | 0.077848 | 5.2278e+01
|  |  |  | bot-5 | 32.5525 | 1.6305 | 6.9731 | 88.9703 | 303.515 | 8292 | 46.547 | 1.6326 | 28.5374 | 3.612 | 1.6696 | 69.1598 | 0.077491 | 5.2278e+01
|  |  |  | 40 | 32.5519 | 1.6343 | 6.9807 | 89.0753 | 303.847 | 9940 | 40.832 | 1.6361 | 28.5364 | 3.6179 | 2.9437 | 70.8696 | 0.075131 | 5.1878e+01
|  |  |  | 35 | 32.5511 | 1.6391 | 6.994 | 89.2547 | 304.425 | 11279 | 35.16 | 1.6407 | 28.5363 | 3.6263 | 1.2197 | 72.8155 | 0.074325 | 5.0100e+01
|  |  |  | 30 | 32.5506 | 1.6426 | 6.9943 | 89.2666 | 304.441 | 12933 | 29.918 | 1.6439 | 28.5357 | 3.6288 | 1.6417 | 73.5481 | 0.074035 | 5.0100e+01
|  |  |  | 25 | 32.5498 | 1.6504 | 7.0122 | 89.5121 | 305.221 | 14268 | 24.949 | 1.6515 | 28.5398 | 3.6392 | 1.3695 | 73.7746 | 0.073667 | 4.7922e+01
|  |  |  | 20 | 32.5487 | 1.6675 | 7.0752 | 90.3542 | 307.961 | 15470 | 20.056 | 1.6684 | 28.5537 | 3.6698 | 1.2777 | 76.5426 | 0.074683 | 4.7833e+01
|  |  |  | 15 | 32.5073 | 2.472 | 7.7944 | 101.5833 | 339.296 | 16984 | 15.401 | 2.4728 | 29.1808 | 4.0521 | 2.5762 | 81.3332 | 0.096996 | 4.7744e+01
|  |  |  | 10 | 32.4909 | 2.8104 | 8.253 | 108.4737 | 359.272 | 18368 | 10.246 | 2.811 | 29.4454 | 4.2882 | 4.919 | 77.5654 | 0.25679 | 4.5744e+01
|  |  |  | 5 | 32.4926 | 2.8584 | 8.3776 | 110.2455 | 364.698 | 19687 | 5.592 | 2.8587 | 29.4788 | 4.3507 | 4.655 | 76.2414 | 1.3189 | 4.5744e+01
| 076 | 076 | UTN-7 | bot-5 | 33.0751 | 1.1223 | 6.4519 | 81.5299 | 280.706 | 8945 | 52.969 | 1.1245 | 28.5343 | 3.3584 | 2.3215 | 49.5552 | 0.10556 | 4.3565e+01
|  |  |  | bot-5 | 33.0745 | 1.1229 | 6.4519 | 81.5309 | 280.706 | 8997 | 52.981 | 1.1251 | 28.5387 | 3.3584 | 2.3629 | 50.2227 | 0.10568 | 4.3565e+01
|  |  |  | bot-5 | 33.0746 | 1.1228 | 6.4488 | 81.4913 | 280.571 | 9044 | 53.041 | 1.125 | 28.5383 | 3.357 | 2.1609 | 49.6936 | 0.10609 | 4.3565e+01 | 912 | 3 | 1732 | 6.881 |  |  |  | 6.881
|  |  |  | 45 | 32.9372 | 1.4905 | 6.9019 | 87.9724 | 300.325 | 10848 | 45.563 | 1.4925 | 28.725 | 3.5789 | 2.0092 | 68.3027 | 0.097386 | 4.4988e+01 | 913 | 4 | 1752 | 7.156 |  |  |  | 7.156
|  |  |  | 40 | 32.9095 | 1.4273 | 6.9084 | 87.8933 | 300.61 | 12382 | 40.654 | 1.4291 | 28.6481 | 3.5783 | 2.6417 | 77.6975 | 0.089023 | 4.5744e+01 | 914 | 5 | 1709 | 7.182 |  |  |  | 7.182
|  |  |  | 35 | 32.8907 | 1.4597 | 6.8917 | 87.7428 | 299.889 | 13808 | 35.717 | 1.4612 | 28.6579 | 3.5751 | 1.9539 | 76.0068 | 0.08194 | 4.5744e+01 | 915 | 6 | 1745 | 7.147 |  |  |  | 7.147
|  |  |  | 30 | 32.8618 | 1.4977 | 7.0778 | 90.1834 | 307.996 | 15175 | 30.612 | 1.4991 | 28.6578 | 3.66 | 4.8508 | 81.1617 | 0.075446 | 4.7833e+01 | 916 | 7 | 1741, 1727 | 7.328 | 7.32 |  |  | 7.328
|  |  |  | 25 | 32.846 | 1.6461 | 8.2897 | 106.0195 | 360.742 | 16498 | 25.734 | 1.6472 | 28.7626 | 4.2001 | 4.937 | 64.1533 | 0.071567 | 4.7922e+01 | 917 | 8 | 1733 | 8.473 |  |  |  | 8.473
|  |  |  | 20 | 32.8155 | 1.655 | 7.6865 | 98.3069 | 334.501 | 18015 | 20.546 | 1.656 | 28.7589 | 3.9402 | 4.5829 | 71.6448 | 0.081229 | 4.7922e+01 | 918 | 9 | 1753 | 8.011 |  |  |  | 8.011
|  |  |  | 15 | 32.7288 | 2.3771 | 8.6429 | 112.5379 | 376.166 | 19401 | 15.517 | 2.3779 | 29.311 | 4.425 | 4.6722 | 71.0205 | 0.32897 | 4.8678e+01 | 919 | 10 | 1700 | 8.808 |  |  | b | 8.808 | stalled burette, OT
|  |  |  | 10 | 32.3367 | 4.747 | 8.7944 | 121.1543 | 382.956 | 20861 | 10.589 | 4.7477 | 30.9409 | 4.7141 | 2.8664 | 83.8426 | 0.82545 | 5.0011e+01 | 920 | 11 | 1696 | 9.08 |  |  |  | 9.08
|  |  |  | 5 | 32.3382 | 4.756 | 8.8596 | 122.0801 | 385.798 | 22460 | 5.604 | 4.7564 | 30.9472 | 4.7483 | 2.2822 | 83.5683 | 2.1701 | 5.0100e+01 | 921 | 12 | 1706 | 9.086 |  |  | b | 9.086 | stalled burette, OT
| 077 | 077 | CCL-1 | 40 | 31.8151 | 3.5159 | 8.0407 | 107.094 | 350.239 | 7299 | 39.736 | 3.5183 | 29.477 | 4.225 | 0.71017 | 79.5745 | 0.13488 | 3.3190e+02 | 922 | 1 | 1336 | 8.314 |  |  |  | 8.314
|  |  |  | 30 | 31.8114 | 3.5691 | 8.0712 | 107.6376 | 351.566 | 8971 | 30.625 | 3.571 | 29.4961 | 4.2477 | 0.81336 | 81.9518 | 0.32942 | 3.1149e+02 | 923 | 2 | 1135, 1094 | 8.321 | 8.318 |  |  | 8.321
|  |  |  | 25 | 31.6878 | 3.8415 | 8.0256 | 107.6679 | 349.623 | 9909 | 24.728 | 3.8431 | 29.6337 | 4.251 | 0.72176 | 81.5027 | 0.7758 | 2.9895e+02 | 924 | 3 | 1293 | 8.267 |  |  | d | 8.267 | leaky Niskin
|  |  |  | 20 | 31.6349 | 4.2482 | 7.9807 | 108.1095 | 347.698 | 10792 | 20.34 | 4.2495 | 29.9214 | 4.2671 | 0.72674 | 83.903 | 1.4866 | 2.9211e+02 | 925 | 4 | 1654 | 8.197 |  |  |  | 8.197
|  |  |  | 15 | 31.5978 | 5.0813 | 7.7651 | 107.3244 | 338.343 | 11682 | 13.901 | 5.0823 | 30.7623 | 4.2415 | 0.94137 | 86.5513 | 3.7865 | 2.8535e+02 | 926 | 5 | 1376 | 7.953 |  |  |  | 7.953
|  |  |  | 10 | 31.0928 | 7.3407 | 7.0881 | 103.0267 | 309.048 | 12788 | 9.492 | 7.3416 | 31.9055 | 4.0921 | 0.56759 | 86.8975 | 6.9501 | 2.8495e+02 | 927 | 6 | 1284 | 7.497 |  |  |  | 7.497
|  |  |  | 5 | 24.3959 | 9.4379 | 6.9033 | 100.8573 | 302.622 | 13712 | 5.231 | 9.4384 | 26.4904 | 4.0154 | 0.94403 | 81.9173 | 16.86 | 2.8535e+02 | 928 | 7 | 1143 | 7.174 |  |  |  | 7.174
| 078 | 078 | CCL-2 | 50 | 32.5806 | 1.6999 | 8.6279 | 110.3037 | 375.538 | 5223 | 50.727 | 1.7023 | 28.6163 | 4.3329 | 0.24335 | 71.0664 | 0.12578 | 6.4544e+02 | 929 | 1 | 1336, 1135 | 8.975 | 8.978 |  | b | 8.974 | stalled burette, 1-OT, 2-add titres
|  |  |  | 45 | 32.5799 | 1.6948 | 8.6224 | 110.2172 | 375.298 | 6614 | 44.65 | 1.6968 | 28.6092 | 4.3331 | 0.24358 | 73.2881 | 0.37193 | 6.3317e+02 | 930 | 2 | 1094, 1293 | 8.974 | 8.983 |  |  | 8.973 | 2-run 4 days later
|  |  |  | 40 | 32.5787 | 1.715 | 8.6245 | 110.2995 | 375.389 | 7957 | 40.253 | 1.7168 | 28.6132 | 4.338 | 0.23007 | 74.9443 | 0.77595 | 6.4450e+02 | 931 | 3 | 1654, 1699 | 8.984 | 8.986 |  |  | 8.983 | 2-run 4 days later
|  |  |  | 35 | 32.563 | 2.0037 | 8.6439 | 111.3595 | 376.245 | 9885 | 35.837 | 2.0055 | 28.8459 | 4.3755 | 0.18619 | 79.5285 | 1.6447 | 6.9291e+02 | 932 | 4 | 1284, 1143 | 9.009 | 9.023* |  | *b | 9.008 | 2-stalled burette, add titres, run 4 days later
|  |  |  | 30 | 32.5609 | 2.0184 | 8.6541 | 111.5311 | 376.692 | 11195 | 30.436 | 2.0198 | 28.8555 | 4.384 | 0.2073 | 79.1479 | 3.3786 | 6.8575e+02 | 933 | 5 | 1127, 1509 | 8.999* | 8.995 |  | *b | 8.998 | 1-stalled burette, OT, 2-run 4 days later
|  |  |  | 25 | 32.56 | 2.4563 | 9.0392 | 117.8034 | 393.469 | 12447 | 24.449 | 2.4576 | 29.2112 | 4.5991 | 0.18068 | 88.6973 | 5.8406 | 6.8615e+02 | 934 | 6 | 879, 1746 | 9.245* | 9.245 |  | *b | 9.245 | 1-stalled burette, add titres
|  |  |  | 20 | 32.5203 | 3.2721 | 8.7437 | 116.2898 | 380.64 | 13554 | 20.214 | 3.2732 | 29.8643 | 4.5477 | 0.18579 | 88.1453 | 8.3258 | 7.1189e+02 | 935 | 7 | 1145, 1748 | 9.081 | 9.091 |  |  | 9.08 | 2-run 4 days later
|  |  |  | 15 | 32.477 | 4.0824 | 8.7136 | 118.2135 | 379.375 | 14785 | 15.476 | 4.0834 | 30.5102 | 4.6133 | 0.22255 | 89.1589 | 12.87 | 7.4613e+02 | 936 | 8 | 1658, 1156 | 9.001 | 9 |  |  | 9 | 2-run 4 days later
|  |  |  | 10 | 31.0357 | 5.6349 | 7.5599 | 105.502 | 329.564 | 15938 | 9.673 | 5.6357 | 30.46 | 4.1802 | 0.59866 | 87.246 | 26.694 | 7.7066e+02 | 937 | 9 | 663, 1137 | 7.84 | 7.849 |  |  | 7.84 | 2-run 4 days later
|  |  |  | 5 | 29.1481 | 6.7644 | 7.2875 | 103.1915 | 318.19 | 16984 | 5.388 | 6.7648 | 29.7327 | 4.101 | 0.5735 | 86.2578 | 68.297 | 8.3312e+02 | 938 | 10 | 1023, 1022 | 7.531 | 7.533 |  |  | 7.53
| 079 | 079 | CCL-3 | 30 | 31.9796 | 3.5104 | 8.0105 | 106.7913 | 348.875 | 11532 | 30.659 | 3.5122 | 29.6077 | 4.2192 | 1.8557 | 73.6717 | 0.15054 | 1.3070e+02 | 939 | 1 | 1666 | 8.314 |  |  |  | 8.314
|  |  |  | 25 | 31.0553 | 4.0421 | 8.1369 | 109.245 | 354.652 | 13284 | 24.465 | 4.0436 | 29.2559 | 4.3041 | 0.69069 | 85.4505 | 0.29968 | 1.3505e+02 | 940 | 2 | 1155 | 8.392 |  |  |  | 8.392
|  |  |  | 20 | 30.7722 | 4.4578 | 7.9909 | 108.1868 | 348.379 | 14748 | 20.978 | 4.4591 | 29.326 | 4.2689 | 0.726 | 86.8587 | 0.46552 | 1.3723e+02 | 941 | 3 | 1136 | 8.195 |  |  |  | 8.195
|  |  |  | 15 | 26.7611 | 7.3779 | 7.1882 | 101.6804 | 314.455 | 16335 | 13.927 | 7.3791 | 27.9692 | 4.0442 | 0.82816 | 86.0801 | 1.6522 | 1.4159e+02 | 942 | 4 | 970, 1169 | 7.502 | 7.441 |  | c | 7.502 | bad reps
|  |  |  | 10 | 26.0074 | 7.8807 | 7.0055 | 99.7786 | 306.658 | 18254 | 10.111 | 7.8816 | 27.6017 | 3.9803 | 0.75445 | 85.8402 | 2.7089 | 1.1931e+02 | 943 | 5 | 1724 | 7.296 |  |  |  | 7.296
|  |  |  | 5 | 25.9962 | 7.8875 | 7.1043 | 101.1941 | 310.988 | 19478 | 4.743 | 7.8879 | 27.5914 | 4.0307 | 0.76006 | 85.7967 | 10.828 | 1.1109e+02 | 944 | 6 | 1089 | 7.28 |  |  |  | 7.28
| 080 | 080 | CCL-4 | b-5 | 790:32 | -1.6987 | 7.1368 | 83.5375 | 310.499 | 8762 | 45.39 | -1.6979 | 26.1219 | 3.4349 | 0.94296 | 84.0624 | 0.097069 | 296.24 | 945 | 1 | 1730 | 7.474 |  |  |  | 7.474
|  |  |  | 40 | 790:30 | -1.6968 | 7.2147 | 84.4538 | 313.892 | 10521 | 39.703 | -1.6961 | 26.1202 | 3.4682 | 1.0654 | 83.6415 | 0.083727 | 285.35 | 946 | 2 | 1742 | 7.461 |  |  |  | 7.461
|  |  |  | 35 | 790:30 | -1.6984 | 7.1925 | 84.1897 | 312.924 | 12533 | 35.619 | -1.6978 | 26.1179 | 3.4608 | 1.0458 | 84.3425 | 0.078035 | 315.85 | 947 | 3 | 1743 | 7.473 |  |  |  | 7.473
|  |  |  | 30 | 790:27 | -1.7033 | 7.1743 | 83.964 | 312.133 | 13959 | 29.947 | -1.7028 | 26.1103 | 3.4554 | 0.93823 | 84.7179 | 0.10484 | 300.6 | 948 | 4 | 1739 | 7.468 |  |  |  | 7.468
|  |  |  | 25 | 790:19 | -1.7055 | 7.209 | 84.3615 | 313.644 | 15288 | 25.222 | -1.7051 | 26.1004 | 3.4707 | 0.90673 | 85.7201 | 0.15507 | 295.71 | 949 | 5 | 1725 | 7.481 |  |  |  | 7.481
|  |  |  | 20 | 786:50 | -1.7224 | 7.4727 | 87.3183 | 325.152 | 16560 | 20.329 | -1.7221 | 25.9818 | 3.5735 | 0.78299 | 87.3274 | 0.21944 | 296.24 | 950 | 6 | 1731, 1193 | 7.779 | 7.778 |  |  | 7.779
|  |  |  | 15 | 782:51 | -1.5594 | 8.5233 | 99.927 | 370.916 | 18104 | 15.089 | -1.5592 | 25.9928 | 4.0055 | 4.9404 | 74.0284 | 0.7563 | 295.84 | 951 | 7 | 1750 | 8.946 |  |  |  | 8.946
|  |  |  | 20 | 736:54 | 0.9119 | 9.9719 | 123.3133 | 434.653 | 19436 | 10.301 | 0.9122 | 26.5043 | 4.7989 | 4.046 | 77.9772 | 8.7855 | 298.42 | 952 | 8 | 1744 | 10.257 |  |  |  | 10.257
|  |  |  | 5 | 732:58 | 1.1289 | 8.8111 | 109.4596 | 384.112 | 20976 | 4.31 | 1.1291 | 26.5129 | 4.3291 | 2.5887 | 81.9489 | 28.751 | 283.48 | 953 | 9 | 1740 | 9.037 |  |  |  | 9.037
|  |  |  | 15 | 784:31 | -1.6269 | 8.1414 | 95.3185 | 354.277 | 22718 | 15.766 | -1.6266 | 25.9855 | 3.8481 | 4.6813 | 81.3241 | 0.33058 | 285.35
|  |  |  | 15 | 784:31 | -1.6272 | 8.1461 | 95.3735 | 354.484 | 22760 | 15.804 | -1.6269 | 25.9857 | 3.8499 | 4.3587 | 81.158 | 0.33496 | 285.35
|  |  |  | 15 | 784:26 | -1.6265 | 8.1411 | 95.3146 | 354.268 | 22806 | 15.821 | -1.6262 | 25.9852 | 3.8479 | 4.3018 | 81.4302 | 0.33868 | 285.35
|  |  | Caro | 15 | 784:13 | -1.6201 | 8.132 | 95.2186 | 353.873 | 22885 | 15.792 | -1.6198 | 25.9834 | 3.8446 | 4.5143 | 81.3882 | 0.35847 | 285.35
|  |  |  | 15 | 32.6748 | -1.6186 | 8.1305 | 95.2042 | 353.807 | 22954 | 15.739 | -1.6183 | 25.9839 | 3.8442 | 4.6538 | 81.2 | 0.39902 | 2.8535e+02
|  |  |  | 15 | 32.6731 | -1.6114 | 8.1511 | 95.4632 | 354.704 | 23018 | 15.682 | -1.6111 | 25.9856 | 3.853 | 4.6173 | 81.3691 | 0.45974 | 2.8535e+02
|  |  |  | 15 | 32.6725 | -1.6066 | 8.1645 | 95.6322 | 355.287 | 23080 | 15.678 | -1.6063 | 25.9888 | 3.8588 | 4.8011 | 81.0081 | 0.51453 | 2.8535e+02
|  |  |  | 15 | 32.6718 | -1.6072 | 8.178 | 95.7883 | 355.875 | 23143 | 15.726 | -1.6069 | 25.9901 | 3.8641 | 4.8075 | 80.9913 | 0.55583 | 2.8535e+02
|  |  |  | 15 | 32.6739 | -1.6114 | 8.1856 | 95.8681 | 356.207 | 23197 | 15.767 | -1.6111 | 25.9894 | 3.8668 | 4.6422 | 80.5321 | 0.57236 | 2.8424e+02
|  |  |  | 15 | 32.6764 | -1.6161 | 8.1947 | 95.9638 | 356.602 | 23256 | 15.787 | -1.6159 | 25.9882 | 3.8701 | 4.503 | 80.8941 | 0.55211 | 2.8321e+02
|  |  |  | 15 | 32.6768 | -1.615 | 8.1906 | 95.9188 | 356.422 | 23318 | 15.748 | -1.6147 | 25.987 | 3.8686 | 4.7003 | 80.8845 | 0.50999 | 2.8317e+02
|  |  |  | 15 | 32.6773 | -1.6169 | 8.1913 | 95.9228 | 356.454 | 23377 | 15.675 | -1.6166 | 25.987 | 3.8687 | 4.7823 | 80.9846 | 0.47975 | 2.8317e+02
|  |  |  | 15 | 32.6746 | -1.6136 | 8.1855 | 95.8619 | 356.203 | 23447 | 15.637 | -1.6133 | 25.9869 | 3.8666 | 4.9094 | 80.7868 | 0.44125 | 2.8317e+02
|  |  |  | 15 | 32.6782 | -1.6168 | 8.1699 | 95.6724 | 355.521 | 23505 | 15.682 | -1.6166 | 25.9869 | 3.8602 | 4.7429 | 80.7543 | 0.39566 | 2.8317e+02
|  |  |  | 15 | 32.6835 | -1.6254 | 8.1703 | 95.6584 | 355.538 | 23566 | 15.771 | -1.6251 | 25.9855 | 3.8597 | 4.4756 | 81.383 | 0.36048 | 2.8317e+02
| 081 | 081 | CCL-4 | 40 | 32.8695 | -1.5749 | 7.222 | 84.7835 | 314.227 | 8317 | 40.093 | -1.5742 | 26.1694 | 3.4791 | 1.5336 | 83.2345 | 0.1008 | 5.1718e+02
|  |  |  | 40 | 32.8707 | -1.5778 | 7.2271 | 84.8364 | 314.445 | 8383 | 40.061 | -1.577 | 26.1695 | 3.4809 | 2.2643 | 83.3265 | 0.10037 | 5.1847e+02
|  |  |  | 30 | 32.8523 | -1.5422 | 7.2708 | 85.4221 | 316.353 | 9734 | 30.683 | -1.5416 | 26.1793 | 3.5044 | 1.9259 | 83.4467 | 0.18999 | 5.1287e+02
|  |  |  | 30 | 32.8536 | -1.5413 | 7.2747 | 85.4708 | 316.523 | 9789 | 30.719 | -1.5408 | 26.1798 | 3.506 | 1.9536 | 83.5702 | 0.1913 | 5.1189e+02
|  |  |  | 15 | 32.3307 | -1.1287 | 9.6115 | 113.7959 | 418.373 | 11097 | 15.362 | -1.1284 | 26.0981 | 4.4774 | 4.9387 | 71.6288 | 5.3634 | 5.3803e+02
|  |  |  | 15 | 32.3488 | -1.1135 | 9.6248 | 114.015 | 418.948 | 11144 | 15.418 | -1.1132 | 26.1024 | 4.4848 | 4.9387 | 72.1925 | 5.2813 | 5.3962e+02
|  |  |  | 15 | 32.3512 | -1.1242 | 9.6437 | 114.2079 | 419.77 | 11212 | 15.474 | -1.1238 | 26.1402 | 4.4914 | 4.9384 | 70.8668 | 5.1975 | 5.4217e+02
|  |  |  | 15 | 32.3675 | -1.16 | 9.6486 | 114.1661 | 419.975 | 11290 | 15.418 | -1.1597 | 26.1433 | 4.49 | 4.9387 | 71.9723 | 5.3776 | 5.4484e+02
|  |  |  | 8 | 29.8524 | 1.7768 | 8.3179 | 104.6036 | 362.82 | 13136 | 7.871 | 1.7771 | 26.4703 | 4.1602 | 1.0788 | 83.404 | 34.146 | 5.1843e+02
|  |  |  | 8 | 29.8637 | 1.7573 | 8.2987 | 104.3174 | 361.979 | 13183 | 7.875 | 1.7576 | 26.4597 | 4.1505 | 1.2295 | 83.2685 | 34.17 | 5.1852e+02
|  |  |  | 4 | 29.8459 | 1.7893 | 8.368 | 105.2623 | 365.007 | 14712 | 4.116 | 1.7894 | 26.4753 | 4.1844 | 1.0761 | 82.7679 | 70.837 | 5.7288e+02
|  |  |  | 4 | 29.8426 | 1.792 | 8.3751 | 105.3565 | 365.316 | 14755 | 4.131 | 1.7922 | 26.474 | 4.1876 | 1.2597 | 83.0672 | 70.074 | 5.7288e+02
|  |  |  | 4 | 29.8355 | 1.7951 | 8.3816 | 105.4422 | 365.604 | 14799 | 4.139 | 1.7952 | 26.4775 | 4.1906 | 0.99273 | 83.1668 | 69.451 | 5.7200e+02
|  |  |  | 4 | 29.8414 | 1.792 | 8.3879 | 105.5169 | 365.876 | 14844 | 4.125 | 1.7922 | 26.4749 | 4.1931 | 0.95474 | 82.9875 | 69.124 | 5.7030e+02
|  |  |  | 2 | 29.8043 | 1.8067 | 8.3724 | 105.3357 | 365.211 | 16450 | 3.384 | 1.8069 | 26.4562 | 4.1872 | 0.68551 | 82.9145 | 78.173 | 5.3532e+02
|  |  |  | 2 | 29.8052 | 1.8063 | 8.3719 | 105.329 | 365.189 | 16493 | 3.325 | 1.8064 | 26.4551 | 4.187 | 0.92443 | 82.9977 | 77.698 | 5.3377e+02
| 082 | 082 | BC-1 | b-x | 32.9237 | -0.7805 | 7.9016 | 94.8312 | 343.788 | 8697 | 81.445 | -0.7784 | 26.8652 | 3.801 | 3.1692 | 80.9937 | 0.08522 | 1.1545e+02 | 970 | 1 | 1752 | 8.23 |  |  |  | 8.23
|  |  |  | 70 | 32.9233 | -0.7723 | 7.9171 | 95.0378 | 344.464 | 10388 | 70.276 | -0.7705 | 26.8694 | 3.8128 | 2.2852 | 81.099 | 0.081075 | 1.2145e+02 | 971 | 2 | 1709 | 8.237 |  |  |  | 8.237
|  |  |  | 60 | 32.9202 | -0.7412 | 7.9742 | 95.8009 | 346.949 | 11598 | 60.082 | -0.7396 | 26.8876 | 3.843 | 2.2335 | 81.2949 | 0.079308 | 1.2198e+02 | 972 | 3 | 1745, 1741 | 8.292 | 8.289 |  |  | 8.292
|  |  |  | 50 | 32.9188 | -0.7269 | 7.9729 | 95.8205 | 346.892 | 12771 | 50.109 | -0.7256 | 26.8907 | 3.8481 | 2.5439 | 81.4452 | 0.078963 | 1.2198e+02 | 973 | 4 | 1727 | 8.297 |  |  |  | 8.297
|  |  |  | 40 | 32.9179 | -0.7218 | 7.9955 | 96.1046 | 347.878 | 14010 | 40.124 | -0.7207 | 26.8927 | 3.8621 | 2.4427 | 81.5941 | 0.078913 | 1.2131e+02 | 974 | 5 | 1733 | 8.315 |  |  |  | 8.315
|  |  |  | 35 | 32.914 | -0.7645 | 7.9716 | 95.7027 | 346.835 | 15058 | 34.896 | -0.7636 | 26.855 | 3.8508 | 2.4124 | 80.9702 | 0.080724 | 1.1980e+02 | 975 | 6 | 1753 | 8.305 |  |  |  | 8.305
|  |  |  | 30 | 32.9091 | -0.7741 | 7.9787 | 95.7603 | 347.148 | 16079 | 29.767 | -0.7734 | 26.829 | 3.855 | 2.444 | 80.7571 | 0.10646 | 1.2416e+02 | 976 | 7 | 1700 | 8.305 |  |  |  | 8.305
|  |  |  | 25 | 32.9131 | -0.743 | 7.987 | 95.942 | 347.505 | 17140 | 24.816 | -0.7423 | 26.8121 | 3.8633 | 2.609 | 80.5465 | 0.20887 | 1.2852e+02 | 977 | 8 | 1696 | 8.331 |  |  |  | 8.331
|  |  |  | 20 | 32.8662 | -0.9304 | 8.1989 | 97.9571 | 356.737 | 18210 | 20.592 | -0.93 | 26.6791 | 3.9343 | 3.5923 | 78.5699 | 0.49279 | 1.2634e+02 | 978 | 9 | 1706 | 8.5 |  |  |  | 8.5
|  |  |  | 15 | 31.5056 | -0.523 | 10.2954 | 123.2122 | 448.444 | 19780 | 15.18 | -0.5226 | 25.9807 | 4.7969 | 4.1659 | 81.5443 | 2.3154 | 1.1545e+02 | 979 | 10 | 1351 | 10.423 |  |  |  | 10.423
|  |  |  | 10 | 29.6103 | 3.4295 | 7.6275 | 99.8886 | 332.808 | 21275 | 10.933 | 3.4301 | 27.5405 | 3.9942 | 0.55354 | 86.7208 | 3.6504 | 1.0238e+02 | 980 | 11 | 1734, 1751 | 8.112 | 8.1 |  |  | 8.112
|  |  |  | 5 | 29.3542 | 2.8001 | 7.6979 | 99.0495 | 335.929 | 22888 | 5.023 | 2.8004 | 26.8427 | 3.9698 | 0.44442 | 86.5369 | 5.6376 | 7.8418e+01 | 981 | 12 | 1423 | 7.942 |  |  |  | 7.942
|  |  |  | chlaM | 32.1702 | -1.2208 | 10.389 | 122.5571 | 452.274 | 25192 | 17.942 | -1.2204 | 25.9363 | 4.775 | 4.9385 | 72.115 | 0.93673 | 6.7526e+01
| 083 | 083 | BC-2 | b- | 33.1687 | -1.4413 | 6.7961 | 80.2452 | 295.624 | 12923 | 120.843 | -1.4387 | 26.5331 | 3.2949 | 19.342 | 80.8055 | 0.10217 | 5.0100e+02 | 982 | 1 | 1737 | 7.051 |  |  |  | 7.051
|  |  |  | 100 | 33.1618 | -1.3975 | 6.8107 | 80.5098 | 296.263 | 15086 | 99.6 | -1.3954 | 26.5533 | 3.3115 | 24.337 | 80.9085 | 0.087683 | 5.0754e+02 | 983 | 2 | 1723 | 7.055 |  |  |  | 7.055
|  |  |  | 75 | 33.1563 | -1.3818 | 6.8165 | 80.6081 | 296.515 | 17087 | 75.461 | -1.3803 | 26.5518 | 3.3236 | 16.616 | 81.8287 | 0.085308 | 5.1189e+02 | 984 | 3 | 1729 | 7.1 |  |  |  | 7.1
|  |  |  | 50 | 33.1534 | -1.4032 | 6.8789 | 81.2959 | 299.231 | 19579 | 49.86 | -1.4022 | 26.5183 | 3.3565 | 20.562 | 82.049 | 0.082538 | 5.1843e+02 | 985 | 4 | 1738, 1735 | 7.112 | 7.115 |  |  | 7.112
|  |  |  | 44 | 33.0878 | -1.5247 | 6.7765 | 79.7832 | 294.792 | 21270 | 43.787 | -1.5239 | 26.4344 | 3.3076 | 39.31 | 78.1615 | 0.11383 | 5.2331e+02 | 986 | 5
|  |  |  | 44 | 33.0794 | -1.5259 | 6.7827 | 79.8484 | 295.061 | 21322 | 43.792 | -1.525 | 26.4276 | 3.3098 | 79.657 | 77.7497 | 0.11807 | 5.2367e+02 | 987 | 6
|  |  |  | 30 | 32.5092 | -1.6857 | 7.2763 | 84.9476 | 316.676 | 23612 | 30.246 | -1.6852 | 25.8172 | 3.4887 | 35.265 | 82.852 | 0.58968 | 5.3150e+02 | 988 | 7
|  |  |  | 30 | 32.5092 | -1.6857 | 7.2716 | 84.8927 | 316.472 | 23653 | 30.241 | -1.6852 | 25.8173 | 3.4869 | 43.386 | 82.749 | 0.59128 | 5.3150e+02 | 989 | 8
|  |  |  | 30 | 32.5104 | -1.686 | 7.2675 | 84.8453 | 316.294 | 23699 | 30.251 | -1.6855 | 25.8178 | 3.4852 | 65.736 | 83.2299 | 0.5904 | 5.3150e+02 | 990 | 9 | 1698 | 6.75 |  |  |  | 6.75
|  |  |  | 20 | 31.9714 | -1.5413 | 9.6601 | 112.807 | 420.606 | 25081 | 20.538 | -1.5409 | 25.5355 | 4.4421 | 89.42 | 81.6626 | 3.374 | 5.3585e+02 | 991 | 10 | 1704 | 9.344 |  |  |  | 9.344
|  |  |  | 18 | 31.8873 | -1.499 | 10.43 | 121.8676 | 454.156 | 26194 | 18.553 | -1.4987 | 25.5066 | 4.752 | 126.71 | 77.3795 | 5.698 | 5.4021e+02 | 992 | 11
|  |  |  | 18 | 31.8867 | -1.4987 | 10.4416 | 122.0046 | 454.665 | 26239 | 18.549 | -1.4984 | 25.5064 | 4.7567 | 116.37 | 78.1876 | 5.6711 | 5.4021e+02 | 993 | 12
|  |  |  | 15 | 31.8022 | -1.4312 | 10.5279 | 123.1686 | 458.451 | 27375 | 15.292 | -1.4309 | 25.4959 | 4.798 | 65.932 | 79.1229 | 11.882 | 5.4457e+02 | 994 | 13 | 1152 | 11.161 |  |  |  | 11.161
|  |  |  | 10 | 30.9722 | -0.417 | 10.2297 | 122.3273 | 445.771 | 28497 | 10.305 | -0.4168 | 25.7459 | 4.7691 | 11.429 | 87.8206 | 26.557 | 5.4892e+02 | 995 | 14
|  |  |  | 10 | 31.0384 | -0.5901 | 10.3499 | 123.2435 | 450.982 | 28806 | 10.293 | -0.5898 | 25.5814 | 4.801 | 11.521 | 86.9248 | 27.094 | 5.5110e+02 | 996 | 15
|  |  |  | 10 | 30.9555 | -0.3943 | 10.3009 | 123.24 | 448.878 | 28937 | 10.297 | -0.394 | 25.6727 | 4.8002 | 29.129 | 87.0057 | 27.567 | 5.5137e+02 | 997 | 16 | 1697 | 11.374 |  |  |  | 11.374
|  |  |  | 5 | 29.6145 | 3.8049 | 7.6635 | 101.3093 | 334.388 | 30167 | 5.848 | 3.8052 | 27.8332 | 4.044 | 8.0027 | 87.5767 | 52.1 | 5.5787e+02 | 998 | 17
|  |  |  | 5 | 29.6089 | 3.8155 | 7.6507 | 101.1621 | 333.828 | 30206 | 5.851 | 3.8159 | 27.8318 | 4.039 | 7.9038 | 87.9925 | 52.863 | 558.51 | 999 | 18
|  |  |  | 5 | 29.6098 | 3.816 | 7.6441 | 101.0767 | 333.54 | 30250 | 5.846 | 3.8163 | 27.8392 | 4.036 | 7.1909 | 87.8715 | 53.818 | 559.13 | 1000 | 19 | 1336 | 8.111 |  |  |  | 8.111
|  |  |  | surface | 29.6048 | 3.8303 | 7.6845 | 101.6437 | 335.306 | 31603 | 2.338 | 3.8305 | 27.8469 | 4.057 | 6.7287 | 87.678 | 105.99 | 566.35 | 1001 | 20
|  |  |  | surface | 29.6049 | 3.8287 | 7.6856 | 101.6543 | 335.354 | 31647 | 2.339 | 3.8289 | 27.8467 | 4.0573 | 7.0925 | 87.7282 | 103.23 | 566.53 | 1002 | 21
| 084 | 084 | BC-3 | b- | 34.8368 | 0.5323 | 6.3645 | 80.1574 | 276.511 | 8152 | 182.304 | 0.5399 | 29.4587 | 3.266 | 5.9969 | 81.5261 | 0.20896 | zz | 1003 | 1 | 663 | 6.584 |  |  |  | 6.584
|  |  |  | 150 | 34.8141 | 0.513 | 6.3534 | 79.9625 | 276.036 | 10560 | 150.768 | 0.5192 | 29.4091 | 3.2709 | 2.7637 | 86.4187 | 0.11986 | 448.41 | 1004 | 2 | 879 | 6.587 |  |  |  | 6.587
|  |  |  | 100 | 33.2355 | -1.3604 | 6.8543 | 81.149 | 298.143 | 13884 | 99.416 | -1.3582 | 26.6366 | 3.333 | 23.614 | 81.9824 | 0.11348 | 452.95 | 1005 | 3 | 970 | 7.108 |  |  |  | 7.108
|  |  |  | 75 | 33.1189 | -1.132 | 7.2942 | 86.8256 | 317.307 | 15838 | 75.516 | -1.1303 | 26.7253 | 3.5335 | 26.112 | 82.9228 | 0.12259 | 452.9 | 1006 | 4 | 1023, 1022 | 7.523 | 7.535 |  | c | 7.523 | reps a bit off
|  |  |  | 50 | 33.0091 | -0.9181 | 7.963 | 95.2666 | 346.435 | 18233 | 51.015 | -0.9168 | 26.8075 | 3.8293 | 25.356 | 83.2082 | 0.10916 | 451.88 | 1007 | 5 | 1089 | 8.265 |  |  |  | 8.265
|  |  |  | 25 | 32.1414 | -1.5995 | 8.8095 | 102.8296 | 383.516 | 20644 | 25.643 | -1.5991 | 25.616 | 4.0996 | 85.111 | 80.9051 | 3.0144 | 442.19 | 1008 | 6 | 1094 | 9.052 |  |  | b | 9.052 | floaties
|  |  |  | 15 | 28.6424 | 2.4819 | 7.9686 | 101.2191 | 347.926 | 23749 | 5.964 | 2.4821 | 25.9915 | 4.0441 | 8.413 | 87.1517 | 41.289 | 432.23 | 1009 | 7 | 1127 | 8.366 |  |  |  | 8.366
|  |  |  | 5` | 28.6554 | 2.5093 | 7.9701 | 101.3182 | 347.989 | 24347 | 5.957 | 2.5096 | 26.114 | 4.0474 | 7.5361 | 87.156 | 43.314 | 429.12 | 1010 | 8 | 1135 | 8.327 |  |  |  | 8.327
| 085 | 085 | BC-4 | b- | 34.847 | 0.5001 | 6.4796 | 81.5841 | 281.51 | 13075 | 582.173 | 0.5264 | 29.6365 | 3.1707 | 1.6237 | 89.5128 | 0.16143 | 259.21 | 1011 | 1 | 1336 | 6.725 |  |  |  | 6.725
|  |  |  | 500 | 34.8457 | 0.5217 | 6.4687 | 81.4832 | 281.035 | 16865 | 500.435 | 0.544 | 29.6141 | 3.196 | 1.5983 | 89.7225 | 0.18272 | 243.88 | 1012 | 2 | 1697, 1152 | 6.733 | 6.732 |  |  | 6.733
|  |  |  | 400 | 34.8387 | 0.5756 | 6.4548 | 81.4102 | 280.436 | 20055 | 400.766 | 0.5932 | 29.6067 | 3.2287 | 1.7201 | 90.5092 | 0.16537 | 224.23 | 1013 | 3 | 1704 | 6.712 |  |  |  | 6.712
|  |  |  | 300 | 34.8181 | 0.6599 | 6.3714 | 80.5158 | 276.82 | 23476 | 300.878 | 0.6731 | 29.6143 | 3.2351 | 1.7695 | 90.3919 | 0.12091 | 200.84 | 1014 | 4 | 1698 | 6.63 |  |  |  | 6.63
|  |  |  | 200 | 34.7572 | 0.5508 | 6.222 | 78.36 | 270.337 | 26721 | 201.411 | 0.5592 | 29.4252 | 3.1993 | 1.9885 | 90.2118 | 0.10356 | 189.51 | 1015 | 5 | 1735 | 6.461 |  |  |  | 6.461
|  |  |  | 150 | 34.5428 | -0.052 | 6.0713 | 75.1375 | 263.824 | 29380 | 150.361 | -0.0466 | 28.7185 | 3.1106 | 9.5395 | 89.7522 | 0.1013 | 191.69 | 1016 | 6 | 1738 | 6.312 |  |  |  | 6.312
|  |  |  | 100 | 33.2455 | -1.3021 | 7.1946 | 85.3193 | 312.942 | 31863 | 100.267 | -1.2999 | 26.6931 | 3.4732 | 9.1463 | 81.4127 | 0.099962 | 198.22 | 1017 | 7 | 1729 | 7.487 |  |  |  | 7.487
|  |  |  | 75 | 33.1578 | -0.9487 | 7.7328 | 92.531 | 336.378 | 33923 | 75.132 | -0.9468 | 26.9221 | 3.7263 | 11.502 | 82.7883 | 0.095592 | 196.04 | 1018 | 8 | 1723 | 8.087 |  |  |  | 8.087
|  |  |  | 50 | 32.963 | -0.766 | 8.1107 | 97.4032 | 352.875 | 37383 | 51.636 | -0.7647 | 26.9072 | 3.9012 | 12.994 | 83.1029 | 0.11249 | 198.22 | 1019 | 9 | 1737 | 7.683 |  |  | c | 7.683 | value is VERY different from CTD profile
|  |  |  | 25 | 31.6922 | -1.5681 | 8.7142 | 101.49 | 379.504 | 39067 | 24.373 | -1.5677 | 25.312 | 4.0545 | 39.487 | 89.1541 | 0.89152 | 200.4 | 1020 | 10 | 1423, 1751 | 9.53 | 9.599* |  | *e | 9.53 | big air bubble in sample
|  |  |  | 15 | 30.3222 | -1.3977 | 9.7534 | 113.052 | 425.224 | 40278 | 13.2 | -1.3975 | 24.4178 | 4.4539 | 26.951 | 88.8434 | 3.1853 | 204.45 | 1021 | 11 | 1734 | 10.148 |  |  |  | 10.148
|  |  |  | 5 | 27.4283 | 1.2167 | 8.4678 | 103.238 | 370.049 | 41608 | 5.609 | 1.2169 | 24.0968 | 4.1161 | 10.875 | 85.099 | 13.414 | 204.76 | 1022 | 12 | 1351 | 8.953 |  |  | b | 8.953 | stalled burette, OT
|  |  |  | 5 | 27.42 | 1.2134 | 8.4819 | 103.3957 | 370.669 | 41709 | 5.61 | 1.2136 | 24.0894 | 4.1215 | 12.701 | 85.0392 | 13.567 | 204.76
|  |  |  | 5 | 27.4242 | 1.224 | 8.4715 | 103.3001 | 370.212 | 41800 | 5.67 | 1.2242 | 24.096 | 4.1182 | 17.385 | 84.9994 | 13.396 | 204.76
| 086 | 086 | BL-1 |  | 31.9967 | -1.5012 | 8.9213 | 104.3119 | 388.428 | 12885 | 22.832 | -1.5008 | 25.5865 | 4.1512 | 8.2392 | 89.8384 | 2.9278 | 625.77
[5 empty] | 31.9964 | -1.5008 | 8.9309 | 104.4253 | 388.846 | 12924 | 22.827 | -1.5004 | 25.5865 | 4.1551 | 25.849 | 89.8484 | 2.9294 | 776.89
[5 empty] | 31.9959 | -1.5002 | 8.934 | 104.4638 | 388.985 | 12956 | 22.823 | -1.4998 | 25.5865 | 4.1564 | 22.964 | 89.8699 | 2.9215 | 822.56
[5 empty] | 31.9961 | -1.5006 | 8.9365 | 104.4916 | 389.092 | 12987 | 22.829 | -1.5002 | 25.5865 | 4.1574 | 24.365 | 89.6293 | 2.9161 | 826.02
[5 empty] | 31.9965 | -1.5008 | 8.9405 | 104.5382 | 389.267 | 13018 | 22.829 | -1.5004 | 25.5863 | 4.159 | 7.4542 | 89.8224 | 2.9162 | 829.37
[5 empty] | 31.9964 | -1.5008 | 8.9455 | 104.5969 | 389.485 | 13048 | 22.829 | -1.5004 | 25.5865 | 4.161 | 8.6336 | 89.7502 | 2.9105 | 831.11
[5 empty] | 31.996 | -1.5004 | 8.9366 | 104.4931 | 389.096 | 13083 | 22.832 | -1.5 | 25.5866 | 4.1574 | 25.177 | 89.9842 | 2.909 | 832.45
[5 empty] | 31.9955 | -1.4994 | 8.9205 | 104.3076 | 388.397 | 13117 | 22.84 | -1.4991 | 25.5867 | 4.1511 | 16.14 | 89.9111 | 2.91 | 833.63
[5 empty] | 31.9967 | -1.5008 | 8.9157 | 104.2478 | 388.184 | 13147 | 22.847 | -1.5004 | 25.5867 | 4.1491 | 3.8875 | 89.6363 | 2.9168 | 834.28
[5 empty] | 31.9961 | -1.5002 | 8.9087 | 104.1678 | 387.882 | 13178 | 22.84 | -1.4998 | 25.5867 | 4.1463 | 30.004 | 89.8743 | 2.9059 | 834.64
[5 empty] | 31.9958 | -1.4998 | 8.9086 | 104.1679 | 387.879 | 13208 | 22.832 | -1.4994 | 25.5867 | 4.1463 | 24.52 | 89.8459 | 2.8953 | 835.33
[5 empty] | 31.9951 | -1.4987 | 8.916 | 104.256 | 388.199 | 13239 | 22.822 | -1.4984 | 25.587 | 4.1493 | 45.676 | 89.7257 | 2.8945 | 836.45
[5 empty] | 31.9947 | -1.4981 | 8.9234 | 104.3444 | 388.522 | 13269 | 22.819 | -1.4978 | 25.5871 | 4.1523 | 10.58 | 89.8579 | 2.9032 | 836.45
[5 empty] | 31.9942 | -1.4973 | 8.9268 | 104.3858 | 388.669 | 13303 | 22.819 | -1.4969 | 25.5871 | 4.1537 | 21.529 | 89.7303 | 2.9021 | 836.45
[5 empty] | 31.9945 | -1.4975 | 8.9308 | 104.4327 | 388.844 | 13345 | 22.811 | -1.4971 | 25.5872 | 4.1554 | 8.1202 | 89.8668 | 2.9026 | 836.45
[5 empty] | 31.993 | -1.4951 | 8.9277 | 104.4027 | 388.712 | 13385 | 22.805 | -1.4947 | 25.5876 | 4.1543 | 11.765 | 89.8519 | 2.907 | 836.32
[5 empty] | 31.9922 | -1.4939 | 8.9242 | 104.3641 | 388.558 | 13418 | 22.807 | -1.4935 | 25.5881 | 4.153 | 11.488 | 89.9798 | 2.9013 | 836.45
[5 empty] | 31.9921 | -1.4935 | 8.9261 | 104.3878 | 388.641 | 13450 | 22.816 | -1.4931 | 25.588 | 4.1538 | 7.4891 | 89.8711 | 2.9034 | 835.92
[5 empty] | 31.993 | -1.4946 | 8.9228 | 104.3456 | 388.494 | 13484 | 22.818 | -1.4942 | 25.588 | 4.1524 | 17.958 | 90.0281 | 2.9063 | 836.45
[5 empty] | 31.9929 | -1.4942 | 8.9233 | 104.353 | 388.517 | 13517 | 22.817 | -1.4938 | 25.5883 | 4.1526 | 6.0165 | 89.8508 | 2.9012 | 836.31
[5 empty] | 31.991 | -1.4916 | 8.9193 | 104.3126 | 388.345 | 13551 | 22.822 | -1.4912 | 25.5882 | 4.1512 | 10.055 | 89.9354 | 2.9022 | 836.1
[5 empty] | 31.9905 | -1.4909 | 8.9247 | 104.3772 | 388.579 | 13588 | 22.824 | -1.4905 | 25.5883 | 4.1534 | 7.4348 | 89.8331 | 2.8985 | 835.91
[5 empty] | 31.9882 | -1.488 | 8.9387 | 104.5475 | 389.189 | 13630 | 22.816 | -1.4876 | 25.5891 | 4.1593 | 24.146 | 89.8349 | 2.8982 | 835.23
[5 empty] | 31.9847 | -1.4832 | 8.9361 | 104.5284 | 389.077 | 13684 | 22.804 | -1.4828 | 25.59 | 4.1586 | 3.5299 | 85.4198 | 2.891 | 835.87
| 087 | 087 | BL-1 | b-5 | 32.4559 | -1.5822 | 8.0648 | 94.3885 | 351.009 | 4305 | 58.67 | -1.5812 | 25.8726 | 3.7977 | 70.412 | 60.1509 | 0.13069 | 489.8 | 1025 | 1 | 1706 | 8.402 |  |  |  | 8.402
|  |  |  | 50 | 32.4534 | -1.584 | 8.0677 | 94.4162 | 351.137 | 6400 | 51.059 | -1.5831 | 25.8654 | 3.8019 | 41.685 | 75.7052 | 0.27243 | 455.7 | 1026 | 2 | 1696 | 8.385 |  |  |  | 8.385
|  |  |  | 40 | 32.2894 | -1.6562 | 8.2798 | 96.5953 | 360.414 | 8488 | 40.382 | -1.6556 | 25.6845 | 3.8808 | 8.5826 | 89.53 | 0.95874 | 454.5 | 1027 | 3 | 1700 | 8.477 |  |  |  | 8.477
|  |  |  | 30 | 32.1449 | -1.6127 | 8.2951 | 96.7931 | 361.123 | 10275 | 30.793 | -1.6122 | 25.6101 | 3.8917 | 4.5189 | 89.6056 | 1.9589 | 459.43 | 1028 | 4 | 1753 | 8.622 |  |  |  | 8.622
|  |  |  | 20 | 31.7053 | -1.2829 | 9.189 | 107.871 | 400.18 | 12434 | 20.914 | -1.2825 | 25.5388 | 4.2729 | 4.2229 | 90.3727 | 4.1385 | 468.86 | 1029 | 5 | 1733, 1727, 1741 | 9.169* | 9.168* | 9.18 | d* | 9.18 | stalled burette, 1,2-added titres, used dosimat values
|  |  |  | 10 | 29.94 | -0.1409 | 9.8496 | 117.8261 | 429.563 | 13993 | 11.463 | -0.1406 | 25.0871 | 4.6142 | 3.8932 | 90.5977 | 8.7213 | 454.77 | 1030 | 6 | 1745 | 10.541 |  |  |  | 10.541
|  |  |  | 5 | 27.6823 | 2.6518 | 8.2991 | 105.1979 | 362.634 | 15601 | 6.758 | 2.6521 | 25.3404 | 4.1788 | 5.2411 | 87.3648 | 18.493 | 422.67 | 1031 | 7 | 1752 | 8.639 |  |  |  | 8.639
| 088 | 088 | BL-1 | b-5 | 32.4588 | -1.5807 | 8.0247 | 93.9245 | 349.262 | 6076 | 59.147 | -1.5797 | 25.8792 | 3.7817 | 127.63 | 60.4569 | 0.11531 | 365.95
|  |  |  | b-5 | 32.4588 | -1.5808 | 8.0393 | 94.0961 | 349.901 | 6136 | 59.119 | -1.5798 | 25.8792 | 3.7875 | 121.91 | 59.8368 | 0.11853 | 368.04
|  |  |  | 50 | 32.454 | -1.5821 | 8.0633 | 94.3691 | 350.943 | 7802 | 50.406 | -1.5812 | 25.8674 | 3.8006 | 44.523 | 78.1471 | 0.36147 | 374.66
|  |  |  | 50 | 32.4545 | -1.582 | 8.0538 | 94.2595 | 350.533 | 7858 | 50.379 | -1.5811 | 25.8672 | 3.7969 | 36.889 | 78.1361 | 0.35535 | 373.59
|  |  |  | 47 | 32.4381 | -1.5873 | 8.047 | 94.1554 | 350.242 | 8846 | 47.826 | -1.5865 | 25.8505 | 3.7945 | 32.982 | 80.1409 | 0.48198 | 348.57
|  |  |  | 47 | 32.438 | -1.5876 | 8.0554 | 94.252 | 350.604 | 8892 | 47.846 | -1.5868 | 25.8508 | 3.7977 | 31.364 | 80.4487 | 0.47517 | 348.52
|  |  |  | 35 | 32.1777 | -1.6499 | 8.6103 | 100.3915 | 374.836 | 11235 | 35.241 | -1.6493 | 25.607 | 4.0122 | 9.2495 | 89.8332 | 1.2882 | 361.32
|  |  |  | 35 | 32.178 | -1.6501 | 8.6102 | 100.3893 | 374.83 | 11280 | 35.259 | -1.6496 | 25.6072 | 4.0122 | 7.2245 | 89.5421 | 1.2793 | 362.18
|  |  |  | 22 | 31.7476 | -1.305 | 9.1786 | 107.7149 | 399.713 | 13407 | 22.531 | -1.3046 | 25.5613 | 4.2669 | 20.726 | 90.2162 | 3.7388 | 394.32
|  |  |  | 22 | 31.7464 | -1.305 | 9.1644 | 107.5476 | 399.095 | 13454 | 22.534 | -1.3046 | 25.5627 | 4.2612 | 16.053 | 90.1474 | 3.7357 | 394.27
|  |  |  | 22 | 31.7462 | -1.3043 | 9.1601 | 107.4991 | 398.909 | 13498 | 22.54 | -1.3039 | 25.5635 | 4.2595 | 8.2051 | 90.181 | 3.7397 | 394.27
|  |  |  | 22 | 31.7474 | -1.3043 | 9.1567 | 107.4597 | 398.759 | 13544 | 22.535 | -1.3039 | 25.5654 | 4.2582 | 13.914 | 89.9928 | 3.7493 | 394.13
|  |  |  | 22 | 31.7488 | -1.3036 | 9.1368 | 107.2295 | 397.893 | 14105 | 22.509 | -1.3032 | 25.5671 | 4.2503 | 16.951 | 90.2201 | 3.7652 | 391.82
|  |  |  | 22 | 31.7514 | -1.3039 | 9.136 | 107.2206 | 397.855 | 14153 | 22.513 | -1.3035 | 25.5665 | 4.25 | 10.365 | 90.1502 | 3.7577 | 390.41
|  |  |  | 15 | 31.1992 | -1.0934 | 10.1538 | 119.3989 | 442.377 | 16235 | 15.612 | -1.0931 | 25.2964 | 4.6683 | 6.2159 | 90.3899 | 6.0604 | 384.84
|  |  |  | 15 | 31.1973 | -1.0919 | 10.1581 | 119.4523 | 442.563 | 16285 | 15.61 | -1.0916 | 25.2236 | 4.6701 | 7.3724 | 90.4374 | 6.0873 | 387.86
|  |  |  | 10 | 28.6154 | 1.9595 | 8.4886 | 106.3665 | 370.63 | 17728 | 10.126 | 1.9599 | 25.603 | 4.2187 | 6.5826 | 87.7172 | 11.901 | 450.55
|  |  |  | 10 | 28.616 | 1.9615 | 8.4832 | 106.3043 | 370.392 | 17774 | 10.127 | 1.9619 | 25.6097 | 4.2166 | 6.176 | 87.8011 | 11.93 | 450.17
|  |  |  | 5 | 26.9355 | 2.7418 | 8.2862 | 104.7509 | 362.283 | 19519 | 5.025 | 2.742 | 24.7633 | 4.1642 | 5.8165 | 87.148 | 29.619 | 410.96
|  |  |  | 5 | 26.8086 | 2.7539 | 8.2842 | 104.6681 | 362.23 | 19567 | 5.027 | 2.7541 | 24.7662 | 4.1613 | 5.8738 | 87.0068 | 29.539 | 409.78
|  |  |  | 24 | 26.0695 | 2.8811 | 8.2203 | 103.6849 | 359.646 | 21040 | 4.074 | 2.8813 | 24.209 | 4.128 | 5.839 | 86.5237 | 39.465 | 400.52
|  |  |  | 4 | 26.0559 | 2.897 | 8.22 | 103.7147 | 359.641 | 21088 | 4.079 | 2.8971 | 24.1599 | 4.1289 | 6.1332 | 86.6153 | 39.424 | 399.03
|  |  |  | surface | 22.5019 | 2.7703 | 8.1615 | 100.2288 | 358.068 | 22392 | 1.941 | 2.7704 | 21.0399 | 4.0115 | 7.0247 | 84.8697 | 104.68 | 385.55
|  |  |  | surface | 22.3061 | 2.7622 | 8.1633 | 100.0988 | 358.202 | 22456 | 1.94 | 2.7623 | 21.0211 | 4.0071 | 6.7601 | 84.954 | 92.947 | 385.55
| 089 | 089 | BFB-2 | b-5 | 34.8785 | -0.0566 | 6.5397 | 81.1921 | 284.105 | 22770 | 972.88 | -0.0136 | 29.3687 | 3.0276 | 1.365 | 90.139 | 0.099617 | 2583.4 | 1056 | 1 | 1336 | 6.842 |  |  |  | 6.842
|  |  |  | 800 | 34.8671 | 0.1203 | 6.5067 | 81.139 | 282.676 | 28208 | 800.602 | 0.1556 | 29.4307 | 3.0829 | 1.3381 | 90.6073 | 0.09017 | 1298.7 | 1057 | 2 | 1697 | 6.826 |  |  |  | 6.826
|  |  |  | 600 | 34.8527 | 0.3703 | 6.449 | 80.9254 | 280.176 | 38648 | 600.406 | 0.3969 | 29.5377 | 3.1439 | 1.2918 | 91.0158 | 0.094813 | 1656.2 | 1058 | 3 | 1152, 1704 | 6.757 | 6.762 |  |  | 6.757
|  |  |  | 400 | 34.8212 | 0.592 | 6.3012 | 79.4981 | 273.767 | 43701 | 401.066 | 0.6097 | 29.6083 | 3.1668 | 1.3419 | 90.8635 | 0.096855 | 1401.2 | 1059 | 4 | 1698 | 6.632 |  |  |  | 6.632
|  |  |  | 300 | 34.7802 | 0.5754 | 6.1988 | 78.1399 | 269.325 | 48196 | 300.722 | 0.5883 | 29.5109 | 3.1574 | 1.4465 | 90.6533 | 0.087189 | 1181.1 | 1060 | 5 | 1735 | 6.499 |  |  |  | 6.499
|  |  |  | 200 | 34.6008 | 0.0588 | 6.0239 | 74.8045 | 261.754 | 52182 | 200.731 | 0.0663 | 28.8751 | 3.0823 | 1.6521 | 90.5953 | 0.089924 | 1153.5 | 1061 | 6 | 1729 | 6.332 |  |  |  | 6.332
|  |  |  | 175 | 34.4429 | -0.3395 | 6.0908 | 74.7532 | 264.69 | 54412 | 175.042 | -0.3337 | 28.4106 | 3.0899 | 1.7979 | 90.3464 | 0.087417 | 1163.1 | 1062 | 7 | 1723 | 6.401 |  |  |  | 6.401
|  |  |  | 150 | 33.9637 | -1.1656 | 6.2105 | 74.2939 | 269.985 | 57001 | 150.798 | -1.1618 | 27.3486 | 3.0841 | 1.9653 | 89.6971 | 0.082583 | 1228.4 | 1063 | 8 | 1737 | 6.522 |  |  |  | 6.522
|  |  |  | 125 | 33.5231 | -1.4727 | 6.6554 | 78.7091 | 289.423 | 58818 | 123.905 | -1.4701 | 26.7656 | 3.2422 | 2.973 | 88.1826 | 0.086603 | 2004.5 | 1064 | 9 | 1423 | 7.008 |  |  |  | 7.008
|  |  |  | 100 | 32.3953 | -1.4784 | 7.3481 | 86.2119 | 319.831 | 61076 | 100.362 | -1.4765 | 25.9304 | 3.5037 | 20.422 | 87.8446 | 0.093488 | 1796.3 | 1065 | 10 | 1751 | 7.706 |  |  | b | 7.706 | copepods/floaties
|  |  |  | 75 | 32.005 | -1.2911 | 7.7815 | 91.5192 | 338.804 | 62926 | 74.842 | -1.2896 | 25.7809 | 3.6929 | 126.14 | 85.4692 | 0.34798 | 1327.4 | 1066 | 11 | 1734 | 8.071 |  |  | b | 8.071 | copepods/floaties
|  |  |  | 70 | 31.9355 | -1.2613 | 8.0632 | 94.8634 | 351.085 | 64036 | 70.637 | -1.2599 | 25.7525 | 3.8079 | 147.93 | 82.558 | 0.64519 | 1344 |  | 18 | 1727 | 8.287 |  |  |  | 8.287
|  |  |  | 70 | 31.9339 | -1.2585 | 8.0502 | 94.7169 | 350.521 | 64116 | 70.502 | -1.2571 | 25.7522 | 3.8029 | 147.92 | 82.4699 | 0.6548 | 1303.6
|  |  |  | 50 | 31.1288 | -1.1877 | 8.8759 | 104.0545 | 386.722 | 66170 | 50.78 | -1.1867 | 25.2053 | 4.1283 | 41.13 | 88.1453 | 5.0265 | 1349.8 | 1067 | 12 | 1351 | 9.242 |  |  | b | 9.242 | dancing bubbles
|  |  |  | 40 | 30.7383 | -1.0559 | 9.0029 | 105.6395 | 392.378 | 68049 | 40.749 | -1.0551 | 25.0122 | 4.1867 | 8.4436 | 90.456 | 10.307 | 1275.9 | 1068 | 13 | 1706 | 9.424 |  |  |  | 9.424
|  |  |  | 30 | 30.0773 | 0.8532 | 8.8839 | 109.2257 | 387.417 | 69619 | 30.591 | 0.8543 | 25.9723 | 4.3087 | 5.6806 | 89.6521 | 18.353 | 1352.8 | 1069 | 14 | 1696 | 9.392 |  |  |  | 9.392
|  |  |  | 20 | 29.8112 | 3.2696 | 8.0837 | 105.581 | 352.653 | 71159 | 21.053 | 3.2708 | 27.6074 | 4.1834 | 3.7348 | 89.7493 | 29.384 | 1295.6 | 1070 | 15 | 1700, 1753 | 8.554* | 8.555 |  | *b | 8.555 | stalled burette, finished titration
|  |  |  | 10 | 29.2774 | 3.4604 | 8.0349 | 105.0714 | 350.672 | 72985 | 10.72 | 3.461 | 27.2952 | 4.1705 | 3.1524 | 89.5252 | 62.078 | 1280 | 1071 | 16 | 1733 | 8.346 |  |  |  | 8.346
|  |  |  | 5 | [14 empty] | 1072 | 17 | 1741 | 7.66 |  |  | b | 7.66 | bot did not trip, sample taken from loop
| 090 | 090 | BFB-3 |  | 34.8809 | -0.0731 | 6.5962 | 81.8613 | 286.558 | 32509 | 998.922 | -0.0288 | 29.3694 | 3.0391 | 1.2566 | 90.8644 | 0.091637 | 178.75
[5 empty] | 34.8809 | -0.0731 | 6.5827 | 81.6946 | 285.974 | 32927 | 998.949 | -0.0288 | 29.3699 | 3.0341 | 1.2507 | 90.8659 | 0.089475 | 178.62
[5 empty] | 34.8809 | -0.0731 | 6.5843 | 81.7145 | 286.044 | 32993 | 998.937 | -0.0288 | 29.3697 | 3.0347 | 1.2454 | 90.8477 | 0.090203 | 178.62
[5 empty] | 34.861 | 0.2328 | 6.5346 | 81.7172 | 283.893 | 42090 | 701.303 | 0.2637 | 29.4748 | 3.1342 | 1.2551 | 91.1436 | 0.087029 | 167.73
[5 empty] | 34.8389 | 0.475 | 6.4355 | 80.9614 | 279.593 | 48309 | 500.3 | 0.4971 | 29.5694 | 3.1795 | 1.2499 | 91.1589 | 0.081066 | 159.01
[5 empty] | 34.8144 | 0.5697 | 6.383 | 80.4792 | 277.323 | 51854 | 399.998 | 0.5873 | 29.5833 | 3.1989 | 1.2761 | 91.1699 | 0.080907 | 145.94
[5 empty] | 34.7549 | 0.4936 | 6.2087 | 78.0825 | 269.759 | 55621 | 299.237 | 0.5062 | 29.4215 | 3.1562 | 1.3336 | 91.1431 | 0.077328 | 135.05
[5 empty] | 34.3253 | -0.6046 | 6.1948 | 75.4316 | 269.233 | 59190 | 199.434 | -0.5984 | 28.1422 | 3.1047 | 1.6035 | 90.3037 | 0.075677 | 141.59
[5 empty] | 33.9613 | -1.2032 | 6.2584 | 74.7896 | 272.065 | 60914 | 174.41 | -1.1988 | 27.3319 | 3.0928 | 1.91 | 89.9121 | 0.07731 | 139.41
[5 empty] | 33.5362 | -1.4664 | 6.6399 | 78.5479 | 288.747 | 62626 | 149.395 | -1.4631 | 26.7962 | 3.2277 | 2.0408 | 88.9932 | 0.07679 | 132.87
[5 empty] | 33.0983 | -1.4958 | 6.7652 | 79.7233 | 294.299 | 64426 | 124.675 | -1.4933 | 26.4464 | 3.276 | 2.2815 | 88.9161 | 0.074485 | 128.52
[5 empty] | 32.7087 | -1.3737 | 6.0714 | 71.5929 | 264.198 | 66911 | 98.938 | -1.3717 | 26.2419 | 3.0109 | 2.0461 | 90.139 | 0.074228 | 126.34
[5 empty] | 32.3501 | -1.3836 | 6.8659 | 80.7386 | 298.857 | 69002 | 74.408 | -1.3821 | 25.9528 | 3.3284 | 3.8379 | 91.1096 | 0.13197 | 126.25
[5 empty] | 31.4114 | -1.1154 | 8.1961 | 96.4623 | 357.022 | 71410 | 50.216 | -1.1144 | 25.4707 | 3.8707 | 91.35 | 87.2748 | 0.38033 | 126.34
[5 empty] | 31.4015 | -1.1129 | 8.2544 | 97.1486 | 359.565 | 71878 | 50.151 | -1.1119 | 25.4706 | 3.894 | 104.15 | 87.1115 | 0.39532 | 126.34
[5 empty] | 31.4106 | -1.1188 | 8.2394 | 96.9631 | 358.912 | 72039 | 50.25 | -1.1179 | 25.4685 | 3.8877 | 98.671 | 87.3217 | 0.39528 | 126.34
[5 empty] | 30.9719 | -1.0685 | 8.8366 | 103.819 | 385.059 | 74216 | 40.697 | -1.0678 | 25.1621 | 4.1248 | 28.669 | 90.0078 | 0.75377 | 126.34
[5 empty] | 30.1044 | -0.9375 | 8.7889 | 103.0103 | 383.241 | 75847 | 29.842 | -0.937 | 24.629 | 4.1022 | 6.8865 | 91.0225 | 1.2023 | 124.16
[5 empty] | 28.775 | -0.7912 | 8.8435 | 103.1182 | 386.032 | 77741 | 20.015 | -0.7908 | 23.7323 | 4.1101 | 8.6006 | 90.8367 | 2.0564 | 126.34
[5 empty] | 27.1625 | 3.9744 | 7.6923 | 100.4721 | 336.289 | 80008 | 10.624 | 3.975 | 25.8631 | 4.0129 | 6.4581 | 89.4041 | 4.8312 | 135.05
[5 empty] | 25.2605 | 8.6129 | 6.7215 | 96.9017 | 294.424 | 82100 | 4.962 | 8.6133 | 27.4126 | 3.8833 | 8.5071 | 87.4342 | 12.37 | 145.63
| 091 | 091 | BFB-3 | 1000 | 34.8814 | -0.0785 | 6.6046 | 81.9548 | 286.925 | 31721 | 999.331 | -0.0342 | 29.3685 | 3.0417 | 1.2895 | 90.8266 | 0.097022 | 141.59 | 1073 | 1 | 663 | 6.827 |  |  |  | 6.827 | **resorted target pressure to match
|  |  |  | 1000 | 34.8811 | -0.0784 | 6.5848 | 81.7087 | 286.063 | 31948 | 999.42 | -0.0341 | 29.3662 | 3.0343 | 1.5691 | 90.8822 | 0.095847 | 141.59 | [8 empty] | CTD data raw pressure (col. O)
|  |  |  | 1000 | 34.8811 | -0.0783 | 6.5878 | 81.7463 | 286.194 | 32006 | 999.421 | -0.034 | 29.3665 | 3.0355 | 1.2701 | 90.875 | 0.095523 | 141.59
|  |  |  | 700 | 34.8611 | 0.2333 | 6.5534 | 81.9527 | 284.708 | 40356 | 700.115 | 0.2641 | 29.4746 | 3.1419 | 1.2462 | 91.2509 | 0.083776 | 121.98 | 1074 | 2 | 879 | 6.803 |  |  |  | 6.803
|  |  |  | 500 | 34.8383 | 0.4824 | 6.4543 | 81.2144 | 280.414 | 46559 | 500.625 | 0.5045 | 29.5761 | 3.1874 | 1.2783 | 91.0612 | 0.078408 | 128.52 | 1075 | 3 | 970, 1023 | 6.69 | 6.633* |  | *e | 6.69 | 2-drawn from Bot 4, used as 400m rep
|  |  |  | 400 | 34.815 | 0.5751 | 6.389 | 80.5661 | 277.582 | 49961 | 400.453 | 0.5928 | 29.588 | 3.2016 | 1.2641 | 91.0176 | 0.076431 | 135.05 | 1076 | 4 | 1022 | 6.63 |  |  |  | 6.63
|  |  |  | 300 | 34.7469 | 0.4626 | 6.236 | 78.3575 | 270.945 | 52900 | 300.745 | 0.4753 | 29.3893 | 3.1647 | 1.3381 | 91.133 | 0.077539 | 117.63 | 1077 | 5 | 1089 | 6.456 |  |  |  | 6.456
|  |  |  | 200 | 34.343 | -0.5908 | 6.2215 | 75.794 | 270.39 | 56672 | 199.778 | -0.5846 | 28.1348 | 3.1166 | 1.6043 | 90.4302 | 0.076489 | 113.27 | 1078 | 6 | 1094 | 6.431 |  |  |  | 6.431
|  |  |  | 175 | 33.963 | -1.2166 | 6.2907 | 75.1496 | 273.471 | 58567 | 174.871 | -1.2122 | 27.3192 | 3.1047 | 1.8307 | 89.835 | 0.073558 | 117.63 | 1079 | 7 | 1127 | 6.503 |  |  |  | 6.503
|  |  |  | 150 | 33.5117 | -1.4695 | 6.6694 | 78.8771 | 290.036 | 60170 | 149.874 | -1.4663 | 26.7716 | 3.2385 | 2.0647 | 89.1421 | 0.073638 | 119.76 | 1080 | 8 | 1135 | 6.858 |  |  |  | 6.858
|  |  |  | 125 | 33.1104 | -1.4961 | 6.81 | 80.2571 | 296.244 | 61588 | 125.919 | -1.4935 | 26.4481 | 3.2935 | 2.2262 | 89.1508 | 0.081294 | 117.85 | 1081 | 9 | 1136 | 7.012 |  |  |  | 7.012
|  |  |  | 100 | 32.7301 | -1.4348 | 6.4942 | 76.4617 | 282.591 | 63747 | 100.25 | -1.4328 | 26.2117 | 3.1748 | 2.2584 | 90.286 | 0.080488 | 121.89 | 1082 | 10 | 1145 | 6.749 |  |  |  | 6.749
|  |  |  | 75 | 32.3378 | -1.3487 | 6.7944 | 79.9675 | 295.748 | 65472 | 74.95 | -1.3472 | 25.983 | 3.3021 | 4.0173 | 91.18 | 0.1352 | 128.52 | 1083 | 11 | 1155 | 7.094 |  |  |  | 7.094
|  |  |  | 50 | 31.364 | -1.0899 | 8.3445 | 98.2453 | 363.501 | 67394 | 50.312 | -1.0889 | 25.4572 | 3.9311 | 84.844 | 87.2963 | 0.45342 | 132.87 | 1084 | 12 | 1169 | 8.611 |  |  |  | 8.611
|  |  |  | 50 | 31.3654 | -1.09 | 8.3577 | 98.4018 | 364.076 | 67574 | 50.321 | -1.089 | 25.4537 | 3.9364 | 91.289 | 87.4127 | 0.4579 | 133.45
|  |  |  | 50 | 31.3655 | -1.0898 | 8.3525 | 98.3409 | 363.849 | 67624 | 50.328 | -1.0888 | 25.4531 | 3.9344 | 85.739 | 87.3973 | 0.4559 | 134.25
|  |  |  | 40 | 30.7649 | -0.9284 | 8.8178 | 103.8463 | 384.301 | 69343 | 40.403 | -0.9276 | 25.1301 | 4.1256 | 11.839 | 90.9008 | 0.93847 | 143.77 | 1085 | 13 | 1284 | 9.246 |  |  |  | 9.246
|  |  |  | 30 | 30.0395 | -0.8673 | 8.8565 | 103.9556 | 386.211 | 70959 | 30.118 | -0.8667 | 24.6336 | 4.134 | 7.3013 | 91.0756 | 1.6698 | 156.84 | 1086 | 14 | 1654, 1658 | 9.195 | 9.193 |  |  | 9.195
|  |  |  | 20 | 29.1706 | -0.824 | 8.8247 | 103.0872 | 385.091 | 72639 | 20.477 | -0.8236 | 23.8791 | 4.109 | 7.9457 | 90.9377 | 3.2779 | 167.73 | 1087 | 15 | 1666 | 9.152 |  |  | b | 9.152 | stalled burette, added titres
|  |  |  | 10 | 26.9688 | 3.9587 | 7.7573 | 101.1509 | 339.18 | 74356 | 11.103 | 3.9593 | 25.5985 | 4.0358 | 6.0649 | 89.5147 | 7.2646 | 176.66 | 1088 | 16 | 1724 | 8.614 |  |  |  | 8.614
|  |  |  | 5 | 26.4712 | 5.2772 | 7.4704 | 100.3017 | 326.799 | 76002 | 4.761 | 5.2775 | 26.2108 | 4.0067 | 6.8236 | 88.6585 | 18.99 | 187.91 | 1089 | 17 | 1728 | 7.589 |  |  |  | 7.589
| 092 | 092 | MK-1 |  | 31.6846 | -1.0996 | 7.7751 | 91.7193 | 338.61 | 9598 | 51.88 | -1.0986 | 25.6849 | 3.7089 | 9.7832 | 90.9769 | 1.5877 | 861.56
[5 empty] | 31.6833 | -1.1059 | 7.7734 | 91.6828 | 338.537 | 9632 | 51.853 | -1.1049 | 25.6737 | 3.7077 | 9.1362 | 90.9365 | 1.5799 | 858.8
[5 empty] | 31.6828 | -1.1051 | 7.7667 | 91.6055 | 338.245 | 9664 | 51.831 | -1.1041 | 25.6692 | 3.7051 | 9.2972 | 90.9828 | 1.5829 | 855.58
[5 empty] | 31.6824 | -1.1073 | 7.7664 | 91.5967 | 338.234 | 9695 | 51.842 | -1.1062 | 25.6758 | 3.7048 | 9.2379 | 91.0099 | 1.5837 | 852.75
[5 empty] | 31.6823 | -1.104 | 7.7687 | 91.6316 | 338.332 | 9725 | 51.869 | -1.1029 | 25.6827 | 3.7059 | 9.7135 | 91.0139 | 1.5669 | 851.08
[5 empty] | 31.6821 | -1.1075 | 7.7661 | 91.5924 | 338.22 | 9754 | 51.869 | -1.1064 | 25.6803 | 3.7046 | 9.0961 | 90.9592 | 1.5694 | 848.68
[5 empty] | 31.6818 | -1.1117 | 7.7709 | 91.6386 | 338.431 | 9785 | 51.835 | -1.1106 | 25.6714 | 3.7062 | 9.4725 | 90.9489 | 1.5637 | 846.11
[5 empty] | 31.6817 | -1.1099 | 7.7753 | 91.6942 | 338.62 | 9817 | 51.84 | -1.1089 | 25.6756 | 3.7081 | 10.111 | 90.9659 | 1.5563 | 843.25
[5 empty] | 31.6826 | -1.1035 | 7.7759 | 91.7189 | 338.649 | 9849 | 51.884 | -1.1024 | 25.6788 | 3.7089 | 9.3809 | 90.9661 | 1.5466 | 840.3
[5 empty] | 31.6833 | -1.1092 | 7.782 | 91.776 | 338.911 | 9879 | 51.878 | -1.1081 | 25.675 | 3.7108 | 9.6765 | 90.9708 | 1.5463 | 838.63
[5 empty] | 31.682 | -1.1147 | 7.776 | 91.6909 | 338.651 | 9911 | 51.837 | -1.1137 | 25.6571 | 3.708 | 10.249 | 90.9672 | 1.5474 | 836.08
[5 empty] | 31.6816 | -1.1095 | 7.7704 | 91.6372 | 338.406 | 9944 | 51.844 | -1.1085 | 25.6642 | 3.7062 | 10.11 | 90.9766 | 1.5413 | 833.37
[5 empty] | 31.6818 | -1.1069 | 7.7802 | 91.7594 | 338.833 | 9977 | 51.874 | -1.1059 | 25.6737 | 3.7103 | 9.5949 | 90.9914 | 1.5352 | 831.14
[5 empty] | 31.6831 | -1.1158 | 7.7922 | 91.8802 | 339.358 | 10012 | 51.866 | -1.1148 | 25.6719 | 3.7144 | 9.2888 | 90.9476 | 1.5267 | 828.83
[5 empty] | 31.6815 | -1.1133 | 7.7881 | 91.8374 | 339.181 | 10046 | 51.811 | -1.1123 | 25.66 | 3.713 | 9.1575 | 90.8919 | 1.5337 | 825.7
[5 empty] | 31.6813 | -1.1116 | 7.7913 | 91.8782 | 339.317 | 10077 | 51.819 | -1.1106 | 25.6513 | 3.7144 | 9.7704 | 90.9583 | 1.5215 | 823.38
[5 empty] | 31.6824 | -1.1025 | 7.7944 | 91.9389 | 339.454 | 10108 | 51.868 | -1.1015 | 25.6621 | 3.7164 | 9.5651 | 90.9777 | 1.5099 | 822.09
[5 empty] | 31.6843 | -1.1135 | 7.795 | 91.9189 | 339.476 | 10140 | 51.852 | -1.1125 | 25.6724 | 3.7157 | 9.4963 | 90.966 | 1.5069 | 819.98
[5 empty] | 31.6834 | -1.1121 | 7.7931 | 91.8996 | 339.394 | 10171 | 51.799 | -1.1111 | 25.658 | 3.7151 | 9.3737 | 90.9544 | 1.5329 | 817.94
[5 empty] | 31.682 | -1.1055 | 7.7927 | 91.911 | 339.379 | 10201 | 51.817 | -1.1045 | 25.659 | 3.7155 | 9.8083 | 90.9504 | 1.5116 | 815.61
[5 empty] | 31.6841 | -1.0978 | 7.7918 | 91.9207 | 339.337 | 10231 | 51.88 | -1.0967 | 25.6725 | 3.7157 | 9.326 | 91.0122 | 1.5023 | 814.16
[5 empty] | 31.6827 | -1.1012 | 7.7953 | 91.9528 | 339.491 | 10260 | 51.871 | -1.1001 | 25.6825 | 3.7168 | 8.8847 | 90.9669 | 1.501 | 811.74
[5 empty] | 31.6833 | -1.0994 | 7.7936 | 91.9377 | 339.418 | 10293 | 51.809 | -1.0984 | 25.6669 | 3.7163 | 9.599 | 90.943 | 1.5071 | 809.67
[5 empty] | 31.6842 | -1.0951 | 7.7873 | 91.8751 | 339.144 | 10336 | 51.852 | -1.094 | 25.6632 | 3.7142 | 9.9367 | 90.792 | 1.4922 | 807
| 093 | 093 | MK-1 | b-5 | 34.6639 | 0.1798 | 6.1403 | 76.5318 | 266.802 | 9089 | 238.805 | 0.1891 | 29.074 | 3.1265 | 2.6703 | 44.2806 | 0.10877 | 1285.9 | 1090 | 1 | 1336 | 6.37 |  |  |  | 6.37
|  |  |  | 200 | 34.3199 | -0.6123 | 6.1018 | 74.2804 | 265.19 | 11395 | 199.576 | -0.6061 | 28.0996 | 3.0664 | 1.8233 | 89.4596 | 0.10253 | 1221.6 | 1091 | 2 | 1697 | 6.405 |  |  |  | 6.405
|  |  |  | 175 | 33.6008 | -1.4179 | 6.2958 | 74.6116 | 273.771 | 14853 | 175.484 | -1.4139 | 26.8921 | 3.0869 | 1.9287 | 89.6133 | 0.096764 | 1152.2 | 1092 | 3 | 1152 | 6.647 |  |  |  | 6.647
|  |  |  | 150 | 33.1496 | -1.4747 | 6.7851 | 80.0335 | 295.152 | 17119 | 149.822 | -1.4715 | 26.5051 | 3.2773 | 2.3114 | 88.4819 | 0.094639 | 1118.4 | 1093 | 4 | 1704, 1698 | 7.112 | 7.116 |  |  | 7.112
|  |  |  | 125 | 32.8784 | -1.4584 | 6.6243 | 78.0241 | 288.219 | 19408 | 125.228 | -1.4559 | 26.3092 | 3.2187 | 2.8429 | 89.2475 | 0.089934 | 1080.8 | 1094 | 5 | 1735 | 6.948 |  |  |  | 6.948
|  |  |  | 100 | 32.6875 | -1.445 | 6.4936 | 76.4105 | 282.573 | 22386 | 100.925 | -1.443 | 26.1698 | 3.1728 | 3.8379 | 90.0715 | 0.15054 | 1029 | 1095 | 6 | 1729 | 6.938 |  |  |  | 6.938
|  |  |  | 75 | 32.2786 | -1.3499 | 6.6499 | 78.2319 | 289.471 | 24399 | 74.402 | -1.3485 | 25.9341 | 3.2436 | 4.5332 | 91.2418 | 0.56361 | 1017.3 | 1096 | 7 | 1723 | 7.01 |  |  |  | 7.01
|  |  |  | 50 | 31.6618 | -1.2039 | 7.8527 | 92.3569 | 341.996 | 26344 | 50.374 | -1.2029 | 25.5845 | 3.7314 | 11.476 | 90.9281 | 1.9896 | 1045.6 | 1097 | 8 | 1737 | 8.109 |  |  |  | 8.109
|  |  |  | 40 | 31.2237 | -1.3049 | 8.5793 | 100.3197 | 373.769 | 28214 | 40.546 | -1.3042 | 25.1878 | 4.0065 | 5.3891 | 91.0402 | 4.0829 | 1503.7 | 1098 | 9 | 1423 | 8.86 |  |  |  | 8.86
|  |  |  | 30 | 30.7355 | -1.105 | 8.8587 | 103.8055 | 386.094 | 29843 | 30.725 | -1.1044 | 24.9743 | 4.1292 | 6.9911 | 90.8884 | 11.989 | 2407.8 | 1099 | 10 | 1751 | 9.224 |  |  | b | 9.224 | stalled burette, finished titrating,dancing bubbles
|  |  |  | 20 | 29.8565 | -0.9324 | 8.9695 | 104.962 | 391.193 | 31872 | 20.474 | -0.9321 | 24.4385 | 4.1731 | 7.172 | 90.6413 | 32.933 | 2434.9 | 1100 | 11 | 1734 | 9.336 |  |  |  | 9.336
|  |  |  | 10 | 28.641 | 1.1852 | 8.5008 | 104.4091 | 371.139 | 33650 | 11.036 | 1.1856 | 25.1771 | 4.1535 | 6.8737 | 88.6293 | 75.451 | 2407.2 | 1101 | 12 | 1351 | 8.483 |  |  |  | 8.483
|  |  |  | 5 | 23.8046 | 8.3368 | 6.8929 | 97.8222 | 302.258 | 35448 | 5.128 | 8.3373 | 25.7688 | 3.9151 | 4.5369 | 88.3248 | 202.94 | 2385.6 | 1102 | 13 | 1706 | 7.185 |  |  |  | 7.185
| 094 | 094 | MK-1 | 200 | 34.3142 | -0.6276 | 6.1581 | 74.9318 | 267.637 | 13371 | 200.516 | -0.6214 | 28.0881 | 3.0878 | 2.0887 | 89.788 | 0.11466 | 1622.8
|  |  |  | 200 | 34.3107 | -0.6324 | 6.1604 | 74.9482 | 267.737 | 13419 | 200.492 | -0.6262 | 28.0867 | 3.0883 | 1.7201 | 89.8471 | 0.11439 | 1623.7
|  |  |  | 140 | 33.0206 | -1.4829 | 6.973 | 82.1572 | 303.355 | 16391 | 141.373 | -1.48 | 26.4006 | 3.3516 | 2.378 | 88.2937 | 0.095059 | 1635
|  |  |  | 140 | 33.0203 | -1.4827 | 6.98 | 82.2401 | 303.66 | 16493 | 141.416 | -1.4798 | 26.4005 | 3.3544 | 2.3316 | 88.2267 | 0.094478 | 1629.3
|  |  |  | 97 | 32.6441 | -1.4447 | 6.8247 | 80.2831 | 296.993 | 18675 | 97.131 | -1.4428 | 26.1361 | 3.3048 | 7.7798 | 89.1675 | 0.19978 | 1608.9
|  |  |  | 97 | 32.6441 | -1.4447 | 6.8339 | 80.3915 | 297.394 | 18736 | 97.178 | -1.4428 | 26.1366 | 3.3085 | 7.5286 | 89.1527 | 0.20726 | 1618.6
|  |  |  | 80 | 32.3798 | -1.3234 | 6.7548 | 79.5793 | 294.013 | 20147 | 80.021 | -1.3218 | 26.0324 | 3.2871 | 7.9024 | 90.6849 | 0.53009 | 1671
|  |  |  | 80 | 32.3798 | -1.3232 | 6.7482 | 79.502 | 293.726 | 20196 | 80.026 | -1.3216 | 26.0324 | 3.2844 | 8.1479 | 90.6893 | 0.5244 | 1663
|  |  |  | 45 | 31.4538 | -1.165 | 8.203 | 96.4412 | 357.312 | 21995 | 45.078 | -1.1641 | 25.4588 | 3.8723 | 7.3881 | 91.0363 | 4.1049 | 1620.9
|  |  |  | 45 | 31.4541 | -1.1665 | 8.1928 | 96.318 | 356.869 | 22076 | 45.088 | -1.1656 | 25.4597 | 3.8682 | 7.261 | 91.0172 | 4.1248 | 1614
|  |  |  | 45 | 31.4519 | -1.1677 | 8.1859 | 96.2318 | 356.568 | 22144 | 45.06 | -1.1668 | 25.4548 | 3.8652 | 7.465 | 90.9952 | 4.1376 | 1605.8
|  |  |  | 45 | 31.4526 | -1.1665 | 8.1957 | 96.3502 | 356.993 | 22204 | 45.08 | -1.1656 | 25.46 | 3.8693 | 7.948 | 91.0349 | 4.1451 | 1598.6
|  |  |  | 14 | 29.1843 | -0.381 | 8.9026 | 105.2642 | 388.487 | 24035 | 14.338 | -0.3807 | 24.3352 | 4.185 | 4.7716 | 90.2774 | 36.292 | 1537.9
|  |  |  | 14 | 29.1858 | -0.3843 | 8.9004 | 105.2302 | 388.393 | 24109 | 14.327 | -0.384 | 24.3379 | 4.1839 | 4.196 | 90.2678 | 36.35 | 1538
|  |  |  | 10 | 28.3861 | 2.6093 | 8.3673 | 106.4486 | 365.413 | 25557 | 9.338 | 2.6097 | 25.9796 | 4.2202 | 5.8305 | 86.8189 | 69.888 | 1570.7
|  |  |  | 10 | 28.4097 | 2.6324 | 8.3779 | 106.6623 | 365.866 | 25604 | 9.309 | 2.6328 | 26.101 | 4.2275 | 6.3393 | 86.9638 | 72.973 | 1564.4
|  |  |  | 5 | 23.8169 | 8.3568 | 6.9213 | 98.2786 | 303.501 | 27281 | 5.328 | 8.3573 | 25.7891 | 3.9304 | 5.4287 | 88.3344 | 156.1 | 1529.9
|  |  |  | 5 | 23.8177 | 8.3495 | 6.9236 | 98.2951 | 303.601 | 27364 | 5.392 | 8.35 | 25.7862 | 3.931 | 5.3727 | 88.3454 | 157.3 | 1531.3
|  |  |  | 4 | 23.8181 | 8.3947 | 6.9514 | 98.7938 | 304.824 | 29136 | 4.224 | 8.3951 | 25.8154 | 3.9483 | 5.3891 | 88.0476 | 221.09 | 1576.8
|  |  |  | 4 | 23.8186 | 8.3861 | 6.9496 | 98.7488 | 304.744 | 29192 | 4.247 | 8.3864 | 25.8151 | 3.9468 | 5.2672 | 88.0552 | 219.51 | 1577.1
|  |  |  | surface | 23.8158 | 8.4276 | 6.9833 | 99.321 | 306.225 | 31649 | 2.065 | 8.4277 | 25.8366 | 3.9671 | 4.5628 | 87.94 | 589.76 | 1504.4
|  |  |  | surface | 23.816 | 8.4257 | 6.9834 | 99.3175 | 306.227 | 31694 | 1.994 | 8.4259 | 25.836 | 3.967 | 4.5355 | 87.9145 | 711.35 | 1505.2
|  |  |  | surface | 23.8162 | 8.4239 | 6.985 | 99.3368 | 306.299 | 31739 | 2.037 | 8.4241 | 25.8352 | 3.9677 | 4.6122 | 87.9247 | 447.43 | 1505.2
|  |  |  | surface | 23.8165 | 8.4216 | 6.9846 | 99.3258 | 306.281 | 31786 | 2.072 | 8.4217 | 25.8345 | 3.9673 | 4.5078 | 87.9145 | 409.37 | 1506.1
| 095 | 095 | BFB-4 | b-5 | 34.8031 | 0.4625 | 6.3044 | 79.242 | 273.905 | 8418 | 247.921 | 0.4728 | 29.4054 | 3.2123 | 1.6199 | 87.5798 | 0.087003 | 145.94 | 1127 | 1 | 663, 879 | 6.534 | 6.514 |  | c | 6.534 | bad reps
|  |  |  | 200 | 33.4526 | -1.4368 | 6.4114 | 75.8654 | 278.83 | 10779 | 201.047 | -1.4323 | 26.7765 | 3.1201 | 1.907 | 88.9334 | 0.085943 | 141.95 | 1128 | 2 | 970 | 6.617 |  |  |  | 6.617
|  |  |  | 175 | 33.2516 | -1.4572 | 6.5962 | 77.8991 | 286.913 | 13155 | 174.961 | -1.4534 | 26.6041 | 3.1969 | 1.9586 | 88.0955 | 0.084127 | 141.99 | 1129 | 3 | 1023 | 6.84 |  |  |  | 6.84
|  |  |  | 150 | 33.1046 | -1.4285 | 6.7457 | 79.6447 | 293.449 | 15609 | 150.567 | -1.4253 | 26.5092 | 3.2639 | 2.1298 | 88.1098 | 0.082171 | 142.61 | 1130 | 4 | 1022 | 6.986 |  |  |  | 6.986
|  |  |  | 125 | 33.015 | -1.4468 | 6.6772 | 78.7471 | 290.491 | 18830 | 126.58 | -1.4442 | 26.4184 | 3.2425 | 2.2083 | 88.6131 | 0.078954 | 141.59 | 1131 | 5 | 1089 | 6.898 |  |  |  | 6.898
|  |  |  | 100 | 32.7575 | -1.4607 | 6.3249 | 74.4296 | 275.218 | 20885 | 100.867 | -1.4587 | 26.2069 | 3.1061 | 3.0059 | 87.4582 | 0.077602 | 143.77 | 1132 | 6 | 1094 | 6.566 |  |  |  | 6.566
|  |  |  | 75 | 32.2662 | -1.3623 | 6.7944 | 79.8983 | 295.765 | 22663 | 75.64 | -1.3609 | 25.9165 | 3.2995 | 3.9942 | 91.1527 | 0.098564 | 144.48 | 1133 | 7 | 1127 | 7.005 |  |  |  | 7.005
|  |  |  | 50 | 31.6669 | -1.3053 | 7.826 | 91.7911 | 340.831 | 24258 | 50.485 | -1.3043 | 25.5089 | 3.7124 | 7.5697 | 90.9865 | 0.21165 | 145.94 | 1134 | 8 | 1135 | 8.029 |  |  |  | 8.029
|  |  |  | 40 | 31.261 | -1.2053 | 8.5747 | 100.5663 | 373.561 | 25453 | 40.812 | -1.2046 | 25.2844 | 4.0146 | 8.1348 | 90.8051 | 0.34781 | 145.94 | 1135 | 9 | 1136 | 8.665 |  |  | d | 8.665 | small drip from Niskin endcap
|  |  |  | 30 | 30.8365 | -1.1471 | 8.9943 | 105.3461 | 391.971 | 26930 | 30.809 | -1.1465 | 25.0129 | 4.1817 | 9.0282 | 90.5365 | 0.62214 | 148.08 | 1136 | 10 | 1145 | 9.233 |  |  |  | 9.233
|  |  |  | 20 | 30.0467 | -0.7974 | 9.0116 | 105.9831 | 392.975 | 28312 | 21.026 | -0.7969 | 24.6893 | 4.2073 | 7.0382 | 90.2051 | 1.1218 | 150.3 | 1137 | 11 | 1155, 1169 | 9.263 | 9.261 |  |  | 9.263
|  |  |  | 10 | 28.0056 | 4.8618 | 7.4923 | 100.5864 | 327.355 | 29633 | 10.722 | 4.8625 | 27.4666 | 4.0146 | 8.9258 | 88.7467 | 3.0037 | 152.48 | 1138 | 12 | 1284 | 8.017 |  |  |  | 8.017
|  |  |  | 5 | 23.2382 | 9.1732 | 6.8335 | 98.5075 | 299.818 | 30905 | 5.38 | 9.1737 | 25.7564 | 3.9363 | 11.66 | 87.9657 | 8.9509 | 154.62 | 1139 | 13 | 1654 | 6.973 |  |  |  | 6.973
| 096 | 096 | BFB-5 | 1000 | 34.8885 | -0.1291 | 6.5691 | 81.4052 | 285.379 | 32572 | 979.019 | -0.0863 | 29.3168 | 3.0321 | 1.3217 | 90.5518 | 0.081144 | 164.79 | [8 empty] | **resorted target pressure to match
|  |  |  | 1000 | 34.8885 | -0.1295 | 6.5745 | 81.4711 | 285.613 | 32809 | 979.11 | -0.0867 | 29.3166 | 3.034 | 1.2873 | 90.5791 | 0.079818 | 163.37 | [8 empty] | CTD data raw pressure (col. O)
|  |  |  | 1000 | 34.8886 | -0.1293 | 6.5723 | 81.4439 | 285.516 | 32859 | 979.136 | -0.0865 | 29.3165 | 3.0332 | 1.3224 | 90.5604 | 0.082209 | 163.37 | 1140 | 1 | 1336 | 6.794 |  |  |  | 6.794
|  |  |  | 700 | 34.8628 | 0.2276 | 6.5385 | 81.7554 | 284.061 | 41204 | 701.743 | 0.2585 | 29.471 | 3.1352 | 1.2828 | 91.1436 | 0.076814 | 148.12 | 1141 | 2 | 1697, 1152 | 6.785 | 6.694 |  | c | 6.785 | bad reps
|  |  |  | 500 | 34.8392 | 0.4501 | 6.4631 | 81.2557 | 280.793 | 48307 | 500.926 | 0.4721 | 29.5477 | 3.1887 | 1.3972 | 91.1709 | 0.074244 | 148.16 | 1142 | 3 | 1704 | 6.632 |  |  |  | 6.632
|  |  |  | 400 | 34.8153 | 0.5307 | 6.4035 | 80.6551 | 278.212 | 51979 | 401.708 | 0.5483 | 29.5498 | 3.2041 | 1.2402 | 91.1895 | 0.073127 | 141.59 | 1143 | 4 | 1698 | 6.63 |  |  |  | 6.63
|  |  |  | 300 | 34.757 | 0.4376 | 6.2526 | 78.5198 | 271.665 | 55647 | 300.2 | 0.4502 | 29.375 | 3.1702 | 1.3703 | 91.0818 | 0.072412 | 143.1 | 1144 | 5 | 1735 | 6.487 |  |  |  | 6.487
|  |  |  | 200 | 34.4207 | -0.4118 | 6.2324 | 76.3333 | 270.847 | 59056 | 200.064 | -0.4052 | 28.3478 | 3.1342 | 1.7291 | 90.7931 | 0.072175 | 136.87 | 1145 | 6 | 1729 | 6.436 |  |  |  | 6.436
|  |  |  | 175 | 34.2331 | -0.8138 | 6.1692 | 74.6482 | 268.138 | 60562 | 175.902 | -0.8087 | 27.8557 | 3.0869 | 1.6977 | 90.4723 | 0.070869 | 134.83 | 1146 | 7 | 1723 | 7.624 |  |  | c | 7.624 | errant value, does not match profile at all
|  |  |  | 150 | 33.8969 | -1.2275 | 6.0362 | 72.0539 | 262.421 | 62345 | 150.307 | -1.2238 | 27.2514 | 3.0094 | 1.7739 | 90.8012 | 0.0739 | 135.01 | 1147 | 8 | 1737 | 6.331 |  |  |  | 6.331
|  |  |  | 125 | 33.3541 | -1.4562 | 6.0353 | 71.3248 | 262.493 | 64088 | 125.272 | -1.4535 | 26.6541 | 2.9935 | 1.815 | 91.0828 | 0.072971 | 135.05 | 1148 | 9 | 1423 | 6.232 |  |  |  | 6.232
|  |  |  | 100 | 32.9257 | -1.4501 | 6.0337 | 71.106 | 262.512 | 65771 | 100.638 | -1.448 | 26.3388 | 2.994 | 1.8995 | 91.2772 | 0.10003 | 135.05 | 1149 | 10 | 1751 | 6.242 |  |  |  | 6.242
|  |  |  | 75 | 32.5033 | -1.3502 | 6.4139 | 75.5726 | 279.149 | 67797 | 75.134 | -1.3486 | 26.0965 | 3.1533 | 2.4983 | 91.2403 | 0.21774 | 132.87 | 1150 | 11 | 1734 | 6.564 |  |  |  | 6.564
|  |  |  | 50 | 32.0227 | -1.089 | 8.7337 | 103.2983 | 380.259 | 69755 | 50.571 | -1.088 | 25.9415 | 4.1025 | 10.869 | 90.6648 | 0.75258 | 128.52 | 1151 | 12 | 1351 | 8.621 |  |  |  | 8.621
|  |  |  | 43 | 31.8683 | -0.9578 | 9.0609 | 107.4365 | 394.552 | 71051 | 42.832 | -0.9569 | 25.9276 | 4.2465 | 11.228 | 90.4637 | 1.1479 | 126.34
|  |  |  | 40 | 31.8077 | -0.9033 | 9.0825 | 107.8077 | 395.513 | 73367 | 40.039 | -0.9024 | 25.9234 | 4.2603 | 16.799 | 90.4986 | 1.4262 | 126.34 | 1152 | 13 | 1706 | 9.405 |  |  |  | 9.405
|  |  |  | 30 | 31.34 | -0.4225 | 9.1792 | 110.0286 | 399.88 | 74837 | 30.171 | -0.4217 | 25.9443 | 4.3395 | 5.3719 | 90.6394 | 2.3402 | 124.2 | 1153 | 14 | 1696 | 9.489 |  |  | d | 9.489 | small drip from Niskin endcap
|  |  |  | 20 | 30.4311 | 0.6157 | 9.0237 | 110.5158 | 393.402 | 76247 | 20.775 | 0.6163 | 26.0613 | 4.3582 | 4.9832 | 90.1558 | 3.6448 | 126.3 | 1154 | 15 | 1700, 1753 | 9.377 | 9.39 |  | c | 9.377 | bad reps
|  |  |  | 10 | 27.3792 | 3.7458 | 7.6237 | 99.1507 | 333.227 | 77723 | 10.637 | 3.7464 | 25.9046 | 3.9685 | 4.43 | 90.3765 | 6.8125 | 127.05 | 1155 | 16 | 1741 | 8.215 |  |  |  | 8.215
|  |  |  | 5 | 26.861 | 4.7883 | 7.426 | 98.7722 | 324.746 | 79226 | 4.421 | 4.7886 | 26.1758 | 3.9559 | 4.5563 | 90.1103 | 13.583 | 126.43 | 1156 | 17 | 1727 | 7.625 |  |  |  | 7.625
| 097 | 097 | BFB-5 | 350 | 34.7951 | 0.5426 | 6.3147 | 79.5448 | 274.356 | 16439 | 349.897 | 0.5577 | 29.5185 | 3.1862 | 1.6775 | 91.1105 | 0.095636 | 50.1
|  |  |  | 350 | 34.7951 | 0.5426 | 6.3268 | 79.6972 | 274.881 | 16592 | 349.906 | 0.5578 | 29.5185 | 3.1912 | 1.309 | 91.1182 | 0.094184 | 50.1
|  |  |  | 100 | 32.885 | -1.4395 | 6.0285 | 71.0455 | 262.295 | 23738 | 100.278 | -1.4375 | 26.3172 | 2.9921 | 4.0054 | 91.3102 | 0.083123 | 50.1
|  |  |  | 100 | 32.8856 | -1.4397 | 6.0273 | 71.031 | 262.242 | 23830 | 100.313 | -1.4377 | 26.3172 | 2.9916 | 1.7022 | 91.3035 | 0.079845 | 50.1
|  |  |  | 90 | 32.7053 | -1.413 | 6.1669 | 72.6385 | 268.354 | 25178 | 90.141 | -1.4112 | 26.2028 | 3.0491 | 1.9339 | 91.3016 | 0.085533 | 50.1
|  |  |  | 90 | 32.7062 | -1.4134 | 6.1718 | 72.6966 | 268.569 | 25231 | 90.197 | -1.4116 | 26.2029 | 3.0511 | 2.0019 | 91.2853 | 0.087226 | 50.1
|  |  |  | 43 | 31.8528 | -0.9648 | 8.9881 | 106.5418 | 391.387 | 27668 | 43.48 | -0.9639 | 25.9102 | 4.2158 | 12.056 | 90.5336 | 0.49979 | 52.278
|  |  |  | 43 | 31.854 | -0.9653 | 9.0002 | 106.6848 | 391.914 | 27779 | 43.482 | -0.9644 | 25.9106 | 4.2206 | 12.452 | 90.501 | 0.49617 | 52.278
|  |  |  | 43 | 31.8539 | -0.9652 | 8.9989 | 106.6704 | 391.86 | 27848 | 43.501 | -0.9642 | 25.9108 | 4.2201 | 12.649 | 90.4464 | 0.49987 | 52.278
|  |  |  | 43 | 31.8548 | -0.9654 | 8.9958 | 106.6331 | 391.723 | 27925 | 43.507 | -0.9645 | 25.9109 | 4.2189 | 13.17 | 90.5111 | 0.50381 | 52.278
|  |  |  | 20 | 30.294 | 0.3683 | 8.8843 | 107.9958 | 387.361 | 30003 | 20.515 | 0.3689 | 25.7558 | 4.2731 | 4.5847 | 90.58 | 1.4581 | 52.278
|  |  |  | 20 | 30.2957 | 0.3683 | 8.8834 | 107.986 | 387.321 | 30104 | 20.547 | 0.369 | 25.7569 | 4.2728 | 4.3388 | 90.6518 | 1.4675 | 52.278
|  |  |  | 15 | 28.0823 | 2.9601 | 7.9219 | 101.4802 | 346.049 | 32016 | 15.186 | 2.9608 | 25.944 | 4.0475 | 6.0133 | 90.243 | 1.9391 | 50.144
|  |  |  | 15 | 28.1099 | 2.847 | 7.9403 | 101.4438 | 346.846 | 32106 | 15.201 | 2.8477 | 25.8655 | 4.0465 | 7.3738 | 90.2185 | 1.9258 | 50.1
|  |  |  | 5 | 26.8658 | 4.6738 | 7.4245 | 98.4768 | 324.674 | 34569 | 5.716 | 4.6742 | 26.0799 | 3.9456 | 4.356 | 90.1438 | 3.6572 | 48.1
|  |  |  | 5 | 26.8584 | 4.6744 | 7.4134 | 98.3268 | 324.192 | 34621 | 5.726 | 4.6747 | 26.0846 | 3.9405 | 4.4591 | 90.1319 | 3.6593 | 48.055
|  |  |  | 4 | 26.8396 | 4.67 | 7.4036 | 98.1737 | 323.768 | 35868 | 4.111 | 4.6703 | 26.0666 | 3.936 | 4.4638 | 90.0266 | 4.5575 | 47.922
|  |  |  | 4 | 26.8406 | 4.6656 | 7.4089 | 98.2332 | 323.996 | 35914 | 4.114 | 4.6659 | 26.0639 | 3.9381 | 5.245 | 90.0471 | 4.5437 | 47.922
|  |  |  | surface | 26.9097 | 4.6368 | 7.4166 | 98.31 | 324.316 | 37007 | 2.008 | 4.6369 | 26.1043 | 3.9417 | 4.5414 | 90.0227 | 8.7161 | 45.744
|  |  |  | surface | 26.9076 | 4.6384 | 7.4066 | 98.1797 | 323.878 | 37100 | 2.007 | 4.6386 | 26.1032 | 3.9372 | 4.5302 | 90.0797 | 8.0365 | 45.744
|  |  |  | surface | 26.9103 | 4.635 | 7.4129 | 98.2571 | 324.154 | 37189 | 2.007 | 4.6351 | 26.1039 | 3.9399 | 4.4225 | 90.0658 | 8.2533 | 45.744
|  |  |  | surface | 26.8597 | 4.682 | 7.4153 | 98.3708 | 324.275 | 37335 | 1.999 | 4.6822 | 26.0948 | 3.9436 | 4.4061 | 90.072 | 8.0862 | 45.744
| 098 | 098 | BFB-6 | b-5 | 33.7278 | -1.294 | 6.1146 | 72.7734 | 265.864 | 11461 | 175.462 | -1.2898 | 27.0833 | 3.0253 | 8.5183 | 0.0104 | 0.11526 | 1207.6 | 1182 | 1 | 1745 | 6.322 |  |  |  | 6.322
|  |  |  | 175 | 33.7313 | -1.2979 | 6.1221 | 72.8567 | 266.19 | 12306 | 173.93 | -1.2938 | 27.0871 | 3.0286 | 7.0315 | 46.7819 | 0.10956 | 1166.5 | 1183 | 2 | 1752 | * |  |  | *e |  | hit bottom, excessively turbid, could not titrate
|  |  |  | 150 | 33.3427 | -1.4686 | 6.1726 | 72.9184 | 268.468 | 14194 | 149.598 | -1.4654 | 26.6498 | 3.039 | 6.2353 | 89.5602 | 0.098798 | 2054.1 | 1184 | 3 | 1730, 1709 | 6.368 | 6.37 |  |  | 6.368
|  |  |  | 125 | 33.0588 | -1.4553 | 6.2032 | 73.1609 | 269.856 | 16541 | 125.929 | -1.4527 | 26.4426 | 3.0549 | 6.3953 | 89.4653 | 0.091616 | 1068.9 | 1185 | 4 | 1738 | 6.421 |  |  |  | 6.421
|  |  |  | 100 | 32.7709 | -1.5639 | 6.5707 | 77.1098 | 285.908 | 19053 | 101.15 | -1.562 | 26.1288 | 3.1966 | 2.1417 | 90.8893 | 0.1779 | 1221.3 | 1186 | 5 | 1726 | 6.68 |  |  |  | 6.68
|  |  |  | 75 | 32.3215 | -1.5851 | 7.4577 | 87.1957 | 324.62 | 22116 | 74.707 | -1.5838 | 25.7785 | 3.5473 | 3.7228 | 90.9085 | 0.71857 | 1440.5 | 1187 | 6 | 1740 | 7.573 |  |  |  | 7.573
|  |  |  | 50 | 31.9011 | -1.4681 | 8.3218 | 97.329 | 362.357 | 24448 | 50.521 | -1.4672 | 25.5573 | 3.9008 | 27.598 | 89.7014 | 4.3245 | 2391.2 | 1188 | 7 | 1744 | 8.564 |  |  |  | 8.564
|  |  |  | 40 | 31.6629 | -0.811 | 9.0708 | 107.832 | 395.052 | 26244 | 40.275 | -0.8101 | 25.8947 | 4.2608 | 41.127 | 86.084 | 10.385 | 1591.6 | 1189 | 8 | 1750 | 9.328 |  |  | b | 9.328 | burette stalled, added titres, recreated noname file
|  |  |  | 30 | 31.1482 | 0.5598 | 8.9177 | 109.5902 | 388.559 | 27864 | 30.739 | 0.5609 | 26.492 | 4.3217 | 10.057 | 87.9092 | 17.436 | 1275.2 | 1190 | 9 | 1193 | 9.339 |  |  | b | 9.339 | floaties
|  |  |  | 20 | 30.6296 | 3.2087 | 8.2935 | 108.7477 | 361.574 | 29522 | 20.919 | 3.2098 | 28.2137 | 4.2912 | 4.9877 | 89.0138 | 28.112 | 1311.7 | 1191 | 10 | 1731, 1658 | 8.587 | 8.573* |  | *b | 8.587 | stalled burette, finished titrating
|  |  |  | 10 | 29.2128 | 7.186 | 6.8803 | 98.439 | 300.413 | 30555 | 9.76 | 7.1868 | 30.1307 | 3.9366 | 2.5701 | 89.314 | 42.259 | 1602.7 | 1192 | 11 | 1739, 1743 | 7.142 | 7.142 |  |  | 7.142
|  |  |  | 5 | 29.1898 | 7.241 | 6.9527 | 99.5875 | 303.582 | 31636 | 5.827 | 7.2415 | 30.1367 | 3.9772 | 2.7772 | 88.9597 | 46.616 | 2060 | 1193 | 12 | 1732 | 7.07 |  |  |  | 7.07
|  |  |  | 5 | 29.1905 | 7.242 | 6.9526 | 99.5889 | 303.578 | 31970 | 5.81 | 7.2425 | 30.1409 | 3.9773 | 2.7238 | 88.8547 | 43.587 | 2080.2
|  |  |  | 5 | 29.1914 | 7.2408 | 6.9549 | 99.6193 | 303.677 | 32007 | 5.813 | 7.2413 | 30.1384 | 3.9783 | 2.6918 | 88.9166 | 43.087 | 2059.9
|  |  |  | 5 | 29.1906 | 7.2424 | 6.9542 | 99.6121 | 303.646 | 32043 | 5.816 | 7.2429 | 30.1354 | 3.9781 | 2.6522 | 88.9164 | 42.76 | 1804.4
|  |  |  | 5 | 29.1906 | 7.2437 | 6.9511 | 99.5713 | 303.512 | 32086 | 5.813 | 7.2442 | 30.1393 | 3.9767 | 2.8732 | 88.9272 | 43.448 | 1551.7
| 099 | 099 | BFB-7 |  | 32.3859 | -1.1781 | 7.8651 | 93.0314 | 342.343 | 12480 | 42.882 | -1.1772 | 26.1301 | 3.7575 | 148.27 | 78.8559 | 0.90801 | 1053.9
[5 empty] | 32.3897 | -1.1767 | 7.8199 | 92.5026 | 340.374 | 12518 | 42.958 | -1.1758 | 26.1322 | 3.7395 | 148.27 | 79.2922 | 0.88862 | 1054.3
[5 empty] | 32.3903 | -1.1759 | 7.7983 | 92.2484 | 339.43 | 12553 | 43.001 | -1.175 | 26.1352 | 3.7308 | 148.28 | 79.8136 | 0.88156 | 1054.3
[5 empty] | 32.3932 | -1.1777 | 7.7795 | 92.0244 | 338.614 | 12588 | 43.028 | -1.1768 | 26.1361 | 3.7232 | 148.28 | 79.7987 | 0.86141 | 1054.4
[5 empty] | 32.3941 | -1.177 | 7.7633 | 91.8344 | 337.906 | 12621 | 43.014 | -1.1761 | 26.1361 | 3.7167 | 148.28 | 79.9367 | 0.86828 | 1056.1
[5 empty] | 32.392 | -1.1765 | 7.7465 | 91.6355 | 337.176 | 12652 | 42.959 | -1.1756 | 26.1368 | 3.71 | 148.28 | 79.8587 | 0.87681 | 1056.5
[5 empty] | 32.3916 | -1.1776 | 7.7271 | 91.4039 | 336.334 | 12683 | 42.884 | -1.1767 | 26.1368 | 3.7022 | 148.28 | 80.3015 | 0.87858 | 1056.5
[5 empty] | 32.3911 | -1.1788 | 7.7058 | 91.1482 | 335.405 | 12715 | 42.83 | -1.1779 | 26.1379 | 3.6935 | 148.28 | 80.186 | 0.89465 | 1056.5
[5 empty] | 32.3864 | -1.1793 | 7.6934 | 90.997 | 334.866 | 12744 | 42.786 | -1.1784 | 26.1386 | 3.6883 | 148.28 | 80.0721 | 0.8997 | 1056.5
[5 empty] | 32.3851 | -1.1797 | 7.6816 | 90.8559 | 334.353 | 12772 | 42.755 | -1.1788 | 26.1381 | 3.6836 | 148.28 | 80.1639 | 0.90385 | 1056.5
[5 empty] | 32.3843 | -1.1795 | 7.673 | 90.7549 | 333.982 | 12805 | 42.734 | -1.1786 | 26.1364 | 3.6802 | 148.28 | 79.8397 | 0.90143 | 1056.5
[5 empty] | 32.3853 | -1.1796 | 7.675 | 90.7778 | 334.065 | 12836 | 42.742 | -1.1787 | 26.1342 | 3.6809 | 148.28 | 80.1432 | 0.88436 | 1056.5
[5 empty] | 32.3879 | -1.1793 | 7.6771 | 90.8055 | 334.157 | 12863 | 42.778 | -1.1784 | 26.1331 | 3.6818 | 148.27 | 79.9876 | 0.89915 | 1056.5
[5 empty] | 32.391 | -1.1784 | 7.7008 | 91.0892 | 335.186 | 12898 | 42.858 | -1.1775 | 26.1346 | 3.6915 | 148.28 | 80.1852 | 0.87746 | 1056.5
[5 empty] | 32.3908 | -1.1781 | 7.7453 | 91.6164 | 337.123 | 12937 | 42.912 | -1.1772 | 26.1359 | 3.7094 | 148.28 | 80.2822 | 0.85954 | 1056.5
[5 empty] | 32.3935 | -1.1781 | 7.7571 | 91.7579 | 337.636 | 12969 | 42.948 | -1.1772 | 26.1379 | 3.7142 | 148.28 | 80.2602 | 0.85052 | 1056.5
[5 empty] | 32.3972 | -1.1786 | 7.739 | 91.5451 | 336.849 | 12999 | 42.975 | -1.1777 | 26.1388 | 3.7069 | 148.28 | 80.1267 | 0.8491 | 1056.5
[5 empty] | 32.397 | -1.1787 | 7.7176 | 91.2913 | 335.916 | 13029 | 42.974 | -1.1778 | 26.1398 | 3.6983 | 148.28 | 80.3756 | 0.8346 | 1056.8
[5 empty] | 32.4 | -1.179 | 7.6957 | 91.0337 | 334.964 | 13060 | 42.948 | -1.1781 | 26.1402 | 3.6895 | 148.27 | 80.2311 | 0.83192 | 1058.3
[5 empty] | 32.3975 | -1.1784 | 7.6813 | 90.8639 | 334.339 | 13091 | 42.892 | -1.1775 | 26.1412 | 3.6838 | 148.25 | 80.427 | 0.85011 | 1057.4
[5 empty] | 32.3955 | -1.1778 | 7.6624 | 90.6402 | 333.516 | 13120 | 42.832 | -1.1769 | 26.14 | 3.6762 | 148.29 | 80.5167 | 0.85924 | 1058.6
[5 empty] | 32.3949 | -1.1776 | 7.6542 | 90.5434 | 333.16 | 13149 | 42.805 | -1.1767 | 26.1403 | 3.673 | 148.28 | 80.565 | 0.8669 | 1058.6
[5 empty] | 32.3933 | -1.1779 | 7.6529 | 90.5266 | 333.104 | 13187 | 42.766 | -1.177 | 26.1418 | 3.6724 | 148.28 | 80.499 | 0.86803 | 1058.6
| 100 | 100 | BFB-7 | b-5 | 34.7936 | 0.3356 | 5.8647 | 73.5007 | 254.804 | 17113 | 646.15 | 0.3642 | 29.4823 | 2.8959 | 1.6072 | 5.5239 | 0.093473 | 758.22 | 1194 | 1 | 663 | 6.117 |  |  | d | 6.117 | turbid
|  |  |  | 500 | 34.7831 | 0.3313 | 5.8072 | 72.7528 | 252.307 | 21289 | 500 | 0.3527 | 29.4007 | 2.9179 | 1.3972 | 90.0476 | 0.089042 | 738.43 | 1195 | 2 | 879, 970 | 6.033 | 6.032 |  |  | 6.033
|  |  |  | 400 | 34.7642 | 0.3185 | 5.6369 | 70.5774 | 244.911 | 24319 | 399.232 | 0.3352 | 29.3263 | 2.8793 | 1.4181 | 90.7931 | 0.085834 | 723.18 | 1196 | 3 | 1023 | 5.872 |  |  |  | 5.872
|  |  |  | 300 | 34.7195 | 0.2409 | 5.3307 | 66.5782 | 231.614 | 27358 | 299.691 | 0.2529 | 29.1757 | 2.7792 | 1.5355 | 90.8084 | 0.08509 | 711.76 | 1197 | 4 | 1022 | 5.531 |  |  |  | 5.531
|  |  |  | 200 | 34.593 | -0.0054 | 5.2371 | 64.9196 | 227.566 | 29932 | 200.003 | 0.002 | 28.8206 | 2.7541 | 1.8622 | 90.7428 | 0.082502 | 703.58 | 1198 | 5 | 1089 | 5.434 |  |  |  | 5.434
|  |  |  | 175 | 34.5206 | -0.1409 | 5.165 | 63.7624 | 224.447 | 31926 | 174.138 | -0.1348 | 28.6372 | 2.7232 | 1.9279 | 90.6696 | 0.082268 | 692.47 | 1199 | 6 | 1094 | 5.342 |  |  |  | 5.342
|  |  |  | 150 | 34.3966 | -0.4246 | 5.2566 | 64.3465 | 228.446 | 34121 | 150.128 | -0.4197 | 28.2921 | 2.7501 | 1.8599 | 90.4814 | 0.080874 | 686.46 | 1200 | 7 | 1127 | 5.472 |  |  |  | 5.472
|  |  |  | 125 | 34.1499 | -0.6702 | 5.2146 | 63.3025 | 226.66 | 35839 | 124.317 | -0.6665 | 27.889 | 2.7229 | 3.0455 | 90.435 | 0.077229 | 675.35 | 1201 | 8 | 1135 | 5.483 |  |  |  | 5.483
|  |  |  | 100 | 33.8827 | -0.9435 | 5.2927 | 63.6603 | 230.103 | 37437 | 100.79 | -0.9408 | 27.4521 | 2.7421 | 5.727 | 90.2362 | 0.077267 | 664.37 | 1202 | 9 | 1136 | 5.481 |  |  |  | 5.481
|  |  |  | 75 | 33.3544 | -1.3284 | 5.6651 | 67.1822 | 246.391 | 39019 | 74.329 | -1.3268 | 26.7349 | 2.8696 | 5.0415 | 90.0519 | 0.094076 | 652.81 | 1203 | 10 | 1145 | 5.855 |  |  |  | 5.855
|  |  |  | 50 | 32.6822 | -1.3399 | 6.2198 | 73.3956 | 270.663 | 40993 | 49.896 | -1.3389 | 26.2259 | 3.0881 | 46.291 | 88.8549 | 0.19927 | 647.66 | 1204 | 11 | 1155 | 6.452 |  |  |  | 6.452
|  |  |  | 40 | 32.4292 | -1.002 | 8.6172 | 102.4476 | 375.066 | 41938 | 41.689 | -1.0011 | 26.3045 | 4.0776 | 148.23 | 78.6168 | 0.62901 | 632.46 | 1205 | 12 | 1169, 1284 | 7.839 | 7.838 |  |  | 7.839
|  |  |  | 30 | 31.5325 | 0.1253 | 9.0889 | 110.7034 | 395.895 | 43515 | 30.189 | 0.1263 | 26.5142 | 4.361 | 14.21 | 88.6921 | 2.5242 | 616.09 | 1206 | 13 | 1654 | 9.604 |  |  | b | 9.604 | burette stalled, added titres, recreated noname file
|  |  |  | 20 | 30.7244 | 7.2686 | 6.7985 | 98.4177 | 296.499 | 44752 | 20.994 | 7.2704 | 31.6172 | 3.9306 | 11.571 | 88.4062 | 4.6505 | 613.42 | 1207 | 14 | 1742 | 8.418 |  |  |  | 8.418
|  |  |  | 10 | 30.409 | 7.8707 | 6.7132 | 98.3468 | 292.873 | 45714 | 10.548 | 7.8717 | 31.7888 | 3.9315 | 9.3324 | 88.2185 | 10.155 | 614.09 | 1208 | 15 | 1666 | 7.008 |  |  |  | 7.008
|  |  |  | 5 | 30.4084 | 7.883 | 6.746 | 98.8542 | 294.306 | 46473 | 6.067 | 7.8835 | 31.8089 | 3.9507 | 9.1956 | 87.9312 | 16.094 | 602.13 | 1209 | 16 | 1724 | 6.995 |  |  |  | 6.995
| 101 | 101 | BFB-8 | 500 | 34.7863 | 0.3316 | 5.8384 | 73.1453 | 253.661 | 20182 | 499.58 | 0.353 | 29.4029 | 2.9305 | 1.3501 | 89.8043 | 0.085276 | 123.09
|  |  |  | 500 | 34.7864 | 0.3316 | 5.8475 | 73.2594 | 254.056 | 20255 | 499.594 | 0.353 | 29.4029 | 2.9342 | 1.368 | 89.8048 | 0.087378 | 124.2
|  |  |  | 250 | 34.6646 | 0.1306 | 5.2962 | 65.9262 | 230.125 | 29003 | 250.836 | 0.1403 | 29.0155 | 2.7723 | 1.6835 | 90.8779 | 0.075639 | 111.71
|  |  |  | 120 | 34.6645 | 0.1304 | 5.3037 | 66.0187 | 230.449 | 29052 | 250.799 | 0.1401 | 29.0154 | 2.7754 | 1.7164 | 90.886 | 0.07642 | 110.78
|  |  |  | 120 | 34.075 | -0.7447 | 5.2003 | 62.9702 | 226.053 | 34315 | 120.454 | -0.7412 | 27.7691 | 2.7129 | 3.628 | 90.3688 | 0.073212 | 106.74
|  |  |  | 60 | 34.0752 | -0.7448 | 5.2077 | 63.0602 | 226.376 | 34362 | 120.517 | -0.7413 | 27.7694 | 2.7159 | 4.5421 | 90.3477 | 0.072698 | 108.46
|  |  |  | 60 | 32.9057 | -1.4055 | 5.9958 | 70.735 | 260.87 | 37168 | 59.889 | -1.4043 | 26.3415 | 2.9946 | 24.55 | 89.1105 | 0.076947 | 96.555
|  |  |  | 60 | 32.9058 | -1.4054 | 5.9952 | 70.7276 | 260.842 | 37227 | 59.933 | -1.4042 | 26.3415 | 2.9943 | 39.9 | 89.2396 | 0.075327 | 98.657
|  |  |  | 60 | 32.9055 | -1.4052 | 6.0037 | 70.8284 | 261.212 | 37275 | 60.032 | -1.404 | 26.3413 | 2.9977 | 33.439 | 89.2318 | 0.074475 | 101.97
|  |  |  | 60 | 32.9056 | -1.4051 | 6.0025 | 70.8145 | 261.16 | 37323 | 60.13 | -1.4039 | 26.3413 | 2.9972 | 35.254 | 88.9947 | 0.074833 | 105.98
|  |  |  | 60 | 32.9065 | -1.4057 | 5.9988 | 70.7696 | 260.997 | 37372 | 60.167 | -1.4044 | 26.3413 | 2.9957 | 47.795 | 89.2577 | 0.074857 | 108.46
|  |  |  | 43 | 32.3046 | -0.9329 | 9.3781 | 111.6077 | 408.225 | 39405 | 42.93 | -0.9319 | 26.2727 | 4.3881 | 148.21 | 74.6388 | 0.2334 | 102.13
|  |  |  | 43 | 32.3093 | -0.9423 | 9.3767 | 111.5669 | 408.165 | 39450 | 42.936 | -0.9414 | 26.2715 | 4.3867 | 148.21 | 74.5966 | 0.23811 | 100.43
|  |  |  | 43 | 32.3083 | -0.9407 | 9.3712 | 111.505 | 407.924 | 39496 | 42.931 | -0.9397 | 26.2699 | 4.3846 | 148.21 | 74.6478 | 0.23746 | 101.29
|  |  |  | 43 | 32.3095 | -0.9427 | 9.3653 | 111.4298 | 407.667 | 39542 | 42.918 | -0.9417 | 26.2696 | 4.3821 | 148.21 | 74.573 | 0.23612 | 102.24
|  |  |  | 43 | 32.3183 | -0.9273 | 9.3708 | 111.5483 | 407.902 | 39637 | 42.958 | -0.9263 | 26.2674 | 4.386 | 148.21 | 74.4789 | 0.23903 | 102.38
|  |  |  | 10 | 30.3936 | 7.7809 | 6.7208 | 98.244 | 293.204 | 41715 | 10.199 | 7.7818 | 31.7128 | 3.9284 | 10.99 | 88.3645 | 2.6174 | 98.022
|  |  |  | 10 | 30.3893 | 7.7863 | 6.7296 | 98.3828 | 293.591 | 41761 | 10.198 | 7.7872 | 31.7306 | 3.9331 | 9.7368 | 88.4008 | 2.762 | 98.249
|  |  |  | 5 | 30.338 | 7.7864 | 6.7346 | 98.4223 | 293.819 | 44883 | 5.611 | 7.787 | 31.6516 | 3.9365 | 10.541 | 87.8992 | 5.0983 | 98.434
|  |  |  | 5 | 30.3344 | 7.7854 | 6.7387 | 98.4773 | 293.998 | 44920 | 5.619 | 7.786 | 31.6482 | 3.9384 | 9.9684 | 87.9027 | 5.3501 | 95.965
|  |  |  | 5 | 30.3334 | 7.7856 | 6.7365 | 98.4451 | 293.903 | 44962 | 5.474 | 7.7861 | 31.6543 | 3.9373 | 9.8931 | 87.9037 | 4.7589 | 92.984
|  |  |  | 5 | 30.3282 | 7.7841 | 6.73 | 98.3444 | 293.623 | 45047 | 5.452 | 7.7846 | 31.6577 | 3.9339 | 10.3 | 87.9346 | 5.0083 | 91.487
|  |  |  | sufrace | 30.0489 | 7.8279 | 6.7795 | 98.9874 | 295.844 | 46594 | 1.925 | 7.828 | 31.4223 | 3.9573 | 8.8338 | 87.9714 | 26.237 | 98.2
|  |  |  | surface | 30.0453 | 7.8286 | 6.7887 | 99.1222 | 296.25 | 46652 | 2.065 | 7.8287 | 31.4215 | 3.9618 | 8.9348 | 87.8632 | 13.602 | 95.444
| 102 | 102 | BFB-8 | b-5 | 34.7726 | 0.3225 | 5.6699 | 71.0038 | 246.345 | 14409 | 412.926 | 0.3398 | 29.3415 | 2.8889 | 1.4188 | 87.5295 | 0.087648 | 1912.5 | 1234 | 1 | 1728 | 5.82 |  |  |  | 5.82
|  |  |  | 400 | 34.7568 | 0.3022 | 5.527 | 69.1681 | 240.137 | 16402 | 401.235 | 0.319 | 29.3068 | 2.8331 | 1.4697 | 89.3289 | 0.081801 | 1871.5 | 1235 | 2 | 1746 | 5.735 |  |  |  | 5.735
|  |  |  | 300 | 34.6598 | 0.0583 | 5.4637 | 67.882 | 237.401 | 21024 | 302.451 | 0.0699 | 28.9762 | 2.8214 | 1.5788 | 90.2798 | 0.073261 | 429.25 | 1236 | 3 | 1733, 1156 | 5.665 | 5.66 |  |  | 5.665
|  |  |  | 200 | 34.1148 | -0.9935 | 5.745 | 69.122 | 249.719 | 25061 | 203.321 | -0.9879 | 27.6299 | 2.8943 | 2.286 | 89.9935 | 0.071643 | 1923.4 | 1237 | 4 | 1509 | 5.941 |  |  |  | 5.941
|  |  |  | 175 | 33.9578 | -1.1758 | 5.8401 | 69.842 | 253.884 | 29033 | 177.603 | -1.1713 | 27.3523 | 2.9266 | 2.2845 | 89.8642 | 0.070299 | 1927.8 | 1238 | 5 | 1699 | 6.051 |  |  |  | 6.051
|  |  |  | 150 | 33.7438 | -1.2584 | 6.0451 | 72.0232 | 262.84 | 31002 | 150.979 | -1.2547 | 27.1155 | 3.0082 | 2.0894 | 89.3308 | 0.072572 | 430.81 | 1239 | 6 | 1137 | 6.202 |  |  |  | 6.202
|  |  |  | 125 | 33.3198 | -1.4459 | 6.2441 | 73.7957 | 271.581 | 32935 | 125.898 | -1.4432 | 26.641 | 3.0763 | 2.8818 | 88.9679 | 0.071036 | 652.95 | 1240 | 7 | 1143 | 6.377 |  |  |  | 6.377
|  |  |  | 100 | 33.0877 | -1.4966 | 6.3569 | 74.9026 | 276.536 | 35087 | 100.701 | -1.4946 | 26.4188 | 3.1221 | 3.2227 | 89.6387 | 0.075962 | 431.43 | 1241 | 8 | 1748 | 6.611 |  |  |  | 6.611
|  |  |  | 75 | 32.7556 | -1.4547 | 6.511 | 76.6299 | 283.316 | 37685 | 74.172 | -1.4532 | 26.1987 | 3.1896 | 2.8661 | 90.4991 | 0.15504 | 1947.3 | 1242 | 9 | 1293 | 6.723 |  |  |  | 6.723
|  |  |  | 50 | 32.1776 | -1.3951 | 7.998 | 93.9076 | 348.179 | 39410 | 50.192 | -1.3941 | 25.8293 | 3.7846 | 38.223 | 88.3459 | 1.0093 | 1945.7 | 1243 | 10 | 1740 | 8.284 |  |  |  | 8.284
|  |  |  | 40 | 32.0177 | -1.2453 | 8.8589 | 104.3282 | 385.709 | 40980 | 39.929 | -1.2445 | 25.8117 | 4.143 | 138.33 | 84.1788 | 4.4293 | 1941.2 | 1244 | 11 | 1726 | 9.123 |  |  | b | 9.123 | burette stalled, added titres, recreated noname file
|  |  |  | 30 | 31.8007 | -0.8649 | 9.1646 | 108.89 | 399.093 | 42920 | 30.492 | -0.8642 | 25.9423 | 4.3018 | 20.54 | 82.0653 | 20.465 | 431.26 | 1245 | 12 | 1744 | 9.498 |  |  |  | 9.498
|  |  |  | 20 | 31.321 | 3.1718 | 8.362 | 110.0497 | 364.36 | 44833 | 20.098 | 3.173 | 28.7829 | 4.3359 | 5.6223 | 87.812 | 54.989 | 1987.8 | 1246 | 13 | 1658, 1731 | 8.761 | 8.761 |  | b | 8.761 | dancing bubbles
|  |  |  | 10 | 30.6413 | 7.8099 | 6.8443 | 100.277 | 298.537 | 46495 | 10.333 | 7.8109 | 31.9729 | 3.9972 | 3.8753 | 88.6925 | 127.21 | 1881.6 | 1247 | 14 | 1739 | 7.08 |  |  |  | 7.08
|  |  |  | 5 | 30.6372 | 8.1198 | 6.7928 | 100.2314 | 296.305 | 48087 | 5.14 | 8.1203 | 32.2238 | 3.9973 | 3.2406 | 88.1859 | 254.94 | 431.3 | 1248 | 15 | 1750 | 7.023 |  |  |  | 7.023
| 103 | 103 | BFB-9 | b-5 | 28.6589 | -0.327 | 6.9497 | 82.0076 | 303.394 | 8467 | 210.817 | -0.3222 | 24.0563 | 3.3187 | 16.051 | 87.8862 | 0.091368 | 152.48
|  |  |  | 200 | 28.6577 | -0.3267 | 6.9474 | 81.9807 | 303.297 | 9944 | 200.73 | -0.3222 | 24.052 | 3.3215 | 19.565 | 88.0241 | 0.090296 | 150.21
|  |  |  | 175 | 28.65 | -0.3268 | 6.9099 | 81.5323 | 301.662 | 11616 | 176.244 | -0.3228 | 24.0357 | 3.3156 | 20.99 | 88.0596 | 0.080987 | 145.94
|  |  |  | 150 | 28.6258 | -0.3444 | 6.8324 | 80.5642 | 298.282 | 13583 | 151.518 | -0.3411 | 23.9943 | 3.2923 | 33.314 | 88.3444 | 0.07711 | 264.95
|  |  |  | 125 | 28.5471 | -0.3701 | 6.6188 | 77.948 | 288.974 | 15719 | 126.586 | -0.3674 | 23.9047 | 3.2137 | 12.719 | 88.9904 | 0.074081 | 139.41
|  |  |  | 100 | 28.4758 | -0.6268 | 5.9942 | 70.0665 | 261.717 | 17658 | 101.523 | -0.6248 | 23.6562 | 2.9574 | 4.0353 | 90.4005 | 0.072323 | 141.59
|  |  |  | 75 | 28.4154 | -0.7004 | 6.2165 | 72.489 | 271.438 | 19670 | 75.855 | -0.6991 | 23.5474 | 3.0476 | 6.485 | 90.5489 | 0.073034 | 144.7
|  |  |  | 50 | 28.346 | -0.725 | 6.4848 | 75.5296 | 283.167 | 21955 | 51.596 | -0.7241 | 23.4703 | 3.1589 | 12.128 | 90.1314 | 0.075744 | 159.15
|  |  |  | 40 | 28.3268 | -0.2849 | 6.7778 | 79.8794 | 295.967 | 24295 | 40.807 | -0.2841 | 23.7641 | 3.3096 | 31.116 | 89.1757 | 0.086786 | 328.15
|  |  |  | 40 | 28.327 | -0.2876 | 6.7869 | 79.9815 | 296.366 | 24334 | 40.845 | -0.2867 | 23.7524 | 3.3131 | 17.606 | 89.0823 | 0.091996 | 272.5
|  |  |  | 40 | 28.3265 | -0.2953 | 6.7878 | 79.9757 | 296.407 | 24373 | 40.888 | -0.2945 | 23.7547 | 3.3128 | 38.72 | 89.0443 | 0.087281 | 201.72
|  |  |  | 40 | 28.326 | -0.3023 | 6.7774 | 79.838 | 295.953 | 24416 | 40.939 | -0.3014 | 23.7428 | 3.3082 | 38.787 | 88.9918 | 0.089112 | 151.66
|  |  |  | 30 | 28.2365 | -0.2022 | 7.0292 | 82.9769 | 306.971 | 26152 | 31.569 | -0.2015 | 23.7473 | 3.4182 | 101.77 | 85.8972 | 0.13837 | 353.99
|  |  |  | 20 | 28.1113 | -0.0961 | 8.053 | 95.2522 | 351.715 | 28098 | 21.133 | -0.0956 | 23.7255 | 3.84 | 148.14 | 68.8696 | 0.94272 | 265.3
|  |  |  | 10 | 25.6018 | 1.8021 | 8.7271 | 106.7194 | 381.94 | 29253 | 10.994 | 1.8025 | 23.0357 | 4.2307 | 21.881 | 86.6921 | 3.192 | 154.66
|  |  |  | 5 | 24.7997 | 2.2371 | 8.7205 | 107.2705 | 381.899 | 30621 | 5.771 | 2.2373 | 22.6739 | 4.2509 | 17.233 | 87.2695 | 6.6924 | 326.74

## profile comparisons
| CTD | depth | WINKLER | offset
NP1 | 1.6284 | 136.604 | 1.669 | 0.0406
| 1.5433 | 99.845 | 1.624 | 0.08066
| 3.4522 | 75.193 | 3.575 | 0.12283
| 4.2955 | 49.999 | 4.431 | 0.13551
| 5.6658 | 25.087 | 5.935 | 0.26919
| 6.5918 | 10.218 | 6.795 | 0.2032
| 6.932 | 5.992 | 7.204 | 0.272
BS5 | CTD | depth | WINKLER | offset
| 5.4373 | 126.223 | 5.633
| 5.4329 | 126.452
| 5.4429 | 119.023 | 5.641
| 5.5033 | 101.136 | 5.707
| 5.5867 | 90.936
| 5.6665 | 76.017 | 5.894
| 6.1548 | 50.433 | 6.349
| 6.4476 | 39.504 | 6.684
| 6.448 | 39.783
| 6.4471 | 39.631
| 6.5525 | 33.492
| 6.5502 | 33.342
| 6.5629 | 30.387 | 6.713
| 6.9985 | 20.843 | 7.242
| 7.014 | 20.862
| 7.0265 | 20.736
| 6.9976 | 13.634 | 7.225
| 7.0073 | 10.383 | 7.227
| 7.1053 | 4.451 | 7.238
| 7.1049 | 4.462
| 7.1013 | 4.316
| 7.1044 | 3.196 | 7.235
BCL2 | 0.43783 | 1001.22 | 0.432
| 0.45224 | 800.528 | 0.447
| 0.59207 | 601.998 | 0.601
| 1.2645 | 400.125 | 1.304
| 2.2367 | 300.737 | 2.32
| 2.7371 | 251.534 | 2.868
| 3.2927 | 200.418 | 3.398
| 3.8065 | 175.575 | 3.962
| 4.3726 | 150.574 | 4.531
| 5.205 | 122.893 | 5.439
| 5.4271 | 100.342 | 5.633
| 5.4432 | 75.272 | 5.665
| 5.6267 | 52.638 | 5.852
| 6.2325 | 25.296 | 6.521 |  | air bubble
| 6.6814 | 11.264 | 6.954
| 6.7014 | 5.036 | 6.931
| CTD | depth | WINKLER | offset
SLIP4 | 5.5491 | 65.918 | 5.741 | 0.19195
| 5.5753 | 60.654 | 5.752 | 0.1767
| 5.5507 | 50.224 | 5.749 | 0.19827
| 6.7883 | 39.102 | 6.995 | 0.20673
| 7.3584 | 35.153 | 7.782 | 0.42357
| 7.6434 | 30.193 | 7.878 | 0.23461
| 7.8001 | 24.288 | 8.117 | 0.31689
| 8.8038 | 20.257 | 9.057 | 0.25325
| 9.141 | 15.425 | 9.357 | 0.21595
| 7.2896 | 10.404 | 7.645 | 0.35535
| 7.4011 | 4.553 | 7.596 | 0.19492
**35m sample seems like bigger offset from CTD, noted turbidity!
BRS1
| 7.9985 | 44.023 | 8.272 | 0.27349
| 8.0074 | 35.511 | 8.275 | 0.26756
| 8.027 | 31.212 | 8.269 | 0.24203
| 7.9947 | 24.569 | 8.219 | 0.22433
| 8.3017 | 19.565 | 8.543 | 0.24131
| 8.4639 | 14.515 | 8.745 | 0.28114
| 8.458 | 9.866 | 8.78 | 0.32199
| 8.6061 | 5.851 | 8.776 | 0.16994
entire water column turbid
BRS4 | 7.6643 | 44.961 | 7.942 | 0.27767
| 7.6592 | 41.122 | 7.944 | 0.2848
| 7.6543 | 35.858 | 7.951 | 0.2967
| 7.7037 | 30.556 | 7.952 | 0.24829
| 7.7248 | 26.261 | 7.975 | 0.25025
| 7.7095 | 20.378 | 7.965 | 0.25548 | flask pulled early
| 7.5969 | 13.842 | 7.866 | 0.26908
| 7.584 | 11.035 | 7.757 | 0.17296
| 6.6596 | 4.275 | 6.791 | 0.13145
UTN7 | 6.4488 | 53.041 | 6.881 | 0.43221
| 6.9019 | 45.563 | 7.156 | 0.25406
| 6.9084 | 40.654 | 7.182 | 0.27364
| 6.8917 | 35.717 | 7.147 | 0.25534
| 7.0778 | 30.612 | 7.328 | 0.25019
| 8.2897 | 25.734 | 8.473 | 0.18326
| 7.6865 | 20.546 | 8.011 | 0.32447
| 8.6429 | 15.517 | 8.808 | 0.16507 | stall OT
| 8.7944 | 10.589 | 9.08 | 0.28563
| 8.8596 | 5.604 | 9.086 | 0.22638 | stall OT
CCL1 | 8.0407 | 39.736 | 8.314 | 0.27325
| 8.0712 | 30.625 | 8.321 | 0.24984
| 8.0256 | 24.728 | 8.267 | 0.24143 | leaky Niskin
| 7.9807 | 20.34 | 8.197 | 0.21625
| 7.7651 | 13.901 | 7.953 | 0.18786
| 7.0881 | 9.492 | 7.497 | 0.40886
| 6.9033 | 5.231 | 7.174 | 0.27069
CCL3 | 8.0105 | 30.659 | 8.314 | 0.30354
| 8.1369 | 24.465 | 8.392 | 0.25514
| 7.9909 | 20.978 | 8.195 | 0.20412
| 7.1882 | 13.927 | 7.502 | 0.31381
| 7.0055 | 10.111 | 7.296 | 0.29054
| 7.1043 | 4.743 | 7.28 | 0.17569
BC1 | 7.9016 | 81.445 | 8.23 | 0.3284
| 7.9171 | 70.276 | 8.237 | 0.31986
| 7.9742 | 60.082 | 8.292 | 0.31777
| 7.9729 | 50.109 | 8.297 | 0.32409
| 7.9955 | 40.124 | 8.315 | 0.31945
| 7.9716 | 34.896 | 8.305 | 0.33343
| 7.9787 | 29.767 | 8.305 | 0.32626
| 7.987 | 24.816 | 8.331 | 0.34405
| 8.1989 | 20.592 | 8.5 | 0.30111
| 10.2954 | 15.18 | 10.423 | 0.12759
| 7.6275 | 10.933 | 8.112 | 0.48446
| 7.6979 | 5.023 | 7.942 | 0.2441
BC4 | 6.4796 | 582.173 | 6.725 | 0.24539
| 6.4687 | 500.435 | 6.733 | 0.26434
| 6.4548 | 400.766 | 6.712 | 0.25718
| 6.3714 | 300.878 | 6.63 | 0.25856
| 6.222 | 201.411 | 6.461 | 0.23902
| 6.0713 | 150.361 | 6.312 | 0.24075
| 7.1946 | 100.267 | 7.487 | 0.29241
| 7.7328 | 75.132 | 8.087 | 0.35424
| 8.1107 | 51.636 | 7.683 | -0.42772 | questionable
| 8.7142 | 24.373 | 9.53 | 0.81579
| 9.7534 | 13.2 | 10.148 | 0.39458
| 8.4678 | 5.609 | 8.953 | 0.48519
BFB3 | 6.6046 | 999.331 | 6.827 | 0.22236
| 6.5534 | 700.115 | 6.803 | 0.24961
| 6.4543 | 500.625 | 6.69 | 0.23566
| 6.389 | 400.453 | 6.63 | 0.24099
| 6.236 | 300.745 | 6.456 | 0.22004
| 6.2215 | 199.778 | 6.431 | 0.20947
| 6.2907 | 174.871 | 6.503 | 0.21231
| 6.6694 | 149.874 | 6.858 | 0.18858
| 6.81 | 125.919 | 7.012 | 0.20199
| 6.4942 | 100.25 | 6.749 | 0.25479
| 6.7944 | 74.95 | 7.094 | 0.29958
| 8.3445 | 50.312 | 8.611 | 0.26651
| 8.8178 | 40.403 | 9.246 | 0.42824
| 8.8565 | 30.118 | 9.195 | 0.33851
| 8.8247 | 20.477 | 9.152 | 0.32726
| 7.7573 | 11.103 | 8.614 | 0.85669
| 7.4704 | 4.761 | 7.589 | 0.11861
BFB5 | 6.5723 | 979.136 | 6.794 | 0.22174
| 6.5385 | 701.743 | 6.694 | 0.1555
| 6.4631 | 500.926 | 6.632 | 0.16892
| 6.4035 | 401.708 | 6.63 | 0.22648
| 6.2526 | 300.2 | 6.487 | 0.23442
| 6.2324 | 200.064 | 6.436 | 0.2036
| 6.1692 | 175.902 | 7.624 | 1.4548
| 6.0362 | 150.307 | 6.331 | 0.29479
| 6.0353 | 125.272 | 6.232 | 0.1967
| 6.0337 | 100.638 | 6.242 | 0.2083
| 6.4139 | 75.134 | 6.564 | 0.15006
| 8.7337 | 50.571 | 8.621 | -0.11272
| 9.0609 | 42.832
| 9.0825 | 40.039 | 9.405 | 0.32251
| 9.1792 | 30.171 | 9.489 | 0.30977
| 9.0237 | 20.775 | 9.377 | 0.3533
| 7.6237 | 10.637 | 8.215 | 0.59128
| 7.426 | 4.421 | 7.625 | 0.19896
BFB6
| 6.1146 | 175.462 | 6.322 | 0.20742
| 6.1726 | 149.598 | 6.368 | 0.19537
| 6.2032 | 125.929 | 6.421 | 0.21784
| 6.5707 | 101.15 | 6.68 | 0.10934
| 7.4577 | 74.707 | 7.573 | 0.11532
| 8.3218 | 50.521 | 8.564 | 0.24216
| 9.0708 | 40.275 | 9.328 | 0.25717
| 8.9177 | 30.739 | 9.339 | 0.42134
| 8.2935 | 20.919 | 8.587 | 0.29345
| 6.8803 | 9.76 | 7.142 | 0.26168
| 6.9527 | 5.827 | 7.07 | 0.11728
BFB7 | 5.8647 | 646.15 | 6.117
| 5.8072 | 500 | 6.033
| 5.6369 | 399.232 | 5.872
| 5.3307 | 299.691 | 5.531
| 5.2371 | 200.003 | 5.434
| 5.165 | 174.138 | 5.342
| 5.2566 | 150.128 | 5.472
| 5.2146 | 124.317 | 5.483
| 5.2927 | 100.79 | 5.481
| 5.6651 | 74.329 | 5.855
| 6.2198 | 49.896 | 6.452
| 8.6172 | 41.689 | 7.839
| 9.0889 | 30.189 | 9.604
| 6.7985 | 20.994 | 8.418
| 6.7132 | 10.548 | 7.008
| 6.746 | 6.067 | 6.995
BFB8 | 5.6699 | 412.926 | 5.82
| 5.527 | 401.235 | 5.735
| 5.4637 | 302.451 | 5.665
| 5.745 | 203.321 | 5.941
| 5.8401 | 177.603 | 6.051
| 6.0451 | 150.979 | 6.202
| 6.2441 | 125.898 | 6.377
| 6.3569 | 100.701 | 6.611
| 6.511 | 74.172 | 6.723
| 7.998 | 50.192 | 8.284
| 8.8589 | 39.929 | 9.123
| 9.1646 | 30.492 | 9.498
| 8.362 | 20.098 | 8.761
| 6.8443 | 10.333 | 7.08
| 6.7928 | 5.14 | 7.023

## chauvenet's criterion
Chauvenet's criterion allows us to determine whether outliers can be removed based on statistical improbability
We take the difference between the replicates and then calculate a mean and standard deviation. The Chauvenet criterion is then calculated for each outlier
The Chauvenet criterion is the difference between the mean and the outlier, divided by the standard deviation. If this value is greater than the critical value for a given n, the outlier can be discarded
Based on Chauvenet's criterion we can reject the two biggest outliers (-0.091, -0.061)
DUPLICATES | [9 empty] | sorted in ascending order
Station | LAT | LONG | depth | OXY-1 | OXY-2 | diff. (di) | test-avg. # (δ) | δ/σ | reject? | diff. (di) | freq.
NP6 | 53.49 | -139.7579 | 25 | 6.998 | 6.997 | -0.001 |  |  |  | -0.091 | 1 | n | Chauvenet
NP7 | 53.0409 | -140.2243 | 400 | 0.919 | 0.922 | 0.003 |  |  |  | -0.061 | 1 | 3 | 1.37
NP7 | 53.0409 | -140.2243 | 50 | 7.25 | 7.245 | -0.005 |  |  |  | -0.02 | 1 | 4 | 1.53
NP9 | 53.2507 | -148.0538 | 600 | 0.299 | 0.299 | 0 |  |  |  | -0.02 | 2 | 5 | 1.64
NP9 | 53.2507 | -148.0538 | 100 | 2.586 | 2.58 | -0.006 |  |  |  | -0.018 | 1 | 6 | 1.73
NP10 | 53.3262 | -150.1838 | 175 | 0.814 | 0.812 | -0.002 |  |  |  | -0.014 | 1 | 7 | 1.8
NP10 | 53.3262 | -150.1838 | 10 | 6.958 | 6.961 | 0.003 |  |  |  | -0.012 | 1 | 8 | 1.87
NP11 | 53.6352 | -153.7618 | 600 | 0.289 | 0.295 | 0.006 |  |  |  | -0.011 | 1 | 9 | 1.91
NP11 | 53.6352 | -153.7618 | 25 | 7.429 | 7.426 | -0.003 |  |  |  | -0.009 | 1 | 10 | 1.97
NP12 | 53.7119 | -156.1256 | 125 | 3.01 | 3.015 | 0.005 |  |  |  | -0.008 | 1 | 20 | 2.24
NP12 | 53.7119 | -156.1256 | 10 | 6.746 | 6.747 | 0.001 |  |  |  | -0.008 | 2 | 30 | 2.39
NP13 | 53.7178 | -159.326 | 1000 | 0.364 | 0.364 | 0 |  |  |  | -0.007 | 1 | 50 | 2.57
NP13 | 53.7178 | -159.326 | 5 | 7.227 | 7.206 |  |  |  |  | -0.006 | 1 | 80 | 2.74
NP14 | 53.7236 | -161.1096 | 250 | 1.5 | 1.502 | 0.002 |  |  |  | -0.006 | 2 | 100 | 2.81
NP14 | 53.7236 | -161.1096 | 50 | 6.434 | 6.431 | -0.003 |  |  |  | -0.005 | 1
UN8 | 53.6412 | -163.7536 | 600 | 0.348 | 0.351 | 0.003 |  |  |  | -0.005 | 2
UN8 | 53.6412 | -163.7536 | 25 | 7.026 | 7.017 | -0.009 |  |  |  | -0.004 | 1
UN7 | 53.7336 | -163.9314 | 400 | 0.735 | 0.736 | 0.001 |  |  |  | -0.004 | 2
UN7 | 53.7336 | -163.9314 | 50 | 6.497 | 6.497 | 0 |  |  |  | -0.004 | 3
UN6 | 53.8073 | -164.0626 | 20 | 7.463 | 7.459 | -0.004 |  |  |  | -0.004 | 4
UN5 | 53.8849 | -164.1991 | 15 | 7.495 | 7.498 | 0.003 |  |  |  | -0.003 | 1
UN4 | 53.9598 | -164.3328 | 40 | 6.765 | 6.764 | -0.001 |  |  |  | -0.003 | 2
UN3 | 54.0379 | -164.4581 | 50 | 6.784 | 6.79 | 0.006 |  |  |  | -0.003 | 3
UN2 | 54.1115 | -164.5974 | 40 | 7.328 | 7.333 | 0.005 |  |  |  | -0.003 | 4
UN1 | 54.188 | -164.7269 | 10 | 7.149 | 7.145 | -0.004 |  |  |  | -0.002 | 1
BCL1 | 54.1754 | -166.6759 | 1000 | 0.543 | 0.543 | 0 |  |  |  | -0.002 | 2
BCL1 | 54.1754 | -166.6759 | 125 | 5.349 | 5.342 | -0.007 |  |  |  | -0.002 | 3
BCL2 | 55.0607 | -170.2157 | 400 | 1.304 | 1.315 | 0.011 |  |  |  | -0.002 | 4
BCL2 | 55.0607 | -170.2157 | 75 | 5.665 | 5.677 | 0.012 |  |  |  | -0.001 | 1
BS0 | 55.7069 | -172.9497 | 1000 | 0.424 | 0.422 | -0.002 |  |  |  | -0.001 | 2
BS0 | 55.7069 | -172.9497 | 100 | 6.714 | 6.694 | -0.02 | -0.018709 | -1.3706 | no | -0.001 | 3
BS1 | 56.3145 | -172.8243 | 80 | 5.536 | 5.543 | 0.007 |  |  |  | -0.001 | 4
BS1 | 56.3145 | -172.8243 | 3.4 | 7.326 | 7.328 | 0.002 |  |  |  | -0.001 | 5
BS5 | 56.6513 | -172.7341 | 126 | 5.633 | 5.634 | 0.001 |  |  |  | -0.001 | 6
BS5 | 56.6513 | -172.7341 | 20 | 7.242 | 7.26 | 0.018 |  |  |  | -0.001 | 7
BCL3 | 57.9603 | -173.0919 | 60 | 6.38 | 6.38 | 0 |  |  |  | -0.001 | 8
BCL4 | 59.3215 | -173.4999 | 30 | 8.754 | 8.723 |  |  |  |  | 0 | 1
BCL5 | 60.6478 | -174.15 | 30 | 8.47 | 8.609 |  |  |  |  | 0 | 2
SLIP1 | 62.0147 | -175.0499 | 60 | 6.336 | 6.325 |  |  |  |  | 0 | 3
SLIP1 | 62.0147 | -175.0499 | 20 | 8.778 | 8.788 | 0.01 |  |  |  | 0 | 4
SLIP2 | 62.0499 | -175.203 | 20 | 9.369 | 9.351 | -0.018 |  |  |  | 0 | 5
SLIP3 | 62.3929 | -174.5711 | 60 | 5.839 | 5.831 | -0.008 |  |  |  | 0 | 6
SLIP3 | 62.3929 | -174.5711 | 15 | 7.925 | 7.929 | 0.004 |  |  |  | 0 | 7
SLIP4 | 63.029 | -173.457 | 50 | 5.749 | 5.759 | 0.01 |  |  |  | 0 | 8
BCL6 | 63.9983 | -171.9907 | 30 | 7.476 | 7.48 | 0.004 |  |  |  | 0 | 9
UTBS5 | 64.6658 | -169.9219 | 25 | 7.109 | 7.105 | -0.004 |  |  |  | 0.001 | 1
UTBS1 | 64.9909 | -169.1371 | 5 | 7.415 | 7.417 | 0.002 |  |  |  | 0.001 | 2
BRS1 | 65.7163 | -168.897 |  | 8.219 | 8.221 | 0.002 |  |  |  | 0.001 | 3
BRS2 | 65.6999 | -168.7173 | 25 | 8.458 | 8.452 | -0.006 |  |  |  | 0.001 | 4
BRS3 | 65.6823 | -168.565 | 25 | 7.988 | 7.977 | -0.011 |  |  |  | 0.002 | 1
BRS4 | 65.6777 | -168.4045 | 30 | 7.952 | 7.951 | -0.001 |  |  |  | 0.002 | 2
BRS5 | 65.6615 | -168.2193 | 25 | 7.567 | 7.57 | 0.003 |  |  |  | 0.002 | 3
UTN1 | 66.7092 | -168.3971 | 10 | 6.952 | 6.952 | 0 |  |  |  | 0.002 | 4
UTN4 | 67.5005 | -168.9075 | 40 | 8.525 | 8.532 | 0.007 |  |  |  | 0.002 | 5
UTN4 | 67.5005 | -168.9075 | 15 | 8.595 | 8.606 | 0.011 |  |  |  | 0.002 | 6
UTN7 | 68.0007 | -165.9324 | 30 | 7.328 | 7.32 | -0.008 |  |  |  | 0.002 | 7
CCL1 | 68.3729 | -167.1776 | 30 | 8.321 | 8.318 | -0.003 |  |  |  | 0.003 | 1
CCL2 | 68.9579 | -168.5236 | 51 | 8.975 | 8.978 | 0.003 |  |  |  | 0.003 | 2
CCL2 | 68.9579 | -168.5236 | 45 | 8.974 | 8.983 | 0.009 |  |  |  | 0.003 | 3
CCL2 | 68.9579 | -168.5236 | 40 | 8.984 | 8.986 | 0.002 |  |  |  | 0.003 | 4
CCL2 | 68.9579 | -168.5236 | 35 | 9.009 | 9.023 | 0.014 |  |  |  | 0.003 | 5
CCL2 | 68.9579 | -168.5236 | 30 | 8.999 | 8.995 | -0.004 |  |  |  | 0.003 | 6
CCL2 | 68.9579 | -168.5236 | 25 | 9.245 | 9.245 | 0 |  |  |  | 0.003 | 7
CCL2 | 68.9579 | -168.5236 | 20 | 9.081 | 9.091 | 0.01 |  |  |  | 0.003 | 8
CCL2 | 68.9579 | -168.5236 | 15 | 9.001 | 9 | -0.001 |  |  |  | 0.004 | 1
CCL2 | 68.9579 | -168.5236 | 10 | 7.84 | 7.849 | 0.009 |  |  |  | 0.004 | 2
CCL2 | 68.9579 | -168.5236 | 5 | 7.531 | 7.533 | 0.002 |  |  |  | 0.004 | 3
CCL3 | 68.9655 | -166.5327 | 15 | 7.502 | 7.441 | -0.061 | -0.059709 | -4.374 | yes | 0.005 | 1
CCL4 | 69.991 | -168.02 | 20 | 7.779 | 7.778 | -0.001 |  |  |  | 0.005 | 2
BC1 | 71.0672 | -159.3526 | 60 | 8.292 | 8.289 | -0.003 |  |  |  | 0.005 | 3
BC1 | 71.0672 | -159.3526 | 10 | 8.112 | 8.1 | -0.012 |  |  |  | 0.006 | 1
BC2 | 71.4137 | -157.4971 | 50 | 7.112 | 7.115 | 0.003 |  |  |  | 0.006 | 2
BC3 | 71.5778 | -156.0236 | 75 | 7.523 | 7.535 | 0.012 |  |  |  | 0.007 | 1
BC4 | 71.933 | -154.8923 | 500 | 6.733 | 6.732 | -0.001 |  |  |  | 0.007 | 2
BC4 | 71.933 | -154.8923 | 25 | 9.53 | 9.599 |  |  |  |  | 0.009 | 1
BFB2 | 71.207 | -148.2912 | 600 | 6.757 | 6.762 | 0.005 |  |  |  | 0.009 | 2
BFB2 | 71.207 | -148.2912 | 20 | 8.554 | 8.555 | 0.001 |  |  |  | 0.01 | 1
BFB3 | 70.6337 | -142.2588 | 400 | 6.63 | 6.633 | 0.003 |  |  |  | 0.01 | 2
BFB3 | 70.6337 | -142.2588 | 30 | 9.195 | 9.193 | -0.002 |  |  |  | 0.01 | 3
MK1 | 70.2298 | -140.004 | 150 | 7.112 | 7.116 | 0.004 |  |  |  | 0.011 | 1
BFB4 | 70.3188 | -137.8368 | 248 | 6.534 | 6.514 | -0.02 |  |  |  | 0.011 | 2
BFB4 | 70.3188 | -137.8368 | 20 | 9.263 | 9.261 | -0.002 |  |  |  | 0.012 | 1
BFB5 | 71.3332 | -133.7479 | 700 | 6.785 | 6.694 | -0.091 | -0.089709 | -6.5717 | yes | 0.012 | 2
BFB5 | 71.3332 | -133.7479 | 20 | 9.377 | 9.39 | 0.013 |  |  |  | 0.013 | 1
BFB6 | 70.8261 | -127.4236 | 150 | 6.368 | 6.37 | 0.002 |  |  |  | 0.014 | 1
BFB6 | 70.8261 | -127.4236 | 20 | 8.587 | 8.573 | -0.014 |  |  |  | 0.018 | 1
BFB6 | 70.8261 | -127.4236 | 10 | 7.142 | 7.142 | 0
BFB7 | 70.5436 | -122.9154 | 500 | 6.033 | 6.032 | -0.001
BFB7 | 70.5436 | -122.9154 | 40 | 7.839 | 7.838 | -0.001
BFB8 | 69.4642 | -118.3323 | 300 | 5.665 | 5.66 | -0.005
BFB8 | 69.4642 | -118.3323 | 20 | 8.761 | 8.761 | 0
[5 empty] | Count | 86
[5 empty] | Average | -0.0012907
[5 empty] | Std dev pop (σ) | 0.013651
| Chauvenet's critical value for n=80 |  |  |  | δ/σ for 80 | 2.74

## precision
[6 empty] | ALL |  | removed commented reds |  | removed outliers
DUPLICATES | [15 empty] | TRIPLICATES
Station | LAT | LONG | depth | OXY-1 | OXY-2 | diff. (di) | di^2 | diff. (di) | di^2 | diff. (di) | di^2 | comments | SUM(di)^2 | Sp |  | Station | LAT | LONG | depth | OXY-1 | OXY-2 | OXY-3 | mean | SS | Sp
NP6 | 53.49 | -139.7579 | 25 | 6.998 | 6.997 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06 |  | 0.041743 | 0.015145 | for all dups | NP1 | 49.1175 | -126.6903 | 10 | 6.795 | 6.798 | 6.796 | 6.7963 | 4.6667e-06 | 0.0025459
NP7 | 53.0409 | -140.2243 | 400 | 0.919 | 0.922 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06 |  |  |  |  | NP2 | 49.6288 | -128.3708 | 125 | 5.15 | 5.149 | 5.149 | 5.1493 | 6.6667e-07
NP7 | 53.0409 | -140.2243 | 50 | 7.25 | 7.245 | -0.005 | 2.5e-05 | -0.005 | 2.5e-05 | -0.005 | 2.5e-05 |  | 0.016138 | 0.0096864 | excluding samples with known problems | NP3 | 50.4673 | -131.4968 | 125 | 3.719 | 3.718 | 3.719 | 3.7187 | 6.6667e-07
NP9 | 53.2507 | -148.0538 | 600 | 0.299 | 0.299 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  |  | NP4 | 50.9829 | -133.6279 | 50 | 7.115 | 7.112 | 7.111 | 7.1127 | 8.6667e-06
NP9 | 53.2507 | -148.0538 | 100 | 2.586 | 2.58 | -0.006 | 3.6e-05 | -0.006 | 3.6e-05 | -0.006 | 3.6e-05 |  | 0.004167 | 0.0049803 | excluding outliers | NP5 | 51.6425 | -136.657 | 2000 | 1.261 | 1.258 | 1.262 | 1.2603 | 8.6667e-06
NP10 | 53.3262 | -150.1838 | 175 | 0.814 | 0.812 | -0.002 | 4e-06 | -0.002 | 4e-06 | -0.002 | 4e-06 |  |  |  |  | NP8 | 52.6198 | -142.4771 | 125 | 3.013 | 3.014 | 3.011 | 3.0127 | 4.6667e-06
NP10 | 53.3262 | -150.1838 | 10 | 6.958 | 6.961 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06 |  |  |  |  | BL1 | 71.3259 | -152.2153 | 20 | 9.169 | 9.168 | 9.18 | 9.1723 | 8.8667e-05 | stalled burette, 1,2-added titres, dosimat discrepancy
NP11 | 53.6352 | -153.7618 | 600 | 0.289 | 0.295 | 0.006 | 3.6e-05 | 0.006 | 3.6e-05 | 0.006 | 3.6e-05
NP11 | 53.6352 | -153.7618 | 25 | 7.429 | 7.426 | -0.003 | 9e-06 | -0.003 | 9e-06 | -0.003 | 9e-06 | [11 empty] | SUM | 0.00011667
NP12 | 53.7119 | -156.1256 | 125 | 3.01 | 3.015 | 0.005 | 2.5e-05 | 0.005 | 2.5e-05 | 0.005 | 2.5e-05
NP12 | 53.7119 | -156.1256 | 10 | 6.746 | 6.747 | 0.001 | 1e-06 | 0.001 | 1e-06 | 0.001 | 1e-06 |  |  |  | Of the 751 unique sample numbers analyzed during this cruise 91 were duplicate pairs and 7 triplicates (total samples run=856)
NP13 | 53.7178 | -159.326 | 1000 | 0.364 | 0.364 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | All replicates were run at the same time except for those highlighted in blue (Stn. CCL2), which were run 6 days after collection
NP13 | 53.7178 | -159.326 | 5 | 7.227 | 7.206 | -0.021 | 0.000441 |  |  |  |  | 2-possibly sampled from wrong bottle |  |  | Pooled standard deviations were calculated separately for duplicates and triplicates.
NP14 | 53.7236 | -161.1096 | 250 | 1.5 | 1.502 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
NP14 | 53.7236 | -161.1096 | 50 | 6.434 | 6.431 | -0.003 | 9e-06 | -0.003 | 9e-06 | -0.003 | 9e-06 |  |  |  | Samples in red that had possible problems noted for one or more of the reps were excluded from the analysis
UN8 | 53.6412 | -163.7536 | 600 | 0.348 | 0.351 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06 |  |  |  | Samples in purple were excluded as outliers based on statistical improbability according to Chauvenet's criterion.
UN8 | 53.6412 | -163.7536 | 25 | 7.026 | 7.017 | -0.009 | 8.1e-05 | -0.009 | 8.1e-05 | -0.009 | 8.1e-05 |  |  |  | The resultant pooled standard deviations for the duplicates and triplicates is 0.005 and 0.003 respectively.
UN7 | 53.7336 | -163.9314 | 400 | 0.735 | 0.736 | 0.001 | 1e-06 | 0.001 | 1e-06 | 0.001 | 1e-06 |  |  |  | Samples where the burette stalled on one or both replicates were not excluded from the analysis as the reproducibility was good.
UN7 | 53.7336 | -163.9314 | 50 | 6.497 | 6.497 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | The pooled standard deviation for these samples is 0.006
UN6 | 53.8073 | -164.0626 | 20 | 7.463 | 7.459 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05
UN5 | 53.8849 | -164.1991 | 15 | 7.495 | 7.498 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06
UN4 | 53.9598 | -164.3328 | 40 | 6.765 | 6.764 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
UN3 | 54.0379 | -164.4581 | 50 | 6.784 | 6.79 | 0.006 | 3.6e-05 | 0.006 | 3.6e-05 | 0.006 | 3.6e-05 |  |  |  |  | Replicates where 1 or both samples experienced a stalled burette
UN2 | 54.1115 | -164.5974 | 40 | 7.328 | 7.333 | 0.005 | 2.5e-05 | 0.005 | 2.5e-05 | 0.005 | 2.5e-05 |  |  |  |  | Station | Sample # | LAT | LONG | depth | OXY-1 | OXY-2 | OXY-3 | mean | SS | Sp
UN1 | 54.188 | -164.7269 | 10 | 7.149 | 7.145 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05 |  |  |  |  | SLIP1 | 681 | 62.0147 | -175.0499 | 20 | 8.778 | 8.788 |  | 8.783 | 5e-05 | 0.0062153 | stalled burette 1-added titres, 2-finished titrating
BCL1 | 54.1754 | -166.6759 | 1000 | 0.543 | 0.543 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  |  | CCL2 | 929 | 68.9579 | -168.5236 | 51 | 8.975 | 8.978 |  | 8.9765 | 4.5e-06 |  | stalled burette, 1-OT, 2-add titres
BCL1 | 54.1754 | -166.6759 | 125 | 5.349 | 5.342 | -0.007 | 4.9e-05 | -0.007 | 4.9e-05 | -0.007 | 4.9e-05 |  |  |  |  | CCL2 | 932 | 68.9579 | -168.5236 | 35 | 9.009 | 9.023 |  | 9.016 | 9.8e-05 |  | 2-stalled burette, add titres, run 4 days later
BCL2 | 55.0607 | -170.2157 | 400 | 1.304 | 1.315 | 0.011 | 0.000121 | 0.011 | 0.000121 | 0.011 | 0.000121 |  |  |  |  | CCL2 | 933 | 68.9579 | -168.5236 | 30 | 8.999 | 8.995 |  | 8.997 | 8e-06 |  | 1-stalled burette, OT, 2-run 4 days later
BCL2 | 55.0607 | -170.2157 | 75 | 5.665 | 5.677 | 0.012 | 0.000144 | 0.012 | 0.000144 | 0.012 | 0.000144 |  |  |  |  | CCL2 | 934 | 68.9579 | -168.5236 | 25 | 9.245 | 9.245 |  | 9.245 | 0 |  | 1-stalled burette, add titres
BS0 | 55.7069 | -172.9497 | 1000 | 0.424 | 0.422 | -0.002 | 4e-06 | -0.002 | 4e-06 | -0.002 | 4e-06 |  |  |  |  | BFB2 | 1070 | 71.207 | -148.2912 | 20 | 8.554 | 8.555 |  | 8.5545 | 5e-07 |  | 1-stalled burette, finished titration
BS0 | 55.7069 | -172.9497 | 100 | 6.714 | 6.694 | -0.02 | 0.0004 | -0.02 | 0.0004 | -0.02 | 0.0004 |  |  |  |  | BFB6 | 1191 | 70.8261 | -127.4236 | 20 | 8.587 | 8.573 |  | 8.58 | 9.8e-05 |  | 2-stalled burette, finished titrating
BS1 | 56.3145 | -172.8243 | 80 | 5.536 | 5.543 | 0.007 | 4.9e-05 | 0.007 | 4.9e-05 | 0.007 | 4.9e-05 |  |  |  |  | BL1 | 1029 | 71.3259 | -152.2153 | 20 | 9.169 | 9.168 | 9.18 | 9.1723 | 8.8667e-05 |  | stalled burette, 1,2-added titres, dosimat discrepancy
BS1 | 56.3145 | -172.8243 | 3.4 | 7.326 | 7.328 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
BS5 | 56.6513 | -172.7341 | 126 | 5.633 | 5.634 | 0.001 | 1e-06 | 0.001 | 1e-06 | 0.001 | 1e-06 | [12 empty] | SUM | 0.00034767
BS5 | 56.6513 | -172.7341 | 20 | 7.242 | 7.26 | 0.018 | 0.000324 | 0.018 | 0.000324 | 0.018 | 0.000324
BCL3 | 57.9603 | -173.0919 | 60 | 6.38 | 6.38 | 0 | 0 | 0 | 0 | 0 | 0
BCL4 | 59.3215 | -173.4999 | 30 | 8.754 | 8.723 | -0.031 | 0.000961 |  |  |  |  | 1-floaties, 2-dancing bubbles
BCL5 | 60.6478 | -174.15 | 30 | 8.47 | 8.609 | 0.139 | 0.019321 |  |  |  |  | very turbid sample
SLIP1 | 62.0147 | -175.0499 | 60 | 6.336 | 6.325 | -0.011 | 0.000121 |  |  |  |  | turbid
SLIP1 | 62.0147 | -175.0499 | 20 | 8.778 | 8.788 | 0.01 | 0.0001 | 0.01 | 0.0001 | 0.01 | 0.0001 | stalled burette 1-added titres, 2-finished titrating
SLIP2 | 62.0499 | -175.203 | 20 | 9.369 | 9.351 | -0.018 | 0.000324 | -0.018 | 0.000324 | -0.018 | 0.000324
SLIP3 | 62.3929 | -174.5711 | 60 | 5.839 | 5.831 | -0.008 | 6.4e-05 | -0.008 | 6.4e-05 | -0.008 | 6.4e-05
SLIP3 | 62.3929 | -174.5711 | 15 | 7.925 | 7.929 | 0.004 | 1.6e-05 | 0.004 | 1.6e-05 | 0.004 | 1.6e-05
SLIP4 | 63.029 | -173.457 | 50 | 5.749 | 5.759 | 0.01 | 0.0001 | 0.01 | 0.0001 | 0.01 | 0.0001
BCL6 | 63.9983 | -171.9907 | 30 | 7.476 | 7.48 | 0.004 | 1.6e-05 | 0.004 | 1.6e-05 | 0.004 | 1.6e-05
UTBS5 | 64.6658 | -169.9219 | 25 | 7.109 | 7.105 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05
UTBS1 | 64.9909 | -169.1371 | 5 | 7.415 | 7.417 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
BRS1 | 65.7163 | -168.897 |  | 8.219 | 8.221 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
BRS2 | 65.6999 | -168.7173 | 25 | 8.458 | 8.452 | -0.006 | 3.6e-05 | -0.006 | 3.6e-05 | -0.006 | 3.6e-05
BRS3 | 65.6823 | -168.565 | 25 | 7.988 | 7.977 | -0.011 | 0.000121 | -0.011 | 0.000121 | -0.011 | 0.000121
BRS4 | 65.6777 | -168.4045 | 30 | 7.952 | 7.951 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
BRS5 | 65.6615 | -168.2193 | 25 | 7.567 | 7.57 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06
UTN1 | 66.7092 | -168.3971 | 10 | 6.952 | 6.952 | 0 | 0 | 0 | 0 | 0 | 0
UTN4 | 67.5005 | -168.9075 | 40 | 8.525 | 8.532 | 0.007 | 4.9e-05 | 0.007 | 4.9e-05 | 0.007 | 4.9e-05
UTN4 | 67.5005 | -168.9075 | 15 | 8.595 | 8.606 | 0.011 | 0.000121 | 0.011 | 0.000121 | 0.011 | 0.000121
UTN7 | 68.0007 | -165.9324 | 30 | 7.328 | 7.32 | -0.008 | 6.4e-05 | -0.008 | 6.4e-05 | -0.008 | 6.4e-05
CCL1 | 68.3729 | -167.1776 | 30 | 8.321 | 8.318 | -0.003 | 9e-06 | -0.003 | 9e-06 | -0.003 | 9e-06
CCL2 | 68.9579 | -168.5236 | 51 | 8.975 | 8.978 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06 | stalled burette, 1-OT, 2-add titres
CCL2 | 68.9579 | -168.5236 | 45 | 8.974 | 8.983 | 0.009 | 8.1e-05 | 0.009 | 8.1e-05 | 0.009 | 8.1e-05
CCL2 | 68.9579 | -168.5236 | 40 | 8.984 | 8.986 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
CCL2 | 68.9579 | -168.5236 | 35 | 9.009 | 9.023 | 0.014 | 0.000196 | 0.014 | 0.000196 | 0.014 | 0.000196 | 2-stalled burette, add titres, run 4 days later
CCL2 | 68.9579 | -168.5236 | 30 | 8.999 | 8.995 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05 | -0.004 | 1.6e-05 | 1-stalled burette, OT, 2-run 4 days later
CCL2 | 68.9579 | -168.5236 | 25 | 9.245 | 9.245 | 0 | 0 | 0 | 0 | 0 | 0 | 1-stalled burette, add titres
CCL2 | 68.9579 | -168.5236 | 20 | 9.081 | 9.091 | 0.01 | 0.0001 | 0.01 | 0.0001 | 0.01 | 0.0001
CCL2 | 68.9579 | -168.5236 | 15 | 9.001 | 9 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
CCL2 | 68.9579 | -168.5236 | 10 | 7.84 | 7.849 | 0.009 | 8.1e-05 | 0.009 | 8.1e-05 | 0.009 | 8.1e-05
CCL2 | 68.9579 | -168.5236 | 5 | 7.531 | 7.533 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
CCL3 | 68.9655 | -166.5327 | 15 | 7.502 | 7.441 | -0.061 | 0.003721 | -0.061 | 0.003721
CCL4 | 69.991 | -168.02 | 20 | 7.779 | 7.778 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
BC1 | 71.0672 | -159.3526 | 60 | 8.292 | 8.289 | -0.003 | 9e-06 | -0.003 | 9e-06 | -0.003 | 9e-06
BC1 | 71.0672 | -159.3526 | 10 | 8.112 | 8.1 | -0.012 | 0.000144 | -0.012 | 0.000144 | -0.012 | 0.000144
BC2 | 71.4137 | -157.4971 | 50 | 7.112 | 7.115 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06
BC3 | 71.5778 | -156.0236 | 75 | 7.523 | 7.535 | 0.012 | 0.000144 | 0.012 | 0.000144 | 0.012 | 0.000144
BC4 | 71.933 | -154.8923 | 500 | 6.733 | 6.732 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
BC4 | 71.933 | -154.8923 | 25 | 9.53 | 9.599 | 0.069 | 0.004761 |  |  |  |  | 2- big air bubble in sample
BFB2 | 71.207 | -148.2912 | 600 | 6.757 | 6.762 | 0.005 | 2.5e-05 | 0.005 | 2.5e-05 | 0.005 | 2.5e-05
BFB2 | 71.207 | -148.2912 | 20 | 8.553 | 8.554 | 0.001 | 1e-06 | 0.001 | 1e-06 | 0.001 | 1e-06 | 1-stalled burette, finished titration
BFB3 | 70.6337 | -142.2588 | 400 | 6.63 | 6.633 | 0.003 | 9e-06 | 0.003 | 9e-06 | 0.003 | 9e-06
BFB3 | 70.6337 | -142.2588 | 30 | 9.195 | 9.193 | -0.002 | 4e-06 | -0.002 | 4e-06 | -0.002 | 4e-06
MK1 | 70.2298 | -140.004 | 150 | 7.112 | 7.116 | 0.004 | 1.6e-05 | 0.004 | 1.6e-05 | 0.004 | 1.6e-05
BFB4 | 70.3188 | -137.8368 | 248 | 6.534 | 6.514 | -0.02 | 0.0004 | -0.02 | 0.0004 | -0.02 | 0.0004
BFB4 | 70.3188 | -137.8368 | 20 | 9.263 | 9.261 | -0.002 | 4e-06 | -0.002 | 4e-06 | -0.002 | 4e-06
BFB5 | 71.3332 | -133.7479 | 700 | 6.785 | 6.694 | -0.091 | 0.008281 | -0.091 | 0.008281
BFB5 | 71.3332 | -133.7479 | 20 | 9.377 | 9.39 | 0.013 | 0.000169 | 0.013 | 0.000169 | 0.013 | 0.000169
BFB6 | 70.8261 | -127.4236 | 150 | 6.368 | 6.37 | 0.002 | 4e-06 | 0.002 | 4e-06 | 0.002 | 4e-06
BFB6 | 70.8261 | -127.4236 | 20 | 8.587 | 8.573 | -0.014 | 0.000196 | -0.014 | 0.000196 | -0.014 | 0.000196 | 2-stalled burette, finished titrating
BFB6 | 70.8261 | -127.4236 | 10 | 7.142 | 7.142 | 0 | 0 | 0 | 0 | 0 | 0
BFB7 | 70.5436 | -122.9154 | 500 | 6.033 | 6.032 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
BFB7 | 70.5436 | -122.9154 | 40 | 7.839 | 7.838 | -0.001 | 1e-06 | -0.001 | 1e-06 | -0.001 | 1e-06
BFB8 | 69.4642 | -118.3323 | 300 | 5.665 | 5.66 | -0.005 | 2.5e-05 | -0.005 | 2.5e-05 | -0.005 | 2.5e-05
BFB8 | 69.4642 | -118.3323 | 20 | 8.761 | 8.761 | 0 | 0 | 0 | 0 | 0 | 0
[5 empty] | Count |  | 91 |  | 86 |  | 84
[5 empty] | Sum of squares |  | 0.041774 |  | 0.016169 |  | 0.004167
[5 empty] | Sp |  | 0.01515 |  | 0.0096957 |  | 0.0049803