## Notes
Notes by: Mary Steel -February 2009
Analysing Equipment
> | Guildline Autosal Model 8400B, serial # 68572
bath temperature of 24OC
accuracy of < +/- 0.0001 conductivity ratio
resolution <0.0001 conductivity ratio and <0.002 equivalent PSU
> | The calibration checks and the standardizations were done using IASO Standard Seawater
K15 Value = 0.99984
1 | Autosal Standby Ratio at beginning of analysing on January 9, 2009 was 5987 and Potentiometer dial was 5.38
| > | samples analysed from # 1 to 67
2 | The Autosal was standardized January 12, 2009 changing the Standby Ratio from 5987 to 5985 and the Potentiomenter dial from 5.38 to 5.35
| > | Samples analysed from # 68 to 189
3 | The Autosal was standardized again January 27, 2009 changing the Standby Ratio from 5987 to 5980 and the Potentiometer Dial from 5.35 to 5.30
| > | Samples analysed from # 214 to 1257
| > | The Autosal remained stable during analysing to the end with a slight drift but within the Guildline Electronic spec's of 0.002 PSU
Data Interpretation Equipment
> | Ocean Scientific International Salinometer Data Logger software was used to log the conductivity readings from the Guildline Autosal calculating salinity to four (4) decimal places or 1/10th part per million
Lab and Sample Condition
> | Analysed in a temperature controlled lab with an ambient temperature changing between 21.3oc-24.3oc overall with a 2.6oc maximum change during a daily sampling session
> | Sample temperatures ranged from 20.3oc to 24.0oc overall with a 2.1oc maximum change during a daily sampling session
> | The sample bottles were inverted and mixed prior to analysis.
> | all samples had minimum 3 flushes and minimum 2 reading
| > | each reading is the average of 5 readings taken in a 5 second period
Analyst Observations and Suggestions
1 | The Autosal was very stable making the analysing process very smooth!!!
| > | Autosal accuracy for electronics is < 0.002 PSU equivalents
| > | the electronics of the Autosal 8400B drifted within the manufactures spec's during analysing
2 | The majority of the samples had the separate insert from cap method of sealing and this helps minimize evaporation and/or leakage
| > | There were a number of samples where the insert was raised and some leaking.
| > | Always make sure you have pushed the insert all the in. This is most important to minimize evaporation
3 | Samples #1077 to #1137 were small bottles with plastic cone inserts in the cap and are marked in the spreadsheet comments section
these are clearly marked in the spreadsheet
4 | Samples #1207 to 1257 were small bottles with separate insert from cap method of sealing.
| > | With in this set samples #1240, 1241, 1242, 1243, 1248 were without inserts and are marked in the spreadsheet comments section
|  | > | check that all inserts are tight before capping to help minimize leakage or evaporation
5 | Some samples were noticeably leaking, had loose inserts or the bottle was cracked etc. These are identified in the spreadsheet comments section
>fill bottles only to the shoulder and insure that the insert is pushed in all the way before storing for transfer
>there may have been evaporation, compromising the accuracy of the salinity values
Timeline of Daily Conditions of Lab and Sample Temeratures and with AutoSalinometer 8400B SN68572
Temperature checks done every 26 to 28 samples
| Lab and Sample Temperatures |  |  |  | AutoSal Information
date | sample | change OC | room | change OC | bath OC | Zero | Standby | Comments | Samples Analysed
2009-01-09 | 21.7 |  | 21.28 |  | 24 | 0.00002 - | 5987 | no standardizing necessary | 1 to 67
| 21.57 | 0.13 | 22 | 0.72 | 24 |  | 5985
2009-01-12 | 21.3 |  | 20.6 |  | 24 | 0.00001 - | 5985 | standardized | 68 to 189
| 20.3 |  | 21.6 |  | 24 |  | 5984
| 21.6 |  | 21.8 |  | 24 |  | 5985
| 22.2 | 1.9 | 22.2 | 1.6 | 24 |  | 5983
2009-01-27 | 21.3 |  | 22.6 |  | 24 | 0.00001 + | 5980 | standardized | 214 to 382
| 21.7 |  | 22.9 |  | 24 |  | 5979
| 22.9 |  | 23.9 |  | 24 |  | 5978
| 23.2 |  | 24.3 |  | 24 |  | 5978
| 23.2 | 1.9 | 24.3 | 1.7 | 24 |  | 5977
2009-01-28 | 21.5 |  | 21.5 |  | 24 | 0.00001 - | 5980 | no standardizing necessary | 383 to 531
| 23.6 |  | 23.2 |  | 24 |  | 5979 | cal check +0.00002
| 22.7 |  | 24 |  | 24 |  | 5978
| 23.4 |  | 23 |  | 24 |  | 5978
| 23.6 |  | 24.1 |  | 24 |  | 5978
| 23.6 | 2.1 | 24.1 | 2.6 | 24 |  | 5978
2009-01-29 | 22.3 |  | 22.3 |  | 24 | 0.00001 + | 5981 | no standardizing necessary | 532 to 658
| 22.2 |  | 23.2 |  | 24 |  | 5979 | cal check +0.000002
| 23.7 |  | 23.8 |  | 24 |  | 5979
| 23.7 | 1.5 | 23.8 | 1.5 | 24 |  | 5979
2009-02-02 | 21.7 |  | 21.7 |  | 24 | 0.00001 + | 5981 | no standardizing necessary | 661 to 776
| 22.8 |  | 23.1 |  | 24 |  | 5980 | cal check -0.000002
| 22.7 |  | 23.8 |  | 24 |  | 5979
| 23.5 | 1.8 | 23.8 | 2.1 | 24 |  | 5978
2009-02-03 | 22 |  | 22 |  | 24 | 0.00001 + | 5981 | no standardizing necessary | 777 to 981
| 22.9 |  | 23.3 |  | 24 |  | 5979 | cal check +0.000002
| 23.9 |  | 23.9 |  | 24 |  | 5979
| 23.7 |  | 24.1 |  | 24 |  | 5978
| 23.8 | 1.9 | 24.1 | 2.1 | 24 |  | 5978
2009-02-04 | 22.2 |  | 22 |  | 24 | 0.00002 - | 5980 | no standardizing necessary | 982 to 1190
| 23 |  | 23.7 |  | 24 |  | 5979 | cal check -0.000004
| 23.4 |  | 24.2 |  | 24 |  | 5978
| 23.4 |  | 23.9 |  | 24 |  | 5978
| 23.6 |  | 23.6 |  | 24 |  | 5977
| 24 | 1.8 | 23.4 | 2.2 | 24 |  | 5978
2009-02-05 | 22.1 |  | 21.9 |  | 24 | 0.00001 + | 5980 | no standardizing necessary
| 22.5 | 0.4 | 23 |  | 24 |  | 5979 | cal check +0.000009
Maximum change in sample temperature in a given day is 2.1 oC
Maximum change in the lab temperature in a given day is 2.6 oC
Autosal accuracy for electronics is < 0.002 PSU equivalents
the electronics of the Autosal 8400B drifted within the manufactures spec's during analysing

## Calibration Timeline2008-02
Calibration Timeline for SWL 2008-02 Salinity Alalysis
| Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
1 | Calibration | K15=0.99984 | 2009-01-09 13:56 | 24C | 0.99984 | 2.7425e-06 | -3.0398e-06 | 0.99984
2 | Calibration | K15=0.99984 | 2009-01-12 11:51 | 24C | 0.99983 | 2.7425e-06 | 6.9737e-06 | 0.99984
3 | Calibration | K15=0.99984 | 2009-01-12 11:52 | 24C | 0.99984 | 6.664e-08 | 0 | 0.99984 |  |  | standarization value on next flush
4 | Calibration | K15=0.99984 | 2009-01-27 10:45 | 24C | 0.99986 | 1.1717e-05 | -1.4961e-05 | 0.99984 |  |  | cal check after 4 flushes
5 | Calibration | K15=0.99984 | 2009-01-27 10:46 | 24C | 0.99984 | 4.189e-06 | -8.9407e-07 | 0.99984 |  |  | standarization value on next flush
6 |  | K15=0.99984 | 2009-01-28 08:57 | 24C | 0.99985 | 5.0791e-06 | -9.5367e-07 | 0.99985 | 34.9942 | 0.00019806 | cal check with in +-0.00001
7 | Calibration | K15=0.99984 | 2009-01-29 09:52 | 24C | 0.99984 | 2.2391e-06 | -9.5367e-07 | 0.99984 |  |  | calibration check right on!!
8 | Calibration | K15=0.99984 | 2009-02-02 11:24 | 24C | 0.99984 | 4.4783e-06 | -1.967e-06 | 0.99984
9 | Calibration | K15=0.99984 | 2009-02-03 10:09 | 24C | 0.99984 | 4.4784e-06 | 1.9073e-06 | 0.99984 |  |  | cal check right on! K15=0.99984
10 | Calibration | K15=0.99984 | 2009-02-04 09:11 | 24C | 0.99984 | 4.1682e-06 | -3.9935e-06 | 0.99984
The Calibration checks and the standardizations were done using IASO Standard Seawater with K15 Value = 0.99984 Salinity Value 34.994
A calibration check was done before starting the daily analysing run with the exception of one day where the check got run as a sample by mistake.
The electronics of the Autosal 8400B drifted within the manufactures spec's during analysing
Autosal accuracy for electronics is < 0.002 PSU equivalents

## 2008-02 salinity Data
2008-02-SWL | [6 empty] | analysts: Mary Steel and assisted by Douglas Sieberg
Data for Transfer | [6 empty] | 2009-02-01
[6 empty] | Count | 764 | 80
Electrontic Rosette Log | [6 empty] | Salinity Value )ppm)
Date (UTC) | Ros Cast # | Stn | Samp # | Depth | ROS BOT # | Samp # | sample | duplicate sample | Comments made at time of analysing
2008-07-02 | 2 | NP1 | 1 | bot-5 | 1 | 1 | 33.9183 | 33.9156
|  |  | 2 | 100 | 2 | 2 | 33.8179
|  |  | 3 | 75 | 3 | 3 | 33.3947
|  |  | 4 | 49.5 | 4 | 4 | 32.813
|  |  | 5 | 25 | 5 | 5 | 32.3389
|  |  | 6 | 10.2 | 6 | 6 | 32.107 | 32.1066
|  |  | 7 | 5.9 | 7 | 7 | 32.0232
2008-07-03 | 3 | NP1 | 8 | 133 | 1
|  |  | 9 | 99 | 2
|  |  | 10 | 79 | 3
|  |  | 11 | 79 | 4
|  |  | 12 | 72 | 5
|  |  | 13 | 72 | 6
|  |  | 14 | 61 | 7
|  |  | 15 | 61 | 8
|  |  | 16 | 49.5 | 9
|  |  | 17 | 49.5 | 10
|  |  | 18 | 25 | 11
|  |  | 19 | 25 | 12
|  |  | 20 | 19 | 13
|  |  | 21 | 19 | 14
|  |  | 22 | 5.5 | 15
|  |  | 23 | 5.5 | 16
|  |  | 24 | 5.5 | 17
|  |  | 25 | 5.5 | 18
|  |  | 26 | 5.5 | 19
|  |  | 27 | 5.5 | 20
|  |  | 28 | 2.4 | 21
|  |  | 29 | 2.4 | 22
2008-07-04 | 4 | NP2 | 30 | 1000 | 1 | 30 | 34.3652 |  | 4 flushes
|  |  | 31 | 800 | 2 | 31 | 34.2473 | 34.2476
|  |  | 32 | 600 | 3 | 32 | 34.111
|  |  | 33 | 400 | 4 | 33 | 33.9649
|  |  | 34 | 300 | 5 | 34 | 33.8881
|  |  | 35 | 250 | 6 | 35 | 33.8793
|  |  | 36 | 200 | 7 | 36 | 33.8453
|  |  | 37 | 175 | 8 | 37 | 33.7943
|  |  | 38 | 150 | 9 | 38 | 33.6774
|  |  | 39 | 125 | 10 | 39 | 33.3021 | 33.2994 | leaking from insert
|  |  | 40 | 100 | 11 | 40 | 32.7049
|  |  | 41 | 75 | 12 | 41 | 32.5489
|  |  | 42 | 50 | 13 | 42 | 32.5006
|  |  | 43 | 25 | 14 | 43 | 32.4262
|  |  | 44 | 10 | 15 | 44 | 32.3834
|  |  | 45 | 5 | 16 | 45 | 32.4086
2008-07-04 | 5 | NP3 | 46 | 1000 | 1 | 46 | 34.3752
|  |  | 47 | 800 | 2 | 47 | 34.2763 | 34.2772
|  |  | 48 | 600 | 3 |  | 34.1836
|  |  | 49 | 400 | 4 |  | 34.0565
|  |  | 50 | 300 | 5 |  | 33.9821
|  |  | 51 | 250 | 6 |  | 33.9459
|  |  | 52 | 200 | 7 |  | 33.8923
|  |  | 53 | 175 | 8 |  | 33.8669 |  | salt on rim is thick
|  |  | 54 | 150 | 9 |  | 33.7554 | 33.7582
|  |  | 55 | 125 | 10 | 55 | 33.5408
|  |  | 56 | 100 | 11 | 56 | 32.967
|  |  | 57 | 75 | 12 | 57 | 32.6022
|  |  | 58 | 50 | 13 | 58 | 32.5943
|  |  | 59 | 25 | 14 | 59 | 32.2912
|  |  | 60 | 10 | 15 | 60 | 32.2642 |  | salt on rim
|  |  | 61 | 5 | 16 | 61 | 32.2502
2008-07-05 | 6 | NP4 | 62 | 1000 | 1 | 62 | 34.3763
|  |  | 63 | 800 | 2 | 63 | 34.3158 | 34.3543
|  |  | 64 | 600 | 3 | 64 | 34.1779
|  |  | 65 | 400 | 4 | 65 | 34.0538
|  |  | 66 | 300 | 5 | 66 | 33.9472
|  |  | 67 | 250 | 6 | 67 | 33.903
|  |  | 68 | 200 | 7 | 68 | 33.8603 |  | new day -cal check 34.99382
|  |  | 69 | 175 | 8 | 69 | 33.8364
|  |  | 70 | 150 | 9 | 70 | 33.7779 | 33.7789
|  |  | 71 | 125 | 10 | 71 | 33.5688
|  |  | 72 | 100 | 11 | 72 | 32.8742
|  |  | 73 | 75 | 12 | 73 | 32.553
|  |  | 74 | 50 | 13 | 74 | 32.5249
|  |  | 75 | 25 | 14 | 75 | 32.5098
|  |  | 76 | 10 | 15 | 76 | 32.4714
|  |  | 77 | 5 | 16
2008-07-05 | 7 | NP_5
2008-07-05 | 8 | NP_5 | 78 | 1000 | 1
|  |  | 79 | 250 | 2
|  |  | 80 | 250 | 3
|  |  | 81 | 110 | 4
|  |  | 82 | 110 | 5
|  |  | 83 | 100 | 6
|  |  | 84 | 100 | 7
|  |  | 85 | 43 | 8
|  |  | 86 | 43 | 9
|  |  | 87 | 40 | 10
|  |  | 88 | 40 | 11
|  |  | 89 | 13 | 12
|  |  | 90 | 13 | 13
|  |  | 91 | 5 | 14
|  |  | 92 | 5 | 15
|  |  | 93 | 5 | 16
|  |  | 94 | 5 | 17
|  |  | 95 | 5 | 18
|  |  | 96 | 5 | 19
|  |  | 97 | 5 | 20
|  |  | 98 | 5 | 21
2008-07-05 | 9 | NP_5 |  | 3000 | 1
|  |  |  | 3000 | 2
|  |  | 99 | 3000 | 3 | 99 | 34.6557 |  | 4 flushes
|  |  | 100 | 2500 | 4 | 100 | 34.6283 | 34.629
|  |  | 101 | 2000 | 5 | 101 | 34.5921
|  |  | 102 | 1500 | 6 | 102 | 34.5038
|  |  | 103 | 1250 | 7 | 103 | 34.4532
|  |  | 104 | 1000 | 8 | 104 | 34.374
|  |  | 105 | 800 | 9 | 105 | 34.2918
|  |  | 106 | 600 | 10 | 106 | 34.1793
|  |  | 107 | 400 | 11 | 107 | 34.028
|  |  | 108 | 300 | 12 | 108 | 33.9368
|  |  | 109 | 250 | 13 | 109 | 33.8875
|  |  | 110 | 200 | 14 | 110 | 33.8448
|  |  | 111 | 175 | 15 | 111 | 33.8276
|  |  | 112 | 150 | 16 | 112 | 33.7768 | 33.7785
|  |  | 113 | 125 | 17 | 113 | 33.5149
|  |  | 114 | 100 | 18 | 114 | 32.7272
|  |  | 115 | 75 | 19 | 115 | 32.6246
|  |  | 116 | 50 | 20 | 116 | 32.5986 |  | loose insert
|  |  | 117 | 100 | 21 | 117 | 32.5499
|  |  | 118 | 75 | 22 | 118 | 32.4913
|  |  | 119 | 50 | 23 | 119 | 32.4891
2008-07-06 | 10 | NP_6
2008-07-06 | 11 | NP_6 | 120 | 1000 | 1 | 120 | 34.3857 |  | loose insert and 4 flushes
|  |  | 121 | 800 | 2 | 121 | 34.3003
|  |  | 122 | 600 | 3 | 122 | 34.197 | 34.1982
|  |  | 123 | 400 | 4 | 123 | 34.0264
|  |  | 124 | 300 | 5 | 124 | 33.9219
|  |  | 125 | 250 | 6 | 125 | 33.8585
|  |  | 126 | 200 | 7 | 126 | 33.8322
|  |  | 127 | 175 | 8 | 127 | 33.8105
|  |  | 128 | 150 | 9 | 128 | 33.7338
|  |  | 129 | 125 | 10 | 129 | 33.3142
|  |  | 130 | 100 | 11 | 130 | 32.7793
|  |  | 131 | 75 | 12 | 131 | 32.6972
|  |  | 132 | 50 | 13 | 132 | 32.6515
|  |  | 133 | 25 | 14 | 133 | 32.5857 | 32.585
|  |  | 134 | 10 | 15 | 134 | 32.5261
|  |  | 135 | 5 | 16 | 135 | 32.5293
2008-07-06 | 12 | NP_6 | 136 | 1000 | 1
|  |  | 137 | 300 | 2
|  |  | 138 | 140 | 3
|  |  | 139 | 140 | 4
|  |  | 140 | 75 | 5
|  |  | 141 | 75 | 6
|  |  | 142 | 71 | 7
|  |  | 143 | 71 | 8
|  |  | 144 | 45 | 9
|  |  | 145 | 45 | 10
|  |  | 146 | 35 | 11
|  |  | 147 | 35 | 12
|  |  | 148 | 7 | 13
|  |  | 149 | 7 | 14
|  |  | 150 | 7 | 15
|  |  | 151 | 7 | 16
|  |  | 152 | 5 | 17
|  |  | 153 | 5 | 18
|  |  | 154 | 3 | 19
|  |  | 155 | 3 | 20
|  |  | 156 | 0 | 21
|  |  | 157 | 0 | 22
2008-07-07 | 13 | NP_7 | 158 | 1000 | 1 | 158 | 34.3681
|  |  | 159 | 800 | 2 | 159 | 34.2955
|  |  | 160 | 600 | 3 | 160 | 34.1774
|  |  | 161 | 400 | 4 | 161 | 33.9982 | 33.9999
|  |  | 162 | 300 | 5 | 162 | 33.8973
|  |  | 163 | 250 | 6 | 163 | 33.8459
|  |  | 164 | 200 | 7 | 164 | 33.7847
|  |  | 165 | 175 | 8 | 165 | 33.7532
|  |  | 166 | 150 | 9 | 166 | 33.598
|  |  | 167 | 125 | 10 | 167 | 33.1561
|  |  | 168 | 100 | 11 | 168 | 32.7388
|  |  | 169 | 75 | 12 | 169 | 32.6942 | 32.6901
|  |  | 170 | 50 | 13 | 170 | 32.6691
|  |  | 171 | 25 | 14 | 171 | 32.6034
|  |  | 172 | 10 | 15 | 172 | 32.5363
|  |  | 173 | 5 | 16 | 173 | 32.5393 |  | sample is only 1/2 full?
2008-07-07 | 14 | NP8
2008-07-07 | 15 | NP8 | 174 | 1000 | 1 | 174 | 34.3804 |  | 4 flushes
|  |  | 175 | 800 | 2 | 175 | 34.3242 | 34.3188
|  |  | 176 | 600 | 3 | 176 | 34.2226
|  |  | 177 | 400 | 4 | 177 | 34.0752
|  |  | 178 | 300 | 5 | 178 | 33.9795
|  |  | 179 | 250 | 6 | 179 | 33.9371
|  |  | 180 | 200 | 7 | 180 | 33.8816 | 33.8765 | reading jumping around
|  |  | 181 | 175 | 8 | 181 | 33.8522
|  |  | 182 | 150 | 9 | 182 | 33.8041
|  |  | 183 | 125 | 10 | 183 | 33.7167
|  |  | 184 | 100 | 11 | 184 | 33.4651
|  |  | 185 | 75 | 12 | 185 | 32.9333
|  |  | 186 | 50 | 13 | 186 | 32.7094
|  |  | 187 | 25 | 14 | 187 | 32.6717
|  |  | 188 | 10 | 15 | 188 | 32.6451
|  |  | 189 | 5 | 16 | 189 | 32.6412
|  |  |  | 1000 | 17
|  |  |  | 1000 | 16
2008-07-07 | 16/17 | NP8 | 190 | 400 | 1
|  |  | 191 | 400 | 2
|  |  | 192 | 120 | 3
|  |  | 193 | 120 | 4
|  |  | 194 | 120 | 5
|  |  | 195 | 120 | 6
|  |  | 196 | 100 | 7
|  |  | 197 | 100 | 8
|  |  | 198 | 75 | 9
|  |  | 199 | 75 | 10
|  |  | 200 | 60 | 11
|  |  | 201 | 60 | 12
|  |  | 202 | 40 | 13
|  |  | 203 | 40 | 14
|  |  | 204 | 31 | 15
|  |  | 205 | 31 | 16
|  |  | 206 | 18 | 17
|  |  | 207 | 18 | 18
|  |  | 208 | 10 | 19
|  |  | 209 | 10 | 20
|  |  | 210 | 5 | 21
|  |  | 211 | 5 | 22
|  |  | 212 | 0 | 23
|  |  | 213 | 0 | 24
2008-07-08 | 18 | NP9 | / |  | 1-21
2008-07-08 | 019 | NP9 | 214 | 1000 | 1 | 214 | 34.4075 |  | algae in tank?
|  |  | 215 | 1000 | 2
|  |  | 216 | 1000 | 3
|  |  | 217 | 800 | 4 | 217 | 34.3501 | 34.3469 | algae in tank?
|  |  | 218 | 600 | 5 | 218 | 34.2659 |  | algae in tank?
|  |  | 219 | 400 | 6 | 219 | 34.1442 |  | algae in tank?
|  |  | 220 | 300 | 7 | 220 | 34.0675 |  | algae in tank?
|  |  | 221 | 250 | 8 | 221 | 34.0366 |  | algae in tank?
|  |  | 222 | 200 | 9 | 222 | 33.9734 |  | algae in tank?
|  |  | 223 | 175 | 10 | 223 | 33.9543 |  | algae in tank?
|  |  | 224 | 150 | 11 | 224 | 33.8936 |  | algae in tank?
|  |  | 225 | 125 | 12 | 225 | 33.8431 |  | algae in tank?
|  |  | 226 | 100 | 13 | 226 | 33.8496 |  | algae in tank?
|  |  | 227 | 100 | 14
|  |  | 228 | 75 | 15 | 228 | 32.9577 | 32.9641 | algae in tank?
|  |  | 229 | 75 | 16
|  |  | 230 | 50 | 17 | 230 | 32.894 |  | algae in tank?
|  |  | 231 | 50 | 18
|  |  | 232 | 25 | 19 | 232 | 32.8677
|  |  | 233 | 25 | 20
|  |  | 234 | 10 | 21 | 234 | 32.8727
|  |  | 235 | 10 | 22
|  |  | 236 | 5 | 23 | 236 | 32.8368
|  |  | 237 | 5 | 24
2008-07-08 | 20 | NP9 | 238 | 300 | 1
|  |  | 239 | 300 | 2
|  |  | 240 | 119 | 3
|  |  | 241 | 119 | 4
|  |  | 242 | 20 | 5
|  |  | 243 | 20 | 6
|  |  | 244 | 70 | 7
|  |  | 245 | 70 | 8
|  |  | 246 | 36 | 9
|  |  | 247 | 36 | 10
|  |  | 248 | 15 | 11
|  |  | 249 | 15 | 12
|  |  | 250 | 11 | 13
|  |  | 251 | 11 | 14
|  |  | 252 | 6 | 15
|  |  | 253 | 6 | 16
|  |  | 254 | 6 | 17
|  |  | 255 | 6 | 18
|  |  | 256 | 4 | 19
|  |  | 257 | 4 | 20
|  |  | 258 | 2 | 21
|  |  | 259 | 2 | 22
2008-07-09 | 021 | NP-10 | 260 | 1000 | 1 | 260 | 34.4046 |  | 4 flushes
|  |  | 261 | 800 | 2 | 261 | 34.3476
|  |  | 262 | 600 | 3 | 262 | 34.2729 | 34.2754
|  |  | 263 | 400 | 4 | 263 | 34.151
|  |  | 264 | 300 | 5 | 264 | 34.0671
|  |  | 265 | 250 | 6 | 265 | 34.0115
|  |  | 266 | 200 | 7 | 266 | 33.9527 |  | see notes -flush cells with CLR and add bleach to tank -JL
|  |  | 267 | 175 | 8 | 267 | 33.9194
|  |  | 268 | 148 | 9 | 268 | 33.8644 | 33.863
|  |  | 269 | 125 | 10 | 269 | 33.7664
|  |  | 270 | 100 | 11 | 270 | 33.5559
|  |  | 271 | 75 | 12 | 271 | 33.2026
|  |  | 272 | 20 | 13 | 272 | 32.9675
|  |  | 273 | 25 | 14 | 273 | 32.7788
|  |  | 274 | 10 | 15 | 274 | 32.7312
|  |  | 275 | 5 | 16 | 275 | 32.7225
2008-07-09 | 022 | NP-11 | 276 | 1000 | 1 | 276 | 34.4225 |  | 4 flushes
|  |  | 277 | 800 | 2 | 277 | 34.3577
|  |  | 278 | 600 | 3 | 278 | 34.2755 | 34.2757
|  |  | 279 | 400 | 4 | 279 | 34.1614
|  |  | 280 | 300 | 5 | 280 | 34.0849
|  |  | 281 | 250 | 6 | 281 | 34.0439
|  |  | 282 | 200 | 7 | 282 | 33.9706
|  |  | 283 | 175 | 8 | 283 | 33.9433
|  |  | 284 | 150 | 9 | 284 | 33.9079
|  |  | 285 | 125 | 10 | 285 | 33.7311
|  |  | 286 | 100 | 11 | 286 | 33.4105
|  |  | 287 | 75 | 12 | 287 | 33.0094
|  |  | 288 | 50 | 13 | 288 | 32.8619
|  |  | 289 | 25 | 14 | 289 | 32.7657 | 32.7654
|  |  | 290 | 10 | 15 | 290 | 32.696
|  |  | 291 | 5 | 16 | 291 | 32.7073
2008-07-10 | 023 | NP-12 | 292 | 1000 | 1 | 292 | 34.4736 |  | 4 flushes
|  |  | 293 | 800 | 2 | 293 | 34.4027 | 34.3965
|  |  | 294 | 600 | 3 | 294 | 34.2759
|  |  | 295 | 400 | 4 | 295 | 34.1633
|  |  | 296 | 300 | 5 | 296 | 34.0649
|  |  | 297 | 250 | 6 | 297 | 34.0133
|  |  | 298 | 200 | 7 | 298 | 33.9456
|  |  | 299 | 175 | 8 | 299 | 33.9145 | 33.9078
|  |  | 300 | 150 | 9 | 300 | 33.8468
|  |  | 301 | 125 | 10 | 301 | 33.5704
|  |  | 302 | 100 | 11 | 302 | 33.2798
|  |  | 303 | 75 | 12 | 303 | 33.0175
|  |  | 304 | 50 | 13 | 304 | 32.7511
|  |  | 305 | 25 | 14 | 305 | 32.6855
|  |  | 306 | 10 | 15 | 306 | 32.6171
|  |  | 307 | 5 | 16 | 307 | 32.6144
|  |  | 308 | 100 | 17
|  |  | 309 | 100 | 18
|  |  | 310 | 50 | 19
|  |  | 311 | 50 | 20
|  |  | 312 | 25 | 21
|  |  | 313 | 25 | 22
|  |  | 314 | 5 | 23
|  |  | 315 | 5 | 24
2008-07-10 | 024 | NP-12 | 316 | 101 | 1 |  |  |  | 4 flushes
|  |  | 317 | 101 | 2
|  |  | 318 | 50 | 3
|  |  | 319 | 50 | 4
|  |  | 320 | 50 | 5
|  |  | 321 | 50 | 6
|  |  | 322 | 50 | 7
|  |  | 323 | 50 | 8
|  |  | 324 | 50 | 9
|  |  | 325 | 50 | 10
|  |  | 326 | 50 | 11
|  |  | 327 | 50 | 12
|  |  | 328 | 50 | 13
|  |  | 329 | 20 | 14
|  |  | 330 | 20 | 15
|  |  | 331 | 2 | 16
|  |  | 332 | 2 | 17
2008-07-10 | 025 | NP-13
2008-07-10 | 026 | NP-13 | 333 | 1000 | 1 | 333 | 34.3445 |  | 4 flushes
|  |  | 334 | 801 | 2 | 334 | 34.2651
|  |  | 335 | 601 | 3 | 335 | 34.1773
|  |  | 336 | 400 | 4 | 336 | 34.048 | 34.0285
|  |  | 337 | 300 | 5 | 337 | 33.9467
|  |  | 338 | 251 | 6 | 338 | 33.851
|  |  | 339 | 200 | 7 | 339 | 33.7893
|  |  | 340 | 175 | 8 | 340 | 33.6689
|  |  | 341 | 150 | 9 | 341 | 33.4529
|  |  | 342 | 125 | 10 | 342 | 33.0649
|  |  | 343 | 100 | 11 | 343 | 32.8124
|  |  | 344 | 75 | 12 | 344 | 32.5626
|  |  | 345 | 50 | 13 | 345 | 32.2816
|  |  | 346 | 26 | 14 | 346 | 32.2248
|  |  | 347 | 10 | 15 | 347 | 32.153 | 32.1623
|  |  | 348 | 5 | 16 | 348 | 32.141
|  |  | 349 | 100 | 17
|  |  | 350 | 100 | 18
|  |  | 351 | 50 | 19
|  |  | 352 | 50 | 20
|  |  | 353 | 30 | 21
|  |  | 354 | 30 | 22
|  |  | 355 | 5 | 23
|  |  | 356 | 5 | 24
2008-07-10 | 027 | NP-13 | 357 | 80 | 1 |  |  |  | 4 flushes
|  |  | 358 | 80 | 2
|  |  | 359 | 30 | 3
|  |  | 360 | 30 | 4
|  |  | 361 | 30 | 5
|  |  | 362 | 30 | 6
|  |  | 363 | 30 | 7
|  |  | 364 | 30 | 8
|  |  | 365 | 30 | 9
|  |  | 366 | 30 | 10
|  |  | 367 | 30 | 11
|  |  | 368 | 30 | 12
|  |  | 369 | 30 | 13
|  |  | 370 | 11 | 14
|  |  | 371 | 11 | 15
|  |  | 372 | 3 | 16
|  |  | 373 | 3 | 17
2008-07-11 | 028 | NP-14 | 374 | 1000 | 1 | 374 | 34.3519 |  | 4 flushes
|  |  | 375 | 800 | 2 | 375 | 34.2644 | 34.2903
|  |  | 376 | 600 | 3 | 376 | 34.1673
|  |  | 377 | 400 | 4 | 377 | 34.0491
|  |  | 378 | 300 | 5 | 378 | 33.9562
|  |  | 379 | 244 | 6 | 379 | 33.8924
|  |  | 380 | 200 | 7 | 380 | 33.8313
|  |  | 381 | 175 | 8 | 381 | 33.7622
|  |  | 382 | 150 | 9 | 382 | 33.6475 |  | cal check with in +-0.00001 34.9942
|  |  | 383 | 125 | 10 | 383 | 33.4718 | 33.4761
|  |  | 384 | 100 | 11 | 384 | 33.0514
|  |  | 385 | 75 | 12 | 385 | 32.7897
|  |  | 386 | 50 | 13 | 386 | 32.6188
|  |  | 387 | 25 | 14 | 387 | 32.4332
|  |  | 388 | 10 | 15 | 388 | 32.2356
|  |  | 389 | 5 | 16 | 389 | 32.1934
|  |  | 390 | 100 | 17
|  |  | 391 | 100 | 18
|  |  | 392 | 50 | 19
|  |  | 393 | 50 | 20
|  |  | 394 | 25 | 21
|  |  | 395 | 25 | 22
|  |  | 396 | 5 | 23
|  |  | 397 | 5 | 24
2008-07-11 | 029 | NP-14 | 398 | 80 | 1
|  |  | 399 | 80 | 2
|  |  | 400 | 25 | 3
|  |  | 401 | 30 | 4
|  |  | 402 | 30 | 5
|  |  | 403 | 30 | 6
|  |  | 404 | 30 | 7
|  |  | 405 | 30 | 8
|  |  | 406 | 30 | 9
|  |  | 407 | 30 | 10
|  |  | 408 | 30 | 11
|  |  | 409 | 30 | 12
|  |  | 410 | 30 | 13
|  |  | 411 | 30 | 14
|  |  | 412 | 30 | 15
|  |  | 413 | 30 | 16
|  |  | 414 | 30 | 17
|  |  | 415 | 30 | 18
|  |  | 416 | 30 | 19
|  |  | 417 | 30 | 20
|  |  | 418 | 30 | 21
2008-07-11 | 030 | UN-8 | 419 | 1000 | 1 | 419 | 34.3508 |  | 4 flushes
|  |  | 420 | 800 | 2 | 420 | 34.2721
|  |  | 421 | 600 | 3 | 421 | 34.1754 | 34.1511
|  |  | 422 | 400 | 4 | 422 | 34.032
|  |  | 423 | 300 | 5 | 423 | 33.9695
|  |  | 424 | 250 | 6 | 424 | 33.9302
|  |  | 425 | 200 | 7 | 425 | 33.8896
|  |  | 426 | 175 | 8 | 426 | 33.8691
|  |  | 427 | 150 | 9 | 427 | 33.8145
|  |  | 428 | 125 | 10 | 428 | 33.6134
|  |  | 429 | 100 | 11 | 429 | 33.2822
|  |  | 430 | 75 | 12 | 430 | 32.9007
|  |  | 431 | 20 | 13 | 431 | 32.7171
|  |  | 432 | 25 | 14 | 432 | 32.4395 | 32.4408
|  |  | 433 | 10 | 15 | 433 | 32.2367
|  |  | 434 | 5 | 16 | 434 | 32.2094
2008-07-11 | 31 | UN-7 | 435 | 1000 | 1 | 435 | 34.3715 |  | 4 flushes
|  |  | 436 | 800 | 2 | 436 | 34.2699
|  |  | 437 | 600 | 3 | 437 | 34.131
|  |  | 438 | 400 | 4 | 438 | 33.9993 | 33.999
|  |  | 439 | 300 | 5 | 439 | 33.9235
|  |  | 440 | 250 | 6 | 440 | 33.8635
|  |  | 441 | 200 | 7 | 441 | 33.7562
|  |  | 442 | 175 | 8 | 442 | 33.4703
|  |  | 443 | 150 | 9 | 443 | 33.1847
|  |  | 444 | 125 | 10 | 444 | 32.9389
|  |  | 445 | 100 | 11 | 445 | 32.7803
|  |  | 446 | 75 | 12 | 446 | 32.5667
|  |  | 447 | 20 | 13 | 447 | 32.2888 | 32.2889
|  |  | 448 | 25 | 14 | 448 | 32.0921
|  |  | 449 | 10 | 15 | 449 | 32.0039
|  |  | 450 | 5 | 16 | 450 | 32.0032
|  |  | 451 | 5 | 17
|  |  | 452 | 5 | 18
2008-07-11 | 32 | UN-6 | 453 | 85 | 1 | 453 | 32.2632 |  | 4 flushes
|  |  | 454 | 75 | 2 | 454 | 32.2279 | 32.2301
|  |  | 455 | 50 | 3 | 455 | 32.1174
|  |  | 456 | 40 | 4 | 456 | 32.1705
|  |  | 457 | 30 | 5 | 457 | 32.082
|  |  | 458 | 20 | 6 | 458 | 32.0401 | 32.041
|  |  | 459 | 10 | 7 | 459 | 32.0409
|  |  | 460 | 5 | 8 | 460 | 32.0411
|  |  | 461 | 2.5 | 9 | 461 | 32.0426
2008-07-11 | 33 | UN-5 | 462 | 75 | 1 | 462 | 32.3378 |  | 4 flushes
|  |  | 463 | 50 | 2 | 463 | 32.2434
|  |  | 464 | 40 | 3 | 464 | 32.1553
|  |  | 465 | 30 | 4 | 465 | 32.1512
|  |  | 466 | 20 | 5 | 466 | 32.1221 | 32.1242
|  |  | 467 | 10 | 6 | 467 | 32.1405
|  |  | 468 | 5 | 7 | 468 | 32.1429
|  |  | 469 | 2.5 | 8 | 469 | 32.144
2008-07-11 | 34 | UN-4 | 470 | 104 | 1 | 470 | 32.4436 |  | 4 flushes
|  |  | 471 | 75 | 2 | 471 | 32.4312
|  |  | 472 | 50 | 3 | 472 | 32.3064
|  |  | 473 | 40 | 4 | 473 | 32.2069 | 32.2071
|  |  | 474 | 30 | 5 | 474 | 32.1669
|  |  | 475 | 15 | 6 | 475 | 32.1196
|  |  | 476 | 10 | 7 | 476 | 32.1211
|  |  | 477 | 5 | 8 | 477 | 32.1214
|  |  | 478 | 3 | 9 | 478 | 32.1203
|  |  | 479a | 3 | 10
|  |  | 479b | 3 | 11
2008-07-11 | 35 | UN-3 | 480 | 95 | 1 | 480 | 32.3023 |  | 4 flushes
|  |  | 481 | 75 | 2 | 481 | 32.3007
|  |  | 482 | 50 | 3 | 482 | 32.2251 | 32.226
|  |  | 483 | 40 | 4 | 483 | 32.1402
|  |  | 484 | 30 | 5 | 484 | 32.1148
|  |  | 485 | 20 | 6 | 485 | 32.092
|  |  | 486 | 10 | 7 | 486 | 31.9357
|  |  | 487 | 5 | 8 | 487 | 31.923
|  |  | 488 | 3 | 9 | 488 | 31.9175
2008-07-11 | 36 | UN-2 | 489 | 83 | 1 | 489 | 32.0864 | 32.0916 | 4 flushes
|  |  | 490 | 50 | 2 | 490 | 32.0548
|  |  | 491 | 40 | 3 | 491 | 31.9097
|  |  | 492 | 30 | 4 | 492 | 31.8753
|  |  | 493 | 20 | 5 | 493 | 31.8305
|  |  | 494 | 10 | 6 | 494 | 31.8043
|  |  | 495 | 5 | 7 | 495 | 31.7983
|  |  | 496 | 3 | 8 | 496 | 31.798
|  |  | 497 | 3 | 9
|  |  | 498 | 3 | 10
2008-07-11 | 37 | UN-1 | 499 | 63 | 1 | 499 | 31.9367 |  | 4 flushes
|  |  | 500 | 50 | 2 | 500 | 31.9303
|  |  | 501 | 40 | 3 | 501 | 31.9261
|  |  | 502 | 30 | 4 | 502 | 31.9256
|  |  | 503 | 20 | 5 | 503 | 31.9111 | 31.9051
|  |  | 504 | 10 | 6 | 504 | 31.9
|  |  | 505 | 5 | 7 | 505 | 31.9024
|  |  | 506 | 2.5 | 8 | 506 | 31.9008
2008-07-12 | 38 | BCL-1 | 507 | 1002 | 1 | 507 | 34.3196 | 34.3185 | 4 flushes
|  |  | 508 | 802 | 2 | 508 | 34.2084
|  |  | 509 | 604 | 3 | 509 | 34.0197
|  |  | 510 | 403 | 4 | 510 | 33.6991
|  |  | 511 | 300 | 5 | 511 | 33.4618
|  |  | 512 | 250 | 6 | 512 | 33.3992
|  |  | 513 | 200 | 7 | 513 | 33.3296
|  |  | 514 | 171 | 8 | 514 | 33.2274
|  |  | 515 | 149 | 9 | 515 | 33.142
|  |  | 516 | 121 | 10 | 516 | 33.0114
|  |  | 517 | 100 | 11 | 517 | 32.8593
|  |  | 518 | 74 | 12 | 518 | 32.7235
|  |  | 519 | 49 | 13 | 519 | 32.6004
|  |  | 520 | 25 | 14 | 520 | 32.5858 | 32.586
|  |  | 521 | 9.8 | 15 | 521 | 32.5472
|  |  | 522 | 3.5 | 16 | 522 | 32.5468
2008-07-13 | 39 | BCL-2 | / | 25 | 1-21
2008-07-13 | 40 | BCL-2 | 523 | 1002 | 1 | 523 | 34.3592 |  | 4 flushes
|  |  | 524 | 800 | 2 | 524 | 34.2905 | 34.2902
|  |  | 525 | 601 | 3 | 525 | 34.1734
|  |  | 526 | 399 | 4 | 526 | 33.9684
|  |  | 527 | 300 | 5 | 527 | 33.794
|  |  | 528 | 251 | 6 | 528 | 33.7012
|  |  | 529 | 200 | 7 | 529 | 33.5882
|  |  | 530 | 175 | 8 | 530 | 33.5071
|  |  | 531 | 150 | 9 | 531 | 33.4337 |  | calibration check right on!!
|  |  | 532 | 123 | 10 | 532 | 33.3225
|  |  | 533 | 100 | 11 | 533 | 33.2745
|  |  | 534 | 75 | 12 | 534 | 33.2456
|  |  | 535 | 52 | 13 | 535 | 33.1544
|  |  | 536 | 25 | 14 | 536 | 32.9645
|  |  | 537 | 10 | 15 | 537 | 32.7141 | 32.7121
|  |  | 538 | 5 | 16 | 538 | 32.7125
2008-07-13 | 41 | BCL-2 | 539 | 600 | 1
|  |  | 540 | 600 | 2
|  |  | 541 | 250 | 3
|  |  | 542 | 250 | 4
|  |  | 543 | 100 | 5
|  |  | 544 | 100 | 6
|  |  | 545 | 100 | 7
|  |  | 546 | 100 | 8
|  |  | 547 | 75 | 9
|  |  | 548 | 75 | 10
|  |  | 549 | 53 | 11
|  |  | 550 | 53 | 12
|  |  | 551 | 25 | 13
|  |  | 552 | 25 | 14
|  |  | 553 | 15 | 15
|  |  | 554 | 15 | 16
|  |  | 555 | 5 | 17
|  |  | 556 | 5 | 18
|  |  | 557 | 5 | 19
|  |  | 558 | 5 | 20
|  |  | 559 | 5 | 21
|  |  | 560 | 5 | 22
|  |  | 561 | 5 | 23
|  |  | 562 | 5 | 24
2008-07-14 | 42 | BS-0 | 563 | 3001 | 1 | 563 | 34.6592 | 34.6584 | 4 flushes
|  |  | 564 | 2499 | 2 | 564 | 34.6341
|  |  | 565 | 1999 | 3 | 565 | 34.5938
|  |  | 566 | 1500 | 4 | 566 | 34.515
|  |  | 567 | 1250 | 5 | 567 | 34.4503
|  |  | 568 | 998 | 6 | 568 | 34.3673 | 34.3698
|  |  | 569 | 800 | 7 | 569 | 34.2881
|  |  | 570 | 601 | 8 | 570 | 34.1533
|  |  | 571 | 400 | 9 | 571 | 33.9161
|  |  | 572 | 299 | 10 | 572 | 33.7403
|  |  | 573 | 249 | 11 | 573 | 33.6524
|  |  | 574 | 199 | 12 | 574 | 33.5176
|  |  | 575 | 175 | 13 | 575 | 33.3935
|  |  | 576 | 150 | 14 | 576 | 33.3263
|  |  | 577 | 125 | 15 | 577 | 33.2645
|  |  | 578 | 100 | 16 | 578 | 33.2225
|  |  | 579 | 75 | 17 | 579 | 33.1834 |  | started fickering -salt floating??
|  |  | 580 | 50 | 18 | 580 | 33.1372
|  |  | 581 | 25 | 19 | 581 | 32.9581
|  |  | 582 | 10 | 20 | 582 | 32.945
|  |  | 583 | 5.7 | 21 | 583 | 32.9501
2008-07-14 | 43 | BS-1 | 584 | 1000 | 1 | 584 | 34.3526 |  | 4 flushes
|  |  | 585 | 150 | 2 | 585 | 33.2573
|  |  | 586 | 100 | 3 | 586 | 33.1604 | 33.1511
|  |  | 587 | 80 | 4 | 587 | 33.0723
|  |  | 588 | 60 | 5 | 588 | 33.009
|  |  | 589 | 50 | 6 | 589 | 32.8934
|  |  | 590 | 40 | 7 | 590 | 32.8397
|  |  | 591 | 30 | 8 | 591 | 32.6209
|  |  | 592 | 20 | 9 | 592 | 32.2646
|  |  | 593 | 15 | 10 | 593 | 32.1819
|  |  | 594 | 10 | 11 | 594 | 32.1698
|  |  | 595 | 5 | 12 | 595 | 32.1699
|  |  | 596 | 3.4 | 13 | 596 | 32.1717
|  |  | 597 | 3 | 14
|  |  | 598 | 3 | 15
2008-07-14 | 45 | BS-3 | 599 | 3 | 1
|  |  | 600 | 3 | 2
2008-07-14 | 46 | BS-4 | /
2008-07-14 | 47 | BS-5 | /
2008-07-14 | 48 | BS-5 | 601 | 126 | 1 | 601 | 33.1005 |  | 4 flushes
|  |  | 602 | 119 | 2 | 602 | 33.0955
|  |  | 603 | 100 | 3 | 603 | 33.0319
|  |  | 604 | 75 | 4 | 604 | 32.8867
|  |  | 605 | 50 | 5 | 605 | 32.6429
|  |  | 606 | 40 | 6 | 606 | 32.4942
|  |  | 607 | 30 | 7 | 607 | 32.4775 | 32.476
|  |  | 608 | 20 | 8 | 608 | 32.2672
|  |  | 609 | 14 | 9 | 609 | 32.1837
|  |  | 610 | 10 | 10 | 610 | 32.1819
|  |  | 611 | 5 | 11 | 611 | 32.1788
|  |  | 612 | 3 | 12 | 612 | 32.1823
|  |  | 613 | 126 | 13
|  |  | 614 | 90 | 14
|  |  | 615 | 40 | 15
|  |  | 616 | 40 | 16
|  |  | 617 | 33 | 17
|  |  | 618 | 33 | 18
|  |  | 619 | 20 | 19
|  |  | 620 | 20 | 20
|  |  | 621 | 5 | 21
|  |  | 622 | 5 | 24
2008-07-14 | 49 | BS-5 | 622 | 128 | 1
|  |  | 623 | 128 | 2
|  |  | 624 | 50.1 | 3
|  |  | 625 | 50.2 | 4
|  |  | 626 | 33.3 | 5
|  |  | 627 | 33.2 | 6
|  |  | 628 | 15.1 | 7
|  |  | 629 | 15.3 | 8
|  |  | 630 | 9.2 | 9
|  |  | 631 | 9.2 | 10
|  |  | 632 | 4.7 | 11
|  |  | 633 | 4.9 | 12
|  |  | 634 | 4.7 | 13
|  |  | 635 | 4.9 | 14
2008-07-14 | 50 | BCL-3 | 636 | 109 | 1
|  |  | 637 | 109 | 2
|  |  | 638 | 109 | 3 | 638 | 32.788 |  | 4 flushes
|  |  | 639 | 100 | 4 | 639 | 32.7954
|  |  | 640 | 81 | 5 | 640 | 32.7812
|  |  | 641 | 60 | 6 | 641 | 32.7417 | 32.7427
|  |  | 642 | 49 | 7 | 642 | 32.634
|  |  | 643 | 40 | 8 | 643 | 32.6309
|  |  | 644 | 30.6 | 9 | 644 | 32.5381
|  |  | 645 | 19.9 | 10 | 645 | 32.2793
|  |  | 646 | 15 | 11 | 646 | 32.2777
|  |  | 647 | 10 | 12 | 647 | 32.2784
|  |  | 648 | 5.3 | 13 | 648 | 32.276
2008-07-15 | 51 | BCL-4 | 649 | 71 | 1 | 649 | 32.5933 |  | 4 flushes
|  |  | 650 | 60 | 2 | 650 | 32.596
|  |  | 651 | 40 | 3 | 651 | 32.4429
|  |  | 652 | 35 | 4 | 652 | 32.4168 | 32.4174
|  |  | 653 | 30 | 5 | 653 | 32.3814
|  |  | 654 | 25 | 6 | 654 | 32.2904
|  |  | 655 | 20 | 7 | 655 | 31.6658 |  | leaking
|  |  | 656 | 15 | 8 | 656 | 31.3095
|  |  | 657 | 10 | 9 | 657 | 31.3094
|  |  | 658 | 4 | 10 | 658 | 31.3236 |  | k value 0.99984 cal is bang on!!
2008-07-15 | 52 | BCL-5 | 659 | 82 | 1
|  |  | 660 | 82 | 2
|  |  | 661 | 82 | 3 | 661 | 32.1653 |  | 4 flushes
|  |  | 662 | 69.8 | 4 | 662 | 32.1524
|  |  | 663 | 60 | 5 | 663 | 32.1465
|  |  | 664 | 48.6 | 6 | 664 | 32.1105
|  |  | 665 | 40.3 | 7 | 665 | 32.0088 |  | leaking
|  |  | 666 | 29.9 | 8 | 666 | 31.8397 | 31.858
|  |  | 667 | 25 | 9 | 667 | 31.7432 |  | leaking
|  |  | 668 | 19.9 | 10 | 668 | 31.6617
|  |  | 669 | 13.3 | 11 | 669 | 31.1864
|  |  | 670 | 9.3 | 12 | 670 | 31.059
|  |  | 671 | 5 | 13 | 671 | 31.0512
2008-07-16 | 53 | SLIP-1 | 672 | 73.6 | 1
|  |  | 673 | 73.6 | 2
|  |  | 674 | 73.6 | 3 | 674 | 32.4014 |  | 4 flushes
|  |  | 675 | 70.8 | 4 | 675 | 32.4058
|  |  | 676 | 59.9 | 5 | 676 | 32.3865
|  |  | 677 | 49.8 | 6 | 677 | 32.2836 | 32.2852
|  |  | 678 | 38.8 | 7 | 678 | 32.1585 |  | leaking
|  |  | 679 | 29.8 | 8 | 679 | 32.1116
|  |  | 680 | 25 | 9 | 680 | 32.0987
|  |  | 681 | 19.4 | 10 | 681 | 31.836 |  | insert raised -popped?
|  |  | 682 | 14.7 | 11 | 682 | 31.7014
|  |  | 683 | 9.7 | 12 | 683 | 31.6972
|  |  | 684 | 4.8 | 13 | 684 | 31.6761
2008-07-16 | 54 | SLIP-2 | 685 | 76 | 1
|  |  | 686 | 76 | 2
|  |  | 687 | 76 | 3 | 687 | 32.4383 |  | 4 flushes
|  |  | 688 | 69 | 4 | 688 | 32.4406
|  |  | 689 | 60 | 5 | 689 | 32.4349
|  |  | 690 | 50 | 6 | 690 | 32.3095
|  |  | 691 | 40 | 7 | 691 | 32.2263 | 32.2239
|  |  | 692 | 35 | 8 | 692 | 32.1852
|  |  | 693 | 30 | 9 | 693 | 32.1275
|  |  | 694 | 25 | 10 | 694 | 32.0554
|  |  | 695 | 19 | 11 | 695 | 31.9345
|  |  | 696 | 14 | 12 | 696 | 31.6793
|  |  | 697 | 9 | 13 | 697 | 31.647
|  |  | 698 | 5 | 14 | 698 | 31.6416
2008-07-15 | 55 | SLIP-3 | / | 68 | 1
|  |  | / | 68 | 2
|  |  | 699 | 68 | 3 | 699 | 32.5429 | 32.5402 | 4 flushes
|  |  | 700 | 58 | 4 | 700 | 32.5382
|  |  | 701 | 50 | 5 | 701 | 32.5302
|  |  | 702 | 40 | 6 | 702 | 32.3581
|  |  | 703 | 33 | 7 | 703 | 32.2578
|  |  | 704 | 30 | 8 | 704 | 32.2252
|  |  | 705 | 24 | 9 | 705 | 32.1926
|  |  | 706 | 20 | 10 | 706 | 32.0707
|  |  | 707 | 15 | 11 | 707 | 31.9108
|  |  | 708 | 10 | 12 | 708 | 31.872
|  |  | 709 | 5 | 13 | 709 | 31.8657
2008-07-16 | 56 | SLIP-5 | / | 61 | 1
|  |  | / | 61 | 2
|  |  | / | 50 | 3
|  |  | / | 40 | 4
|  |  | / | 25 | 5
|  |  | / | 10 | 6
|  |  | / | 5 | 7
2008-07-16 | 57 | SLIP-4 | 710 | 66 | 1 | 710 | 32.7578 |  | 4 flushes
|  |  | 711 | 60 | 2 | 711 | 32.7443 |  | leaking
|  |  | 712 | 50 | 3 | 712 | 32.6811 | 32.6809
|  |  | 713 | 39.3 | 4 | 713 | 32.4745
|  |  | 714 | 35 | 5 | 714 | 32.4276
|  |  | 715 | 30.2 | 6 | 715 | 32.3997
|  |  | 716 | 24.4 | 7 | 716 | 32.3361
|  |  | 717 | 20.3 | 8 | 717 | 32.2432
|  |  | 718 | 15 | 9 | 718 | 32.2301
|  |  | 719 | 10 | 10 | 719 | 31.9861
|  |  | 720 | 4.8 | 11 | 720 | 31.9758
|  |  | / | 20.2 | 12
2008-07-16 | 58 | SLIP-4 | 721 | 65 | 1
|  |  | 722 | 59 | 2
|  |  | 723 | 59 | 3
|  |  | 724 | 39 | 4
|  |  | 725 | 29.9 | 5
|  |  | 726 | 29.8 | 6
|  |  | 727 | 28.6 | 7
|  |  | 728 | 28.6 | 8
|  |  | 729 | 28.4 | 9 | [5 empty] | 4 flushes
|  |  | 730 | 28.8 | 10
|  |  | 731 | 15.2 | 11
|  |  | 732 | 15.3 | 12
|  |  | 733 | 15.3 | 13
|  |  | 734 | 15.1 | 14
|  |  | 735 | 5.1 | 15
|  |  | 736 | 5.1 | 16
|  |  | 737 | 4.9 | 17
|  |  | 738 | 4.6 | 18
|  |  | 739 | 4.3 | 19
|  |  | 740 | 4.3 | 20
|  |  | 741 | 3.4 | 21
|  |  | 742 | 3.2 | 22
2008-07-17 | 59 | BCL-6 | / | 48 | 1
|  |  | / | 48 | 2
|  |  | 743 | 48 | 3 | 743 | 32.7158 |  | 4 flushes
|  |  | 744 | 40 | 4 | 744 | 32.7184 | 32.7183
|  |  | 745 | 35.5 | 5 | 745 | 32.7182
|  |  | 746 | 31 | 6 | 746 | 32.7121
|  |  | 747 | 25 | 7 | 747 | 32.3703
|  |  | 748 | 21.3 | 8 | 748 | 32.2896
|  |  | 749 | 15 | 9 | 749 | 32.1085
|  |  | 750 | 10 | 10 | 750 | 32.0639
|  |  | 751 | 5 | 11 | 751 | 32.0628
|  |  | 752 | 5 | 12
|  |  | 753 | 5 | 13
2008-07-17 | 60 | UTBS-5 | / | 41 | 1
|  |  | / | 41 | 2
|  |  | 754 | 41 | 3 | 754 | 32.6576 | 32.6542 | 4 flushes
|  |  | 755 | 34 | 4 | 755 | 32.6574 |  | # 754 to #761 were written over the typed number and increased by 1
|  |  | 756 | 30 | 5 | 756 | 32.653
|  |  | 757 | 24 | 6 | 757 | 32.6544
|  |  | 758 | 20 | 7 | 758 | 32.6565
|  |  | 759 | 15 | 8 | 759 | 32.6548
|  |  | 760 | 10 | 9 | 760 | 32.6386
|  |  | 761 | 5 | 10 | 761 | 32.5928
2008-07-17 | 61 | UTBS-4 | / | 44 | 1
|  |  | / | 44 | 2
|  |  | 762 | 44 | 3 | 762 | 32.8837 |  | 4 flushes
|  |  | 763 | 35 | 4
|  |  | 764 | 30 | 5
|  |  | 765 | 25 | 6
|  |  | 766 | 20 | 7
|  |  | 767 | 15 | 8
|  |  | 768 | 10 | 9
|  |  | 769 | 5 | 10
2008-07-17 | 62 | UTBS-1 | / | 41.3 | 1
|  |  | / | 41.7 | 2
|  |  | 770 | 41.6 | 3 | 770 | 32.7686 |  | 4 flushes
|  |  | 771 | 40.3 | 4 | 771 | 32.7908
|  |  | 772 | 35.3 | 5 | 772 | 32.7901
|  |  | 773 | 30 | 6 | 773 | 32.798
|  |  | 774 | 25 | 7 | 774 | 32.7646
|  |  | 775 | 20 | 8 | 775 | 32.7487
|  |  | 776 | 15.3 | 9 | 776 | 32.7266
|  |  | 777 | 10.5 | 10 | 777 | 32.6986
|  |  | 778 | 5.5 | 11 | 778 | 32.6081 | 32.6005
|  |  | / | 15.3 | 12-24
2008-07-17 | 63 | UTBS-1 | 779 | 44.2 | 1
|  |  | 780 | 44.2 | 2
|  |  | 781 | 40.3 | 3
|  |  | 782 | 40.3 | 4
|  |  | 783 | 32.3 | 5
|  |  | 784 | 32.3 | 6
|  |  | 785 | 29.9 | 7
|  |  | 786 | 29.9 | 8
|  |  | 787 | 20.3 | 9
|  |  | 788 | 20.3 | 10
|  |  | 789 | 20.3 | 11
|  |  | 790 | 20.3 | 12
|  |  | 791 | 15.1 | 13
|  |  | 792 | 15.1 | 14
|  |  | 793 | 15 | 15
|  |  | 794 | 15 | 16
|  |  | 795 | 6.1 | 17
|  |  | 796 | 6.3 | 18
|  |  | 797 | 5.5 | 19
|  |  | 798 | 5.5 | 20
|  |  | 799 | 5.5 | 21
|  |  | 800 | 5.5 | 22
|  |  | 801 | 3.3 | 23
|  |  | 802 | 3.3 | 24 |  |  |  | cal check right on! K15=0.99984
2008-07-18 | 64 | BRS-1 | 803 | 44 | 1 | 803 | 32.455 |  | 4 flushes
|  |  | 804 | 35 | 2 | 804 | 32.4599
|  |  | 805 | 30 | 3 | 805 | 32.4519 | 32.4488
|  |  | 806 | 24.7 | 4 | 806 | 32.451
|  |  | 807 | 20 | 5 | 807 | 32.3729
|  |  | 808 | 14.5 | 6 | 808 | 32.3475
|  |  | 809 | 9.7 | 7 | 809 | 32.3481
|  |  | 810 | 5.8 | 8 | 810 | 32.3472
718/08 | 65 | BRS-2 | 811 | 45.4 | 1 | 811 | 32.3843 |  | 4 flushes
|  |  | 812 | 40.1 | 2 | 812 | 32.3815
|  |  | 813 | 34.4 | 3 | 813 | 32.3842
|  |  | 814 | 30.5 | 4 | 814 | 32.3751 | 32.3745
|  |  | 815 | 25.1 | 5 | 815 | 32.3748
|  |  | 816 | 20 | 6 | 816 | 32.3746
|  |  | 817 | 15.7 | 7 | 817 | 32.2694
|  |  | 818 | 10.4 | 8 | 818 | 32.1439
|  |  | 819 | 5.7 | 9 | 819 | 32.1432
2008-07-18 | 66 | BRS-3 | 820 | 48 | 1 | 820 | 32.0768
|  |  | 821 | 42 | 2 | 821 | 32.0902
|  |  | 822 | 35 | 3 | 822 | 32.094
|  |  | 823 | 30 | 4 | 823 | 32.1003 | 32.0999
|  |  | 824 | 25 | 5 | 824 | 32.1003
|  |  | 825 | 20 | 6 | 825 | 32.0743
|  |  | 826 | 15 | 7 | 826 | 31.9945
|  |  | 827 | 10 | 8 | 827 | 31.9916
|  |  | 828 | 5 | 9 | 828 | 31.9305
2008-07-18 | 67 | BRS-4 | 829 | 44 | 1 | 829 | 31.8149 |  | 4 flushs
|  |  | 830 | 40 | 2 | 830 | 31.8084
|  |  | 831 | 35 | 3 | 831 | 31.8035 | 31.8017
|  |  | 832 | 30 | 4 | 832 | 31.809
|  |  | 833 | 25 | 5 | 833 | 31.8038
|  |  | 834 | 20 | 6 | 834 | 31.7973
|  |  | 835 | 15 | 7 | 835 | 31.5533
|  |  | 836 | 10 | 8 | 836 | 31.1603
|  |  | 837 | 5 | 9 | 837 | 28.0647
2008-07-18 | 68 | BRS-5 | 838 | 37 | 1 | 838 | 30.9396 |  | 4 flushes
|  |  | 839 | 35 | 2 | 839 | 30.9064
|  |  | 840 | 30 | 3 | 840 | 30.8172 | 30.809
|  |  | 841 | 25 | 4 | 841 | 30.7497
|  |  | 842 | 20 | 5 | 842 | 30.3318
|  |  | 843 | 15 | 6 | 843 | 29.9092
|  |  | 844 | 10 | 7 | 844 | 28.773
|  |  | 845 | 5 | 8 | 845 | 27.4971
2008-07-18 | 69 | UTN-1 | / | 28 | 1
|  |  | / | 28 | 2
|  |  | 846 | 28 | 3 | 846 | 30.0408 |  | 4 flushes
|  |  | 847 | 25 | 4 | 847 | 29.9989
|  |  | 848 | 20 | 5 | 848 | 29.8455 | 29.8489
|  |  | 849 | 15 | 6 | 849 | 29.4623
|  |  | 850 | 10 | 7 | 850 | 29.2955
|  |  | 851 | 6 | 8 | 851 | 29.1164
2008-07-18 | 70 | UTN-2 | / | 40 | 1
|  |  | / | 40 | 2
|  |  | 852 | 40 | 3 | 852 | 32.3073
|  |  | 853 | 34.3 | 4
|  |  | 854 | 29.7 | 5
|  |  | 855 | 24 | 6
|  |  | 856 | 19.9 | 7
|  |  | 857 | 14.7 | 8
|  |  | 858 | 10.4 | 9
|  |  | 859 | 5.1 | 10
2008-07-19 | 71 | UTN-3 | / | 49 | 1
|  |  | / | 49 | 2
|  |  | 860 | 49 | 3
|  |  | 861 | 40 | 4 | 861 | 32.4232 |  | 4 flushes
|  |  | 862 | 35 | 5
|  |  | 863 | 29.4 | 6
|  |  | 864 | 24.9
|  |  | 865 | 20.6 | 7
|  |  | 866 | 15.6 | 8
|  |  | 867 | 10.6 | 9
|  |  | 868 | 5.3 | 10
2008-07-19 | 72 | UTN-4 | 869 | 49 | 1 | 869 | 32.4566 |  | 4 flushes
|  |  | 870 | 40 | 2 | 870 | 32.4569
|  |  | 871 | 33.6 | 3 | 871 | 32.4557
|  |  | 872 | 29.4 | 4 | 872 | 32.4569
|  |  | 873 | 25 | 5 | 873 | 32.4578
|  |  | 874 | 20.4 | 6 | 874 | 32.4575
|  |  | 875 | 15 | 7 | 875 | 32.4253
|  |  | 876 | 9.8 | 8 | 876 | 32.3966
|  |  | 877 | 5.2 | 9 | 877 | 32.3927
|  |  | / | 23.2 | 10
|  |  | / | 40 | 23
|  |  | / | 40 | 24
2008-07-19 | 73 | UTN-4 | 878 | 42 | 1
|  |  | 879 | 42 | 2
|  |  | 880 | 40 | 3
|  |  | 881 | 40 | 4
|  |  | 882 | 24.7 | 5
|  |  | 883 | 24.8 | 6
|  |  | 884 | 24.8 | 7
|  |  | 885 | 24.8 | 8
|  |  | 886 | 11.8 | 9
|  |  | 887 | 11.8 | 10
|  |  | 888 | 5.2 | 11
|  |  | 889 | 5.2 | 12
|  |  | 890 | 5.3 | 13
|  |  | 891 | 5.3 | 14
|  |  | 892 | 2.4 | 15
|  |  | 893 | 2.4 | 16
2008-07-19 | 74 | UTN-5 | / | 48 | 1
|  |  | / | 48 | 2
|  |  | 894 | 48 | 3 | 894 | 32.6176 |  | 4 flushes
|  |  | 895 | 40 | 4
|  |  | 896 | 35 | 5
|  |  | 897 | 30 | 6
|  |  | 898 | 25 | 7
|  |  | 899 | 20 | 8
|  |  | 900 | 13 | 9
|  |  | 901 | 10 | 10
|  |  | 902 | 5 | 11
2008-07-19 | 75 | UTN-6 | / | 46 | 1
|  |  | / | 46 | 2
|  |  | 903 | 46 | 3 | 903 | 32.5631 |  | 4 flushes
|  |  | 904 | 40 | 4
|  |  | 905 | 35 | 5
|  |  | 906 | 30 | 6
|  |  | 907 | 25 | 7
|  |  | 908 | 20 | 8
|  |  | 909 | 15 | 9
|  |  | 910 | 10 | 10
|  |  | 911 | 5 | 11
2008-07-19 | 76 | UTN-7 | / | 53 | 1
|  |  | / | 53 | 2
|  |  | 912 | 53 | 3 | 912 | 32.999 |  | 4 flushes
|  |  | 913 | 45 | 4 | 913 | 32.9566
|  |  | 914 | 40 | 5 | 914 | 32.9283
|  |  | 915 | 35 | 6 | 915 | 32.9075
|  |  | 916 | 30 | 7 | 916 | 32.8829
|  |  | 917 | 25 | 8 | 917 | 32.878
|  |  | 918 | 20 | 9 | 918 | 32.8303
|  |  | 919 | 15 | 10 | 919 | 32.6866 | 32.7052 | not very close
|  |  | 920 | 10 | 11 | 920 | 32.3933
|  |  | 921 | 5 | 12 | 921 | 32.3555
2008-07-19 | 77 | CCL-1 | 922 | 39 | 1 | 922 | 31.8248 |  | 4 flushes
|  |  | 923 | 30 | 2 | 923 | 31.8266
|  |  | 924 | 24.7 | 3 | 924 | 31.7086
|  |  | 925 | 20.3 | 4 | 925 | 31.6505
|  |  | 926 | 13.9 | 5 | 926 | 31.5094
|  |  | 927 | 9.6 | 6 | 927 | 31.099 | 31.1021
|  |  | 928 | 5.2 | 7 | 928 | 26.8937
2008-07-20 | 78 | CCL-2 | 929 | 51 | 1 | 929 | 32.5905 |  | 4 flushes
|  |  | 930 | 44.9 | 2 | 930 | 32.5905
|  |  | 931 | 39.9 | 3 | 931 | 32.5882
|  |  | 932 | 36 | 4 | 932 | 32.5729
|  |  | 933 | 30.4 | 5 | 933 | 32.5729
|  |  | 934 | 24.6 | 6 | 934 | 32.5667
|  |  | 935 | 20.7 | 7 | 935 | 32.541
|  |  | 936 | 15.4 | 8 | 936 | 32.4836
|  |  | 937 | 9.4 | 9 | 937 | 30.5513
|  |  | 938 | 5.4 | 10 | 938 | 29.5299
2008-07-20 | 79 | CCL-3 | 939 | 31 | 1 | 939 | 31.9336 |  | loose insert
|  |  | 940 | 25 | 2 | 940 | 31.0564
|  |  | 941 | 20 | 3 | 941 | 30.692
|  |  | 942 | 15 | 4 | 942 | 26.6612
|  |  | 943 | 10 | 5 | 943 | 26.0265
|  |  | 944 | 5 | 6 | 944 | 26.0222 |  | loose insert
2008-07-20 | 80 | CCL-4 | 945 | 45 | 1 | 945 | 32.9491
|  |  | 946 | 40 | 2 | 946 | 32.9502
|  |  | 947 | 35 | 3 | 947 | 32.9463
|  |  | 948 | 30 | 4 | 948 | 32.9884
|  |  | 949 | 25 | 5 | 949 | 32.948
|  |  | 950 | 20 | 6 | 950 | 32.7944 | 32.7972 | loose insert
|  |  | 951 | 15 | 7 | 951 | 32.6292
|  |  | 952 | 10 | 8 | 952 | 31.8805
|  |  | 953 | 4 | 9 | 953 | 30.1652
|  |  | / | 15 | 10-24
2008-07-20 | 81 | CCL-4 | 954 | 40 | 1
|  |  | 955 | 40 | 2
|  |  | 956 | 30 | 3
|  |  | 957 | 30 | 4
|  |  | 958 | 15 | 5
|  |  | 959 | 15 | 6
|  |  | 960 | 15 | 7
|  |  | 961 | 15 | 8
|  |  | 962 | 8 | 9
|  |  | 963 | 8 | 10
|  |  | 964 | 4 | 11
|  |  | 965 | 4 | 12
|  |  | 966 | 4 | 13
|  |  | 967 | 4 | 14
|  |  | 968 | 3.5 | 15
|  |  | 969 | 3.5 | 16
2008-07-21 | 82 | BC-1 | 970 | 81.4 | 1 | 970 | 32.8977 |  | 4 flushes and loose insert
|  |  | 971 | 70 | 2 | 971 | 32.9372
|  |  | 972 | 60 | 3 | 972 | 32.9628
|  |  | 973 | 50 | 4 | 973 | 32.9306 |  | ** #973 to #1006 no caps justs inserts
|  |  | 974 | 40 | 5 | 974 | 32.9724 |  | **
|  |  | 975 | 35 | 6 | 975 | 32.9435 |  | **
|  |  | 976 | 30 | 7 | 976 | 32.9295 |  | **
|  |  | 977 | 25 | 8 | 977 | 32.9483 |  | **
|  |  | 978 | 20 | 9 | 978 | 32.874 |  | **
|  |  | 979 | 15 | 10 | 979 | 31.6017 | 31.6166 | **
|  |  | 980 | 10 | 11 | 980 | 29.7297 |  | **
|  |  | 981 | 5 | 12 | 981 | 29.3545 |  | **
|  |  | 982 | 18 | 13 |  |  |  | ?????? Number repeated
2008-07-21 | 83 | BC-2 | 982 | 120.8 | 1 | 982 | 33.1786 |  | ** and 4 flushes
|  |  | 983 | 99.5 | 2 | 983 | 33.1665 |  | **
|  |  | 984 | 75.4 | 3 | 984 | 33.1616 |  | **
|  |  | 985 | 49.8 | 4 | 985 | 33.1591 |  | **
|  |  | 986 | 43.7 | 5 |  |  |  | **
|  |  | 987 | 43.8 | 6 |  |  |  | **
|  |  | 988 | 30.2 | 7 |  |  |  | **
|  |  | 989 | 30.2 | 8 |  |  |  | **
|  |  | 990 | 30.2 | 9 | 990 | 32.6556 |  | **
|  |  | 991 | 20.5 | 10 | 991 | 32.0557 |  | **
|  |  | 992 | 18.5 | 11 |  |  |  | **
|  |  | 993 | 18.5 | 12 |  |  |  | **
|  |  | 994 | 15.2 | 13 | 994 | 31.8656 |  | **
|  |  | 995 | 10.3 | 14 |  |  |  | **
|  |  | 996 | 10.3 | 15 |  |  |  | **
|  |  | 997 | 10.3 | 16 | 997 | 31.3845 |  | **
|  |  | 998 | 5.8 | 17 |  |  |  | **
|  |  | 999 | 5.8 | 18 |  |  |  | **
|  |  | 1000 | 5.8 | 19 | 1000 | 30.0539 |  | **
|  |  | 1001 | 2.3 | 20 |  |  |  | **
|  |  | 1002 | 2.3 | 21 |  |  |  | **
2008-07-22 | 84 | BC-3 | 1003 | 182 | 1 | 1003 | 34.8478 |  | ** and 4 flushes
|  |  | 1004 | 150.1 | 2 | 1004 | 34.8272 |  | **
|  |  | 1005 | 99.3 | 3 | 1005 | 33.2478 |  | **
|  |  | 1006 | 75.4 | 4 | 1006 | 33.1405 | 33.137 | **
|  |  | 1007 | 50.9 | 5 | 1007 | 33.0173
|  |  | 1008 | 25.3 | 6 | 1008 | 32.1726
|  |  | 1009 | 16.8 | 7 | 1009 | 29.0704
|  |  | 1010 | 5.9 | 8 | 1010 | 29.0095
2008-07-22 | 85 | BC-4 | 1011 | 582 | 1 | 1011 | 34.8556 |  | ** and 4 flushes
|  |  | 1012 | 500 | 2 | 1012 | 34.8522 |  | ** and 4 flushes
|  |  | 1013 | 400 | 3 | 1013 | 34.8479
|  |  | 1014 | 300 | 4 | 1014 | 34.8236
|  |  | 1015 | 201 | 5 | 1015 | 34.7647
|  |  | 1016 | 150 | 6 | 1016 | 34.5542
|  |  | 1017 | 100 | 7 | 1017 | 33.2513
|  |  | 1018 | 75 | 8 | 1018 | 33.151
|  |  | 1019 | 51.2 | 9 | 1019 | 32.8752
|  |  | 1020 | 24.1 | 10 | 1020 | 31.783
|  |  | 1021 | 13.3 | 11 | 1021 | 30.5865
|  |  | 1022 | 5.5 | 12 | 1022 | 27.6055
|  |  | 1023 | 5.6 | 13
|  |  | 1024 | 5.7 | 14
2008-07-23 | 86 | BL-1 | / | 22.9 | 1
2008-07-23 | 87 | BL-1 | 1025 | 58.6 | 1 | 1025 | 32.4228 |  | 4 flushes
|  |  | 1026 | 51 | 2 | 1026 | 32.4602
|  |  | 1027 | 40.3 | 3 | 1027 | 32.3377
|  |  | 1028 | 30.7 | 4 | 1028 | 32.18
|  |  | 1029 | 20 | 5 | 1029 | 31.805 | 31.8292 | dup not even close
|  |  | 1030 | 11.4 | 6 | 1030 | 30.3815
|  |  | 1031 | 6.7 | 7 | 1031 | 27.9516
2008-07-23 | 88 | BL-1 | 1032 | 59.2 | 1
|  |  | 1033 | 59.2 | 2
|  |  | 1034 | 50.3 | 3
|  |  | 1035 | 50.3 | 4
|  |  | 1036 | 47.8 | 5
|  |  | 1037 | 47.8 | 6
|  |  | 1038 | 35.2 | 7
|  |  | 1039 | 35.2 | 8
|  |  | 1040 | 22.5 | 9
|  |  | 1041 | 22.5 | 10
|  |  | 1042 | 22.5 | 11
|  |  | 1043 | 22.5 | 12
|  |  | 1044 | 22.4 | 13
|  |  | 1045 | 22.5 | 14
|  |  | 1046 | 15.5 | 15
|  |  | 1047 | 15.5 | 16
|  |  | 1048 | 10 | 17
|  |  | 1049 | 10 | 18
|  |  | 1050 | 5 | 19
|  |  | 1051 | 4.9 | 20
|  |  | 1052 | 4 | 21
|  |  | 1053 | 4 | 22
|  |  | 1054 | 1.9 | 23
|  |  | 1055 | 1.9 | 24
2008-07-23 | 89 | BFB-2 | 1056 | 973 | 1 | 1056 | 34.8893
|  |  | 1057 | 800 | 2 | 1057 | 34.8773
|  |  | 1058 | 600 | 3 | 1058 | 34.8609
|  |  | 1059 | 401 | 4 | 1059 | 34.8324 | 34.829
|  |  | 1060 | 300.7 | 5 | 1060 | 34.7922
|  |  | 1061 | 200.7 | 6 | 1061 | 34.6149
|  |  | 1062 | 175 | 7 | 1062 | 34.4536
|  |  | 1063 | 150 | 8 | 1063 | 33.9747
|  |  | 1064 | 123.8 | 9 | 1064 | 33.527
|  |  | 1065 | 100 | 10 | 1065 | 32.4038
|  |  | 1066 | 74.8 | 11 | 1066 | 32.0448
|  |  | 1067 | 50.7 | 12 | 1067 | 31.1979
|  |  | 1068 | 40.7 | 13 | 1068 | 30.7922
|  |  | 1069 | 30 | 14 | 1069 | 30.149
|  |  | 1070 | 21 | 15 | 1070 | 29.8227 | 29.83
|  |  | 1071 | 10 | 16 | 1071 | 29.3439
|  |  | 1072 | 5 | 17 | 1072 |  |  | no insert -sample lost
|  |  | / | 70.4 | 21
|  |  | / | 70.5 | 22
2008-07-24 | 91 | BFB-3 | 1073 | 999.4 | 1 | 1073 | 34.889 |  | 4 flushes
|  |  | 1074 | 700.1 | 2 | 1074 | 34.868
|  |  | 1075 | 500.6 | 3 | 1075 | 34.8459 | 34.8522
|  |  | 1076 | 400 | 4 | 1076 | 34.8252
|  |  | 1077 | 300 | 5 | 1077 | 34.7589 |  | caps with cone inserts 1077 to 1137
|  |  | 1078 | 200 | 6 | 1078 | 34.3656 |  | cone insert
|  |  | 1079 | 174.9 | 7 | 1079 | 34.0931 |  | " "
|  |  | 1080 | 149.8 | 8 | 1080 | 33.554 |  | " "
|  |  | 1081 | 125.8 | 9 | 1081 | 33.196 |  | " "
|  |  | 1082 | 106 | 10 | 1082 | 32.7605 |  | " "
|  |  | 1083 | 74.9 | 11 | 1083 | 32.3951 | 32.3921 | " " | " "
|  |  | 1084 | 50 | 12 | 1084 | 31.553 |  | " "
|  |  | 1085 | 40 | 13 | 1085 | 30.9802 |  | " "
|  |  | 1086 | 30 | 14 | 1086 | 30.3094 |  | " "
|  |  | 1087 | 20.4 | 15 | 1087 | 29.2892 |  | " "
|  |  | 1088 | 10.9 | 16 | 1088 | 27.5666 |  | " "
|  |  | 1089 | 4.6 | 17 | 1089 | 26.3406 |  | " "
|  |  | / | 50 | 18
|  |  | / | 50 | 19
|  |  | / | 999.4 | 21
|  |  | / | 999.4 | 22
2008-07-25 | 92 | MK-1 | / | 51.8 | 1-21
2008-07-25 | 93 | MK-1 | 1090 | 238 | 1 | 1090 | 34.6758 |  | cone insert
|  |  | 1091 | 199 | 2 | 1091 | 34.3346 |  | " "
|  |  | 1092 | 175 | 3 | 1092 | 33.7368 |  | " "
|  |  | 1093 | 149 | 4 | 1093 | 33.195 |  | " "
|  |  | 1094 | 125 | 5 | 1094 | 32.8947 |  | " "
|  |  | 1095 | 100 | 6 | 1095 | 32.6956 |  | " "
|  |  | 1096 | 74 | 7 | 1096 | 32.2912 |  | " " | " "
|  |  | 1097 | 50 | 8 | 1097 | 31.8247 |  | " "
|  |  | 1098 | 40.6 | 9 | 1098 | 31.3428 |  | " "
|  |  | 1099 | 30.7 | 10 | 1099 | 30.845 |  | " "
|  |  | 1100 | 20 | 11 | 1100 | 29.9287 | 29.9305 | " "
|  |  | 1101 | 11 | 12 | 1101 | 27.6412 |  | " "
|  |  | 1102 | 5.1 | 13 | 1102 | 24.0785 |  | " "
2008-07-25 | 94 | MK-1 | 1103 | 200.5 | 1
|  |  | 1104 | 200.5 | 2
|  |  | 1105 | 141 | 3
|  |  | 1106 | 141 | 4
|  |  | 1107 | 97.1 | 5
|  |  | 1108 | 97.1 | 6
|  |  | 1109 | 79.9 | 7
|  |  | 1110 | 79.9 | 8
|  |  | 1111 | 44.9 | 9
|  |  | 1112 | 44.9 | 10
|  |  | 1113 | 44.9 | 11
|  |  | 1114 | 44.9 | 12
|  |  | 1115 | 14.3 | 13
|  |  | 1116 | 14.3 | 14
|  |  | 1117 | 9.3 | 15
|  |  | 1118 | 9.3 | 16
|  |  | 1119 | 5.4 | 17
|  |  | 1120 | 5.4 | 18
|  |  | 1121 | 4.2 | 19
|  |  | 1122 | 4.2 | 20
|  |  | 1123 | 2.1 | 21
|  |  | 1124 | 2.1 | 22
|  |  | 1125 | 2.1 | 23
|  |  | 1126 | 2.1 | 24
2008-07-25 | 95 | BFB-4 | 1127 | 248 | 1 | 1127 | 34.7247 |  | cone insert
|  |  | 1128 | 201 | 2 | 1128 | 33.5596 |  | " "
|  |  | 1129 | 175 | 3 | 1129 | 33.2584 |  | " "
|  |  | 1130 | 150 | 4 | 1130 | 33.116 |  | " "
|  |  | 1131 | 126 | 5 | 1131 | 33.419 |  | " "
|  |  | 1132 | 100 | 6 | 1132 | 32.7682 |  | " "
|  |  | 1133 | 75.1 | 7 | 1133 | 32.3076 |  | " "
|  |  | 1134 | 50.4 | 8 | 1134 | 31.7154 |  | " "
|  |  | 1135 | 40.7 | 9 | 1135 | 31.3621 |  | " "
|  |  | 1136 | 30.1 | 10 | 1136 | 30.9248 |  | " "
|  |  | 1137 | 20.3 | 11 | 1137 | 30.304 | 30.3789 | " " and duplicate not even close
|  |  | 1138 | 10.1 | 12 | 1138 | 28.3061
|  |  | 1139 | 5.3 | 13 | 1139 | 23.382
2008-07-26 | 96 | BFB-5 | 1140 | 978 | 1 | 1140 | 34.892
|  |  | 1141 | 701.7 | 2 | 1141 | 34.8678
|  |  | 1142 | 500.1 | 3 | 1142 | 34.8431 | 34.8442
|  |  | 1143 | 401.6 | 4 | 1143 | 34.8235
|  |  | 1144 | 300.1 | 5 | 1144 | 34.7641
|  |  | 1145 | 200 | 6 | 1145 | 34.4307
|  |  | 1146 | 175.8 | 7 | 1146 | 26.9102
|  |  | 1147 | 150.3 | 8 | 1147 | 33.9363
|  |  | 1148 | 125.2 | 9 | 1148 | 33.3634
|  |  | 1149 | 100.5 | 10 | 1149 | 32.9868
|  |  | 1150 | 75.1 | 11 | 1150 | 32.5666
|  |  | 1151 | 50.5 | 12 | 1151 | 32.1
|  |  | 1152 | 40 | 13 | 1152 | 31.8313
|  |  | 1153 | 30.2 | 14 | 1153 | 31.5909
|  |  | 1154 | 20.8 | 15 | 1154 | 30.7396
|  |  | 1155 | 10.6 | 16 | 1155 | 27.9061
|  |  | 1156 | 4.3 | 17 | 1156 | 26.9274
1157
|  |  | 1158 | 1000 | 19
|  |  | 1159 | 1000 | 20
|  |  | / | 42.8 | 22
| 97 | BFB-5 | 1160 | 349.7 | 1
|  |  | 1161 | 349.7 | 2
|  |  | 1162 | 100 | 3
|  |  | 1163 | 100 | 4
|  |  | 1164 | 90 | 5
|  |  | 1165 | 90 | 6
|  |  | 1166 | 43.5 | 7
|  |  | 1167 | 43.5 | 8
|  |  | 1168 | 43.5 | 9
|  |  | 1169 | 43.5 | 10
|  |  | 1170 | 20 | 11
|  |  | 1171 | 20 | 12
|  |  | 1172 | 15.2 | 13
|  |  | 1173 | 15.2 | 14
|  |  | 1174 | 5.7 | 15
|  |  | 1175 | 5.7 | 16
|  |  | 1176 | 4.1 | 17
|  |  | 1177 | 4.1 | 18
|  |  | 1178 | 2 | 19
|  |  | 1179 | 1.9 | 20
|  |  | 1180 | 2 | 21
|  |  | 1181 | 1.9 | 22
2008-07-26 | 98 | BFB-6 | 1182 | 177 | 1 | 1182 | 33.7083
|  |  | 1183 | 174 | 2 | 1183 | 33.7213
|  |  | 1184 | 149.8 | 3 | 1184 | 33.3713
|  |  | 1185 | 125 | 4 | 1185 | 33.0758
|  |  | 1186 | 101 | 5 | 1186 | 32.7994
|  |  | 1187 | 74.3 | 6 | 1187 | 32.3706 | 32.3665
|  |  | 1188 | 50 | 7 | 1188 | 31.9538
|  |  | 1189 | 40 | 8 | 1189 | 31.7191
|  |  | 1190 | 30 | 9 | 1190 | 31.3075
|  |  | 1191 | 20 | 10 | 1191 | 30.7007
|  |  | 1192 | 9.7 | 11 | 1192 | 29.3508
|  |  | 1193 | 5.8 | 12 | 1193 | 29.2001
|  |  | / |  | 13
|  |  | / | 5.8 | 14
|  |  | / | 5.8 | 15
|  |  | / | 5.8 | 16
|  |  | / | 5.8 | 17
2008-07-27 | 99 | BFB-7 | / | 70 | 1-21
7/27/2008 | 100 | BFB-7 | 1194 | 646 | 1 | 1194 | 34.8001
|  |  | 1195 | 499 | 2 | 1195 | 34.7864
|  |  | 1196 | 399 | 3 | 1196 | 34.7675
|  |  | 1197 | 299 | 4 | 1197 | 34.7243
|  |  | 1198 | 200 | 5 | 1198 | 34.5992
|  |  | 1199 | 174 | 6 | 1199 | 34.5248
|  |  | 1200 | 150 | 7 | 1200 | 34.4089 | 34.41
|  |  | 1201 | 124 | 8 | 1201 | 34.169
|  |  | 1202 | 100 | 9 | 1202 | 33.9192
|  |  | 1203 | 74 | 10 | 1203 | 33.4093
|  |  | 1204 | 49 | 11 | 1204 | 32.7113
|  |  | 1205 | 41 | 12 | 1205 | 32.5404
|  |  | 1206 | 30 | 13 | 1206 | 31.7835
|  |  | 1207 | 20.9 | 14 | 1207 | 31.6274 |  | samples 1207 to 1257 are small bottles with inserts
|  |  | 1208 | 10 | 15 | 1208 | 30.5306
|  |  | 1209 | 6 | 16 | 1209 | 30.4788
7/27/2008 | 101 | BFB-7 | 1210 | 499.5 | 1
|  |  | 1211 | 499.5 | 2
|  |  | 1212 | 250.5 | 3
|  |  | 1213 | 250.7 | 4
|  |  | 1214 | 119.2 | 5
|  |  | 1215 | 119.2 | 6
|  |  | 1216 | 60.1 | 7
|  |  | 1217 | 60.1 | 8
|  |  | 1218 | 60.1 | 9
|  |  | 1219 | 60.1 | 10
|  |  | 1220 | 60.1 | 11
|  |  | 1221 | 42.9 | 12
|  |  | 1222 | 42.9 | 13
|  |  | 1223 | 42.9 | 14
|  |  | 1224 | 42.9 | 15
|  |  | 1225 | 42.9 | 16
|  |  | 1226 | 10.2 | 17
|  |  | 1227 | 10.2 | 18
|  |  | 1228 | 5.3 | 19
|  |  | 1229 | 5.3 | 20
|  |  | 1230 | 5.3 | 21
|  |  | 1231 | 5.3 | 22
|  |  | 1232 | 1.7 | 23
|  |  | 1233 | 1.7 | 24
7/27/2008 | 102 | BFB-8 | 1234 | 412 | 1 | 1234 | 34.8861
|  |  | 1235 | 401 | 2 | 1235 | 34.863
|  |  | 1236 | 302 | 3 | 1236 | 34.8237
|  |  | 1237 | 203 | 4 | 1237 | 34.2324
|  |  | 1238 | 177 | 5 | 1238 | 34.1428
|  |  | 1239 | 151 | 6 | 1239 | 33.964
|  |  | 1240 | 125 | 7 | 1240 | 33.4624 | 33.481 | no insert
|  |  | 1241 | 100 | 8 | 1241 | 33.2599 |  | no insert
|  |  | 1242 | 74 | 9 | 1242 | 32.7812 |  | no insert
|  |  | 1243 | 50 | 10 | 1243 | 32.2201 |  | no insert
|  |  | 1244 | 39.7 | 11 | 1244 | 32.1152
|  |  | 1245 | 30.1 | 12 | 1245 | 31.9297
|  |  | 1246 | 20.1 | 13 | 1246 | 31.5258
|  |  | 1247 | 10.3 | 14 | 1247 | 30.8044
|  |  | 1248 | 5.2 | 15 | 1248 | 30.6456 |  | no insert
7/28/2008 | 103 | BFB-9 | 1249 | 210 | 1 | 1249 | 28.6903
|  |  | 1250 | 200 | 2 | 1250 | 28.686
|  |  | 1251 | 176 | 3 | 1251 | 28.7567
|  |  | 1252 | 151 | 4 | 1252 | 28.657 | 28.7248 | not close at all
|  |  | 1253 | 126 | 5 | 1253 | 28.5562
|  |  | 1254 | 101.5 | 6 | 1254 | 28.5782
|  |  | 1255 | 75 | 7 | 1255 | 28.5
|  |  | 1256 | 51 | 8 | 1256 | 28.3835
|  |  | 1257 | 40.8 | 9 | 1257 | 28.4474
|  |  | 1258 | 40.8 | 10
|  |  | 1259 | 40.8 | 11
|  |  | 1260 | 40.8 | 12
|  |  | 1261 | 31 | 13
|  |  | 1262 | 21.1 | 14
|  |  | 1263 | 11 | 15
|  |  | 1264 | 5.7 | 16

## #1-189 RAW DATA 2008-02
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | K15 | 2009-01-09 13:56 | 24C | 0.99984 | 2.7425e-06 | -3.0398e-06 | 0.99984
2008-02SWL#1 | check | 2009-01-09 13:57 | 24C | 0.99984 | 2.4911e-06 | -3.0398e-06 | 0.99984 | 34.9935 | 0.00010182
2008-02SWL#2 | 1 | 2009-01-09 14:04 | 24C | 0.97245 | 3.8598e-06 | -3.0398e-06 | 0.97245 | 33.9183 | 0.00014981
2008-02SWL#3 | 1 d | 2009-01-09 14:06 | 24C | 0.97238 | 3.6105e-06 | -3.0398e-06 | 0.97238 | 33.9156 | 0.00014155
2008-02SWL#4 | 2 | 2009-01-09 14:08 | 24C | 0.96989 | 1.1203e-06 | -3.0398e-06 | 0.96988 | 33.8179 | 4.4364e-05
2008-02SWL#5 | 3 | 2009-01-09 14:10 | 24C | 0.95907 | 4.3186e-06 | -3.0398e-06 | 0.95906 | 33.3947 | 0.00016797
2008-02SWL#6 | 4 | 2009-01-09 14:12 | 24C | 0.94416 | 3.1414e-06 | -2.9206e-06 | 0.94416 | 32.813 | 0.00011967
2008-02SWL#7 | 5 | 2009-01-09 14:14 | 24C | 0.93198 | 2.2401e-06 | -3.0398e-06 | 0.93198 | 32.3389 | 8.9564e-05
2008-02SWL#8 | 6 | 2009-01-09 14:17 | 24C | 0.92601 | 3.1415e-06 | -2.9802e-06 | 0.92601 | 32.107 | 0.00012312
2008-02SWL#9 | 6d | 2009-01-09 14:19 | 24C | 0.926 | 2.4915e-06 | -3.0398e-06 | 0.92599 | 32.1066 | 9.4928e-05
2008-02SWL#10 | 7 | 2009-01-09 14:21 | 24C | 0.92385 | 3.4659e-06 | -2.9802e-06 | 0.92385 | 32.0232 | 0.00013595
2008-02SWL#11 | 30 | 2009-01-09 14:23 | 24C | 0.98385 | 3.1417e-06 | -2.9802e-06 | 0.98385 | 34.3652 | 0.00012313 | 4 flushes
2008-02SWL#12 | 31 | 2009-01-09 14:25 | 24C | 0.98085 | 3.9871e-06 | -3.0398e-06 | 0.98084 | 34.2473 | 0.00015604
2008-02SWL#13 | 31d | 2009-01-09 14:27 | 24C | 0.98085 | 2.2397e-06 | -2.9802e-06 | 0.98085 | 34.2476 | 8.792e-05
2008-02SWL#14 | 32 | 2009-01-09 14:29 | 24C | 0.97737 | 1.1196e-06 | -3.0398e-06 | 0.97737 | 34.111 | 4.4364e-05
2008-02SWL#15 | 33 | 2009-01-09 14:31 | 24C | 0.97364 | 1.3882e-06 | -3.0398e-06 | 0.97364 | 33.9649 | 5.4326e-05
2008-02SWL#16 | 34 | 2009-01-09 14:34 | 24C | 0.97168 | 2.7425e-06 | -2.9802e-06 | 0.97168 | 33.8881 | 0.00010451
2008-02SWL#17 | 35 | 2009-01-09 14:35 | 24C | 0.97146 | 3.6109e-06 | -3.0398e-06 | 0.97145 | 33.8793 | 0.00014263
2008-02SWL#18 | 36 | 2009-01-09 14:37 | 24C | 0.97059 | 1.3712e-06 | -2.9802e-06 | 0.97059 | 33.8453 | 5.2235e-05
2008-02SWL#19 | 37 | 2009-01-09 14:39 | 24C | 0.96929 | 3.1417e-06 | -2.9802e-06 | 0.96928 | 33.7943 | 0.00012105
2008-02SWL#20 | 38 | 2009-01-09 14:41 | 24C | 0.9663 | 2.8901e-06 | -3.0398e-06 | 0.9663 | 33.6774 | 0.00011401
2008-02SWL#21 | 39 | 2009-01-09 14:44 | 24C | 0.9567 | 4.3194e-06 | -3.0994e-06 | 0.95669 | 33.3021 | 0.00016576 | leaking from insert
2008-02SWL#22 | 39d | 2009-01-09 14:46 | 24C | 0.95663 | 1.4046e-06 | -2.9206e-06 | 0.95663 | 33.2994 | 5.5371e-05
2008-02SWL#23 | 40 | 2009-01-09 14:48 | 24C | 0.94138 | 1.3549e-06 | -2.9206e-06 | 0.94138 | 32.7049 | 5.4326e-05
2008-02SWL#24 | 41 | 2009-01-09 14:51 | 24C | 0.93738 | 3.2141e-06 | -3.0398e-06 | 0.93737 | 32.5489 | 0.00012382
2008-02SWL#25 | 42 | 2009-01-09 14:53 | 24C | 0.93614 | 2.7425e-06 | -2.9802e-06 | 0.93613 | 32.5006 | 0.00010656
2008-02SWL#26 | 43 | 2009-01-09 14:55 | 24C | 0.93422 | 2.7425e-06 | -3.0398e-06 | 0.93422 | 32.4262 | 0.00010865
2008-02SWL#27 | 44 | 2009-01-09 14:57 | 24C | 0.93312 | 6.0358e-06 | -3.0398e-06 | 0.93312 | 32.3834 | 0.00023387
2008-02SWL#28 | 45 | 2009-01-09 14:59 | 24C | 0.93377 | 3.057e-06 | -3.0398e-06 | 0.93377 | 32.4086 | 0.00012034
2008-02SWL#29 | 46 | 2009-01-09 15:02 | 24C | 0.98411 | 2.743e-06 | -2.9802e-06 | 0.9841 | 34.3752 | 0.00010795
2008-02SWL#30 | 47 | 2009-01-09 15:05 | 24C | 0.98159 | 4.3047e-06 | -3.0398e-06 | 0.98158 | 34.2763 | 0.00016888
2008-02SWL#31 | 47d | 2009-01-09 15:41 | 24C | 0.98161 | 2.2394e-06 | -2.9802e-06 | 0.98161 | 34.2772 | 8.7013e-05
2008-02SWL#32 | 48 | 2009-01-09 15:44 | 24C | 0.97922 | 2.4908e-06 | -2.9206e-06 | 0.97922 | 34.1836 | 0.00010059
2008-02SWL#33 | 49 | 2009-01-09 15:46 | 24C | 0.97598 | 4.1892e-06 | -2.9206e-06 | 0.97598 | 34.0565 | 0.0001676
2008-02SWL#34 | 50 | 2009-01-09 15:48 | 24C | 0.97408 | 3.8406e-06 | -2.9206e-06 | 0.97408 | 33.9821 | 0.00015044
2008-02SWL#35 | 51 | 2009-01-09 15:50 | 24C | 0.97316 | 3.2034e-06 | -2.9802e-06 | 0.97316 | 33.9459 | 0.00012807
2008-02SWL#36 | 52 | 2009-01-09 15:52 | 24C | 0.97179 | 4.0096e-06 | -2.9802e-06 | 0.97179 | 33.8923 | 0.00015611
2008-02SWL#37 | 53 | 2009-01-09 15:54 | 24C | 0.97114 | 6.6785e-06 | -3.0398e-06 | 0.97114 | 33.8669 | 0.00026074 | salt on rim is thick
2008-02SWL#38 | 54 | 2009-01-09 15:56 | 24C | 0.96829 | 5.7779e-06 | -2.861e-06 | 0.96829 | 33.7554 | 0.00022662
2008-02SWL#39 | 54d | 2009-01-09 15:59 | 24C | 0.96836 | 3.1414e-06 | -2.9802e-06 | 0.96836 | 33.7582 | 0.00012002
2008-02SWL#40 | 55 | 2009-01-09 16:01 | 24C | 0.96281 | 3.4494e-06 | -2.9802e-06 | 0.9628 | 33.5408 | 0.00013412
2008-02SWL#41 | 56 | 2009-01-09 16:03 | 24C | 0.94811 | 3.4657e-06 | -2.9802e-06 | 0.94811 | 32.967 | 0.0001339
2008-02SWL#42 | 57 | 2009-01-09 16:06 | 24C | 0.93875 | 2.7425e-06 | -3.0994e-06 | 0.93874 | 32.6022 | 0.00010447
2008-02SWL#43 | 58 | 2009-01-09 16:08 | 24C | 0.93854 | 2.7425e-06 | -2.9802e-06 | 0.93854 | 32.5943 | 0.00010865
2008-02SWL#44 | 59 | 2009-01-09 16:10 | 24C | 0.93075 | 3.4659e-06 | -3.0398e-06 | 0.93075 | 32.2912 | 0.00013319
2008-02SWL#45 | 60 | 2009-01-09 16:12 | 24C | 0.93006 | 1.1064e-06 | -3.0398e-06 | 0.93005 | 32.2642 | 4.5227e-05 | salt on rim
2008-02SWL#46 | 61 | 2009-01-09 16:14 | 24C | 0.9297 | 2.2391e-06 | -3.0398e-06 | 0.92969 | 32.2502 | 8.7447e-05
2008-02SWL#47 | 62 | 2009-01-09 16:39 | 24C | 0.98413 | 3.8436e-06 | -2.9802e-06 | 0.98413 | 34.3763 | 0.00014951
2008-02SWL#48 | 63 | 2009-01-09 16:41 | 24C | 0.98259 | 3.8598e-06 | -3.0994e-06 | 0.98259 | 34.3158 | 0.00014953
2008-02SWL#49 | 63d | 2009-01-09 16:43 | 24C | 0.98357 | 2.5036e-06 | -2.9802e-06 | 0.98357 | 34.3543 | 9.8238e-05
2008-02SWL#50 | 64 | 2009-01-09 16:46 | 24C | 0.97908 | 2.2391e-06 | -3.0994e-06 | 0.97907 | 34.1779 | 9.0443e-05
2008-02SWL#51 | 65 | 2009-01-09 16:48 | 24C | 0.97591 | 5.5696e-06 | -3.0994e-06 | 0.97591 | 34.0538 | 0.00021722
2008-02SWL#52 | 66 | 2009-01-09 16:50 | 24C | 0.97319 | 3.2142e-06 | -2.9206e-06 | 0.97319 | 33.9472 | 0.00012891
2008-02SWL#53 | 67 | 2009-01-09 16:52 | 24C | 0.97206 | 1.1064e-06 | -2.9802e-06 | 0.97206 | 33.903 | 4.4364e-05
Calibration#2 | K15 | 2009-01-12 11:51 | 24C | 0.99983 | 2.7425e-06 | 6.9737e-06 | 0.99984
Calibration#3 | K15 | 2009-01-12 11:52 | 24C | 0.99984 | 6.664e-08 | 0 | 0.99984
2008-02SWL#54 | check | 2009-01-12 11:53 | 24C | 0.99984 | 3.4553e-06 | 0 | 0.99984 | 34.9938 | 0.00013203
2008-02SWL#55 | 68 | 2009-01-12 12:04 | 24C | 0.97097 | 4.322e-06 | 0 | 0.97097 | 33.8603 | 0.00016981
2008-02SWL#56 | 69 | 2009-01-12 12:07 | 24C | 0.97036 | 3.594e-06 | 0 | 0.97036 | 33.8364 | 0.00014259
2008-02SWL#57 | 70 | 2009-01-12 12:09 | 24C | 0.96886 | 3.4658e-06 | 0 | 0.96886 | 33.7779 | 0.00013791
2008-02SWL#58 | 70d | 2009-01-12 12:11 | 24C | 0.96889 | 1.3712e-06 | 0 | 0.96889 | 33.7789 | 5.5395e-05
2008-02SWL#59 | 71 | 2009-01-12 12:13 | 24C | 0.96352 | 3.4658e-06 | 0 | 0.96352 | 33.5688 | 0.00013487
2008-02SWL#60 | 72 | 2009-01-12 12:15 | 24C | 0.94573 | 4.8797e-06 | 0 | 0.94573 | 32.8742 | 0.00018947
2008-02SWL#61 | 73 | 2009-01-12 12:16 | 24C | 0.93748 | 2.7425e-06 | 0 | 0.93748 | 32.553 | 0.00010553
2008-02SWL#62 | 74 | 2009-01-12 12:18 | 24C | 0.93676 | 5.0936e-06 | 0 | 0.93676 | 32.5249 | 0.00020181
2008-02SWL#63 | 75 | 2009-01-12 12:21 | 24C | 0.93637 | 2.4916e-06 | 0 | 0.93637 | 32.5098 | 9.7462e-05
2008-02SWL#64 | 76 | 2009-01-12 12:23 | 24C | 0.93538 | 4.4783e-06 | 0 | 0.93538 | 32.4714 | 0.00017361
2008-02SWL#65 | 99 | 2009-01-12 13:13 | 24C | 0.99125 | 1.3882e-06 | 0 | 0.99125 | 34.6557 | 5.2235e-05
2008-02SWL#66 | 100 | 2009-01-12 13:15 | 24C | 0.99055 | 3.332e-08 | 0 | 0.99055 | 34.6283 | 0
2008-02SWL#67 | 100d | 2009-01-12 13:17 | 24C | 0.99057 | 5.6348e-06 | 0 | 0.99057 | 34.629 | 0.00021944
2008-02SWL#68 | 101 | 2009-01-12 13:20 | 24C | 0.98963 | 2.7425e-06 | 0 | 0.98963 | 34.5921 | 0.00010865
2008-02SWL#69 | 102 | 2009-01-12 13:22 | 24C | 0.98738 | 2.4776e-06 | 0 | 0.98738 | 34.5038 | 9.7169e-05
2008-02SWL#70 | 103 | 2009-01-12 13:24 | 24C | 0.98609 | 3.466e-06 | 0 | 0.98609 | 34.4532 | 0.00013633
2008-02SWL#71 | 104 | 2009-01-12 13:26 | 24C | 0.98407 | 3.4657e-06 | 0 | 0.98407 | 34.374 | 0.00013341
2008-02SWL#72 | 105 | 2009-01-12 13:28 | 24C | 0.98198 | 5.1484e-06 | 0 | 0.98198 | 34.2918 | 0.00020363
2008-02SWL#73 | 106 | 2009-01-12 13:31 | 24C | 0.97911 | 3.4658e-06 | 0 | 0.97911 | 34.1793 | 0.00013203
2008-02SWL#74 | 107 | 2009-01-12 13:33 | 24C | 0.97525 | 3.1414e-06 | 0 | 0.97525 | 34.028 | 0.00012312
2008-02SWL#75 | 108 | 2009-01-12 13:35 | 24C | 0.97292 | 3.4658e-06 | 0 | 0.97292 | 33.9368 | 0.00013478
2008-02SWL#76 | 109 | 2009-01-12 13:38 | 24C | 0.97166 | 2.4915e-06 | 0 | 0.97166 | 33.8875 | 9.869e-05
2008-02SWL#77 | 110 | 2009-01-12 13:40 | 24C | 0.97057 | 3.1901e-06 | 0 | 0.97057 | 33.8448 | 0.00012891
2008-02SWL#78 | 111 | 2009-01-12 13:42 | 24C | 0.97013 | 3.4352e-06 | 0 | 0.97013 | 33.8276 | 0.0001341
2008-02SWL#79 | 112 | 2009-01-12 13:44 | 24C | 0.96883 | 5.7486e-06 | 0 | 0.96883 | 33.7768 | 0.00022591
2008-02SWL#80 | 112d | 2009-01-12 13:47 | 24C | 0.96888 | 2.2391e-06 | 0 | 0.96888 | 33.7785 | 8.7024e-05
2008-02SWL#81 | 113 | 2009-01-12 13:49 | 24C | 0.96214 | 3.5954e-06 | 0 | 0.96214 | 33.5149 | 0.00013864
2008-02SWL#82 | 114 | 2009-01-12 13:51 | 24C | 0.94195 | 2.71e-06 | 0 | 0.94195 | 32.7272 | 9.8233e-05
2008-02SWL#83 | 115 | 2009-01-12 13:53 | 24C | 0.93932 | 3.4494e-06 | 0 | 0.93932 | 32.6246 | 0.0001334
2008-02SWL#84 | 116 | 2009-01-12 13:55 | 24C | 0.93865 | 3.6108e-06 | 0 | 0.93865 | 32.5986 | 0.00014042
2008-02SWL#85 | 117 | 2009-01-12 13:57 | 24C | 0.9374 | 4.6356e-06 | 0 | 0.9374 | 32.5499 | 0.00018028
2008-02SWL#86 | 118 | 2009-01-12 13:59 | 24C | 0.93589 | 1.1196e-06 | 0 | 0.93589 | 32.4913 | 4.265e-05
2008-02SWL#87 | 119 | 2009-01-12 14:02 | 24C | 0.93584 | 1.7995e-06 | 0 | 0.93584 | 32.4891 | 6.9839e-05
2008-02SWL#88 | 120 | 2009-01-12 14:05 | 24C | 0.98437 | 5.8085e-06 | 0 | 0.98437 | 34.3857 | 0.00023041
2008-02SWL#89 | 121 | 2009-01-12 14:07 | 24C | 0.98219 | 3.4657e-06 | 0 | 0.98219 | 34.3003 | 0.00013824
2008-02SWL#90 | 122 | 2009-01-12 14:10 | 24C | 0.97956 | 3.2144e-06 | 0 | 0.97956 | 34.197 | 0.00013016
2008-02SWL#91 | 122d | 2009-01-12 14:12 | 24C | 0.97959 | 5.9204e-06 | 0 | 0.97959 | 34.1982 | 0.00023154
2008-02SWL#92 | 123 | 2009-01-12 14:14 | 24C | 0.97521 | 3.1414e-06 | 0 | 0.97521 | 34.0264 | 0.00012001
2008-02SWL#93 | 124 | 2009-01-12 14:17 | 24C | 0.97254 | 4.6354e-06 | 0 | 0.97254 | 33.9219 | 0.00018076
2008-02SWL#94 | 125 | 2009-01-12 14:19 | 24C | 0.97092 | 4.9203e-06 | 0 | 0.97092 | 33.8585 | 0.0001938
2008-02SWL#95 | 126 | 2009-01-12 14:21 | 24C | 0.97025 | 4.0241e-06 | 0 | 0.97025 | 33.8322 | 0.0001582
2008-02SWL#96 | 127 | 2009-01-12 14:24 | 24C | 0.9697 | 3.594e-06 | 0 | 0.9697 | 33.8105 | 0.00013903
2008-02SWL#97 | 128 | 2009-01-12 14:26 | 24C | 0.96774 | 1.3712e-06 | 0 | 0.96774 | 33.7338 | 5.2235e-05
2008-02SWL#98 | 129 | 2009-01-12 14:28 | 24C | 0.957 | 3.4515e-06 | 0 | 0.957 | 33.3142 | 0.0001339
2008-02SWL#99 | 130 | 2009-01-12 14:30 | 24C | 0.94329 | 4.9188e-06 | 0 | 0.94329 | 32.7793 | 0.00019532
2008-02SWL#100 | 131 | 2009-01-12 14:33 | 24C | 0.94118 | 3.323e-06 | 0 | 0.94118 | 32.6972 | 0.00013031
2008-02SWL#101 | 132 | 2009-01-12 14:36 | 24C | 0.94001 | 3.9739e-06 | 0 | 0.94001 | 32.6515 | 0.00015205
2008-02SWL#102 | 133 | 2009-01-12 14:38 | 24C | 0.93832 | 3.2153e-06 | 0 | 0.93832 | 32.5857 | 0.00012284
2008-02SWL#103 | 133d | 2009-01-12 14:41 | 24C | 0.9383 | 3.2145e-06 | 0 | 0.9383 | 32.585 | 0.00012427
2008-02SWL#104 | 134 | 2009-01-12 14:43 | 24C | 0.93679 | 2.7425e-06 | 0 | 0.93679 | 32.5261 | 0.00010865
2008-02SWL#105 | 135 | 2009-01-12 14:46 | 24C | 0.93687 | 3.3592e-06 | 0 | 0.93687 | 32.5293 | 0.00012614
2008-02SWL#106 | 158 | 2009-01-12 15:38 | 24C | 0.98392 | 4.3341e-06 | 0 | 0.98392 | 34.3681 | 0.00017057
2008-02SWL#107 | 159 | 2009-01-12 15:40 | 24C | 0.98207 | 2.7097e-06 | 0 | 0.98207 | 34.2955 | 0.00010447
2008-02SWL#108 | 160 | 2009-01-12 15:43 | 24C | 0.97906 | 4.3336e-06 | 0 | 0.97906 | 34.1774 | 0.00017133
2008-02SWL#109 | 161 | 2009-01-12 15:45 | 24C | 0.97449 | 4.8416e-06 | 0 | 0.97449 | 33.9982 | 0.00019175
2008-02SWL#110 | 161d | 2009-01-12 15:48 | 24C | 0.97453 | 2.0873e-06 | 0 | 0.97453 | 33.9999 | 8.2082e-05
2008-02SWL#111 | 162 | 2009-01-12 15:50 | 24C | 0.97191 | 2.2398e-06 | 0 | 0.97191 | 33.8973 | 8.9617e-05
2008-02SWL#112 | 163 | 2009-01-12 15:53 | 24C | 0.9706 | 4.2257e-06 | 0 | 0.9706 | 33.8459 | 0.00016282
2008-02SWL#113 | 164 | 2009-01-12 15:55 | 24C | 0.96904 | 3.2492e-06 | 0 | 0.96904 | 33.7847 | 0.00012633
2008-02SWL#114 | 165 | 2009-01-12 15:57 | 24C | 0.96823 | 1.1196e-06 | 0 | 0.96823 | 33.7532 | 4.1831e-05
2008-02SWL#115 | 166 | 2009-01-12 16:00 | 24C | 0.96427 | 2.7425e-06 | 0 | 0.96427 | 33.598 | 0.00010972
2008-02SWL#116 | 167 | 2009-01-12 16:02 | 24C | 0.95295 | 2.1278e-06 | 0 | 0.95295 | 33.1561 | 8.1169e-05
2008-02SWL#117 | 168 | 2009-01-12 16:05 | 24C | 0.94225 | 2.4909e-06 | 0 | 0.94225 | 32.7388 | 9.869e-05
2008-02SWL#118 | 169 | 2009-01-12 16:07 | 24C | 0.94111 | 5.3585e-06 | 0 | 0.94111 | 32.6942 | 0.00020897
2008-02SWL#119 | 169d | 2009-01-12 16:12 | 24C | 0.941 | 2.89e-06 | 0 | 0.941 | 32.6901 | 0.00011145
2008-02SWL#120 | 170 | 2009-01-12 16:15 | 24C | 0.94046 | 5.2783e-06 | 0 | 0.94046 | 32.6691 | 0.00020815
2008-02SWL#121 | 171 | 2009-01-12 16:17 | 24C | 0.93877 | 3.4657e-06 | 0 | 0.93877 | 32.6034 | 0.00013386
2008-02SWL#122 | 172 | 2009-01-12 16:20 | 24C | 0.93705 | 2.1429e-06 | 0 | 0.93705 | 32.5363 | 8.4736e-05
2008-02SWL#123 | 173 | 2009-01-12 16:22 | 24C | 0.93713 | 2.4615e-06 | 0 | 0.93713 | 32.5393 | 8.9214e-05
2008-02SWL#124 | 174 | 2009-01-12 16:24 | 24C | 0.98424 | 2.8901e-06 | 0 | 0.98424 | 34.3804 | 0.00011346
2008-02SWL#125 | 175 | 2009-01-12 16:27 | 24C | 0.9828 | 2.4916e-06 | 0 | 0.9828 | 34.3242 | 9.6599e-05
2008-02SWL#126 | 175d | 2009-01-12 16:29 | 24C | 0.98267 | 3.8746e-06 | 0 | 0.98267 | 34.3188 | 0.00015449
2008-02SWL#127 | 176 | 2009-01-12 16:32 | 24C | 0.98021 | 2.4908e-06 | 0 | 0.98021 | 34.2226 | 9.4885e-05
2008-02SWL#128 | 177 | 2009-01-12 16:34 | 24C | 0.97645 | 3.9873e-06 | 0 | 0.97645 | 34.0752 | 0.00015763
2008-02SWL#129 | 178 | 2009-01-12 16:36 | 24C | 0.97401 | 3.2141e-06 | 0 | 0.97401 | 33.9795 | 0.00012601
2008-02SWL#130 | 179 | 2009-01-12 16:38 | 24C | 0.97293 | 2.4911e-06 | 0 | 0.97293 | 33.9371 | 9.8715e-05
2008-02SWL#131 | 180 | 2009-01-12 16:41 | 24C | 0.97151 | 3.6814e-06 | 0 | 0.97151 | 33.8816 | 0.00014595
2008-02SWL#132 | 180d | 2009-01-12 16:44 | 24C | 0.97138 | 1.8035e-06 | 0 | 0.97138 | 33.8765 | 6.8803e-05
2008-02SWL#133 | 181 | 2009-01-12 16:47 | 24C | 0.97076 | 9.1415e-07 | 0 | 0.97076 | 33.8522 | 3.6914e-05
2008-02SWL#134 | 182 | 2009-01-12 16:50 | 24C | 0.96953 | 1.1199e-06 | 0 | 0.96953 | 33.8041 | 4.1799e-05
2008-02SWL#135 | 183 | 2009-01-12 16:53 | 24C | 0.9673 | 3.3587e-06 | 0 | 0.9673 | 33.7167 | 0.00013117
2008-02SWL#136 | 184 | 2009-01-12 16:55 | 24C | 0.96087 | 2.4911e-06 | 0 | 0.96087 | 33.4651 | 9.9543e-05
2008-02SWL#137 | 185 | 2009-01-12 16:57 | 24C | 0.94724 | 3.6112e-06 | 0 | 0.94724 | 32.9333 | 0.00014323
2008-02SWL#138 | 186 | 2009-01-12 17:00 | 24C | 0.9415 | 6.4282e-06 | 0 | 0.9415 | 32.7094 | 0.00024796
2008-02SWL#139 | 187 | 2009-01-12 17:02 | 24C | 0.94053 | 3.3588e-06 | 0 | 0.94053 | 32.6717 | 0.0001316
2008-02SWL#140 | 188 | 2009-01-12 17:05 | 24C | 0.93984 | 3.8437e-06 | 0 | 0.93984 | 32.6451 | 0.00015203
2008-02SWL#141 | 189 | 2009-01-12 17:07 | 24C | 0.93974 | 2.4909e-06 | 0 | 0.93974 | 32.6412 | 9.3856e-05

## #214-981 RAW DATA 2008-02
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
2008-02#1 | check | 2009-01-27 10:39 | 24C | 0.99987 | 3.1417e-06 | 0 | 0.99987 | 34.9948 | 0.00012315 | reading after 5 flushes with new standard
Calibration#1 | K15 | 2009-01-27 10:45 | 24C | 0.99986 | 1.1717e-05 | -1.4961e-05 | 0.99984 |  |  | cal check after 4 flushes
Calibration#2 | K15 | 2009-01-27 10:46 | 24C | 0.99984 | 4.189e-06 | -8.9407e-07 | 0.99984 |  |  | standarization value on next flush
2008-02#2 | 214 | 2009-01-27 10:54 | 24C | 0.98493 | 4.5839e-06 | -9.5367e-07 | 0.98493 | 34.4075 | 0.00018096 | algae in tank?
2008-02#3 | 217 | 2009-01-27 10:57 | 24C | 0.98347 | 3.8277e-06 | -9.5367e-07 | 0.98346 | 34.3501 | 0.00015033 | "
2008-02#4 | 217d | 2009-01-27 11:00 | 24C | 0.98338 | 2.475e-06 | -9.5367e-07 | 0.98338 | 34.3469 | 9.6599e-05 | "
2008-02#5 | 218 | 2009-01-27 11:04 | 24C | 0.98132 | 4.7143e-06 | -9.5367e-07 | 0.98132 | 34.2659 | 0.00018541 | "
2008-02#6 | 219 | 2009-01-27 11:06 | 24C | 0.97822 | 4.1822e-06 | -9.5367e-07 | 0.97821 | 34.1442 | 0.00016462 | "
2008-02#7 | 220 | 2009-01-27 11:10 | 24C | 0.97626 | 4.1892e-06 | -1.0133e-06 | 0.97626 | 34.0675 | 0.00016326 | "
2008-02#8 | 221 | 2009-01-27 11:31 | 24C | 0.97547 | 3.7635e-06 | -8.3447e-07 | 0.97547 | 34.0366 | 0.00014882 | "
2008-02#9 | 222 | 2009-01-27 11:34 | 24C | 0.97386 | 4.2152e-06 | -9.5367e-07 | 0.97386 | 33.9734 | 0.00017082 | "
2008-02#10 | 223 | 2009-01-27 11:37 | 24C | 0.97337 | 3.1417e-06 | -8.9407e-07 | 0.97337 | 33.9543 | 0.00012621 | "
2008-02#11 | 224 | 2009-01-27 11:39 | 24C | 0.97182 | 3.8647e-06 | -9.5367e-07 | 0.97182 | 33.8936 | 0.00015448 | "
2008-02#12 | 225 | 2009-01-27 11:41 | 24C | 0.97053 | 2.2241e-06 | -9.5367e-07 | 0.97053 | 33.8431 | 8.5299e-05 | "
2008-02#13 | 226 | 2009-01-27 11:43 | 24C | 0.9707 | 2.2592e-06 | -8.3447e-07 | 0.9707 | 33.8496 | 8.7862e-05 | "
2008-02#14 | 228 | 2009-01-27 11:45 | 24C | 0.94787 | 2.7425e-06 | -9.5367e-07 | 0.94787 | 32.9577 | 0.00010761 | "
2008-02#15 | 228d | 2009-01-27 11:47 | 24C | 0.94803 | 4.6353e-06 | -9.5367e-07 | 0.94803 | 32.9641 | 0.00018445 | "
2008-02#16 | 230 | 2009-01-27 11:49 | 24C | 0.94624 | 3.2141e-06 | -9.5367e-07 | 0.94623 | 32.894 | 0.00012692 | "
2008-02#17 | 232 | 2009-01-27 11:52 | 24C | 0.94556 | 3.9635e-06 | -9.5367e-07 | 0.94556 | 32.8677 | 0.00015421
2008-02#18 | 234 | 2009-01-27 11:54 | 24C | 0.94569 | 3.5825e-06 | -9.5367e-07 | 0.94569 | 32.8727 | 0.00014293
2008-02#19 | 236 | 2009-01-27 11:56 | 24C | 0.94477 | 2.4911e-06 | -9.5367e-07 | 0.94477 | 32.8368 | 9.3005e-05
2008-02#20 | 260 | 2009-01-27 11:59 | 24C | 0.98485 | 3.4906e-06 | -9.5367e-07 | 0.98485 | 34.4046 | 0.00013786
2008-02#21 | 261 | 2009-01-27 12:01 | 24C | 0.9834 | 3.4665e-06 | -1.0133e-06 | 0.9834 | 34.3476 | 0.00013801
2008-02#22 | 262 | 2009-01-27 12:05 | 24C | 0.9815 | 3.6105e-06 | -8.3447e-07 | 0.9815 | 34.2729 | 0.00014157
2008-02#23 | 262d | 2009-01-27 12:07 | 24C | 0.98156 | 3.1415e-06 | -1.0729e-06 | 0.98156 | 34.2754 | 0.00012304
2008-02#24 | 263 | 2009-01-27 12:09 | 24C | 0.97839 | 2.7426e-06 | -9.5367e-07 | 0.97839 | 34.151 | 0.00010865
2008-02#25 | 264 | 2009-01-27 12:11 | 24C | 0.97625 | 2.2391e-06 | -1.0729e-06 | 0.97625 | 34.0671 | 8.7077e-05
2008-02#26 | 265 | 2009-01-27 12:13 | 24C | 0.97483 | 3.618e-06 | -9.5367e-07 | 0.97483 | 34.0115 | 0.00014216
2008-02#27 | 266 | 2009-01-27 13:33 | 24C | 0.97333 | 5.2934e-06 | -9.5367e-07 | 0.97333 | 33.9527 | 0.00020962 | see notes -flush cells with CLR and add bleach to tank -JL
2008-02#28 | 267 | 2009-01-27 13:36 | 24C | 0.97248 | 2.7426e-06 | -1.0729e-06 | 0.97248 | 33.9194 | 0.00010865
2008-02#29 | 268 | 2009-01-27 13:38 | 24C | 0.97107 | 2.4912e-06 | -9.5367e-07 | 0.97107 | 33.8644 | 9.9734e-05
2008-02#30 | 268d | 2009-01-27 13:40 | 24C | 0.97104 | 2.2394e-06 | -9.5367e-07 | 0.97104 | 33.863 | 8.8748e-05
2008-02#31 | 269 | 2009-01-27 13:42 | 24C | 0.96857 | 3.2141e-06 | -1.0729e-06 | 0.96857 | 33.7664 | 0.00012553
2008-02#32 | 270 | 2009-01-27 13:44 | 24C | 0.96319 | 6.1782e-06 | -9.5367e-07 | 0.96319 | 33.5559 | 0.00024656
2008-02#33 | 271 | 2009-01-27 13:47 | 24C | 0.95415 | 4.287e-06 | -8.9407e-07 | 0.95415 | 33.2026 | 0.00016863
2008-02#34 | 272 | 2009-01-27 13:50 | 24C | 0.94812 | 2.4912e-06 | -9.5367e-07 | 0.94812 | 32.9675 | 9.802e-05
2008-02#35 | 273 | 2009-01-27 13:53 | 24C | 0.94328 | 1.1536e-06 | -9.5367e-07 | 0.94328 | 32.7788 | 4.265e-05
2008-02#36 | 274 | 2009-01-27 13:55 | 24C | 0.94206 | 3.2148e-06 | -9.5367e-07 | 0.94206 | 32.7312 | 0.00012382
2008-02#37 | 275 | 2009-01-27 13:58 | 24C | 0.94183 | 3.6105e-06 | -1.0133e-06 | 0.94183 | 32.7225 | 0.00013903
2008-02#38 | 276 | 2009-01-27 14:01 | 24C | 0.98531 | 2.7425e-06 | -9.5367e-07 | 0.98531 | 34.4225 | 0.00010765
2008-02#39 | 277 | 2009-01-27 14:04 | 24C | 0.98366 | 4.6251e-06 | -9.5367e-07 | 0.98366 | 34.3577 | 0.00018436
2008-02#40 | 278 | 2009-01-27 14:07 | 24C | 0.98156 | 5.0935e-06 | -9.5367e-07 | 0.98156 | 34.2755 | 0.00019958
2008-02#41 | 278d | 2009-01-27 14:10 | 24C | 0.98157 | 4.2256e-06 | -9.5367e-07 | 0.98157 | 34.2757 | 0.00016807
2008-02#42 | 279 | 2009-01-27 14:12 | 24C | 0.97866 | 4.5981e-06 | -1.0133e-06 | 0.97865 | 34.1614 | 0.00017516
2008-02#43 | 280 | 2009-01-27 14:16 | 24C | 0.9767 | 5.3107e-06 | -9.5367e-07 | 0.9767 | 34.0849 | 0.0002084
2008-02#44 | 281 | 2009-01-27 14:19 | 24C | 0.97566 | 2.8901e-06 | -8.9407e-07 | 0.97565 | 34.0439 | 0.00011317
2008-02#45 | 282 | 2009-01-27 14:22 | 24C | 0.97379 | 2.1278e-06 | -9.5367e-07 | 0.97378 | 33.9706 | 8.1169e-05
2008-02#46 | 283 | 2009-01-27 14:26 | 24C | 0.97309 | 2.7425e-06 | -9.5367e-07 | 0.97309 | 33.9433 | 0.0001097
2008-02#47 | 284 | 2009-01-27 14:30 | 24C | 0.97218 | 1.6496e-06 | -9.5367e-07 | 0.97218 | 33.9079 | 6.4686e-05
2008-02#48 | 285 | 2009-01-27 14:33 | 24C | 0.96767 | 2.7262e-06 | -8.9407e-07 | 0.96767 | 33.7311 | 0.00010656
2008-02#49 | 286 | 2009-01-27 14:36 | 24C | 0.95947 | 1.1532e-06 | -9.5367e-07 | 0.95947 | 33.4105 | 4.4364e-05
2008-02#50 | 287 | 2009-01-27 14:38 | 24C | 0.9492 | 3.5946e-06 | -9.5367e-07 | 0.94919 | 33.0094 | 0.00014191
2008-02#51 | 288 | 2009-01-27 14:41 | 24C | 0.94541 | 3.3587e-06 | -8.9407e-07 | 0.94541 | 32.8619 | 0.0001349
2008-02#52 | 289 | 2009-01-27 14:44 | 24C | 0.94294 | 2.1282e-06 | -9.5367e-07 | 0.94294 | 32.7657 | 8.2558e-05
2008-02#53 | 289d | 2009-01-27 14:47 | 24C | 0.94293 | 3.3591e-06 | -9.5367e-07 | 0.94293 | 32.7654 | 0.00013116
2008-02#54 | 290 | 2009-01-27 14:51 | 24C | 0.94115 | 4.1786e-06 | -8.3447e-07 | 0.94115 | 32.696 | 0.00016096
2008-02#55 | 291 | 2009-01-27 14:53 | 24C | 0.94144 | 1.1196e-06 | -9.5367e-07 | 0.94144 | 32.7073 | 4.4364e-05
2008-02#56 | 292 | 2009-01-27 14:57 | 24C | 0.98661 | 2.8897e-06 | -1.0133e-06 | 0.98661 | 34.4736 | 0.00011145
2008-02#57 | 293 | 2009-01-27 14:59 | 24C | 0.9848 | 2.2397e-06 | -1.0133e-06 | 0.9848 | 34.4027 | 8.7024e-05
2008-02#58 | 293d | 2009-01-27 15:02 | 24C | 0.98465 | 3.8756e-06 | -9.5367e-07 | 0.98465 | 34.3965 | 0.00015551
2008-02#59 | 294 | 2009-01-27 15:31 | 24C | 0.98157 | 2.2394e-06 | -8.3447e-07 | 0.98157 | 34.2759 | 8.7077e-05
2008-02#60 | 295 | 2009-01-27 15:34 | 24C | 0.9787 | 3.5406e-06 | -1.0729e-06 | 0.9787 | 34.1633 | 0.00014162
2008-02#61 | 296 | 2009-01-27 15:37 | 24C | 0.97619 | 1.3712e-06 | -9.5367e-07 | 0.97619 | 34.0649 | 5.2235e-05
2008-02#62 | 297 | 2009-01-27 15:39 | 24C | 0.97488 | 2.4776e-06 | -9.5367e-07 | 0.97488 | 34.0133 | 9.7452e-05
2008-02#63 | 298 | 2009-01-27 15:42 | 24C | 0.97315 | 3.3595e-06 | -8.9407e-07 | 0.97315 | 33.9456 | 0.00013139
2008-02#64 | 299 | 2009-01-27 15:44 | 24C | 0.97235 | 3.466e-06 | -8.3447e-07 | 0.97235 | 33.9145 | 0.00013826
2008-02#65 | 299d | 2009-01-27 15:47 | 24C | 0.97218 | 2.4911e-06 | -8.9407e-07 | 0.97218 | 33.9078 | 9.869e-05
2008-02#66 | 300 | 2009-01-27 15:49 | 24C | 0.97063 | 2.8901e-06 | -1.0133e-06 | 0.97062 | 33.8468 | 0.0001115
2008-02#67 | 301 | 2009-01-27 15:52 | 24C | 0.96356 | 1.1199e-06 | -1.0133e-06 | 0.96356 | 33.5704 | 4.5227e-05
2008-02#68 | 302 | 2009-01-27 15:55 | 24C | 0.95612 | 1.4046e-06 | -9.5367e-07 | 0.95612 | 33.2798 | 5.2235e-05
2008-02#69 | 303 | 2009-01-27 15:57 | 24C | 0.9494 | 3.4658e-06 | -8.9407e-07 | 0.9494 | 33.0175 | 0.00013799
2008-02#70 | 304 | 2009-01-27 16:00 | 24C | 0.94257 | 3.4355e-06 | -9.5367e-07 | 0.94257 | 32.7511 | 0.00013596
2008-02#71 | 305 | 2009-01-27 16:03 | 24C | 0.94088 | 2.2392e-06 | -9.5367e-07 | 0.94088 | 32.6855 | 8.7473e-05
2008-02#72 | 306 | 2009-01-27 16:06 | 24C | 0.93913 | 5.2672e-06 | -9.5367e-07 | 0.93913 | 32.6171 | 0.0002076
2008-02#73 | 307 | 2009-01-27 16:10 | 24C | 0.93906 | 3.1414e-06 | -9.5367e-07 | 0.93906 | 32.6144 | 0.00012417
2008-02#74 | 333 | 2009-01-27 16:12 | 24C | 0.98332 | 5.8494e-06 | -9.5367e-07 | 0.98332 | 34.3445 | 0.00022973
2008-02#75 | 334 | 2009-01-27 16:15 | 24C | 0.9813 | 4.0931e-06 | -8.3447e-07 | 0.9813 | 34.2651 | 0.00015909
2008-02#76 | 335 | 2009-01-27 16:19 | 24C | 0.97906 | 2.2394e-06 | -1.0729e-06 | 0.97906 | 34.1773 | 8.958e-05
2008-02#77 | 336 | 2009-01-27 16:22 | 24C | 0.97576 | 3.4657e-06 | -9.5367e-07 | 0.97576 | 34.048 | 0.00013654
2008-02#78 | 338 | 2009-01-27 16:25 | 24C | 0.97073 | 2.4911e-06 | -8.3447e-07 | 0.97073 | 33.851 | 9.869e-05 | incorrect order
2008-02#79 | 337 | 2009-01-27 16:30 | 24C | 0.97318 | 3.359e-06 | -9.5367e-07 | 0.97318 | 33.9467 | 0.00013255 | incorrect order
2008-02#80 | 336d | 2009-01-27 16:33 | 24C | 0.97526 | 2.4747e-06 | -9.5367e-07 | 0.97526 | 34.0285 | 9.9734e-05 | incorrect order
2008-02#81 | 339 | 2009-01-27 16:36 | 24C | 0.96916 | 4.3337e-06 | -8.9407e-07 | 0.96916 | 33.7893 | 0.00016844
2008-02#82 | 340 | 2009-01-27 16:39 | 24C | 0.96608 | 3.4657e-06 | -9.5367e-07 | 0.96608 | 33.6689 | 0.00013633
2008-02#83 | 341 | 2009-01-27 16:41 | 24C | 0.96055 | 2.7425e-06 | -8.3447e-07 | 0.96055 | 33.4529 | 0.00010658
2008-02#84 | 342 | 2009-01-27 16:43 | 24C | 0.95062 | 3.4657e-06 | -9.5367e-07 | 0.95062 | 33.0649 | 0.00013904
2008-02#85 | 343 | 2009-01-27 16:46 | 24C | 0.94414 | 3.1417e-06 | -9.5367e-07 | 0.94414 | 32.8124 | 0.00012313
2008-02#86 | 344 | 2009-01-27 16:48 | 24C | 0.93773 | 2.4909e-06 | -8.9407e-07 | 0.93773 | 32.5626 | 9.4938e-05
2008-02#87 | 345 | 2009-01-27 16:50 | 24C | 0.9305 | 2.2262e-06 | -9.5367e-07 | 0.9305 | 32.2816 | 8.5299e-05
2008-02#88 | 346 | 2009-01-27 16:53 | 24C | 0.92904 | 3.1418e-06 | -8.9407e-07 | 0.92904 | 32.2248 | 0.00012419
2008-02#89 | 347 | 2009-01-27 16:56 | 24C | 0.92719 | 2.4911e-06 | -8.3447e-07 | 0.92719 | 32.153 | 9.8725e-05
2008-02#90 | 347d | 2009-01-27 16:58 | 24C | 0.92743 | 2.4909e-06 | -9.5367e-07 | 0.92743 | 32.1623 | 9.4885e-05
2008-02#91 | 348 | 2009-01-27 17:01 | 24C | 0.92688 | 2.2392e-06 | -9.5367e-07 | 0.92688 | 32.141 | 8.5299e-05
2008-02#92 | 374 | 2009-01-27 17:13 | 24C | 0.98351 | 2.4751e-06 | -9.5367e-07 | 0.98351 | 34.3519 | 9.7495e-05
2008-02#93 | 375 | 2009-01-27 17:16 | 24C | 0.98128 | 2.0945e-06 | -9.5367e-07 | 0.98128 | 34.2644 | 8.2783e-05
2008-02#94 | 375d | 2009-01-27 17:18 | 24C | 0.98194 | 1.9265e-06 | -9.5367e-07 | 0.98194 | 34.2903 | 7.4302e-05
2008-02#95 | 376 | 2009-01-27 17:22 | 24C | 0.9788 | 3.1415e-06 | -9.5367e-07 | 0.9788 | 34.1673 | 0.00012419
2008-02#96 | 377 | 2009-01-27 17:25 | 24C | 0.97579 | 3.9871e-06 | -9.5367e-07 | 0.97579 | 34.0491 | 0.00015603
2008-02#97 | 378 | 2009-01-27 17:27 | 24C | 0.97342 | 2.7433e-06 | -8.9407e-07 | 0.97342 | 33.9562 | 0.00010447
2008-02#98 | 379 | 2009-01-27 17:29 | 24C | 0.97179 | 1.8038e-06 | -8.3447e-07 | 0.97179 | 33.8924 | 6.8797e-05
2008-02#99 | 380 | 2009-01-27 17:32 | 24C | 0.97023 | 1.1199e-06 | -9.5367e-07 | 0.97023 | 33.8313 | 4.2703e-05
2008-02#100 | 381 | 2009-01-27 17:35 | 24C | 0.96846 | 2.4753e-06 | -9.5367e-07 | 0.96846 | 33.7622 | 9.8045e-05
2008-02#101 | 382 | 2009-01-27 17:39 | 24C | 0.96553 | 3.3587e-06 | -9.5367e-07 | 0.96553 | 33.6475 | 0.00012972
2008-02#102 | cal check | 2009-01-28 08:57 | 24C | 0.99985 | 5.0791e-06 | -9.5367e-07 | 0.99985 | 34.9942 | 0.00019806 | cal check with in +-0.00001
2008-02#103 | 383 | 2009-01-28 09:06 | 24C | 0.96104 | 6.0066e-06 | -8.9407e-07 | 0.96104 | 33.4718 | 0.00023729
2008-02#104 | 383d | 2009-01-28 09:09 | 24C | 0.96115 | 4.4783e-06 | -9.5367e-07 | 0.96115 | 33.4761 | 0.00017692
2008-02#105 | 384 | 2009-01-28 09:12 | 24C | 0.95027 | 2.5036e-06 | -1.0729e-06 | 0.95027 | 33.0514 | 9.8238e-05
2008-02#106 | 385 | 2009-01-28 09:17 | 24C | 0.94356 | 3.4659e-06 | -1.0133e-06 | 0.94356 | 32.7897 | 0.00013345
2008-02#107 | 386 | 2009-01-28 09:20 | 24C | 0.93917 | 4.2263e-06 | -9.5367e-07 | 0.93917 | 32.6188 | 0.00016232
2008-02#108 | 387 | 2009-01-28 09:23 | 24C | 0.9344 | 2.7425e-06 | -8.3447e-07 | 0.9344 | 32.4332 | 0.00010344
2008-02#109 | 388 | 2009-01-28 09:25 | 24C | 0.92932 | 1.3712e-06 | -9.5367e-07 | 0.92932 | 32.2356 | 5.5371e-05
2008-02#110 | 389 | 2009-01-28 09:28 | 24C | 0.92823 | 3.974e-06 | -1.0133e-06 | 0.92823 | 32.1934 | 0.00015175
2008-02#111 | 419 | 2009-01-28 09:40 | 24C | 0.98348 | 4.2151e-06 | -9.5367e-07 | 0.98348 | 34.3508 | 0.00016417
2008-02#112 | 420 | 2009-01-28 09:42 | 24C | 0.98148 | 5.3586e-06 | -1.0729e-06 | 0.98148 | 34.2721 | 0.00020941
2008-02#113 | 421 | 2009-01-28 09:45 | 24C | 0.97901 | 3.3587e-06 | -9.5367e-07 | 0.97901 | 34.1754 | 0.00012795
2008-02#114 | 421d | 2009-01-28 09:47 | 24C | 0.97839 | 1.3712e-06 | -9.5367e-07 | 0.97839 | 34.1511 | 5.4326e-05
2008-02#115 | 422 | 2009-01-28 09:50 | 24C | 0.97535 | 2.7426e-06 | -8.9407e-07 | 0.97535 | 34.032 | 0.00011074
2008-02#116 | 423 | 2009-01-28 09:53 | 24C | 0.97376 | 6.2828e-06 | -9.5367e-07 | 0.97376 | 33.9695 | 0.0002448
2008-02#117 | 424 | 2009-01-28 09:56 | 24C | 0.97276 | 4.6354e-06 | -9.5367e-07 | 0.97275 | 33.9302 | 0.0001836
2008-02#118 | 425 | 2009-01-28 09:58 | 24C | 0.97172 | 4.7138e-06 | -9.5367e-07 | 0.97172 | 33.8896 | 0.000186
2008-02#119 | 426 | 2009-01-28 10:01 | 24C | 0.9712 | 2.8901e-06 | -9.5367e-07 | 0.97119 | 33.8691 | 0.00011317
2008-02#120 | 427 | 2009-01-28 10:04 | 24C | 0.9698 | 2.2397e-06 | -1.0133e-06 | 0.9698 | 33.8145 | 8.5352e-05
2008-02#121 | 428 | 2009-01-28 10:06 | 24C | 0.96466 | 2.4912e-06 | -1.0729e-06 | 0.96466 | 33.6134 | 9.4885e-05
2008-02#122 | 429 | 2009-01-28 10:33 | 24C | 0.95619 | 4.3337e-06 | -9.5367e-07 | 0.95619 | 33.2822 | 0.00016797
2008-02#123 | 430 | 2009-01-28 10:36 | 24C | 0.94641 | 3.3588e-06 | -1.0133e-06 | 0.94641 | 32.9007 | 0.00013192
2008-02#124 | 431 | 2009-01-28 10:38 | 24C | 0.9417 | 3.9515e-06 | -9.5367e-07 | 0.94169 | 32.7171 | 0.00015428
2008-02#125 | 432 | 2009-01-28 10:42 | 24C | 0.93456 | 3.4658e-06 | -9.5367e-07 | 0.93456 | 32.4395 | 0.00013635
2008-02#126 | 432d | 2009-01-28 10:44 | 24C | 0.93459 | 1.1196e-06 | -1.0133e-06 | 0.93459 | 32.4408 | 4.4374e-05
2008-02#127 | 433 | 2009-01-28 10:47 | 24C | 0.92935 | 5.5668e-06 | -9.5367e-07 | 0.92935 | 32.2367 | 0.00021459
2008-02#128 | 434 | 2009-01-28 10:50 | 24C | 0.92864 | 4.1144e-06 | -8.3447e-07 | 0.92864 | 32.2094 | 0.00016016
2008-02#129 | 435 | 2009-01-28 10:53 | 24C | 0.98401 | 2.0949e-06 | -9.5367e-07 | 0.98401 | 34.3715 | 8.3668e-05
2008-02#130 | 436 | 2009-01-28 10:55 | 24C | 0.98142 | 4.5021e-06 | -1.0133e-06 | 0.98142 | 34.2699 | 0.00017592
2008-02#131 | 437 | 2009-01-28 10:57 | 24C | 0.97788 | 2.8904e-06 | -9.5367e-07 | 0.97788 | 34.131 | 0.00011062
2008-02#132 | 438 | 2009-01-28 11:00 | 24C | 0.97452 | 3.4494e-06 | -8.9407e-07 | 0.97452 | 33.9993 | 0.0001334
2008-02#133 | 438d | 2009-01-28 11:02 | 24C | 0.97451 | 3.6232e-06 | -9.5367e-07 | 0.97451 | 33.999 | 0.00013993
2008-02#134 | 439 | 2009-01-28 11:07 | 24C | 0.97258 | 4.276e-06 | -9.5367e-07 | 0.97258 | 33.9235 | 0.00016653
2008-02#135 | 440 | 2009-01-28 11:13 | 24C | 0.97105 | 2.8794e-06 | -9.5367e-07 | 0.97105 | 33.8635 | 0.00011316
2008-02#136 | 441 | 2009-01-28 11:16 | 24C | 0.96831 | 2.4921e-06 | -1.0133e-06 | 0.96831 | 33.7562 | 9.7246e-05
2008-02#137 | 442 | 2009-01-28 11:19 | 24C | 0.961 | 2.7262e-06 | -8.9407e-07 | 0.961 | 33.4703 | 0.00010659
2008-02#138 | 443 | 2009-01-28 11:21 | 24C | 0.95369 | 2.7426e-06 | -9.5367e-07 | 0.95369 | 33.1847 | 0.00010865
2008-02#139 | 444 | 2009-01-28 11:26 | 24C | 0.94739 | 1.1529e-06 | -8.3447e-07 | 0.94739 | 32.9389 | 4.265e-05
2008-02#140 | 445 | 2009-01-28 11:34 | 24C | 0.94332 | 2.8661e-06 | -9.5367e-07 | 0.94332 | 32.7803 | 0.00011606
2008-02#141 | 446 | 2009-01-28 11:37 | 24C | 0.93783 | 3.5957e-06 | -9.5367e-07 | 0.93783 | 32.5667 | 0.00014113
2008-02#142 | 447 | 2009-01-28 11:39 | 24C | 0.93069 | 2.2391e-06 | -1.0729e-06 | 0.93069 | 32.2888 | 8.7024e-05
2008-02#143 | 447d | 2009-01-28 11:41 | 24C | 0.93069 | 3.2142e-06 | -9.5367e-07 | 0.93069 | 32.2889 | 0.00012555
2008-02#144 | 448 | 2009-01-28 11:44 | 24C | 0.92562 | 2.4777e-06 | -9.5367e-07 | 0.92562 | 32.0921 | 9.7029e-05
2008-02#145 | 449 | 2009-01-28 11:47 | 24C | 0.92335 | 2.7426e-06 | -9.5367e-07 | 0.92335 | 32.0039 | 0.00010447
2008-02#146 | 450 | 2009-01-28 11:53 | 24C | 0.92333 | 2.4912e-06 | -8.3447e-07 | 0.92333 | 32.0032 | 9.6642e-05
2008-02#147 | 453 | 2009-01-28 11:56 | 24C | 0.93003 | 3.4658e-06 | -9.5367e-07 | 0.93003 | 32.2632 | 0.0001334
2008-02#148 | 454 | 2009-01-28 11:58 | 24C | 0.92912 | 3.466e-06 | -9.5367e-07 | 0.92912 | 32.2279 | 0.00013448
2008-02#149 | 454d | 2009-01-28 12:00 | 24C | 0.92918 | 2.4912e-06 | -8.9407e-07 | 0.92918 | 32.2301 | 0.00010061
2008-02#150 | 455 | 2009-01-28 12:02 | 24C | 0.92628 | 2.1278e-06 | -8.9407e-07 | 0.92627 | 32.1174 | 8.2541e-05
2008-02#151 | 456 | 2009-01-28 12:04 | 24C | 0.92764 | 2.7262e-06 | -9.5367e-07 | 0.92764 | 32.1705 | 0.00010447
2008-02#152 | 457 | 2009-01-28 12:07 | 24C | 0.92536 | 1.1064e-06 | -9.5367e-07 | 0.92536 | 32.082 | 4.2703e-05
2008-02#153 | 458 | 2009-01-28 12:09 | 24C | 0.92428 | 2.4909e-06 | -1.0729e-06 | 0.92428 | 32.0401 | 9.7452e-05
2008-02#154 | 458d | 2009-01-28 12:11 | 24C | 0.92431 | 1.4046e-06 | -8.3447e-07 | 0.92431 | 32.041 | 5.1197e-05
2008-02#155 | 459 | 2009-01-28 12:14 | 24C | 0.9243 | 3.3588e-06 | -8.9407e-07 | 0.9243 | 32.0409 | 0.00012969
2008-02#156 | 460 | 2009-01-28 12:16 | 24C | 0.92431 | 2.2608e-06 | -9.5367e-07 | 0.92431 | 32.0411 | 8.7551e-05
2008-02#157 | 461 | 2009-01-28 12:18 | 24C | 0.92435 | 2.7261e-06 | -9.5367e-07 | 0.92435 | 32.0426 | 0.00010343
2008-02#158 | 462 | 2009-01-28 12:20 | 24C | 0.93195 | 2.8799e-06 | -9.5367e-07 | 0.93195 | 32.3378 | 0.0001115
2008-02#159 | 463 | 2009-01-28 12:22 | 24C | 0.92952 | 2.4916e-06 | -9.5367e-07 | 0.92952 | 32.2434 | 9.6599e-05
2008-02#160 | 464 | 2009-01-28 12:24 | 24C | 0.92725 | 1.7597e-06 | -9.5367e-07 | 0.92725 | 32.1553 | 6.7435e-05
2008-02#161 | 465 | 2009-01-28 12:26 | 24C | 0.92715 | 1.3713e-06 | -9.5367e-07 | 0.92715 | 32.1512 | 5.4326e-05
2008-02#162 | 466 | 2009-01-28 12:28 | 24C | 0.9264 | 2.1278e-06 | -1.0133e-06 | 0.92639 | 32.1221 | 8.1174e-05
2008-02#163 | 466d | 2009-01-28 12:30 | 24C | 0.92645 | 3.2147e-06 | -8.3447e-07 | 0.92645 | 32.1242 | 0.00012416
2008-02#164 | 467 | 2009-01-28 12:33 | 24C | 0.92687 | 3.6108e-06 | -1.0729e-06 | 0.92687 | 32.1405 | 0.00014293
2008-02#165 | 468 | 2009-01-28 12:35 | 24C | 0.92693 | 3.3587e-06 | -8.9407e-07 | 0.92693 | 32.1429 | 0.0001316
2008-02#166 | 469 | 2009-01-28 12:37 | 24C | 0.92696 | 2.4912e-06 | -1.0729e-06 | 0.92696 | 32.144 | 9.7029e-05
2008-02#167 | 470 | 2009-01-28 12:40 | 24C | 0.93467 | 2.7097e-06 | -9.5367e-07 | 0.93467 | 32.4436 | 0.00010447
2008-02#168 | 471 | 2009-01-28 12:43 | 24C | 0.93435 | 4.5129e-06 | -8.3447e-07 | 0.93435 | 32.4312 | 0.00017589
2008-02#169 | 472 | 2009-01-28 12:45 | 24C | 0.93114 | 1.1529e-06 | -9.5367e-07 | 0.93114 | 32.3064 | 4.265e-05
2008-02#170 | 473 | 2009-01-28 12:47 | 24C | 0.92858 | 1.1196e-06 | -1.0133e-06 | 0.92858 | 32.2069 | 4.6915e-05
2008-02#171 | 473d | 2009-01-28 12:49 | 24C | 0.92858 | 2.2392e-06 | -1.0133e-06 | 0.92858 | 32.2071 | 8.7009e-05
2008-02#172 | 474 | 2009-01-28 12:56 | 24C | 0.92755 | 4.1892e-06 | -8.3447e-07 | 0.92755 | 32.1669 | 0.00016096
2008-02#173 | 475 | 2009-01-28 12:59 | 24C | 0.92633 | 5.47e-06 | -9.5367e-07 | 0.92633 | 32.1196 | 0.00021231
2008-02#174 | 476 | 2009-01-28 13:42 | 24C | 0.92637 | 3.8332e-06 | -9.5367e-07 | 0.92637 | 32.1211 | 0.00014717
2008-02#175 | 477 | 2009-01-28 13:44 | 24C | 0.92638 | 3.4551e-06 | -9.5367e-07 | 0.92638 | 32.1214 | 0.00013203
2008-02#176 | 478 | 2009-01-28 13:47 | 24C | 0.92635 | 5.4274e-06 | -1.0133e-06 | 0.92635 | 32.1203 | 0.00021247
2008-02#177 | 480 | 2009-01-28 13:52 | 24C | 0.93103 | 5.1158e-06 | -9.5367e-07 | 0.93103 | 32.3023 | 0.000199
2008-02#178 | 481 | 2009-01-28 13:54 | 24C | 0.93099 | 2.7425e-06 | -9.5367e-07 | 0.93099 | 32.3007 | 0.00010447
2008-02#179 | 482 | 2009-01-28 13:57 | 24C | 0.92905 | 3.125e-06 | -9.5367e-07 | 0.92905 | 32.2251 | 0.00012213
2008-02#180 | 482d | 2009-01-28 14:00 | 24C | 0.92907 | 2.4915e-06 | -1.0729e-06 | 0.92907 | 32.226 | 9.6599e-05
2008-02#181 | 483 | 2009-01-28 14:03 | 24C | 0.92686 | 1.1196e-06 | -9.5367e-07 | 0.92686 | 32.1402 | 4.4374e-05
2008-02#182 | 484 | 2009-01-28 14:05 | 24C | 0.92621 | 3.4659e-06 | -9.5367e-07 | 0.92621 | 32.1148 | 0.00013724
2008-02#183 | 485 | 2009-01-28 14:08 | 24C | 0.92562 | 2.7425e-06 | -8.3447e-07 | 0.92562 | 32.092 | 0.00010656
2008-02#184 | 486 | 2009-01-28 14:11 | 24C | 0.92159 | 1.3713e-06 | -9.5367e-07 | 0.92159 | 31.9357 | 5.2235e-05
2008-02#185 | 487 | 2009-01-28 14:13 | 24C | 0.92127 | 3.2148e-06 | -9.5367e-07 | 0.92127 | 31.923 | 0.00012553
2008-02#186 | 488 | 2009-01-28 14:15 | 24C | 0.92113 | 2.8901e-06 | -1.0133e-06 | 0.92112 | 31.9175 | 0.00011392
2008-02#187 | 489 | 2009-01-28 14:18 | 24C | 0.92548 | 2.4909e-06 | -9.5367e-07 | 0.92548 | 32.0864 | 9.9792e-05
2008-02#188 | 489d | 2009-01-28 14:20 | 24C | 0.92561 | 2.7425e-06 | -8.3447e-07 | 0.92561 | 32.0916 | 0.0001097
2008-02#189 | 490 | 2009-01-28 14:24 | 24C | 0.92466 | 1.1064e-06 | -9.5367e-07 | 0.92466 | 32.0548 | 4.265e-05
2008-02#190 | 491 | 2009-01-28 14:26 | 24C | 0.92093 | 3.2145e-06 | -9.5367e-07 | 0.92092 | 31.9097 | 0.00012398
2008-02#191 | 492 | 2009-01-28 14:29 | 24C | 0.92004 | 3.2141e-06 | -9.5367e-07 | 0.92004 | 31.8753 | 0.000126
2008-02#192 | 493 | 2009-01-28 14:33 | 24C | 0.91888 | 2.4916e-06 | -1.0729e-06 | 0.91888 | 31.8305 | 9.7081e-05
2008-02#193 | 494 | 2009-01-28 14:35 | 24C | 0.91821 | 3.131e-06 | -8.9407e-07 | 0.91821 | 31.8043 | 0.00012365
2008-02#194 | 495 | 2009-01-28 14:37 | 24C | 0.91805 | 1.3712e-06 | -9.5367e-07 | 0.91805 | 31.7983 | 5.3301e-05
2008-02#195 | 496 | 2009-01-28 14:39 | 24C | 0.91804 | 2.4908e-06 | -9.5367e-07 | 0.91804 | 31.798 | 9.7017e-05
2008-02#196 | 499 | 2009-01-28 14:44 | 24C | 0.92162 | 3.1414e-06 | -8.9407e-07 | 0.92162 | 31.9367 | 0.00012228 | missing 497 and 498 ?
2008-02#197 | 500 | 2009-01-28 14:46 | 24C | 0.92146 | 1.1064e-06 | -9.5367e-07 | 0.92145 | 31.9303 | 4.543e-05
2008-02#198 | 501 | 2009-01-28 14:48 | 24C | 0.92135 | 1.3713e-06 | -1.0729e-06 | 0.92135 | 31.9261 | 5.5913e-05
2008-02#199 | 502 | 2009-01-28 14:51 | 24C | 0.92134 | 2.1278e-06 | -9.5367e-07 | 0.92133 | 31.9256 | 8.2127e-05
2008-02#200 | 503 | 2009-01-28 14:53 | 24C | 0.92096 | 2.7425e-06 | -1.0729e-06 | 0.92096 | 31.9111 | 0.00010708
2008-02#201 | 503d | 2009-01-28 14:56 | 24C | 0.92081 | 3.1415e-06 | -1.0133e-06 | 0.92081 | 31.9051 | 0.00012139
2008-02#202 | 504 | 2009-01-28 14:59 | 24C | 0.92067 | 2.4909e-06 | -9.5367e-07 | 0.92067 | 31.9 | 9.9652e-05
2008-02#203 | 505 | 2009-01-28 15:01 | 24C | 0.92074 | 1.3713e-06 | -9.5367e-07 | 0.92074 | 31.9024 | 5.4847e-05
2008-02#204 | 506 | 2009-01-28 15:03 | 24C | 0.9207 | 1.4046e-06 | -8.9407e-07 | 0.9207 | 31.9008 | 5.4847e-05
2008-02#205 | 507 | 2009-01-28 15:43 | 24C | 0.98269 | 2.1136e-06 | -9.5367e-07 | 0.98269 | 34.3196 | 8.1169e-05
2008-02#206 | 507d | 2009-01-28 15:45 | 24C | 0.98266 | 2.4909e-06 | -1.0729e-06 | 0.98266 | 34.3185 | 9.869e-05
2008-02#207 | 508 | 2009-01-28 15:47 | 24C | 0.97985 | 4.6245e-06 | -9.5367e-07 | 0.97985 | 34.2084 | 0.00018324
2008-02#208 | 509 | 2009-01-28 15:50 | 24C | 0.97504 | 1.3712e-06 | -9.5367e-07 | 0.97504 | 34.0197 | 5.5395e-05
2008-02#209 | 510 | 2009-01-28 15:52 | 24C | 0.96685 | 3.8603e-06 | -9.5367e-07 | 0.96685 | 33.6991 | 0.00015265
2008-02#210 | 511 | 2009-01-28 15:54 | 24C | 0.96078 | 4.2353e-06 | -9.5367e-07 | 0.96078 | 33.4618 | 0.00016512
2008-02#211 | 512 | 2009-01-28 15:57 | 24C | 0.95918 | 4.7525e-06 | -9.5367e-07 | 0.95918 | 33.3992 | 0.00018397
2008-02#212 | 513 | 2009-01-28 16:00 | 24C | 0.9574 | 1.3884e-06 | -9.5367e-07 | 0.9574 | 33.3296 | 5.5371e-05
2008-02#213 | 514 | 2009-01-28 16:03 | 24C | 0.95478 | 2.4912e-06 | -1.0729e-06 | 0.95478 | 33.2274 | 9.7029e-05
2008-02#214 | 515 | 2009-01-28 16:05 | 24C | 0.95259 | 3.6104e-06 | -9.5367e-07 | 0.95259 | 33.142 | 0.00014259
2008-02#215 | 516 | 2009-01-28 16:10 | 24C | 0.94925 | 3.2037e-06 | -9.5367e-07 | 0.94924 | 33.0114 | 0.00012554
2008-02#216 | 517 | 2009-01-28 16:12 | 24C | 0.94534 | 5.8378e-06 | -1.0133e-06 | 0.94534 | 32.8593 | 0.00022909
2008-02#217 | 518 | 2009-01-28 16:14 | 24C | 0.94186 | 3.1418e-06 | -9.5367e-07 | 0.94186 | 32.7235 | 0.00012313
2008-02#218 | 519 | 2009-01-28 16:17 | 24C | 0.9387 | 2.4911e-06 | -9.5367e-07 | 0.9387 | 32.6004 | 9.87e-05
2008-02#219 | 520 | 2009-01-28 16:19 | 24C | 0.93832 | 3.8647e-06 | -9.5367e-07 | 0.93832 | 32.5858 | 0.00015042
2008-02#220 | 520d | 2009-01-28 16:21 | 24C | 0.93833 | 2.7425e-06 | -1.0133e-06 | 0.93833 | 32.586 | 0.00010451
2008-02#221 | 521 | 2009-01-28 16:24 | 24C | 0.93733 | 2.7425e-06 | -8.9407e-07 | 0.93733 | 32.5472 | 0.00010658
2008-02#222 | 522 | 2009-01-28 16:26 | 24C | 0.93732 | 2.8901e-06 | -8.9407e-07 | 0.93732 | 32.5468 | 0.00011402
2008-02#223 | 523 | 2009-01-28 16:30 | 24C | 0.9837 | 4.4633e-06 | -9.5367e-07 | 0.9837 | 34.3592 | 0.0001736
2008-02#224 | 524 | 2009-01-28 16:34 | 24C | 0.98195 | 4.5979e-06 | -9.5367e-07 | 0.98195 | 34.2905 | 0.00018167
2008-02#225 | 524d | 2009-01-28 16:36 | 24C | 0.98194 | 3.5792e-06 | -9.5367e-07 | 0.98194 | 34.2902 | 0.00013756
2008-02#226 | 525 | 2009-01-28 16:39 | 24C | 0.97896 | 3.2011e-06 | -9.5367e-07 | 0.97896 | 34.1734 | 0.00012366
2008-02#227 | 526 | 2009-01-28 16:42 | 24C | 0.97373 | 2.4912e-06 | -9.5367e-07 | 0.97373 | 33.9684 | 9.7029e-05
2008-02#228 | 527 | 2009-01-28 16:45 | 24C | 0.96928 | 2.8904e-06 | -9.5367e-07 | 0.96928 | 33.794 | 0.00011147
2008-02#229 | 528 | 2009-01-28 16:48 | 24C | 0.9669 | 2.2394e-06 | -9.5367e-07 | 0.9669 | 33.7012 | 8.8738e-05
2008-02#230 | 529 | 2009-01-28 16:51 | 24C | 0.96402 | 1.4046e-06 | -8.3447e-07 | 0.96402 | 33.5882 | 5.4326e-05
2008-02#231 | 530 | 2009-01-28 16:53 | 24C | 0.96194 | 2.7426e-06 | -9.5367e-07 | 0.96194 | 33.5071 | 0.00010972
2008-02#232 | 531 | 2009-01-28 16:55 | 24C | 0.96006 | 3.359e-06 | -8.9407e-07 | 0.96006 | 33.4337 | 0.00013404
Calibration#3 | K15 | 2009-01-29 09:52 | 24C | 0.99984 | 2.2391e-06 | -9.5367e-07 | 0.99984 |  |  | calibration check right on!!
2008-02#233 | 532 | 2009-01-29 10:00 | 24C | 0.95722 | 2.2391e-06 | -8.9407e-07 | 0.95722 | 33.3225 | 8.792e-05
2008-02#234 | 533 | 2009-01-29 10:03 | 24C | 0.95599 | 2.9855e-06 | -9.5367e-07 | 0.95599 | 33.2745 | 0.00011492
2008-02#235 | 534 | 2009-01-29 10:07 | 24C | 0.95525 | 2.4909e-06 | -9.5367e-07 | 0.95525 | 33.2456 | 9.4885e-05
2008-02#236 | 535 | 2009-01-29 10:09 | 24C | 0.95291 | 3.9742e-06 | -9.5367e-07 | 0.95291 | 33.1544 | 0.00015323
2008-02#237 | 536 | 2009-01-29 10:12 | 24C | 0.94804 | 3.8622e-06 | -9.5367e-07 | 0.94804 | 32.9645 | 0.00015126
2008-02#238 | 537 | 2009-01-29 10:14 | 24C | 0.94162 | 4.2257e-06 | -1.0133e-06 | 0.94162 | 32.7141 | 0.00016355
2008-02#239 | 537d | 2009-01-29 10:16 | 24C | 0.94157 | 4.6353e-06 | -9.5367e-07 | 0.94157 | 32.7121 | 0.00017994
2008-02#240 | 538 | 2009-01-29 10:18 | 24C | 0.94158 | 3.8749e-06 | -1.0729e-06 | 0.94158 | 32.7125 | 0.00015239
2008-02#241 | 563 | 2009-01-29 10:21 | 24C | 0.99134 | 3.3439e-06 | -8.9407e-07 | 0.99134 | 34.6592 | 0.00012993
2008-02#242 | 563d | 2009-01-29 10:23 | 24C | 0.99132 | 3.1415e-06 | -9.5367e-07 | 0.99132 | 34.6584 | 0.00012104
2008-02#243 | 564 | 2009-01-29 10:25 | 24C | 0.9907 | 2.7426e-06 | -9.5367e-07 | 0.9907 | 34.6341 | 0.00010447
2008-02#244 | 565 | 2009-01-29 10:27 | 24C | 0.98967 | 4.0097e-06 | -9.5367e-07 | 0.98967 | 34.5938 | 0.00015604
2008-02#245 | 566 | 2009-01-29 10:29 | 24C | 0.98767 | 1.8038e-06 | -9.5367e-07 | 0.98766 | 34.515 | 6.8803e-05
2008-02#246 | 567 | 2009-01-29 10:32 | 24C | 0.98602 | 3.5987e-06 | -9.5367e-07 | 0.98602 | 34.4503 | 0.00014284
2008-02#247 | 568 | 2009-01-29 10:34 | 24C | 0.9839 | 1.3713e-06 | -9.5367e-07 | 0.9839 | 34.3673 | 5.4326e-05
2008-02#248 | 568d | 2009-01-29 10:36 | 24C | 0.98397 | 2.7425e-06 | -9.5367e-07 | 0.98397 | 34.3698 | 0.00010241
2008-02#249 | 569 | 2009-01-29 10:41 | 24C | 0.98188 | 1.3712e-06 | -8.3447e-07 | 0.98188 | 34.2881 | 5.5371e-05
2008-02#250 | 570 | 2009-01-29 10:43 | 24C | 0.97845 | 3.2141e-06 | -1.0729e-06 | 0.97845 | 34.1533 | 0.000126
2008-02#251 | 571 | 2009-01-29 10:45 | 24C | 0.9724 | 2.5369e-06 | -1.0133e-06 | 0.97239 | 33.9161 | 9.7279e-05
2008-02#252 | 572 | 2009-01-29 10:46 | 24C | 0.9679 | 3.3587e-06 | -9.5367e-07 | 0.9679 | 33.7403 | 0.0001272
2008-02#253 | 573 | 2009-01-29 10:48 | 24C | 0.96566 | 3.5795e-06 | -8.9407e-07 | 0.96566 | 33.6524 | 0.00013903
2008-02#254 | 574 | 2009-01-29 10:50 | 24C | 0.96221 | 1.3713e-06 | -1.0729e-06 | 0.96221 | 33.5176 | 5.5395e-05
2008-02#255 | 575 | 2009-01-29 10:52 | 24C | 0.95904 | 1.8038e-06 | -1.0133e-06 | 0.95903 | 33.3935 | 7.081e-05
2008-02#256 | 576 | 2009-01-29 10:54 | 24C | 0.95731 | 3.4517e-06 | -9.5367e-07 | 0.95731 | 33.3263 | 0.00013386
2008-02#257 | 577 | 2009-01-29 10:57 | 24C | 0.95573 | 1.6608e-06 | -9.5367e-07 | 0.95573 | 33.2645 | 6.649e-05
2008-02#258 | 578 | 2009-01-29 10:59 | 24C | 0.95466 | 2.4912e-06 | -1.0729e-06 | 0.95465 | 33.2225 | 9.7001e-05
2008-02#259 | 579 | 2009-01-29 11:07 | 24C | 0.95366 | 9.3824e-06 | -9.5367e-07 | 0.95365 | 33.1834 | 0.00036689 | started fickering -salt floating??
2008-02#260 | 580 | 2009-01-29 11:11 | 24C | 0.95247 | 2.7425e-06 | -8.9407e-07 | 0.95247 | 33.1372 | 0.00011074
2008-02#261 | 581 | 2009-01-29 11:12 | 24C | 0.94788 | 3.6088e-06 | -9.5367e-07 | 0.94788 | 32.9581 | 0.00013993
2008-02#262 | 582 | 2009-01-29 11:15 | 24C | 0.94754 | 2.4908e-06 | -1.0133e-06 | 0.94754 | 32.945 | 9.4885e-05
2008-02#263 | 583 | 2009-01-29 11:17 | 24C | 0.94767 | 3.3587e-06 | -1.0729e-06 | 0.94767 | 32.9501 | 0.00013151
2008-02#264 | 584 | 2009-01-29 11:19 | 24C | 0.98353 | 3.4657e-06 | -9.5367e-07 | 0.98353 | 34.3526 | 0.00013904
2008-02#265 | 585 | 2009-01-29 11:21 | 24C | 0.95555 | 2.7425e-06 | -8.9407e-07 | 0.95555 | 33.2573 | 0.00010451 | ?
2008-02#266 | 586 | 2009-01-29 11:24 | 24C | 0.95307 | 2.4908e-06 | -9.5367e-07 | 0.95307 | 33.1604 | 0.00010059
2008-02#267 | 586d | 2009-01-29 11:27 | 24C | 0.95283 | 2.7425e-06 | -9.5367e-07 | 0.95283 | 33.1511 | 0.00010765
2008-02#268 | 587 | 2009-01-29 11:29 | 24C | 0.95081 | 3.4658e-06 | -8.9407e-07 | 0.95081 | 33.0723 | 0.00013595
2008-02#269 | 588 | 2009-01-29 11:31 | 24C | 0.94919 | 2.0945e-06 | -8.9407e-07 | 0.94918 | 33.009 | 8.1621e-05
2008-02#270 | 589 | 2009-01-29 11:34 | 24C | 0.94622 | 2.4909e-06 | -9.5367e-07 | 0.94622 | 32.8934 | 9.6609e-05
2008-02#271 | 590 | 2009-01-29 11:36 | 24C | 0.94484 | 3.8748e-06 | -9.5367e-07 | 0.94484 | 32.8397 | 0.00014765
2008-02#272 | 591 | 2009-01-29 11:38 | 24C | 0.93922 | 2.7426e-06 | -8.9407e-07 | 0.93922 | 32.6209 | 0.00010972
2008-02#273 | 592 | 2009-01-29 11:40 | 24C | 0.93006 | 2.4613e-06 | -9.5367e-07 | 0.93006 | 32.2646 | 0.00010059
2008-02#274 | 593 | 2009-01-29 11:43 | 24C | 0.92794 | 4.9165e-06 | -9.5367e-07 | 0.92793 | 32.1819 | 0.00019288
2008-02#275 | 594 | 2009-01-29 11:45 | 24C | 0.92763 | 0 | -1.0729e-06 | 0.92762 | 32.1698 | 0
2008-02#276 | 595 | 2009-01-29 11:48 | 24C | 0.92763 | 2.7426e-06 | -9.5367e-07 | 0.92763 | 32.1699 | 0.00010447
2008-02#277 | 596 | 2009-01-29 11:50 | 24C | 0.92767 | 2.8897e-06 | -9.5367e-07 | 0.92767 | 32.1717 | 0.00010858
2008-02#278 | 601 | 2009-01-29 11:52 | 24C | 0.95153 | 0 | -9.5367e-07 | 0.95153 | 33.1005 | 0 | 4 flushes
2008-02#279 | 602 | 2009-01-29 11:54 | 24C | 0.9514 | 1.4046e-06 | -9.5367e-07 | 0.9514 | 33.0955 | 5.5371e-05
2008-02#280 | 603 | 2009-01-29 11:57 | 24C | 0.94977 | 2.4908e-06 | -9.5367e-07 | 0.94977 | 33.0319 | 9.4928e-05
2008-02#281 | 604 | 2009-01-29 11:59 | 24C | 0.94605 | 3.2141e-06 | -1.0133e-06 | 0.94605 | 32.8867 | 0.00012434
2008-02#282 | 605 | 2009-01-29 12:02 | 24C | 0.93979 | 2.89e-06 | -9.5367e-07 | 0.93979 | 32.6429 | 0.00011008
2008-02#283 | 606 | 2009-01-29 12:06 | 24C | 0.93597 | 1.3713e-06 | -8.3447e-07 | 0.93597 | 32.4942 | 5.2235e-05
2008-02#284 | 607 | 2009-01-29 12:08 | 24C | 0.93554 | 2.4909e-06 | -9.5367e-07 | 0.93554 | 32.4775 | 9.5571e-05
2008-02#285 | 607d | 2009-01-29 12:10 | 24C | 0.9355 | 3.2002e-06 | -8.9407e-07 | 0.9355 | 32.476 | 0.00012617
2008-02#286 | 608 | 2009-01-29 12:13 | 24C | 0.93013 | 3.1414e-06 | -8.9407e-07 | 0.93013 | 32.2672 | 0.00012108
2008-02#287 | 609 | 2009-01-29 12:15 | 24C | 0.92798 | 2.7426e-06 | -9.5367e-07 | 0.92798 | 32.1837 | 0.00010447
2008-02#288 | 610 | 2009-01-29 12:18 | 24C | 0.92794 | 2.4908e-06 | -8.3447e-07 | 0.92794 | 32.1819 | 9.8045e-05
2008-02#289 | 611 | 2009-01-29 13:13 | 24C | 0.92786 | 3.3591e-06 | -1.0133e-06 | 0.92786 | 32.1788 | 0.00013202
2008-02#290 | 612 | 2009-01-29 13:16 | 24C | 0.92795 | 2.4777e-06 | -9.5367e-07 | 0.92795 | 32.1823 | 9.7029e-05
2008-02#291 | 638 | 2009-01-29 13:20 | 24C | 0.94351 | 4.189e-06 | -9.5367e-07 | 0.94351 | 32.788 | 0.00016191 | 4 flushes
2008-02#292 | 639 | 2009-01-29 13:23 | 24C | 0.9437 | 2.7425e-06 | -8.3447e-07 | 0.9437 | 32.7954 | 0.00010656
2008-02#293 | 640 | 2009-01-29 13:25 | 24C | 0.94334 | 2.4908e-06 | -9.5367e-07 | 0.94334 | 32.7812 | 9.6599e-05
2008-02#294 | 641 | 2009-01-29 13:28 | 24C | 0.94233 | 2.7425e-06 | -9.5367e-07 | 0.94232 | 32.7417 | 0.00010656
2008-02#295 | 641d | 2009-01-29 13:30 | 24C | 0.94235 | 4.1723e-06 | -9.5367e-07 | 0.94235 | 32.7427 | 0.00016508
2008-02#296 | 642 | 2009-01-29 13:44 | 24C | 0.93956 | 2.7426e-06 | -9.5367e-07 | 0.93956 | 32.634 | 0.00010656
2008-02#297 | 643 | 2009-01-29 13:46 | 24C | 0.93948 | 1.7702e-06 | -9.5367e-07 | 0.93948 | 32.6309 | 6.8803e-05
2008-02#298 | 644 | 2009-01-29 13:49 | 24C | 0.9371 | 3.8651e-06 | -1.0729e-06 | 0.9371 | 32.5381 | 0.00014931
2008-02#299 | 645 | 2009-01-29 13:51 | 24C | 0.93044 | 1.4046e-06 | -9.5367e-07 | 0.93044 | 32.2793 | 5.2235e-05
2008-02#300 | 646 | 2009-01-29 14:04 | 24C | 0.9304 | 2.0947e-06 | -8.9407e-07 | 0.9304 | 32.2777 | 7.9819e-05
2008-02#301 | 647 | 2009-01-29 14:06 | 24C | 0.93042 | 3.3472e-06 | -9.5367e-07 | 0.93042 | 32.2784 | 0.00012994
2008-02#302 | 648 | 2009-01-29 14:09 | 24C | 0.93036 | 3.5406e-06 | -8.3447e-07 | 0.93036 | 32.276 | 0.00013759
2008-02#303 | 649 | 2009-01-29 14:12 | 24C | 0.93851 | 4.4783e-06 | -9.5367e-07 | 0.93851 | 32.5933 | 0.00017406 | 4 flushes
2008-02#304 | 650 | 2009-01-29 14:15 | 24C | 0.93858 | 3.2034e-06 | -9.5367e-07 | 0.93858 | 32.596 | 0.00012432
2008-02#305 | 651 | 2009-01-29 14:17 | 24C | 0.93465 | 2.4908e-06 | -8.3447e-07 | 0.93465 | 32.4429 | 9.6599e-05
2008-02#306 | 652 | 2009-01-29 14:19 | 24C | 0.93398 | 1.3712e-06 | -1.0133e-06 | 0.93398 | 32.4168 | 5.2278e-05
2008-02#307 | 652d | 2009-01-29 14:22 | 24C | 0.934 | 3.4659e-06 | -9.5367e-07 | 0.93399 | 32.4174 | 0.00013701
2008-02#308 | 653 | 2009-01-29 14:24 | 24C | 0.93307 | 3.974e-06 | -1.0133e-06 | 0.93307 | 32.3814 | 0.00015449
2008-02#309 | 654 | 2009-01-29 14:26 | 24C | 0.93073 | 3.2641e-06 | -9.5367e-07 | 0.93073 | 32.2904 | 0.00012611
2008-02#310 | 655 | 2009-01-29 14:28 | 24C | 0.91463 | 2.4909e-06 | -1.0133e-06 | 0.91463 | 31.6658 | 9.4885e-05
2008-02#311 | 656 | 2009-01-29 14:30 | 24C | 0.90543 | 4.4252e-06 | -9.5367e-07 | 0.90543 | 31.3095 | 0.000171
2008-02#312 | 657 | 2009-01-29 14:32 | 24C | 0.90543 | 2.8452e-06 | -9.5367e-07 | 0.90543 | 31.3094 | 0.00011094
2008-02#313 | 658 | 2009-01-29 14:34 | 24C | 0.9058 | 3.1414e-06 | -9.5367e-07 | 0.9058 | 31.3236 | 0.00011934
Calibration#4 | K15 | 2009-02-02 11:24 | 24C | 0.99984 | 4.4783e-06 | -1.967e-06 | 0.99984 |  |  | k value 0.99984 cal is bang on!!
2008-02#314 | 661 | 2009-02-02 11:27 | 24C | 0.92751 | 2.555e-06 | -2.0862e-06 | 0.92751 | 32.1653 | 9.9526e-05 | 4 flushes
2008-02#315 | 662 | 2009-02-02 11:30 | 24C | 0.92718 | 1.4046e-06 | -1.9073e-06 | 0.92717 | 32.1524 | 5.5371e-05
2008-02#316 | 663 | 2009-02-02 11:32 | 24C | 0.92703 | 4.3203e-06 | -1.967e-06 | 0.92702 | 32.1465 | 0.0001695
2008-02#317 | 664 | 2009-02-02 11:35 | 24C | 0.9261 | 4.4045e-06 | -1.9073e-06 | 0.9261 | 32.1105 | 0.00016962
2008-02#318 | 665 | 2009-02-02 11:40 | 24C | 0.92348 | 3.1999e-06 | -2.0266e-06 | 0.92348 | 32.0088 | 0.00012415
2008-02#319 | 666 | 2009-02-02 11:42 | 24C | 0.91912 | 4.3338e-06 | -1.967e-06 | 0.91912 | 31.8397 | 0.00016978
2008-02#320 | 666d | 2009-02-02 13:28 | 24C | 0.91959 | 3.3588e-06 | -1.967e-06 | 0.91959 | 31.858 | 0.0001325
2008-02#321 | 667 | 2009-02-02 13:30 | 24C | 0.91663 | 1.7934e-06 | -1.9073e-06 | 0.91663 | 31.7432 | 6.8177e-05
2008-02#322 | 668 | 2009-02-02 13:32 | 24C | 0.91453 | 5.5032e-06 | -2.0266e-06 | 0.91453 | 31.6617 | 0.00021326
2008-02#323 | 669 | 2009-02-02 13:35 | 24C | 0.90225 | 3.6105e-06 | -2.0266e-06 | 0.90225 | 31.1864 | 0.00014202
2008-02#324 | 670 | 2009-02-02 13:37 | 24C | 0.89895 | 1.1199e-06 | -1.967e-06 | 0.89895 | 31.059 | 4.2663e-05
2008-02#325 | 671 | 2009-02-02 13:40 | 24C | 0.89875 | 2.4908e-06 | -2.0266e-06 | 0.89875 | 31.0512 | 9.9639e-05
2008-02#326 | 674 | 2009-02-02 13:43 | 24C | 0.93358 | 2.7425e-06 | -1.9073e-06 | 0.93358 | 32.4014 | 0.00010765 | 4 flushes
2008-02#327 | 675 | 2009-02-02 13:46 | 24C | 0.9337 | 1.1196e-06 | -1.9073e-06 | 0.9337 | 32.4058 | 4.4374e-05
2008-02#328 | 676 | 2009-02-02 13:50 | 24C | 0.9332 | 3.8746e-06 | -2.0862e-06 | 0.9332 | 32.3865 | 0.00014951
2008-02#329 | 677 | 2009-02-02 13:52 | 24C | 0.93055 | 2.7425e-06 | -2.0266e-06 | 0.93055 | 32.2836 | 0.00010656
2008-02#330 | 677d | 2009-02-02 13:55 | 24C | 0.9306 | 3.9978e-06 | -1.9073e-06 | 0.93059 | 32.2852 | 0.00015672
2008-02#331 | 678 | 2009-02-02 13:59 | 24C | 0.92733 | 5.0715e-06 | -1.967e-06 | 0.92733 | 32.1585 | 0.00019989 | leaking
2008-02#332 | 679 | 2009-02-02 14:01 | 24C | 0.92613 | 2.7426e-06 | -1.967e-06 | 0.92613 | 32.1116 | 0.00010447
2008-02#333 | 680 | 2009-02-02 14:03 | 24C | 0.9258 | 2.4908e-06 | -2.0266e-06 | 0.92579 | 32.0987 | 9.8131e-05
2008-02#334 | 681 | 2009-02-02 14:06 | 24C | 0.91903 | 2.4908e-06 | -2.0266e-06 | 0.91902 | 31.836 | 9.4898e-05
2008-02#335 | 682 | 2009-02-02 14:08 | 24C | 0.91555 | 5.8496e-06 | -1.9073e-06 | 0.91555 | 31.7014 | 0.00022617
2008-02#336 | 683 | 2009-02-02 14:10 | 24C | 0.91545 | 4.8105e-06 | -1.9073e-06 | 0.91544 | 31.6972 | 0.00018537
2008-02#337 | 684 | 2009-02-02 14:13 | 24C | 0.9149 | 1.1199e-06 | -1.967e-06 | 0.9149 | 31.6761 | 4.2663e-05
2008-02#338 | 687 | 2009-02-02 14:15 | 24C | 0.93453 | 3.1417e-06 | -2.0266e-06 | 0.93453 | 32.4383 | 0.00011796 | 4 flushes
2008-02#339 | 688 | 2009-02-02 14:17 | 24C | 0.93459 | 3.2141e-06 | -1.9073e-06 | 0.93459 | 32.4406 | 0.00012777
2008-02#340 | 689 | 2009-02-02 14:19 | 24C | 0.93445 | 2.7261e-06 | -2.0266e-06 | 0.93444 | 32.4349 | 0.00010447
2008-02#341 | 690 | 2009-02-02 14:22 | 24C | 0.93122 | 2.7425e-06 | -2.0266e-06 | 0.93122 | 32.3095 | 0.00010451
2008-02#342 | 691 | 2009-02-02 14:23 | 24C | 0.92908 | 4.2257e-06 | -2.0266e-06 | 0.92908 | 32.2263 | 0.00016348
2008-02#343 | 691d | 2009-02-02 14:25 | 24C | 0.92902 | 3.1311e-06 | -1.9073e-06 | 0.92901 | 32.2239 | 0.00012314
2008-02#344 | 692 | 2009-02-02 14:27 | 24C | 0.92802 | 5.7072e-06 | -1.967e-06 | 0.92802 | 32.1852 | 0.00022072
2008-02#345 | 693 | 2009-02-02 14:30 | 24C | 0.92654 | 2.4921e-06 | -1.9073e-06 | 0.92653 | 32.1275 | 0.00010059
2008-02#346 | 694 | 2009-02-02 14:32 | 24C | 0.92468 | 4.3337e-06 | -1.967e-06 | 0.92468 | 32.0554 | 0.00016512
2008-02#347 | 695 | 2009-02-02 14:34 | 24C | 0.92156 | 2.8901e-06 | -1.9073e-06 | 0.92156 | 31.9345 | 0.00011104
2008-02#348 | 696 | 2009-02-02 14:37 | 24C | 0.91498 | 2.4911e-06 | -2.0266e-06 | 0.91498 | 31.6793 | 9.4898e-05
2008-02#349 | 697 | 2009-02-02 14:42 | 24C | 0.91415 | 3.6355e-06 | -2.0266e-06 | 0.91415 | 31.647 | 0.00014283
2008-02#350 | 698 | 2009-02-02 14:45 | 24C | 0.91401 | 3.4658e-06 | -1.9073e-06 | 0.91401 | 31.6416 | 0.0001317
2008-02#351 | 699 | 2009-02-02 14:50 | 24C | 0.93722 | 3.4659e-06 | -1.967e-06 | 0.93722 | 32.5429 | 0.00013596 | 4 flushes
2008-02#352 | 699d | 2009-02-02 14:53 | 24C | 0.93715 | 3.1417e-06 | -1.9073e-06 | 0.93715 | 32.5402 | 0.00012417
2008-02#353 | 700 | 2009-02-02 14:55 | 24C | 0.9371 | 2.4908e-06 | -2.0266e-06 | 0.9371 | 32.5382 | 9.7029e-05
2008-02#354 | 701 | 2009-02-02 14:58 | 24C | 0.93689 | 2.0947e-06 | -2.0266e-06 | 0.93689 | 32.5302 | 8.1169e-05
2008-02#355 | 702 | 2009-02-02 15:00 | 24C | 0.93247 | 3.9873e-06 | -2.0266e-06 | 0.93247 | 32.3581 | 0.00015466
2008-02#356 | 703 | 2009-02-02 15:41 | 24C | 0.92989 | 2.2398e-06 | -1.967e-06 | 0.92989 | 32.2578 | 8.7024e-05
2008-02#357 | 704 | 2009-02-02 15:42 | 24C | 0.92905 | 2.7097e-06 | -1.9073e-06 | 0.92905 | 32.2252 | 0.00010451
2008-02#358 | 705 | 2009-02-02 15:45 | 24C | 0.92821 | 1.4046e-06 | -1.9073e-06 | 0.92821 | 32.1926 | 5.2235e-05
2008-02#359 | 706 | 2009-02-02 15:47 | 24C | 0.92507 | 3.1417e-06 | -2.0266e-06 | 0.92507 | 32.0707 | 0.00012312
2008-02#360 | 707 | 2009-02-02 15:49 | 24C | 0.92095 | 2.4909e-06 | -2.0266e-06 | 0.92095 | 31.9108 | 9.7078e-05
2008-02#361 | 708 | 2009-02-02 15:51 | 24C | 0.91995 | 2.2391e-06 | -2.0862e-06 | 0.91995 | 31.872 | 8.4463e-05
2008-02#362 | 709 | 2009-02-02 15:53 | 24C | 0.91979 | 3.1867e-06 | -1.9073e-06 | 0.91979 | 31.8657 | 0.00012244
2008-02#363 | 710 | 2009-02-02 15:55 | 24C | 0.94274 | 3.4517e-06 | -2.0266e-06 | 0.94274 | 32.7578 | 0.00013392 | 4 flushes
2008-02#364 | 711 | 2009-02-02 15:57 | 24C | 0.9424 | 3.1533e-06 | -2.0266e-06 | 0.94239 | 32.7443 | 0.00012087
2008-02#365 | 712 | 2009-02-02 16:00 | 24C | 0.94077 | 4.3079e-06 | -2.0266e-06 | 0.94077 | 32.6811 | 0.00016842
2008-02#366 | 712d | 2009-02-02 16:02 | 24C | 0.94077 | 2.8897e-06 | -1.967e-06 | 0.94076 | 32.6809 | 0.00011145
2008-02#367 | 713 | 2009-02-02 16:04 | 24C | 0.93546 | 2.8346e-06 | -1.9073e-06 | 0.93546 | 32.4745 | 0.00010967
2008-02#368 | 714 | 2009-02-02 16:07 | 24C | 0.93426 | 2.7439e-06 | -1.967e-06 | 0.93426 | 32.4276 | 0.00010456
2008-02#369 | 715 | 2009-02-02 16:10 | 24C | 0.93354 | 2.2394e-06 | -1.967e-06 | 0.93354 | 32.3997 | 8.7013e-05
2008-02#370 | 716 | 2009-02-02 16:12 | 24C | 0.9319 | 3.4392e-06 | -2.0266e-06 | 0.9319 | 32.3361 | 0.00013386
2008-02#371 | 717 | 2009-02-02 16:14 | 24C | 0.92951 | 1.1203e-06 | -1.967e-06 | 0.92951 | 32.2432 | 4.4374e-05
2008-02#372 | 718 | 2009-02-02 16:16 | 24C | 0.92918 | 2.2267e-06 | -1.9073e-06 | 0.92918 | 32.2301 | 8.5299e-05
2008-02#373 | 719 | 2009-02-02 16:19 | 24C | 0.92289 | 1.1529e-06 | -2.0862e-06 | 0.92289 | 31.9861 | 4.2232e-05
2008-02#374 | 720 | 2009-02-02 16:21 | 24C | 0.92263 | 2.8897e-06 | -2.0266e-06 | 0.92263 | 31.9758 | 0.00011103
2008-02#375 | 743 | 2009-02-02 16:23 | 24C | 0.94166 | 2.4909e-06 | -1.967e-06 | 0.94166 | 32.7158 | 9.7029e-05 | 4 flushes
2008-02#376 | 744 | 2009-02-02 16:25 | 24C | 0.94173 | 2.7425e-06 | -1.9073e-06 | 0.94173 | 32.7184 | 0.00010447
2008-02#377 | 744d | 2009-02-02 16:27 | 24C | 0.94173 | 1.4046e-06 | -2.0266e-06 | 0.94172 | 32.7183 | 5.2235e-05
2008-02#378 | 745 | 2009-02-02 16:30 | 24C | 0.94172 | 2.4909e-06 | -2.0266e-06 | 0.94172 | 32.7182 | 9.6976e-05
2008-02#379 | 746 | 2009-02-02 16:33 | 24C | 0.94157 | 3.3588e-06 | -2.0266e-06 | 0.94157 | 32.7121 | 0.00013159
2008-02#380 | 747 | 2009-02-02 16:35 | 24C | 0.93278 | 2.2126e-06 | -1.967e-06 | 0.93278 | 32.3703 | 8.5299e-05
2008-02#381 | 748 | 2009-02-02 16:37 | 24C | 0.93071 | 2.4912e-06 | -1.9073e-06 | 0.93071 | 32.2896 | 9.4928e-05
2008-02#382 | 749 | 2009-02-02 16:39 | 24C | 0.92605 | 3.7045e-06 | -2.0862e-06 | 0.92604 | 32.1085 | 0.00014372
2008-02#383 | 750 | 2009-02-02 16:42 | 24C | 0.9249 | 2.1278e-06 | -1.9073e-06 | 0.9249 | 32.0639 | 8.1638e-05
2008-02#384 | 751 | 2009-02-02 16:44 | 24C | 0.92487 | 2.2391e-06 | -1.967e-06 | 0.92487 | 32.0628 | 8.7862e-05
2008-02#385 | 754 | 2009-02-02 16:47 | 24C | 0.94017 | 3.1535e-06 | -1.9073e-06 | 0.94017 | 32.6576 | 0.00012394 | 4 flushes
2008-02#386 | 754d | 2009-02-02 16:50 | 24C | 0.94008 | 2.0947e-06 | -2.0266e-06 | 0.94008 | 32.6542 | 8.1174e-05
2008-02#387 | 755 | 2009-02-02 16:53 | 24C | 0.94016 | 4.928e-06 | -1.9073e-06 | 0.94016 | 32.6574 | 0.0001922
2008-02#388 | 756 | 2009-02-02 16:55 | 24C | 0.94005 | 1.3712e-06 | -2.0862e-06 | 0.94005 | 32.653 | 5.1197e-05
2008-02#389 | 757 | 2009-02-02 16:57 | 24C | 0.94009 | 2.4908e-06 | -2.0266e-06 | 0.94008 | 32.6544 | 9.9759e-05
2008-02#390 | 758 | 2009-02-02 16:59 | 24C | 0.94014 | 3.2141e-06 | -2.0862e-06 | 0.94014 | 32.6565 | 0.00012781
2008-02#391 | 759 | 2009-02-02 17:01 | 24C | 0.9401 | 3.2144e-06 | -2.1458e-06 | 0.94009 | 32.6548 | 0.00012617
2008-02#392 | 760 | 2009-02-02 17:04 | 24C | 0.93968 | 3.8647e-06 | -1.967e-06 | 0.93968 | 32.6386 | 0.00015135
2008-02#393 | 761 | 2009-02-02 17:06 | 24C | 0.9385 | 2.2725e-06 | -2.0266e-06 | 0.9385 | 32.5928 | 8.7009e-05
2008-02#394 | 762 | 2009-02-02 17:09 | 24C | 0.94597 | 3.1414e-06 | -2.0266e-06 | 0.94597 | 32.8837 | 0.00012383 | 4 flushes
2008-02#395 | 770 | 2009-02-02 17:11 | 24C | 0.94302 | 2.7434e-06 | -1.967e-06 | 0.94301 | 32.7686 | 0.00011074
2008-02#396 | 771 | 2009-02-02 17:13 | 24C | 0.94359 | 5.6983e-06 | -2.0266e-06 | 0.94359 | 32.7908 | 0.00022219
2008-02#397 | 772 | 2009-02-02 17:16 | 24C | 0.94357 | 4.384e-06 | -1.9073e-06 | 0.94357 | 32.7901 | 0.00017179
2008-02#398 | 773 | 2009-02-02 17:18 | 24C | 0.94377 | 4.6249e-06 | -2.0862e-06 | 0.94377 | 32.798 | 0.00018209
2008-02#399 | 774 | 2009-02-02 17:20 | 24C | 0.94292 | 2.4908e-06 | -1.9073e-06 | 0.94291 | 32.7646 | 9.6414e-05
2008-02#400 | 775 | 2009-02-02 17:22 | 24C | 0.94251 | 1.1199e-06 | -2.0266e-06 | 0.94251 | 32.7487 | 4.265e-05
2008-02#401 | 776 | 2009-02-02 17:24 | 24C | 0.94194 | 4.189e-06 | -1.967e-06 | 0.94194 | 32.7266 | 0.0001635
2008-02#402 | 777 | 2009-02-02 17:26 | 24C | 0.94122 | 2.4909e-06 | -1.9073e-06 | 0.94122 | 32.6986 | 9.8715e-05
2008-02#403 | 778 | 2009-02-02 17:29 | 24C | 0.9389 | 4.189e-06 | -2.0266e-06 | 0.93889 | 32.6081 | 0.00016324
2008-02#404 | 778d | 2009-02-02 17:31 | 24C | 0.9387 | 3.9992e-06 | -1.967e-06 | 0.9387 | 32.6005 | 0.00015444
Calibration#5 | K15 | 2009-02-03 10:09 | 24C | 0.99984 | 4.4784e-06 | 1.9073e-06 | 0.99984 |  |  | cal check right on! K15=0.99984
2008-02#405 | 803 | 2009-02-03 10:12 | 24C | 0.93496 | 4.0094e-06 | 1.9073e-06 | 0.93496 | 32.455 | 0.00015782 | 4 flushes
2008-02#406 | 804 | 2009-02-03 10:16 | 24C | 0.93508 | 2.4908e-06 | 1.967e-06 | 0.93509 | 32.4599 | 9.6423e-05
2008-02#407 | 805 | 2009-02-03 10:18 | 24C | 0.93488 | 1.7705e-06 | 2.0266e-06 | 0.93488 | 32.4519 | 6.8797e-05
2008-02#408 | 805d | 2009-02-03 10:20 | 24C | 0.9348 | 4.3702e-06 | 2.0266e-06 | 0.9348 | 32.4488 | 0.00017032
2008-02#409 | 806 | 2009-02-03 10:22 | 24C | 0.93486 | 3.4908e-06 | 2.0266e-06 | 0.93486 | 32.451 | 0.00013678
2008-02#410 | 807 | 2009-02-03 10:25 | 24C | 0.93285 | 3.2145e-06 | 2.0266e-06 | 0.93285 | 32.3729 | 0.00012599
2008-02#411 | 808 | 2009-02-03 10:27 | 24C | 0.93219 | 2.7425e-06 | 2.0266e-06 | 0.9322 | 32.3475 | 0.00010656
2008-02#412 | 809 | 2009-02-03 10:43 | 24C | 0.93221 | 2.5747e-06 | 2.0862e-06 | 0.93221 | 32.3481 | 0.00010087
2008-02#413 | 810 | 2009-02-03 10:46 | 24C | 0.93219 | 3.9742e-06 | 1.967e-06 | 0.93219 | 32.3472 | 0.00015279
2008-02#414 | 811 | 2009-02-03 10:49 | 24C | 0.93314 | 3.8724e-06 | 2.0266e-06 | 0.93314 | 32.3843 | 0.00015151 | 4 flushes
2008-02#415 | 812 | 2009-02-03 10:52 | 24C | 0.93307 | 3.6105e-06 | 1.967e-06 | 0.93307 | 32.3815 | 0.00014259
2008-02#416 | 813 | 2009-02-03 10:55 | 24C | 0.93314 | 1.4046e-06 | 2.0266e-06 | 0.93314 | 32.3842 | 5.1206e-05
2008-02#417 | 814 | 2009-02-03 10:57 | 24C | 0.9329 | 3.359e-06 | 1.967e-06 | 0.93291 | 32.3751 | 0.00013159
2008-02#418 | 814d | 2009-02-03 10:59 | 24C | 0.93289 | 2.7426e-06 | 2.0266e-06 | 0.93289 | 32.3745 | 0.00010451
2008-02#419 | 815 | 2009-02-03 11:01 | 24C | 0.9329 | 3.8647e-06 | 1.967e-06 | 0.9329 | 32.3748 | 0.00015135
2008-02#420 | 816 | 2009-02-03 11:03 | 24C | 0.93289 | 3.2141e-06 | 1.967e-06 | 0.93289 | 32.3746 | 0.00012708
2008-02#421 | 817 | 2009-02-03 11:06 | 24C | 0.93018 | 1.1196e-06 | 2.0266e-06 | 0.93019 | 32.2694 | 4.2703e-05
2008-02#422 | 818 | 2009-02-03 11:08 | 24C | 0.92695 | 2.2392e-06 | 2.0266e-06 | 0.92696 | 32.1439 | 8.7239e-05
2008-02#423 | 819 | 2009-02-03 11:10 | 24C | 0.92693 | 2.7425e-06 | 2.1458e-06 | 0.92694 | 32.1432 | 0.00010766
2008-02#424 | 820 | 2009-02-03 11:13 | 24C | 0.92523 | 3.6104e-06 | 2.0266e-06 | 0.92523 | 32.0768 | 0.00013904
2008-02#425 | 821 | 2009-02-03 11:15 | 24C | 0.92557 | 2.2398e-06 | 2.0266e-06 | 0.92557 | 32.0902 | 8.7067e-05
2008-02#426 | 822 | 2009-02-03 11:18 | 24C | 0.92567 | 2.4785e-06 | 2.0266e-06 | 0.92567 | 32.094 | 9.4938e-05
2008-02#427 | 823 | 2009-02-03 11:20 | 24C | 0.92583 | 2.7426e-06 | 1.967e-06 | 0.92583 | 32.1003 | 0.00010865
2008-02#428 | 823d | 2009-02-03 11:23 | 24C | 0.92582 | 3.6105e-06 | 2.0266e-06 | 0.92582 | 32.0999 | 0.00013979
2008-02#429 | 824 | 2009-02-03 11:25 | 24C | 0.92583 | 2.4912e-06 | 1.9073e-06 | 0.92583 | 32.1003 | 9.87e-05
2008-02#430 | 825 | 2009-02-03 11:27 | 24C | 0.92516 | 1.1196e-06 | 2.0266e-06 | 0.92516 | 32.0743 | 4.2703e-05
2008-02#431 | 826 | 2009-02-03 11:30 | 24C | 0.92311 | 2.1433e-06 | 1.967e-06 | 0.92311 | 31.9945 | 8.4615e-05
2008-02#432 | 827 | 2009-02-03 11:33 | 24C | 0.92303 | 1.1196e-06 | 2.0266e-06 | 0.92303 | 31.9916 | 4.3729e-05
2008-02#433 | 828 | 2009-02-03 11:35 | 24C | 0.92146 | 2.8339e-06 | 2.0266e-06 | 0.92146 | 31.9305 | 0.00011087
2008-02#434 | 829 | 2009-02-03 11:39 | 24C | 0.91848 | 2.5638e-06 | 1.967e-06 | 0.91848 | 31.8149 | 9.7457e-05 | 4 flushs
2008-02#435 | 830 | 2009-02-03 11:42 | 24C | 0.91831 | 2.8897e-06 | 1.967e-06 | 0.91831 | 31.8084 | 0.00011361
2008-02#436 | 831 | 2009-02-03 11:45 | 24C | 0.91818 | 2.2398e-06 | 2.0862e-06 | 0.91818 | 31.8035 | 8.4887e-05
2008-02#437 | 831d | 2009-02-03 11:47 | 24C | 0.91814 | 1.1196e-06 | 1.9073e-06 | 0.91814 | 31.8017 | 4.2663e-05
2008-02#438 | 832 | 2009-02-03 11:50 | 24C | 0.91833 | 1.7702e-06 | 1.967e-06 | 0.91833 | 31.809 | 6.9858e-05
2008-02#439 | 833 | 2009-02-03 11:52 | 24C | 0.91819 | 3.6105e-06 | 2.0266e-06 | 0.91819 | 31.8038 | 0.00013942
2008-02#440 | 834 | 2009-02-03 11:54 | 24C | 0.91802 | 1.1199e-06 | 1.967e-06 | 0.91803 | 31.7973 | 4.3716e-05
2008-02#441 | 835 | 2009-02-03 11:57 | 24C | 0.91173 | 3.4658e-06 | 2.0862e-06 | 0.91173 | 31.5533 | 0.00013279
2008-02#442 | 836 | 2009-02-03 11:59 | 24C | 0.90157 | 3.4659e-06 | 2.0266e-06 | 0.90157 | 31.1603 | 0.0001366
2008-02#443 | 837 | 2009-02-03 12:01 | 24C | 0.8209 | 2.4908e-06 | 1.967e-06 | 0.8209 | 28.0647 | 9.3948e-05
2008-02#444 | 838 | 2009-02-03 12:04 | 24C | 0.89586 | 3.3591e-06 | 1.967e-06 | 0.89586 | 30.9396 | 0.00012754 | 4 flushes
2008-02#445 | 839 | 2009-02-03 12:07 | 24C | 0.895 | 4.1139e-06 | 1.967e-06 | 0.895 | 30.9064 | 0.00015619
2008-02#446 | 840 | 2009-02-03 12:09 | 24C | 0.89269 | 2.4908e-06 | 2.0266e-06 | 0.89269 | 30.8172 | 9.4364e-05
2008-02#447 | 840d | 2009-02-03 12:11 | 24C | 0.89248 | 1.1203e-06 | 1.967e-06 | 0.89248 | 30.809 | 4.543e-05
2008-02#448 | 841 | 2009-02-03 12:13 | 24C | 0.89094 | 1.1196e-06 | 1.967e-06 | 0.89094 | 30.7497 | 4.4782e-05
2008-02#449 | 842 | 2009-02-03 12:16 | 24C | 0.8801 | 2.7425e-06 | 2.0266e-06 | 0.8801 | 30.3318 | 0.00010186
2008-02#450 | 843 | 2009-02-03 12:18 | 24C | 0.86911 | 2.2401e-06 | 2.0266e-06 | 0.86912 | 29.9092 | 8.5313e-05
2008-02#451 | 844 | 2009-02-03 13:08 | 24C | 0.83946 | 3.4657e-06 | 1.967e-06 | 0.83947 | 28.773 | 0.00013412
2008-02#452 | 845 | 2009-02-03 13:10 | 24C | 0.80597 | 2.4909e-06 | 2.0266e-06 | 0.80597 | 27.4971 | 9.4885e-05
2008-02#453 | 846 | 2009-02-03 13:12 | 24C | 0.87254 | 3.4662e-06 | 2.1458e-06 | 0.87254 | 30.0408 | 0.0001343 | 4 flushes
2008-02#454 | 847 | 2009-02-03 13:14 | 24C | 0.87145 | 2.2391e-06 | 2.0266e-06 | 0.87145 | 29.9989 | 8.5326e-05
2008-02#455 | 848 | 2009-02-03 13:17 | 24C | 0.86746 | 2.7425e-06 | 2.0266e-06 | 0.86746 | 29.8455 | 0.00010447
2008-02#456 | 848d | 2009-02-03 13:19 | 24C | 0.86754 | 4.319e-06 | 1.967e-06 | 0.86755 | 29.8489 | 0.00016535
2008-02#457 | 849 | 2009-02-03 13:22 | 24C | 0.85747 | 4.1917e-06 | 2.0862e-06 | 0.85747 | 29.4623 | 0.00016042
2008-02#458 | 850 | 2009-02-03 13:25 | 24C | 0.85312 | 3.6081e-06 | 2.0862e-06 | 0.85312 | 29.2955 | 0.00013754
2008-02#459 | 851 | 2009-02-03 13:29 | 24C | 0.84844 | 2.1278e-06 | 2.0266e-06 | 0.84845 | 29.1164 | 8.04e-05
2008-02#460 | 852 | 2009-02-03 13:31 | 24C | 0.93116 | 1.7597e-06 | 2.0862e-06 | 0.93116 | 32.3073 | 6.7435e-05
2008-02#461 | 861 | 2009-02-03 13:34 | 24C | 0.93414 | 3.805e-06 | 2.0862e-06 | 0.93414 | 32.4232 | 0.00015005 | 4 flushes
2008-02#462 | 869 | 2009-02-03 13:38 | 24C | 0.935 | 2.4911e-06 | 2.0266e-06 | 0.935 | 32.4566 | 9.4885e-05 | 4 flushes
2008-02#463 | 870 | 2009-02-03 13:40 | 24C | 0.93501 | 4.4374e-06 | 2.0266e-06 | 0.93501 | 32.4569 | 0.00017298
2008-02#464 | 871 | 2009-02-03 13:42 | 24C | 0.93498 | 3.6104e-06 | 2.0266e-06 | 0.93498 | 32.4557 | 0.00014217
2008-02#465 | 872 | 2009-02-03 13:44 | 24C | 0.93501 | 2.4909e-06 | 2.0862e-06 | 0.93501 | 32.4569 | 0.00010061
2008-02#466 | 873 | 2009-02-03 13:47 | 24C | 0.93503 | 2.4912e-06 | 2.0266e-06 | 0.93503 | 32.4578 | 9.7001e-05
2008-02#467 | 874 | 2009-02-03 13:50 | 24C | 0.93502 | 3.579e-06 | 2.1458e-06 | 0.93502 | 32.4575 | 0.00013756
2008-02#468 | 875 | 2009-02-03 13:52 | 24C | 0.9342 | 3.865e-06 | 2.1458e-06 | 0.9342 | 32.4253 | 0.00015134
2008-02#469 | 876 | 2009-02-03 13:54 | 24C | 0.93346 | 2.7425e-06 | 2.1458e-06 | 0.93346 | 32.3966 | 0.00010656
2008-02#470 | 877 | 2009-02-03 13:57 | 24C | 0.93336 | 5.1342e-06 | 2.1458e-06 | 0.93336 | 32.3927 | 0.00019966
2008-02#471 | 894 | 2009-02-03 13:59 | 24C | 0.93914 | 2.2391e-06 | 2.0862e-06 | 0.93914 | 32.6176 | 8.5352e-05 | 4 flushes
2008-02#472 | 903 | 2009-02-03 14:01 | 24C | 0.93774 | 2.8901e-06 | 2.1458e-06 | 0.93774 | 32.5631 | 0.0001147 | 4 flushes
2008-02#473 | 912 | 2009-02-03 14:04 | 24C | 0.94893 | 1.1199e-06 | 2.0862e-06 | 0.94893 | 32.999 | 4.2703e-05 | 4 flushes
2008-02#474 | 913 | 2009-02-03 14:06 | 24C | 0.94784 | 1.3713e-06 | 2.1458e-06 | 0.94784 | 32.9566 | 5.2235e-05
2008-02#475 | 914 | 2009-02-03 14:08 | 24C | 0.94711 | 3.6105e-06 | 2.1458e-06 | 0.94711 | 32.9283 | 0.00013979
2008-02#476 | 915 | 2009-02-03 14:11 | 24C | 0.94658 | 2.9783e-06 | 2.0266e-06 | 0.94658 | 32.9075 | 0.00011345
2008-02#477 | 916 | 2009-02-03 14:15 | 24C | 0.94595 | 4.8105e-06 | 2.0862e-06 | 0.94595 | 32.8829 | 0.00018375
2008-02#478 | 917 | 2009-02-03 14:17 | 24C | 0.94582 | 3.8296e-06 | 2.1458e-06 | 0.94582 | 32.878 | 0.00014981
2008-02#479 | 918 | 2009-02-03 14:20 | 24C | 0.9446 | 2.89e-06 | 2.1458e-06 | 0.9446 | 32.8303 | 0.00011145
2008-02#480 | 919 | 2009-02-03 14:23 | 24C | 0.94091 | 2.407e-06 | 1.967e-06 | 0.94091 | 32.6866 | 9.3401e-05
2008-02#481 | 919d | 2009-02-03 14:25 | 24C | 0.94139 | 3.1309e-06 | 2.1458e-06 | 0.94139 | 32.7052 | 0.00012312 | not very close
2008-02#482 | 920 | 2009-02-03 14:28 | 24C | 0.93337 | 3.4665e-06 | 2.0266e-06 | 0.93337 | 32.3933 | 0.00013725
2008-02#483 | 921 | 2009-02-03 14:31 | 24C | 0.9324 | 2.7425e-06 | 2.1458e-06 | 0.9324 | 32.3555 | 0.00010657
2008-02#484 | 922 | 2009-02-03 14:33 | 24C | 0.91873 | 2.2259e-06 | 2.1458e-06 | 0.91874 | 31.8248 | 8.5299e-05 | 4 flushes
2008-02#485 | 923 | 2009-02-03 14:36 | 24C | 0.91878 | 2.4744e-06 | 2.0862e-06 | 0.91878 | 31.8266 | 9.7089e-05
2008-02#486 | 924 | 2009-02-03 14:38 | 24C | 0.91574 | 1.1532e-06 | 2.0266e-06 | 0.91574 | 31.7086 | 4.5848e-05
2008-02#487 | 925 | 2009-02-03 14:41 | 24C | 0.91424 | 1.3713e-06 | 2.1458e-06 | 0.91424 | 31.6505 | 5.2783e-05
2008-02#488 | 926 | 2009-02-03 14:44 | 24C | 0.91059 | 3.6105e-06 | 2.0862e-06 | 0.9106 | 31.5094 | 0.00013753
2008-02#489 | 927 | 2009-02-03 14:47 | 24C | 0.89999 | 1.3712e-06 | 2.1458e-06 | 0.89999 | 31.099 | 5.28e-05
2008-02#490 | 927d | 2009-02-03 14:49 | 24C | 0.90007 | 3.6105e-06 | 2.0266e-06 | 0.90007 | 31.1021 | 0.00013706
2008-02#491 | 928 | 2009-02-03 14:51 | 24C | 0.79006 | 1.8176e-06 | 2.1458e-06 | 0.79006 | 26.8937 | 7.0357e-05
2008-02#492 | 929 | 2009-02-03 14:54 | 24C | 0.93844 | 2.4615e-06 | 2.0862e-06 | 0.93844 | 32.5905 | 0.0001006 | 4 flushes
2008-02#493 | 930 | 2009-02-03 14:56 | 24C | 0.93844 | 2.7262e-06 | 2.1458e-06 | 0.93844 | 32.5905 | 0.00010663
2008-02#494 | 931 | 2009-02-03 14:59 | 24C | 0.93838 | 2.7425e-06 | 2.0266e-06 | 0.93838 | 32.5882 | 0.00010657
2008-02#495 | 932 | 2009-02-03 15:00 | 24C | 0.93799 | 1.1532e-06 | 2.0266e-06 | 0.93799 | 32.5729 | 4.4364e-05
2008-02#496 | 933 | 2009-02-03 15:03 | 24C | 0.93799 | 2.4909e-06 | 2.0266e-06 | 0.93799 | 32.5729 | 9.4938e-05
2008-02#497 | 934 | 2009-02-03 15:06 | 24C | 0.93783 | 1.6607e-06 | 2.0266e-06 | 0.93783 | 32.5667 | 6.6528e-05
2008-02#498 | 935 | 2009-02-03 15:08 | 24C | 0.93717 | 1.1199e-06 | 2.0266e-06 | 0.93717 | 32.541 | 4.4374e-05
2008-02#499 | 936 | 2009-02-03 15:11 | 24C | 0.93569 | 2.6729e-06 | 2.1458e-06 | 0.9357 | 32.4836 | 0.00010736
2008-02#500 | 937 | 2009-02-03 15:14 | 24C | 0.8858 | 9.148e-07 | 2.0862e-06 | 0.8858 | 30.5513 | 3.7639e-05
2008-02#501 | 938 | 2009-02-03 15:46 | 24C | 0.85923 | 5.0673e-06 | 2.0862e-06 | 0.85924 | 29.5299 | 0.00019478
2008-02#502 | 939 | 2009-02-03 15:50 | 24C | 0.92154 | 1.1536e-06 | 2.0266e-06 | 0.92154 | 31.9336 | 4.4795e-05
2008-02#503 | 940 | 2009-02-03 15:52 | 24C | 0.89888 | 1.4046e-06 | 2.0862e-06 | 0.89889 | 31.0564 | 5.3312e-05
2008-02#504 | 941 | 2009-02-03 15:54 | 24C | 0.88945 | 2.4909e-06 | 2.0266e-06 | 0.88945 | 30.692 | 9.4909e-05
2008-02#505 | 942 | 2009-02-03 15:55 | 24C | 0.78391 | 7.3337e-06 | 2.1458e-06 | 0.78391 | 26.6612 | 0.00027681
2008-02#506 | 943 | 2009-02-03 15:57 | 24C | 0.7671 | 0 | 2.0862e-06 | 0.7671 | 26.0265 | 0
2008-02#507 | 944 | 2009-02-03 15:59 | 24C | 0.76698 | 4.4373e-06 | 2.0862e-06 | 0.76698 | 26.0222 | 0.00016358
2008-02#508 | 945 | 2009-02-03 16:02 | 24C | 0.94765 | 3.1534e-06 | 2.1458e-06 | 0.94765 | 32.9491 | 0.0001229
2008-02#509 | 946 | 2009-02-03 16:05 | 24C | 0.94767 | 3.2142e-06 | 2.1458e-06 | 0.94768 | 32.9502 | 0.00012411
2008-02#510 | 947 | 2009-02-03 16:06 | 24C | 0.94757 | 3.1309e-06 | 2.0266e-06 | 0.94758 | 32.9463 | 0.00012313
2008-02#511 | 948 | 2009-02-03 16:08 | 24C | 0.94865 | 5.5968e-06 | 2.0266e-06 | 0.94866 | 32.9884 | 0.00021663
2008-02#512 | 949 | 2009-02-03 16:10 | 24C | 0.94762 | 3.4354e-06 | 2.0862e-06 | 0.94762 | 32.948 | 0.00013596
2008-02#513 | 950 | 2009-02-03 16:12 | 24C | 0.94368 | 4.37e-06 | 2.0266e-06 | 0.94368 | 32.7944 | 0.00017331
2008-02#514 | 950d | 2009-02-03 16:14 | 24C | 0.94375 | 2.2263e-06 | 2.0266e-06 | 0.94375 | 32.7972 | 8.5352e-05
2008-02#515 | 951 | 2009-02-03 16:16 | 24C | 0.93943 | 5.3086e-06 | 2.0862e-06 | 0.93944 | 32.6292 | 0.00020548
2008-02#516 | 952 | 2009-02-03 16:18 | 24C | 0.92017 | 1.1196e-06 | 2.1458e-06 | 0.92017 | 31.8805 | 4.265e-05
2008-02#517 | 953 | 2009-02-03 16:21 | 24C | 0.87577 | 8.1819e-06 | 2.0862e-06 | 0.87577 | 30.1652 | 0.00031559
2008-02#518 | 970 | 2009-02-03 16:23 | 24C | 0.94633 | 2.4908e-06 | 2.0266e-06 | 0.94633 | 32.8977 | 9.8725e-05 | 4 flushes
2008-02#519 | 971 | 2009-02-03 16:26 | 24C | 0.94734 | 8.0459e-06 | 2.0266e-06 | 0.94734 | 32.9372 | 0.0003138
2008-02#520 | 972 | 2009-02-03 16:30 | 24C | 0.948 | 4.8365e-06 | 2.0266e-06 | 0.948 | 32.9628 | 0.00018716
2008-02#521 | 973 | 2009-02-03 16:32 | 24C | 0.94717 | 1.1532e-06 | 2.0266e-06 | 0.94717 | 32.9306 | 4.4374e-05
2008-02#522 | 974 | 2009-02-03 16:34 | 24C | 0.94824 | 3.9739e-06 | 2.0862e-06 | 0.94825 | 32.9724 | 0.00015273
2008-02#523 | 975 | 2009-02-03 16:36 | 24C | 0.9475 | 4.1137e-06 | 2.0862e-06 | 0.9475 | 32.9435 | 0.00016377
2008-02#524 | 976 | 2009-02-03 16:38 | 24C | 0.94714 | 2.7425e-06 | 2.1458e-06 | 0.94714 | 32.9295 | 0.00010763
2008-02#525 | 977 | 2009-02-03 16:41 | 24C | 0.94763 | 1.4046e-06 | 2.0266e-06 | 0.94763 | 32.9483 | 5.5371e-05
2008-02#526 | 978 | 2009-02-03 16:43 | 24C | 0.94572 | 2.4909e-06 | 2.0862e-06 | 0.94572 | 32.874 | 9.7495e-05
2008-02#527 | 979 | 2009-02-03 16:45 | 24C | 0.91298 | 3.2145e-06 | 2.1458e-06 | 0.91298 | 31.6017 | 0.00012687
2008-02#528 | 979d | 2009-02-03 16:47 | 24C | 0.91336 | 3.466e-06 | 2.0266e-06 | 0.91336 | 31.6166 | 0.00013324
2008-02#529 | 980 | 2009-02-03 16:49 | 24C | 0.86444 | 3.1312e-06 | 2.0862e-06 | 0.86444 | 29.7297 | 0.00012399
2008-02#530 | 981 | 2009-02-03 16:51 | 24C | 0.85466 | 4.3195e-06 | 2.0266e-06 | 0.85466 | 29.3545 | 0.00016662
2008-02#531 | LSSL1313 | 2009-02-03 16:55 | 24C | 0.99912 | 4.8864e-06 | 2.1458e-06 | 0.99912 | 34.9655 | 0.00019364 | 4 flushes
2008-02#532 | LSSL1313-2 | 2009-02-03 16:58 | 24C | 0.99937 | 5.6583e-06 | 2.1458e-06 | 0.99937 | 34.9753 | 0.00022407
2008-02#533 | LSSL1313-3 | 2009-02-03 17:02 | 24C | 0.99935 | 4.1858e-06 | 2.0862e-06 | 0.99935 | 34.9745 | 0.00016506
2008-02#534 | LSSL1313-4 | 2009-02-03 17:04 | 24C | 0.99935 | 2.4909e-06 | 2.0862e-06 | 0.99935 | 34.9744 | 0.0001006

## #982-1257 RAW DATA2008-02
Sample ID | Bottle Label | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
Calibration#1 | K15 | 2009-02-04 09:11 | 24C | 0.99984 | 4.1682e-06 | -3.9935e-06 | 0.99984
SWL2008-02#1 | 982 | 2009-02-04 09:18 | 24C | 0.95354 | 3.4764e-06 | -3.9935e-06 | 0.95353 | 33.1786 | 0.000138 | 4 flushes
SWL2008-02#2 | 983 | 2009-02-04 09:20 | 24C | 0.95323 | 4.9385e-06 | -3.9935e-06 | 0.95322 | 33.1665 | 0.0001904
SWL2008-02#3 | 984 | 2009-02-04 09:23 | 24C | 0.9531 | 2.7593e-06 | -3.9935e-06 | 0.9531 | 33.1616 | 0.00010656
SWL2008-02#4 | 985 | 2009-02-04 09:26 | 24C | 0.95303 | 2.4912e-06 | -3.9339e-06 | 0.95303 | 33.1591 | 9.4885e-05
SWL2008-02#5 | 990 | 2009-02-04 09:39 | 24C | 0.94012 | 1.4046e-06 | -3.9339e-06 | 0.94011 | 32.6556 | 5.2235e-05
SWL2008-02#6 | 991 | 2009-02-04 09:42 | 24C | 0.92469 | 2.1278e-06 | -3.9339e-06 | 0.92469 | 32.0557 | 8.1866e-05
SWL2008-02#7 | 994 | 2009-02-04 09:44 | 24C | 0.91979 | 2.4775e-06 | -3.9339e-06 | 0.91979 | 31.8656 | 9.703e-05
SWL2008-02#8 | 997 | 2009-02-04 09:47 | 24C | 0.90737 | 1.3549e-06 | -3.9339e-06 | 0.90737 | 31.3845 | 5.2246e-05
SWL2008-02#9 | 1000 | 2009-02-04 09:50 | 24C | 0.87288 | 2.2393e-06 | -3.9935e-06 | 0.87288 | 30.0539 | 8.5037e-05
SWL2008-02#10 | 1003 | 2009-02-04 09:54 | 24C | 0.99614 | 2.4908e-06 | -3.8147e-06 | 0.99613 | 34.8478 | 9.9543e-05 | 4 flushes
SWL2008-02#11 | 1004 | 2009-02-04 09:57 | 24C | 0.99561 | 4.0093e-06 | -3.9935e-06 | 0.99561 | 34.8272 | 0.00015628
SWL2008-02#12 | 1005 | 2009-02-04 09:59 | 24C | 0.95531 | 3.1421e-06 | -3.8743e-06 | 0.9553 | 33.2478 | 0.00012108
SWL2008-02#13 | 1006 | 2009-02-04 10:03 | 24C | 0.95256 | 3.4906e-06 | -3.9935e-06 | 0.95255 | 33.1405 | 0.00013503
SWL2008-02#14 | 1006d | 2009-02-04 10:06 | 24C | 0.95247 | 3.1415e-06 | -3.8743e-06 | 0.95247 | 33.137 | 0.00012103
SWL2008-02#15 | 1007 | 2009-02-04 10:09 | 24C | 0.9494 | 3.1417e-06 | -3.9339e-06 | 0.9494 | 33.0173 | 0.00012103
SWL2008-02#16 | 1008 | 2009-02-04 10:11 | 24C | 0.9277 | 3.347e-06 | -3.9339e-06 | 0.9277 | 32.1726 | 0.00012795
SWL2008-02#17 | 1009 | 2009-02-04 10:14 | 24C | 0.84725 | 4.9165e-06 | -3.9339e-06 | 0.84724 | 29.0704 | 0.00018975
SWL2008-02#18 | 1010 | 2009-02-04 10:16 | 24C | 0.84565 | 3.2034e-06 | -3.9339e-06 | 0.84565 | 29.0095 | 0.00012212
SWL2008-02#19 | 1011 | 2009-02-04 10:19 | 24C | 0.99633 | 3.1309e-06 | -3.8743e-06 | 0.99633 | 34.8556 | 0.00012488 | 4 flushes
SWL2008-02#20 | 1012 | 2009-02-04 10:55 | 24C | 0.99625 | 3.4352e-06 | -3.9339e-06 | 0.99624 | 34.8522 | 0.0001341 | 4 flushes
SWL2008-02#21 | 1013 | 2009-02-04 10:58 | 24C | 0.99614 | 4.334e-06 | -3.8147e-06 | 0.99613 | 34.8479 | 0.00017313
SWL2008-02#22 | 1014 | 2009-02-04 11:00 | 24C | 0.99552 | 4.1859e-06 | -3.8743e-06 | 0.99552 | 34.8236 | 0.00016341
SWL2008-02#23 | 1015 | 2009-02-04 11:02 | 24C | 0.99402 | 3.5406e-06 | -3.9339e-06 | 0.99402 | 34.7647 | 0.00013828
SWL2008-02#24 | 1016 | 2009-02-04 11:05 | 24C | 0.98867 | 2.4916e-06 | -3.9339e-06 | 0.98866 | 34.5542 | 9.8121e-05
SWL2008-02#25 | 1017 | 2009-02-04 11:07 | 24C | 0.9554 | 2.2398e-06 | -3.9339e-06 | 0.95539 | 33.2513 | 8.959e-05
SWL2008-02#26 | 1018 | 2009-02-04 11:10 | 24C | 0.95283 | 2.4908e-06 | -3.9339e-06 | 0.95282 | 33.151 | 9.6599e-05
SWL2008-02#27 | 1019 | 2009-02-04 11:13 | 24C | 0.94576 | 3.0573e-06 | -3.8743e-06 | 0.94575 | 32.8752 | 0.00011723
SWL2008-02#28 | 1020 | 2009-02-04 11:16 | 24C | 0.91766 | 2.4915e-06 | -3.8743e-06 | 0.91766 | 31.783 | 9.6599e-05
SWL2008-02#29 | 1021 | 2009-02-04 11:18 | 24C | 0.88672 | 3.4657e-06 | -3.8743e-06 | 0.88671 | 30.5865 | 0.00013341
SWL2008-02#30 | 1022 | 2009-02-04 11:20 | 24C | 0.80883 | 1.4046e-06 | -3.9935e-06 | 0.80882 | 27.6055 | 5.2246e-05
SWL2008-02#31 | 1025 | 2009-02-04 11:23 | 24C | 0.93414 | 2.4912e-06 | -3.9339e-06 | 0.93413 | 32.4228 | 9.6976e-05
SWL2008-02#32 | 1026 | 2009-02-04 11:26 | 24C | 0.9351 | 3.8651e-06 | -3.8147e-06 | 0.93509 | 32.4602 | 0.00015066
SWL2008-02#33 | 1027 | 2009-02-04 11:27 | 24C | 0.93195 | 2.226e-06 | -3.9339e-06 | 0.93194 | 32.3377 | 8.5352e-05
SWL2008-02#34 | 1028 | 2009-02-04 11:31 | 24C | 0.92789 | 3.7931e-06 | -3.9339e-06 | 0.92789 | 32.18 | 0.00014614
SWL2008-02#35 | 1029 | 2009-02-04 11:36 | 24C | 0.91823 | 1.3713e-06 | -3.8147e-06 | 0.91823 | 31.805 | 5.1717e-05 | dup not even close
SWL2008-02#36 | 1029d | 2009-02-04 11:40 | 24C | 0.91885 | 3.2141e-06 | -3.9339e-06 | 0.91885 | 31.8292 | 0.0001263 | dup not even close
SWL2008-02#37 | 1030 | 2009-02-04 11:44 | 24C | 0.8814 | 3.8749e-06 | -3.9935e-06 | 0.88139 | 30.3815 | 0.00015214
SWL2008-02#38 | 1031 | 2009-02-04 11:47 | 24C | 0.81793 | 1.1532e-06 | -3.9339e-06 | 0.81793 | 27.9516 | 4.2663e-05
SWL2008-02#39 | 1056 | 2009-02-04 11:50 | 24C | 0.99719 | 2.7425e-06 | -3.9339e-06 | 0.99719 | 34.8893 | 0.00010972
SWL2008-02#40 | 1057 | 2009-02-04 11:52 | 24C | 0.99689 | 4.2256e-06 | -3.9339e-06 | 0.99688 | 34.8773 | 0.00016693
SWL2008-02#41 | 1058 | 2009-02-04 11:55 | 24C | 0.99647 | 2.4615e-06 | -3.9339e-06 | 0.99646 | 34.8609 | 0.00010061
SWL2008-02#42 | 1059 | 2009-02-04 11:57 | 24C | 0.99574 | 3.1418e-06 | -3.9339e-06 | 0.99574 | 34.8324 | 0.00012381
SWL2008-02#43 | 1059d | 2009-02-04 11:59 | 24C | 0.99566 | 1.3713e-06 | -3.8743e-06 | 0.99565 | 34.829 | 5.4326e-05
SWL2008-02#44 | 1060 | 2009-02-04 12:02 | 24C | 0.99472 | 5.0342e-06 | -3.9339e-06 | 0.99472 | 34.7922 | 0.00019882
SWL2008-02#45 | 1061 | 2009-02-04 12:05 | 24C | 0.99021 | 3.4658e-06 | -3.8743e-06 | 0.99021 | 34.6149 | 0.00013929
SWL2008-02#46 | 1062 | 2009-02-04 12:07 | 24C | 0.98611 | 4.9398e-06 | -4.0531e-06 | 0.9861 | 34.4536 | 0.00019155
SWL2008-02#47 | 1063 | 2009-02-04 12:10 | 24C | 0.97389 | 3.4657e-06 | -3.9339e-06 | 0.97389 | 33.9747 | 0.00013655
SWL2008-02#48 | 1064 | 2009-02-04 12:12 | 24C | 0.96245 | 3.3455e-06 | -3.9339e-06 | 0.96245 | 33.527 | 0.00012993
SWL2008-02#49 | 1065 | 2009-02-04 12:15 | 24C | 0.93365 | 2.8904e-06 | -3.9339e-06 | 0.93364 | 32.4038 | 0.00011145
SWL2008-02#50 | 1066 | 2009-02-04 12:17 | 24C | 0.92441 | 1.3722e-06 | -3.8147e-06 | 0.9244 | 32.0448 | 5.2235e-05
SWL2008-02#51 | 1067 | 2009-02-04 12:20 | 24C | 0.90255 | 2.7425e-06 | -3.9339e-06 | 0.90255 | 31.1979 | 0.0001071
SWL2008-02#52 | 1068 | 2009-02-04 12:22 | 24C | 0.89205 | 2.7425e-06 | -3.9935e-06 | 0.89204 | 30.7922 | 0.00010448
SWL2008-02#53 | 1069 | 2009-02-04 13:05 | 24C | 0.87536 | 3.8747e-06 | -4.0531e-06 | 0.87535 | 30.149 | 0.00014661
SWL2008-02#54 | 1070 | 2009-02-04 13:08 | 24C | 0.86687 | 3.6232e-06 | -4.0531e-06 | 0.86686 | 29.8227 | 0.00014042
SWL2008-02#55 | 1070 | 2009-02-04 13:10 | 24C | 0.86706 | 1.4046e-06 | -3.9339e-06 | 0.86705 | 29.83 | 5.4847e-05
SWL2008-02#56 | 1071 | 2009-02-04 13:13 | 24C | 0.85439 | 2.2271e-06 | -3.8743e-06 | 0.85438 | 29.3439 | 8.5326e-05
SWL2008-02#57 | 1073 | 2009-02-04 13:15 | 24C | 0.99718 | 2.2394e-06 | -3.8743e-06 | 0.99718 | 34.889 | 8.3674e-05 | 4 flushes
SWL2008-02#58 | 1074 | 2009-02-04 13:18 | 24C | 0.99665 | 3.2037e-06 | -3.9339e-06 | 0.99664 | 34.868 | 0.00012452
SWL2008-02#59 | 1075 | 2009-02-04 13:20 | 24C | 0.99609 | 3.1418e-06 | -3.9339e-06 | 0.99608 | 34.8459 | 0.00012618
SWL2008-02#60 | 1075d | 2009-02-04 13:22 | 24C | 0.99625 | 5.8633e-06 | -3.9339e-06 | 0.99624 | 34.8522 | 0.00023109
SWL2008-02#61 | 1076 | 2009-02-04 13:25 | 24C | 0.99556 | 3.0571e-06 | -3.8147e-06 | 0.99556 | 34.8252 | 0.00012128
SWL2008-02#62 | 1077 | 2009-02-04 13:28 | 24C | 0.99387 | 5.3055e-06 | -3.8743e-06 | 0.99387 | 34.7589 | 0.00020764 | caps with cone inserts 1077 to 1137
SWL2008-02#63 | 1078 | 2009-02-04 13:31 | 24C | 0.98386 | 3.466e-06 | -3.9339e-06 | 0.98386 | 34.3656 | 0.00013207
SWL2008-02#64 | 1079 | 2009-02-04 13:33 | 24C | 0.97692 | 9.1415e-07 | -3.9339e-06 | 0.97691 | 34.0931 | 3.4823e-05
SWL2008-02#65 | 1080 | 2009-02-04 13:36 | 24C | 0.96315 | 3.6104e-06 | -3.9339e-06 | 0.96314 | 33.554 | 0.00014134
SWL2008-02#66 | 1081 | 2009-02-04 13:38 | 24C | 0.95398 | 1.3882e-06 | -3.9935e-06 | 0.95398 | 33.196 | 5.1197e-05
SWL2008-02#67 | 1082 | 2009-02-04 13:40 | 24C | 0.94281 | 3.332e-08 | -3.8147e-06 | 0.94281 | 32.7605 | 0
SWL2008-02#68 | 1083 | 2009-02-04 13:42 | 24C | 0.93342 | 4.1137e-06 | -3.8743e-06 | 0.93342 | 32.3951 | 0.00016273
SWL2008-02#69 | 1083d | 2009-02-04 13:44 | 24C | 0.93335 | 1.3712e-06 | -3.9339e-06 | 0.93334 | 32.3921 | 5.2235e-05
SWL2008-02#70 | 1084 | 2009-02-04 13:47 | 24C | 0.91173 | 1.1529e-06 | -3.8147e-06 | 0.91172 | 31.553 | 4.2663e-05
SWL2008-02#71 | 1085 | 2009-02-04 13:49 | 24C | 0.89692 | 3.131e-06 | -3.9339e-06 | 0.89691 | 30.9802 | 0.00012195
SWL2008-02#72 | 1086 | 2009-02-04 13:51 | 24C | 0.87952 | 3.5635e-06 | -3.8743e-06 | 0.87952 | 30.3094 | 0.00013676
SWL2008-02#73 | 1087 | 2009-02-04 13:54 | 24C | 0.85296 | 3.2148e-06 | -3.9339e-06 | 0.85296 | 29.2892 | 0.000122
SWL2008-02#74 | 1088 | 2009-02-04 13:57 | 24C | 0.8078 | 2.7425e-06 | -3.8147e-06 | 0.8078 | 27.5666 | 0.00010448
SWL2008-02#75 | 1089 | 2009-02-04 14:00 | 24C | 0.77543 | 2.8792e-06 | -3.9339e-06 | 0.77543 | 26.3406 | 0.00010768
SWL2008-02#76 | 1090 | 2009-02-04 14:03 | 24C | 0.99176 | 1.8176e-06 | -3.9935e-06 | 0.99176 | 34.6758 | 7.2435e-05
SWL2008-02#77 | 1091 | 2009-02-04 14:06 | 24C | 0.98307 | 6.2734e-06 | -3.9935e-06 | 0.98307 | 34.3346 | 0.00024657
SWL2008-02#78 | 1092 | 2009-02-04 14:11 | 24C | 0.96782 | 4.6251e-06 | -3.9935e-06 | 0.96781 | 33.7368 | 0.00018004
SWL2008-02#79 | 1093 | 2009-02-04 14:14 | 24C | 0.95395 | 3.8746e-06 | -3.8743e-06 | 0.95395 | 33.195 | 0.00014956
SWL2008-02#80 | 1094 | 2009-02-04 14:16 | 24C | 0.94626 | 3.6232e-06 | -3.9339e-06 | 0.94625 | 32.8947 | 0.00013973
SWL2008-02#81 | 1095 | 2009-02-04 14:19 | 24C | 0.94115 | 4.0096e-06 | -3.9339e-06 | 0.94114 | 32.6956 | 0.00015604
SWL2008-02#82 | 1096 | 2009-02-04 14:22 | 24C | 0.93075 | 2.0945e-06 | -3.9339e-06 | 0.93075 | 32.2912 | 8.1169e-05
SWL2008-02#83 | 1097 | 2009-02-04 14:24 | 24C | 0.91874 | 2.7425e-06 | -4.0531e-06 | 0.91873 | 31.8247 | 0.00010969
SWL2008-02#84 | 1098 | 2009-02-04 14:26 | 24C | 0.9063 | 3.8746e-06 | -3.9339e-06 | 0.90629 | 31.3428 | 0.00015213
SWL2008-02#85 | 1099 | 2009-02-04 14:29 | 24C | 0.89342 | 2.7262e-06 | -3.9339e-06 | 0.89341 | 30.845 | 0.00010395
SWL2008-02#86 | 1100 | 2009-02-04 14:31 | 24C | 0.86963 | 2.2263e-06 | -4.0531e-06 | 0.86962 | 29.9287 | 8.7027e-05
SWL2008-02#87 | 1100d | 2009-02-04 14:34 | 24C | 0.86967 | 2.4911e-06 | -3.9935e-06 | 0.86967 | 29.9305 | 9.3938e-05
SWL2008-02#88 | 1101 | 2009-02-04 14:36 | 24C | 0.80977 | 3.3596e-06 | -3.9339e-06 | 0.80977 | 27.6412 | 0.00012716
SWL2008-02#89 | 1102 | 2009-02-04 14:40 | 24C | 0.71514 | 3.2144e-06 | -3.9339e-06 | 0.71514 | 24.0785 | 0.00011998
SWL2008-02#90 | 1127 | 2009-02-04 14:46 | 24C | 0.99301 | 3.8747e-06 | -3.9339e-06 | 0.993 | 34.7247 | 0.00015447
SWL2008-02#91 | 1128 | 2009-02-04 14:48 | 24C | 0.96329 | 3.4352e-06 | -3.8743e-06 | 0.96328 | 33.5596 | 0.00013207
SWL2008-02#92 | 1129 | 2009-02-04 14:51 | 24C | 0.95558 | 1.1199e-06 | -3.9339e-06 | 0.95558 | 33.2584 | 4.4374e-05
SWL2008-02#93 | 1130 | 2009-02-04 14:53 | 24C | 0.95193 | 4.2256e-06 | -3.9339e-06 | 0.95193 | 33.116 | 0.00016204
SWL2008-02#94 | 1131 | 2009-02-04 14:56 | 24C | 0.95969 | 5.1463e-06 | -3.9339e-06 | 0.95969 | 33.419 | 0.00020049
SWL2008-02#95 | 1132 | 2009-02-04 14:58 | 24C | 0.94301 | 4.6496e-06 | -3.8743e-06 | 0.94301 | 32.7682 | 0.00018212
SWL2008-02#96 | 1133 | 2009-02-04 15:00 | 24C | 0.93117 | 3.1415e-06 | -3.8147e-06 | 0.93117 | 32.3076 | 0.00012306
SWL2008-02#97 | 1134 | 2009-02-04 15:02 | 24C | 0.91592 | 2.4908e-06 | -3.9339e-06 | 0.91591 | 31.7154 | 9.6499e-05
SWL2008-02#98 | 1135 | 2009-02-04 15:05 | 24C | 0.9068 | 3.8598e-06 | -3.9339e-06 | 0.90679 | 31.3621 | 0.00014947
SWL2008-02#99 | 1136 | 2009-02-04 15:07 | 24C | 0.89548 | 1.8035e-06 | -3.9935e-06 | 0.89548 | 30.9248 | 6.9858e-05
SWL2008-02#100 | 1137 | 2009-02-04 15:10 | 24C | 0.87938 | 6.7868e-06 | -3.8743e-06 | 0.87938 | 30.304 | 0.00026235 | not even close
SWL2008-02#101 | 1137d | 2009-02-04 15:12 | 24C | 0.88133 | 7.3903e-06 | -3.8743e-06 | 0.88132 | 30.3789 | 0.00028552 | not even close
SWL2008-02#102 | 1138 | 2009-02-04 16:05 | 24C | 0.82724 | 2.4776e-06 | -3.9339e-06 | 0.82723 | 28.3061 | 9.4459e-05
SWL2008-02#103 | 1139 | 2009-02-04 16:07 | 24C | 0.69644 | 2.8794e-06 | -3.9339e-06 | 0.69643 | 23.382 | 0.00010604
SWL2008-02#104 | 1140 | 2009-02-04 16:11 | 24C | 0.99726 | 3.84e-06 | -3.9935e-06 | 0.99726 | 34.892 | 0.00014997
SWL2008-02#105 | 1141 | 2009-02-04 16:14 | 24C | 0.99664 | 2.4777e-06 | -3.9339e-06 | 0.99664 | 34.8678 | 9.7194e-05
SWL2008-02#106 | 1142 | 2009-02-04 16:17 | 24C | 0.99602 | 3.8749e-06 | -3.9339e-06 | 0.99601 | 34.8431 | 0.00015554
SWL2008-02#107 | 1142d | 2009-02-04 16:20 | 24C | 0.99605 | 2.2401e-06 | -3.8743e-06 | 0.99604 | 34.8442 | 8.958e-05
SWL2008-02#108 | 1143 | 2009-02-04 16:22 | 24C | 0.99552 | 4.1892e-06 | -3.8147e-06 | 0.99551 | 34.8235 | 0.00016688
SWL2008-02#109 | 1144 | 2009-02-04 16:24 | 24C | 0.99401 | 3.1867e-06 | -3.9339e-06 | 0.994 | 34.7641 | 0.00012279
SWL2008-02#110 | 1145 | 2009-02-04 16:27 | 24C | 0.98552 | 3.1421e-06 | -3.9935e-06 | 0.98552 | 34.4307 | 0.000119
SWL2008-02#111 | 1146 | 2009-02-04 16:29 | 24C | 0.7905 | 2.4908e-06 | -3.9339e-06 | 0.79049 | 26.9102 | 9.4379e-05
SWL2008-02#112 | 1147 | 2009-02-04 16:32 | 24C | 0.97291 | 1.3712e-06 | -3.9339e-06 | 0.97291 | 33.9363 | 5.5371e-05
SWL2008-02#113 | 1148 | 2009-02-04 16:34 | 24C | 0.95827 | 2.4908e-06 | -3.9339e-06 | 0.95826 | 33.3634 | 9.9734e-05
SWL2008-02#114 | 1149 | 2009-02-04 16:36 | 24C | 0.94862 | 2.7425e-06 | -3.9339e-06 | 0.94861 | 32.9868 | 0.00010451
SWL2008-02#115 | 1150 | 2009-02-04 16:39 | 24C | 0.93783 | 3.4658e-06 | -3.9935e-06 | 0.93783 | 32.5666 | 0.00013553
SWL2008-02#116 | 1151 | 2009-02-04 16:43 | 24C | 0.92583 | 3.2144e-06 | -3.9935e-06 | 0.92583 | 32.1 | 0.00012382
SWL2008-02#117 | 1152 | 2009-02-04 16:45 | 24C | 0.91891 | 3.5971e-06 | -3.9339e-06 | 0.9189 | 31.8313 | 0.00013753
SWL2008-02#118 | 1153 | 2009-02-04 16:48 | 24C | 0.9127 | 1.4046e-06 | -3.9339e-06 | 0.9127 | 31.5909 | 5.2783e-05
SWL2008-02#119 | 1154 | 2009-02-04 16:51 | 24C | 0.89068 | 2.4912e-06 | -3.9339e-06 | 0.89068 | 30.7396 | 9.6591e-05
SWL2008-02#120 | 1155 | 2009-02-04 16:53 | 24C | 0.81674 | 2.4916e-06 | -3.9935e-06 | 0.81673 | 27.9061 | 9.4909e-05
SWL2008-02#121 | 1156 | 2009-02-04 16:56 | 24C | 0.79095 | 5.4556e-06 | -3.8147e-06 | 0.79095 | 26.9274 | 0.00020537
SWL2008-02#122 | 1182 | 2009-02-04 16:59 | 24C | 0.96709 | 3.1415e-06 | -3.8743e-06 | 0.96708 | 33.7083 | 0.00012001
SWL2008-02#123 | 1183 | 2009-02-04 17:02 | 24C | 0.96742 | 3.7966e-06 | -3.9935e-06 | 0.96742 | 33.7213 | 0.00014791
SWL2008-02#124 | 1184 | 2009-02-04 17:05 | 24C | 0.95847 | 2.5034e-06 | -3.9339e-06 | 0.95847 | 33.3713 | 9.7279e-05
SWL2008-02#125 | 1185 | 2009-02-04 17:08 | 24C | 0.9509 | 4.2257e-06 | -3.9339e-06 | 0.9509 | 33.0758 | 0.00016204
SWL2008-02#126 | 1186 | 2009-02-04 17:10 | 24C | 0.94381 | 4.3836e-06 | -3.9935e-06 | 0.94381 | 32.7994 | 0.00017382
SWL2008-02#127 | 1187 | 2009-02-04 17:13 | 24C | 0.93279 | 2.0945e-06 | -3.8743e-06 | 0.93279 | 32.3706 | 7.8884e-05
SWL2008-02#132 | 1187 | 2009-02-04 17:46 | 24C | 0.93269 | 2.8774e-06 | -3.9935e-06 | 0.93268 | 32.3665 | 0.00011145
SWL2008-02#133 | 1188 | 2009-02-04 17:48 | 24C | 0.92207 | 3.8459e-06 | -3.9339e-06 | 0.92206 | 31.9538 | 0.00015134
SWL2008-02#134 | 1189 | 2009-02-04 17:51 | 24C | 0.91601 | 1.3713e-06 | -3.9339e-06 | 0.91601 | 31.7191 | 5.3301e-05
SWL2008-02#135 | 1190 | 2009-02-04 17:53 | 24C | 0.90538 | 3.466e-06 | -3.9339e-06 | 0.90538 | 31.3075 | 0.00013592
SWL2008-02#136 | 1191 | 2009-02-05 09:24 | 24C | 0.88968 | 2.2401e-06 | -3.8743e-06 | 0.88967 | 30.7007 | 8.7432e-05
SWL2008-02#137 | 1192 | 2009-02-05 09:27 | 24C | 0.85457 | 3.8437e-06 | -3.9935e-06 | 0.85456 | 29.3508 | 0.00015
SWL2008-02#138 | 1193 | 2009-02-05 09:29 | 24C | 0.85064 | 1.1529e-06 | -3.8147e-06 | 0.85063 | 29.2001 | 4.329e-05
SWL2008-02#139 | 1194 | 2009-02-05 09:32 | 24C | 0.99492 | 3.1417e-06 | -3.8743e-06 | 0.99492 | 34.8001 | 0.00012306
SWL2008-02#140 | 1195 | 2009-02-05 09:35 | 24C | 0.99457 | 2.4784e-06 | -3.9339e-06 | 0.99457 | 34.7864 | 9.9543e-05
SWL2008-02#141 | 1196 | 2009-02-05 09:37 | 24C | 0.99409 | 3.2034e-06 | -3.9339e-06 | 0.99409 | 34.7675 | 0.00012709
SWL2008-02#142 | 1197 | 2009-02-05 09:40 | 24C | 0.99299 | 4.716e-06 | -3.8147e-06 | 0.99299 | 34.7243 | 0.00018643
SWL2008-02#143 | 1198 | 2009-02-05 09:42 | 24C | 0.98981 | 1.1196e-06 | -3.9339e-06 | 0.98981 | 34.5992 | 4.5227e-05
SWL2008-02#144 | 1199 | 2009-02-05 09:44 | 24C | 0.98792 | 3.466e-06 | -3.9339e-06 | 0.98792 | 34.5248 | 0.00013908
SWL2008-02#145 | 1200 | 2009-02-05 09:46 | 24C | 0.98496 | 2.5219e-06 | -3.9339e-06 | 0.98496 | 34.4089 | 9.7279e-05
SWL2008-02#146 | 1200d | 2009-02-05 09:49 | 24C | 0.985 | 4.1892e-06 | -3.8743e-06 | 0.98499 | 34.41 | 0.000158
SWL2008-02#147 | 1201 | 2009-02-05 09:52 | 24C | 0.97885 | 1.1199e-06 | -3.8147e-06 | 0.97885 | 34.169 | 4.5217e-05
SWL2008-02#148 | 1202 | 2009-02-05 09:56 | 24C | 0.97248 | 2.4908e-06 | -3.9339e-06 | 0.97247 | 33.9192 | 9.7029e-05
SWL2008-02#149 | 1203 | 2009-02-05 09:58 | 24C | 0.95944 | 5.3589e-06 | -3.9339e-06 | 0.95944 | 33.4093 | 0.00020821
SWL2008-02#150 | 1204 | 2009-02-05 10:01 | 24C | 0.94155 | 3.5298e-06 | -3.9339e-06 | 0.94155 | 32.7113 | 0.00013829
SWL2008-02#151 | 1205 | 2009-02-05 10:05 | 24C | 0.93716 | 2.4915e-06 | -3.8743e-06 | 0.93715 | 32.5404 | 9.4885e-05
SWL2008-02#152 | 1206 | 2009-02-05 10:07 | 24C | 0.91767 | 3.2144e-06 | -3.9339e-06 | 0.91767 | 31.7835 | 0.00012687
SWL2008-02#153 | 1207 | 2009-02-05 10:42 | 24C | 0.91365 | 3.579e-06 | -3.9339e-06 | 0.91364 | 31.6274 | 0.00013654
SWL2008-02#154 | 1208 | 2009-02-05 10:44 | 24C | 0.88527 | 3.6239e-06 | -3.8147e-06 | 0.88526 | 30.5306 | 0.00013995
SWL2008-02#155 | 1209 | 2009-02-05 10:47 | 24C | 0.88392 | 4.1786e-06 | -3.8743e-06 | 0.88392 | 30.4788 | 0.0001636
SWL2008-02#156 | 1234 | 2009-02-05 10:52 | 24C | 0.99711 | 1.1203e-06 | -3.9339e-06 | 0.99711 | 34.8861 | 4.5217e-05
SWL2008-02#157 | 1235 | 2009-02-05 10:54 | 24C | 0.99652 | 1.1532e-06 | -3.8147e-06 | 0.99652 | 34.863 | 4.5217e-05
SWL2008-02#158 | 1236 | 2009-02-05 10:57 | 24C | 0.99552 | 2.4918e-06 | -3.8147e-06 | 0.99552 | 34.8237 | 9.6642e-05
SWL2008-02#159 | 1237 | 2009-02-05 10:59 | 24C | 0.98047 | 1.1196e-06 | -3.9339e-06 | 0.98046 | 34.2324 | 4.5227e-05
SWL2008-02#160 | 1238 | 2009-02-05 11:02 | 24C | 0.97818 | 2.2391e-06 | -3.8743e-06 | 0.97818 | 34.1428 | 8.9601e-05
SWL2008-02#161 | 1239 | 2009-02-05 11:05 | 24C | 0.97362 | 3.6233e-06 | -3.9339e-06 | 0.97362 | 33.964 | 0.00014
SWL2008-02#162 | 1240 | 2009-02-05 11:07 | 24C | 0.9608 | 2.2725e-06 | -3.9339e-06 | 0.9608 | 33.4624 | 8.6805e-05 | no insert
SWL2008-02#163 | 1240d | 2009-02-05 11:10 | 24C | 0.96128 | 1.3713e-06 | -3.9339e-06 | 0.96127 | 33.481 | 5.4326e-05 | no insert
SWL2008-02#164 | 1241 | 2009-02-05 11:13 | 24C | 0.95562 | 3.6235e-06 | -3.8147e-06 | 0.95561 | 33.2599 | 0.00014164 | no insert
SWL2008-02#165 | 1242 | 2009-02-05 11:16 | 24C | 0.94334 | 5.0651e-06 | -3.9339e-06 | 0.94334 | 32.7812 | 0.00019734 | no insert
SWL2008-02#166 | 1243 | 2009-02-05 11:19 | 24C | 0.92892 | 2.4744e-06 | -3.9339e-06 | 0.92892 | 32.2201 | 9.6609e-05 | no insert
SWL2008-02#167 | 1244 | 2009-02-05 11:21 | 24C | 0.92622 | 3.332e-08 | -3.9339e-06 | 0.92622 | 32.1152 | 0
SWL2008-02#168 | 1245 | 2009-02-05 11:24 | 24C | 0.92144 | 2.0947e-06 | -3.9339e-06 | 0.92144 | 31.9297 | 7.9337e-05
SWL2008-02#169 | 1246 | 2009-02-05 11:28 | 24C | 0.91102 | 2.2392e-06 | -3.9339e-06 | 0.91102 | 31.5258 | 8.8244e-05
SWL2008-02#170 | 1247 | 2009-02-05 11:31 | 24C | 0.89236 | 2.8768e-06 | -3.9339e-06 | 0.89236 | 30.8044 | 0.00011104
SWL2008-02#171 | 1248 | 2009-02-05 11:33 | 24C | 0.88825 | 2.4777e-06 | -3.9339e-06 | 0.88824 | 30.6456 | 9.751e-05
SWL2008-02#172 | 1249 | 2009-02-05 11:37 | 24C | 0.83731 | 1.4046e-06 | -3.9339e-06 | 0.8373 | 28.6903 | 5.2781e-05
SWL2008-02#173 | 1250 | 2009-02-05 11:40 | 24C | 0.83719 | 3.4658e-06 | -3.9935e-06 | 0.83719 | 28.686 | 0.00013341
SWL2008-02#174 | 1251 | 2009-02-05 11:43 | 24C | 0.83904 | 4.3337e-06 | -3.8743e-06 | 0.83904 | 28.7567 | 0.00016438
SWL2008-02#175 | 1252 | 2009-02-05 11:45 | 24C | 0.83643 | 3.8748e-06 | -3.9339e-06 | 0.83643 | 28.657 | 0.00014711
SWL2008-02#176 | 1252d | 2009-02-05 11:48 | 24C | 0.83821 | 3.1414e-06 | -3.9339e-06 | 0.8382 | 28.7248 | 0.00012104
SWL2008-02#177 | 1253 | 2009-02-05 11:50 | 24C | 0.83379 | 2.8897e-06 | -3.9339e-06 | 0.83379 | 28.5562 | 0.0001098
SWL2008-02#178 | 1254 | 2009-02-05 11:54 | 24C | 0.83437 | 1.3712e-06 | -3.9339e-06 | 0.83437 | 28.5782 | 5.2781e-05
SWL2008-02#179 | 1255 | 2009-02-05 11:56 | 24C | 0.83232 | 1.1063e-06 | -3.9935e-06 | 0.83232 | 28.5 | 4.3716e-05
SWL2008-02#180 | 1256 | 2009-02-05 11:59 | 24C | 0.82927 | 1.7702e-06 | -3.9339e-06 | 0.82926 | 28.3835 | 6.7494e-05
SWL2008-02#181 | 1257 | 2009-02-05 12:02 | 24C | 0.83094 | 1.1196e-06 | -3.9339e-06 | 0.83094 | 28.4474 | 4.265e-05
SWL2008-02#250 | cal check | 2009-02-09 09:20 | 24C | 0.99985 | 4.3688e-06 | -3.9339e-06 | 0.99985 | 34.9941 | 0.00017061