## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2008-05

Agency: OSD

Location: WCVI

Project: La Perouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: May 20, 2008 – May 28, 2008

Processed by: Germaine Gatien

Date of Processing: 9 February 2009 – 25 March 2009

Number of original CTD casts: 	60  	Number of CTD casts processed:  58 (2 missing)

Number of original bottle casts: 	29	Number of bottle casts processed: 26 (1 raw file missing, 2         unintended firings of bottles – no sampling)

Number of original TSG files: Unknown No files available at time of processing – unknown if any exist.

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1176), a Seapoint Fluorometer (#2228) with a 10X cable and an altimeter (#1252). The deck unit was an SBE model 11, s/n 0424 and the logging computer was an HP COMPAQ.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log lacked information about personnel and the TSG; it had a list of CTD equipment and a good record of changes of sensors. The rosette logs were generally in good order.

Cast #59 had bad temperature, salinity, dissolved oxygen and fluorescence data so those channels were removed. That left only pressure and transmissivity in the CTD file. The CHE file for that cast contains pressure, transmissivity and the chemistry data.

The data for casts #7-34 were mislabelled as being from cruise 2008-36.

Many of the salinity samples were identified incorrectly on the analysis spreadsheet. Best guesses were made based on rosette log information and comparison with CTD values. It appears that the correct sample number assignments have been made, but flags were assigned where any doubt remains, and in all cases where samples numbers were changed notes were entered.

Some raw data files are missing: casts files for events #34 and #128 and the thermosalinograph files.

The dissolved oxygen data for event #2 were thought to be lost until late in the processing when it was discovered to be “hiding” in the file for event #1. Such concatenation of data is fairly common in the oxygen analysis output.

Different combinations of T and C sensors were used for casts #1-59 and #60-112 and 114-128. The primary sensors were the same for casts #60-128. The primary salinity data were somewhat noisier for casts #60-126 than for #1-59. Based on post-cruise calibrations and bottle comparisons, the salinity channel was recalibrated for the first pair only.

The configuration files were updated to add the parameters needed to run the latest Sea-Bird algorithm for dissolved oxygen concentration.

The dissolved oxygen data in the CTD files should be considered roughly:

- ±0.8 ml/l from 0 - 50db

- ±0.5 ml/l from 50 – 300db

- ±0.15 ml/l below 300db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. Some files (7-34) were mislabelled as being from 2008-34. They were renamed. Raw data were missing for casts #34 and 128. And there were no TSG data, though there are loop samples, so the TSG may have been operating.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets. There was a list of equipment (except for the TSG), but no list of personnel. There were several instrument changes which are well documented in the log.

Salinity, extracted chlorophyll and nutrient data were available in spreadsheet format.

Dissolved oxygen data were available in individual OXY files without flags or comments.

The cruise summary sheet was completed.

There were 3 configurations used during the cruise. Those from casts 1, 61 and 114 were used to create the following configuration files:
2008-05-ctd1.con – for events #1-59

2008-05-ctd3.con – for events #60-112

2008-05-ctd3.con – for events #114-126

The oxygen calibrations had to be updated to include the extra parameters needed for the new Sea-Bird algorithm for calculating dissolved oxygen concentration. All parameters were checked and no other changes were needed.

The histories of the conductivity, DO and pressure sensors were obtained. There was a post-cruise calibration available for 3 of the 4 conductivity sensors. The primary sensors used for casts 60 -126 showed little drift since they were last recalibrated; that conductivity sensor has been used on a lot of cruises for 2008, so it is expected the primary sensors will be selected for those casts. The other primary sensor has a post-calibration drift estimate that suggests only a small error. One of the secondary conductivity sensors died during the cruise and a post-cruise drift check was impossible and the other was probably low by ~0.004 based on a post-cruise drift check.

##### Conversion of Raw Data

Data were converted using the configuration files given above.

A few casts were examined and all expected channels are present and look reasonable when plotted. As usual the upcast temperature and conductivity traces are noisier and further apart than in the downcasts.

For the deepest casts, the fluorescence is about 0.08ug/l below 1000db.

The transmissivity has small spikes at depth; downcasts and upcasts are close.

The dissolved oxygen shows the usual offset, but the response time does not look too bad.

The descent rate was very noisy for many casts.

The altimetry generally looks useful.

Rosette files were converted using a start time of -5s and duration of 10s. The stops were at least 30s long. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT. Station names were fixed in the headers of some files.

All BOT files were plotted and the only problem noted was in cast #108 when the secondary salinity is clearly bad. The sensor was replaced at that point. No editing was done.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels.

Parameters used were: Pass 1 Std Dev = 2        Pass 2 Std Dev = 5           Points per block = 50

##### CELLTM

The descent rate was very noisy for most of the few casts without stops for bottles, making the tests for selecting CELLTM parameters difficult to assess. Only two suitable casts could be found, one from each of the conductivity sensors. Settings tried were (α = 0.01, β=7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9), (0.04, 7), (0.04, 9) and (0.0245, 9.5). From cast #5 sensor #2399 looked best with (0.02, 7) as it did for 2008-12 when the same equipment was used. Sensor #3396 looked slightly better with (0.03, 7) than the 2008-12 choice for this sensor of (0.0245, 9.5); that channel is unlikely to be archived, so this is not a critical choice – either will improve the data. For sensor #2102 the best choice for was (0.0245, 9.5) and for #2102 the best is (0.03, 7).

Based on those tests CELLTM was run using the following parameters:
Casts 1-59: Primary - (α = 0.02, β = 7); Secondary - (α = 0.03, β = 7)

Casts 60-112: Primary - (α = 0.0245, β = 9.5); Secondary - (α = 0.03, β = 9)

Casts 114-126: Primary - (α = 0.0245, β = 9.5); Secondary - (α = 0.03, β = 7)

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 5 | 100 | ~0 XN | -0.0007 N | -0.008 XN | Noisy/moderate |
| 39 | 1400 | +0.0004 | -0.0007 | -0.009 N | Noisy/moderate |
| 41 | 1200 | +0.0004 XN | -0.0007 N | -0.008 | Noisy/low |
| 79 | 1200 1750 | +0.0004 +0.0002 | ~0 ~0 | -0.0004 N -0.0002 N | Moderate/high Moderate/high |
| 89 | 1000 | ~0 | +0.0008 | +0.0008 | Moderate/high |

The first three casts had different sensors from those for the second two. The differences for the first 3 are similar but slightly higher, than for 2008-05 when the same sensors were used. The differences for the second two are small, but they differ from each other, which may be a sign of the problem that developed later with the secondary conductivity sensor.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and a cross-reference listing were produced. Errors and format inconsistencies were fixed in some station names. Two cast files were misnamed: files 2008-05-0016 and 2008-05-0017 were renamed as 2008-05-19 and 2005-05-0021.

The cruise track was plotted and no problems found.

The average surface pressure is 2.6db which is reasonable for the Tully.

The altimeter readings from the headers of CLN files were exported to a spreadsheet and a random sample of entries were checked against plots of altimetry against pressure. No problems were found.

##### BOTTLE FILE PREPARATION

##### Errors were found in the event numbers recorded on the rosette sheets: 109 should be 108 and 113 should be 114.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Problems included:

For events #13 and 102 there are bottle files, but no sample numbers and no indication that any sampling was done, so those files will not be processed further. For #102 there is a note in the log that this firing was not done by the operator.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

Extracted Chlorophyll

The chlorophyll data were delivered in spreadsheet format in file QF2008-05CHL.xls. This file was edited to change channel names to standard format and to remove some unnecessary lines and columns and it was saved as 2008-05-CHL.csv. Event numbers were changed from 109 to 108 and from 113 to 114. The loop sample data were moved to file 2008-05-loop-bottles.csv. There are no TSG data available at this time, but the file will be saved in case it is needed later to compare TSG and loop data. For information on duplicates see the original file. The file was converted to individual CHL files.

Nutrients

The chlorophyll data were delivered in spreadsheet format in file QF2008-05nuts.xls. This file was edited to change channel names to standard format, to reorder on sample number and to remove some unnecessary lines and columns and it was saved as 2008-05-nuts.csv. The event #s were changed from 109 to 108 and 113 to 114. The spreadsheet was converted to individual NUT files. The loop results were added to file 2008-05-loop-bottles.csv.

SALINITY

The salinity data were delivered in spreadsheet format 2008-05 salinity.xls with flag and comment columns. That file contained both loop and rosette samples; the rosette data were saved as 2008-05-sal.csv and the loop data were copied to 2008-05-loop-bottles.csv. Headers were changed to standard formats and the spreadsheet was simplified. There were many errors in the station name entries and sample #s and some comments look like they should have been with other samples. It looks as if some columns got shifted relative to others. The best guesses were made as to where data belonged based on sampling notes in the rosette sheets and values that looked like they were from the surface;  where sample numbers were changed, a “c” flag was entered and a comment was entered with the details. (After COMPARE was run a few of the flags were removed because there was little doubt about the sample #s and the fit was excellent.) The duplicates were copied to file 2008-05-sal-duplicates.xls. In the main spreadsheet the two values were replaced by a single averaged value and flag “f” was entered. This file was converted into individual SAL files.

There were 5 duplicates all from one cast. The differences were all <0.002 with only one of them >0.001. Using all 5 pairs Sp = 0.0006 and if one pair are excluded it is 0.0003 where Sp is defined as:

Sp = Square Root (sum of squares of differences / 2 * number of pairs)

These results look good. During 2008-26 Sp=0.0012 using 7 pairs and 0.0005 when 1 pair were excluded.

DISSOLVED OXGYEN

The dissolved oxygen files were delivered as *.OXY files with no flags or comments. Flag and COMMENTS channels were added. (Output:*.ADD) The file names were changed from 2008-05-0109 to 2008-05-0108 and 2008-05-0113 to 2008-05-0114.  Every ADD file was examined and flags were attached to a few samples due to bubbles, or confusion about flask numbers. There were no DO duplicates.

There were no DO data for cast #2 though sampling is indicated on the rosette log sheet. (The data were later found for this cast – they were concatenated with cast #1. They are included in the bottle file but were not included in COMPARE. A quick check shows values are reasonably close to CTD bottles.)

The SAL, CHL, ADD and NUT files were merged with CST files in four steps, put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

COMPARE was run using pressure as the reference channel. There were some major outliers:

- For cast #27, sample 27 the nutrients had been flagged “c” and the analyst felt this was a case of a misfire. The chlorophyll, salinity and DO confirm this. All should be flagged “d”.

- For cast #31, sample #54 differs from both CTD channels by 0.5 – probably mislabelled. The value was recorded in a header comment and then replaced with a pad value and flag “e” was added.

- For cast #31, sample #45 had already been flagged “c” because the sample number was uncertain – the one in the raw salinity file was marked both 455 and 45 and there was another sample marked 45 but with station name of another cast. The flag was changed to “d”

- For cast #59 all samples are extreme outliers in COMPARE, but the CTD data look bad, so no flags added.

- For cast #108, the surface bottle is an extreme outlier in comparison to the secondary CTD salinity only, and the CTD secondary salinity looks bad. No flag attached.

COMPARE was rerun and this time the plots were divided according to conductivity sensor:

1. Primary Sensor #2399: The only severe outliers were identified above, the worst being from cast #59 and cast #27, sample #27. A few other minor outliers came from the surface and/or places with high standard deviation in the CTD data. The data in the top 200db are very noisy. When only bottles below 200db are included and one other outlier excluded, the salinity is high by 0.0015. It is known that there were some Autosal linearity problems in early 2008 leading to the Autosal reading significantly high for low salinity samples; the issue was probably resolved before these samples were analyzed. Selecting data below 200db will minimize the effect of such non-linearity  if it exists.

2. Primary Sensor #1766: The results for this sensor are harder to understand. There are 4 outliers, of which 3 involved Niskin #2. It would be easy to imagine a misfire from the salinity data, but the other samples don’t suggest that. For two of the samples #176 (cast #61) and #274 (cast 90) the salinity is a little out of line, but not so far as would be explained by drawing from the wrong bottle. The CTD data are a little noisy, so flags will not be added. Two 1500db samples during casts #76 and 83 are harder to explain. The CTD data look quiet. The dissolved oxygen values seem slightly little out of line. The nutrients are harder to judge. Only 6 points are left below 200db after these outliers are excluded; they indicate the CTD is within 0.0001 of the bottles.

3. Secondary Sensor #3396: All outliers had already been noted elsewhere. There were only 11 bottles below 200db and they indicate the CTD is low by ~0.007, but there is more scatter than in the primary comparison.

4. Secondary Sensor #2102: The 1500db samples show up as outliers again. There are only 5 bottles in the fit below 200db and there is a lot of scatter. They indicate the CTD is within 0.001 of the bottles, but the scatter limits confidence in this comparison.

Some more flags were added after the rerun of COMPARE:

- Sample #6 from cast #2 was flagged “c”. This bottle is a minor outlier in COMPARE, but taken with the fact that there is no such sample mentioned on the rosette log sheet, it must be treated with caution.

- Sample #209 from cast #76, all samples will be flagged “c” with a comment that salinity and dissolved oxygen seem a little out of line and that Niskin #2 frequently misfired, so it is possible it closed in water deeper than indicated.

- Sample #229 from cast #83, all samples will be flagged “c” with a comment that salinity and dissolved oxygen seem a little out of line and that Niskin #2 frequently misfired, so it is possible it closed in water deeper than indicated.

A detailed study of the linearity of the Autosal data was not pursued since there are so many sensors and so many outliers that it is unlikely to prove anything. Moreover, there are post-cruise drift reports for some of the sensors so recalibration will not be as dependent on the bottle comparisons as usual.

Extracted CHL vs CTD Fluorescence

COMPARE was run using pressure as the reference channel. On average CTD fluorescence was about 56% of the extracted chlorophyll. The ratio CTD FL / Extracted CHL was plotted against event # and against CHL. For CHL<0.1 there were some very high ratios, but small errors could be larger than the differences, so those were excluded. A pattern emerges that the ratio is low close to shore with values mostly ~0.4 or 0.5. Offshore the values are ~0.8 in the south and ~ 1.5 at the LG and CS lines. This probably just reflects that the ratio is higher for lower chlorophyll regions. For 0.1<CHL<2 the average ratio is 1.

Dissolved Oxygen

COMPARE was run using pressure as the reference channel. The plot of differences against DO concentration shows a lot more scatter than usual. So a plot was made against file pair number. From that outliers were identified. Large differences at the surface are not unusual, but only 2 of the outliers fit that description. Cast #59 stands out as bad; that was the cast with bad CTD conductivity. Whatever caused the trouble appears to have affected dissolved oxygen as well. The fluorescence also looks unreliable, so all pumped channels are affected. When the outliers identified above were excluded, the fit against DO concentration looks a little tighter, but still somewhat scattered.

The following outliers were checked:

- Cast 4, 4db, Sample #7, Niskin #7 – All bottle data look like they came from much deeper. The CTD data look ok, so it looks like firing went ok, but that the bottle had actually closed earlier. All samples should be flagged “d”.

- Cast 27, 125db, sample #27 – all samples already flagged “d” after salinity COMPARE.

- Cast 33, 151db, sample #61, Niskin #7 – only dissolved oxygen stands out as a misfit – CTD ok. Flag DO “d”.

- Cast 42, 4db, sample #127, Niskin #16 – DO already flagged “c” due to bubble – change to “d”.

- Cast 50, 74db, sample #140, Niskin #2 – Nutrients already flagged “c”, DO clearly out of line, looks like all samples from above 50db. Flag all as “d”

- Cast 58, 76db, sample #157, Niskin #2 – Again, nutrients already flagged “c”, DO clearly out of line, looks like bottle closed above 50db. Flag all as “d”.

- Cast 83, 50db, sample #243, Niskin #16 – Nutrients look fine. The DO looks a little odd, but the CTD temperature and salinity suggest there is a reversal in gradient below 50db, so the CTD DO may not have been able to record this successfully. The DO sample is probably fine.

- Cast 90, 150db, sample #267, Niskin #2 - Again, nutrients already flagged “c”, DO clearly out of line, looks like bottle closed above 75db. Flag all as “d”.

Flags were added to all the basic files, SAL, CHL, ADD and NUT, as appropriate, but where all samples were involved, a general comment was added only for the first occurrence in the order above, so the same comment does not appear repeatedly in the final files. For example, if only nutrients and dissolved oxygen were sampled, then the comment was added to the ADD file.

Various fits were tried to see if the source of noisy data could be identified. Some bottles from casts #29-36 seemed to have a lot of minor outliers, but the CTD data do not look particularly noisy. However, there is a temperature reversal around 75db for casts 31-39 which may account for the noise level. There was also more vertical motion than usual for some bottle stops. But there is no suggestion of stops being too short. Bad bottles were removed from the fit as well as outliers thought to be related to the temperature reversal. Residuals were used to exclude a few more outliers. The fit found was then:

CTD-BOT = 1.0625 DOX-CTD + 0.042

The slope lies between those found for the cruises before and after 2008-05 which were:

CTD-BOT = 1.0379 DOX-CTD - 0.026 (2008-32 Late April 2008)

CTD-BOT = 1.0708 DOX-CTD - 0.01   (2008-26 Late May 2008)

(See 2008-05-dox-comp1.xls.)

Plots were made of salinity versus CTD DO and Titrated DO. Nothing was found that needs flagging beyond what has already been assigned.

##### 13. Shift

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2 (since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The usual shift of +24 records (1s) was found to be appropriate and was applied. (Output: SHFFL)

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. For this sensor SHIFTS of from +100 to +110 records have been used in the past. However, Sea-Bird have made several recent changes to the algorithm for calculating dissolved oxygen and the SHIFT values have changed. For 2008-61 a shift of +70 records worked well. Tests were run using values from +50 to +110 records and the best results were with +70.

SHIFT was run using +70 records for all casts.

Conductivity

Tests were run on 2 casts for each primary sensor. The data are noisy so the tests were not clear. For sensor #2399 the tests were especially difficult, but a setting of -0.4s looks reasonable. For 2008-32 which preceded this cruise a setting of -0.4s was also found best for that sensor. For sensor #1766 a setting of -0.3 or -0.4s looked best, so since -0.3s was used for that sensor during 2008-26 which followed this cruise, and had more suitable data for testing, it will be used for this cruise as well. All casts were put through two runs of SHIFT using those settings.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

- Primary sensor #2399 has been used for a lot of recent cruises, but there is little confidence in the bottle comparison due to linearity problem in the Autosal when some of the analyses were run. An rough correction made to allow for that error led to estimates that the salinity was generally high by <0.001.

- Primary sensor #1766 was used on a few cruises after this one, and seemed to be reading a little low, but the bottle comparisons were limited. 2008-05 was the first use after recalibration.

- Secondary sensor #3396 was used on 2 other cruises in 2008 and the CTD salinity was low by 0.0005 and 0.003 when allowance was made for linearity errors.

- Sensor #2102 has not been used since it was last recalibrated just before this cruise.

2. Dissolved Oxygen – This sensor has been used for many cruises since it was last recalibrated. The slope for fits are found to be increasing fairly steadily, though there is a significant difference between the COMPARE fits for cruises just before and after this one.

3. Pressure – This sensor has been drifting significantly in recent years. An offset of +6.5db was used from March 2008 until August 2008. For cruises since Sept. 2008 an offset of 6.5db has been used.

Post-cruise report – Based on drift reports when T and C sensors sent for recalibration

- Sensors #2339/4054 salinity likely high by ~0.0015

- Sensors #1766/4054 salinity looked low by ~0.001 seven months after this cruise, so is presumed to have very small error for this cruise

- Sensors #3396/4700 salinity low by ~0.004 shortly after this cruise

- Sensors #2102/4700 – no post-cruise report available

Historic ranges –There were a few excursions from the historic ranges:

- There were minor excursions in shallow water in both temperature and salinity with no systematic nature, sometimes high, sometimes low. The profiles of the climatology look odd in most of these places, so the excursions are probably are due to inadequacies in the range data.

- The temperature at the bottom of some of the deeper casts in the northern half of the cruise are slightly above the maxima. These excursions look believable and were also noted during 2008-32 in late April. The local climatology is limited and out-of-date.

None of these excursions are considered evidence of calibration problems.

Repeat Casts – There were no repeat casts.

##### DETAILED EDITING

The primary sensors look like the best choice based on both the bottle comparison and the post-cruise drift reports.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. All casts required some editing. The near-surface salinity data were noisier for the second T/C sensor pair. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

Based on COMPARE and the post-cruise calibration the primary salinity is believed to be high by 0.0015 for the sensors used for casts 1-59 and low by <<0.001 for the sensors used for the second half of the cruise.

File 2008-05-ctd.ccf was prepared to subtract 0.0015 from the primary salinity for casts 1-59 and to apply the following equation to the CTD Dissolved Oxygen channel in the SAM and MRG files for all casts:

CTD-BOT = 1.0625 DOX-CTD + 0.042

COMPARE was then rerun with the SAL and DO data to check that the results were as expected and they were. However, as for the first run, there is a lot of noise. (See 2008-05-sal-comp2.xls and 2008-05-dox-comp2.xls.)

The same DO calibration was applied to the edited CTD files.

##### Special Fluorometer Processing

The COR1 files were clipped to 120db and processed separately for A. Peña. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. A further correction will be applied to at least partly correct for remaining response time errors. To do this we compare downcast data to bottle data from the same pressure.

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When outliers are excluded, the fit against DO looked most useful. The CTD was higher than the bottles except at depth, where there were few bottles and 2 of those were flagged. The following correction is found:

DOX_CTD (Corrected) = 1.0049 * DOX_CTD -0.0455

(See 2008-05-dox-comp3.xls.)

The thinned files were recalibrated by applying the above equation using file 2008-05-recal2.ccf.and the comparison was rerun. The average difference was +0.0001 and there was little pressure dependence, showing that the recalibration had been applied correctly. (See 2008-05-comp4.xls.)

Recalibration using file 2008-05-recal2.ccf was applied to the downcast files only as this correction is not appropriate for data acquired while stopped since the sensor has time to equilibrate. (Output: COR2)

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

For cast #59 all channels were removed except Pressure and Transmissivity.

A second SBE DO channel was added; REORDER was run to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered roughly:

  - ±0.8 ml/l from 0 - 50db

  - ±0.5 ml/l from 50 – 300db

  - ±0.15 ml/l below 300db

The Standards Check routine was run and no problems were found. The final files were named CTD.

A header check turned up no problems.

Profile plots were made and no problems were found.

The track plot looks ok.  The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The values ranged from 85% to 160%, with the highest values in the Cape Scott area and the lowest near the coast in the south. Most values were between 100 and 120%. There was little DO sampling at the casts with the highest saturation, but a check of one such cast does suggest that the surface CTD DO may be too high.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

For cast #59 all CTD channels were removed except Pressure and Transmissivity. The rosette sample data were kept.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

13. Bottle fired at surface – no sampling, so no CHE file prepared.

25. Touched Bottom

27. CHLa and NH4 taken from bottles 1,2,3 instead of 7,8,9

29. Nutrient samples mislabelled

36. Changed computer times – off by one week up to this point

60. Change of equipment

100. No altimeter trace

102. Bottle fired at bottom, not by operator. No CHE file prepared.

108. Secondary conductivity noisy

114. Change of equipment

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Jan08 | Factory “ |  |  |
| Conductivity (1-59) | 2399 | 15Jan08 | “ | 13June08 | Factory |
| Conductivity(60-114) | 1766 |  |  | 24Dec08 | Factory |
| Secondary Temp. | 4700 | 16Jan08 | “ | 24Dec08 | Factory |
| Secondary Cond. (1-59 & 114-126) | 3396 | 21Dec07 | “ | 13June08 | Factory |
| Secondary Cond.(60-112) | 2102 | 07May08 | “ | 24Dec08 | Factory |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 14Feb2007 | Factory |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |