## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 2 April 2025 | Updated channel names & formats in TOB files. GG |

## PROCESSING NOTES

Cruise: 2008-07

Agency: OSD

Location: Broughton Archipelago

Project: Broughton

Party Chief: Stucchi D.

Platform: Vector

Date: March 24, 2008 – March 31, 2008

Processed by: Germaine Gatien

Date of Processing: 19 May 2008 – 28 June 2008

Number of original CTD casts: 59		Number of CTD casts processed: 59

Number of bottle casts: 	 15			Number of bottle casts processed: 15

Number of original TSG files: 2			Number of TSG files processed: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1176), on the primary pump), a Seapoint Fluorometer (#2228) with a 10X cable (on the secondary pump) and an altimeter (#43281). The deck unit was a model 911+ (#0424) and the logging computer was an HP Compaq-II.

A thermosalinograph (SeaBird 21 S/N 2487) was used without an intake temperature sensor; there were no external sensors mounted on it.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order and there was a post-cruise report available.

The dissolved oxygen data in the CTD files should be considered

- ±0.3ml/l from 0 –100db

- ±0.2ml/l below 100db

There are 2 thermosalinograph files with a 1-day gap between them. There were no positions in the TSG files. Position data were obtained where possible from text files that contained times and positions. There was a 1-hour discrepancy between those times and the log book times when positions were matched. The log book is believed to be correct. There are pad values for many of the latitude and longitude entries because there was no match available. A small error in the match of time and positions could lead to significant errors.

The recalibration of the thermosalinograph data is based on a comparison to CTD data and on history, but there was little of that available. Given the likelihood of errors in positions and the paucity of recalibration information, the data should be considered of lower quality than usual.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets and a post-cruise report.

Salinity data were available in spreadsheet format.

Nutrient data were not available at the beginning of the job, but later arrived in spreadsheet format.

Dissolved oxygen data were available in individual files.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors were found.

All CTD con files were the same; one was saved as 2008-07-ctd.con.

##### Conversion of Raw Data

Data were converted using configuration file 2008-07-ctd.con.

Cast file #44 was named 2008-07-0044hex.*, so the name was changed to 2008-07-0044.*.

A few casts were examined and all expected channels are present and look reasonable when plotted. As usual the upcast temperature and conductivity traces are noisier and further apart than in the downcasts. The fluorescence has small spikes, but no major problems; for the deepest cast the dark value is ~0.11ug/l.

The transmissivity also has small spikes at depth, but not large ones.

The dissolved oxygen shows the usual offset, but the response time does not look too bad.

The descent rate is usually very steady.

The altimetry looks reasonable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and no editing was found necessary.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on a few casts to determine the best choice of parameters for CELLTM. Settings tried were (α = 0.01, β=7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9), (0.04, 7), (0.04, 9) and (0.0245, 9.5). The differences were slight because the temperature gradients were low and the best choice varied from one area to another. The choice of (0.0245, 9.5) looked reasonable for all cases, so was applied to all casts for both channels.

CELLTM was run using (α = 0.0245, β = 9.5).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy, especially in the shallow temperature, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 19 | 220 240 400 | +0.001VN +0.0001 +0.0002 | -0.0003 -0.0004 -0.0004 | -0.0040 -0.0043 -0.0044 | Very Steady ~.8 |
| 46 | 220 | -0.0001 | -0.0004 | -0.0038 | Very Steady ~.8 |
| 55 | 220 | -0.0002 | -0.0003 | -0.0040 | Very steady ~.9 |

The differences are reasonably consistent given that the casts are shallow and there is no evidence of time dependence.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and a cross-reference listing were produced. Errors were found in two station names and in some of the dates. Times and positions look fine. The NMEA date does not agree with the log book or computer dates for casts #6-10, 19-24, 28-33 and 50-54. Changes were not made at this stage because the chief scientist and CTD technician were unavailable to confirm which dates were correct. Processing continued, but after the DELETE stage changes will be made to headers as needed, and the header checks will be rerun. Changes will be applied to the DEL files.

The cruise track was plotted and no problems found.

The average surface pressure is 2.9db, which is fairly high for the Vector. It is possible that the offset is slightly too high, but SEABIRD say that the drift in pressure sensors always leads to increasingly large offsets, this is not likely to be a significant source of error. It is more likely that the CTD was just started deeper than usual.

The altimeter readings from the headers of CLN files were exported to a spreadsheet and some casts were checked. There is a lot of noise in the data for many casts but the algorithm appears to have worked well for all except cast #27 where the header entry was based on just a few spikes to values <15. All other values are higher than 15 so the entry was removed from the header. The log book entries indicate that the CTD was about 20db above the bottom at its deepest point for that cast. All altimetry header entries >10db were checked and no other problems were noted. The bottle files had no altimetry header for cast #27 so no correction was needed there.

Cast #45 had an initial lowering to 10db followed by a return to 2.8db and then a full cast. A text editor was used to remove records from that initial lowering to ensure that DELETE would select the best data, and not have a patch of the two. A note of this was made in the header.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

SALINITY

The salinity data were delivered in spreadsheet format 2008-07sal.xls. Headers were changed to standard formats and the spreadsheet was simplified and saved as2008-07-sal.csv. This was converted into individual SAL files. There were no duplicate salinity samples.

DISSOLVED OXGYEN

Dissolved oxygen files (*.add) were provided with a flag channel and comments entered in the headers. A few changes were made to the comments to make it clear which property and which sample are being discussed. For cast #17 the data were reordered on sample #. Where there were duplicate samples, both were replaced with a single averaged value and “f” flags were attached. It is noted that there is a larger difference between duplicates for cast #30 where the two samples from the bottom bottle at 260db differed by 0.055ml/l; one of those samples was very close to the bottle from 250db, so a note of both values was put in a header comment and both “f” and “c” flags were attached. There are other casts with a DO rise at the bottom, so perhaps this rosette bottle was not as homogeneous as most. It will be re-examined after running COMPARE.

All duplicate information was transferred to document 2008-07-DO-dups.xls where analysis was done of the differences for the 8 pairs. The average difference was 0.015ml/l if all data are included and 0.009 if one pair are excluded. Similarly when expressed as a % difference (absolute difference *100 / average concentration) the average is 0.27% using all data and 0.17% excluding cast #30.

Plots were made of CTD Salinity versus CTD DO and Bottle DO – no outliers were found.

NUTRIENTS

The nutrient spreadsheet was simplified and saved as 2008-07-nuts.csv. Extraneous columns were removed, header names were changed to standard format and lines were removed for which there was no sampling for all 3. “Nutrients:” was placed before all comments. Data were sorted on sample number. File 2008-07nuts.csv was then converted to NUT files.

The SAL, ADD and NUT files were merged with CST files in three steps. (Output: MRG1, MRG3, and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

Data were exported to a spreadsheet 2008-07-bottles.xls and compared to the rosette sheets to ensure all expected data are present. No problems were found.

##### Compare

Salinity

COMPARE was run using pressure as the reference channel. When 6 outliers are removed flat fits are found showing that the primary salinity was low by an average of 0.0007 and the secondary by 0.0047. This is consistent with the salinity differences found between channels in section 7. The variations with time do not look significant. The most extreme outliers are records with large standard deviation in the CTD data. There is a lot of noise in the comparison, and there are some concerns about the reliability of the Autosal results, particularly in the lower range of salinity as sampled during this cruise. In fact, based on some other cruises it seems likely that the bottle values are high. A preliminary linearity test on the Autosal suggests that for salinity of ~32 the results are likely to be high by ~0.005 to 0.006psu in the salinity range for this cruise which would imply that the secondary channel has the more accurate salinity.

Dissolved Oxygen

COMPARE was run using pressure as the reference channel. When standard deviations in the CTD data >0.01 were removed plus a few outliers identified by residuals the fit was:

CTD-BOT = 1.0254 DOX-CTD + 0.0081

The only significant outlier was associated with noisy CTD data. (See 2008-07-dox-comp1.xls.)

The fit is reasonably close to the fit from 2007-43 given the distribution of DO values was somewhat different: 	CTD-BOT = 1.0351 DOX-CTD + 0.0098

The analyst was concerned about the bottle DO for cast #49 since the computer output was not complete. It looks fine in the comparison and plots of bottle DO versus salinity show a very similar pattern to other casts so there does not seem to be a problem.

At this point the duplicates from cast #30 were re-examined. The duplicates from sample 82 had values 5.761 and 5.706. In the comparison with CTD the larger value looks like an outlier and the lower value is in line with other differences. It also looks most like the samples from the bottle 10m above it. So the lower value was selected for the bottle file, and a note was put in the header to explain the choice.

##### 13. Shift

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The usual shift of +24 records (1s) was found to be appropriate and was applied. (Output: SHFFL)

Conductivity

Tests were run on 6 casts using shifts between -1s and +1s and T-S plots were prepared to compare the results. The range of values was small, so the results were not clear, but a setting of -0.3s looked best overall for both sensors. All casts were put through two runs of SHIFT using those settings. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on 3 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly.

SHIFTS of from +100 to +130 records were tested and +110 looked best overall, though the differences were not large.

SHIFT was run using +110 records for all casts.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

At this point the dates were fixed in the headers for casts 6-10, 19-24, 28-33 and 50-54. The station names were fixed for casts #1 and 36. The same changes were made in the BOT, SAM and SAMAVG files. The log book is presumed to be correct.

An examination of a few casts shows no significant difference in the noise level in the two salinity channels. COMPARE shows the primary closer to the bottles, but it is likely that the Autosal results are high so that the secondary is likely more accurate. All sensors were recalibrated shortly before this cruise and the report on drift between calibrations was similar for both temperature and conductivity sensors. There is no difference in the variations with time for this cruise, but the secondary is slightly flatter with pressure. So choosing the secondary looks wisest. If future use of the sensor indicates that the salinity really is low, it can be recalibrated later. The secondary channels were selected for editing.

There were a few casts which had an initial lowering of the CTD after which the package returned to the surface after which a full cast was run. Cast #45 was edited earlier to remove the initial lowering so that DELETE did not patch parts of the two lowerings. In running CTDEDIT it was found that many other casts had such partial lowerings, but for many DELETE picked out only records from the full cast. For the others a text editor was used to remove the initial sections from the SHFO file and DELETE was rerun. The files so edited were: 6, 8, 9, 32, 40, 41, 43, 44 and 46.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were used to guide the editing. Most of the editing was in the top 10db or the bottom 1db. For many casts the CTD had been lowered to 10db then returned to the surface and lowered again after a short stop. It is unknown why this was done; sometimes low surface conductivity leads to the pump not starting, but that was not the case for these casts since the pump was on for the first lowering. There is a lot of noise in the top 10db and some of it is probably due to the stirring by the initial lowering.

All casts required light editing. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

The secondary salinity is probably very close to the bottle values, though the bottle calibration shows it to be low by 0.0047; no recalibration of salinity will be applied at this point, though information from subsequent cruises might lead to a future recalibration.

File 2008-07-ctd.ccf was prepared to apply the following equation to the CTD Dissolved Oxygen channel in the SAM and MRG files.:

CTD-BOT = 1.0254 DOX-CTD + 0.0081

COMPARE was then rerun with the DO data to check that the results were as expected and they were.

(See 2008-07-dox-comp2.xls.)

The same DO calibration was applied to the edited CTD files.

The above comparison was done using temporary files before the nutrients were ready, so a test comparison was done to ensure that the results were the same when the final files were used, and they were.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: Both conductivity sensors have been recalibrated since last use.

2. Dissolved Oxygen – This sensor was used for 2007-46 and 2007-43 in May and June 2007.

3. Pressure – This sensor has been drifting significantly in recent years. When used in Oct. 2007 an offset of +6.3db was found appropriate.

Historic ranges –Local climatology is not available.

Repeat Casts – There were no repeat casts.

Comparison of Nearby Casts – Local variability was high, so this did not prove useful.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. A further correction will be applied to at least partly correct for response time. To do this we compare downcast data to bottle data from the same pressure.

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When outliers are excluded, the differences were quite flat when plotted against pressure, but not quite so flat against DO. There are few low values, so the fit is quite arbitrary, depending on what points are included. The fit against pressure looks a little more convincing with CTD DO being high by between 0.026 and 0.035 depending on the choice of what to include. (See 2008-07-dox-comp3.xls.)

When used during 2008-46 the final correction was a simple offset of -0.05ml/l; there were some complicated structures sampled during that cruise and some bottle stops were rather short. During 2007-43 a pressure-dependent fit was selected with corrections of about -0.12ml/l at the surface and ~0 at 300db; many bottle stops were short during that cruise.

The thinned files were recalibrated by subtracting 0.03ml/l and the comparison was rerun. That showed that the results were good. (See 2008-07-comp4.xls.)

Recalibration using file 2008-07-recal2.ccf was applied to the downcast files only to subtract 0.030ml/l. (Output: COR2)

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added; REORDER was run to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

- ±0.3ml/l from 0 –100db

- ±0.2ml/l below 100db

The Standards Check routine was run and no problems were found. The final files were named CTD.

A header check turned up no problems.

Profile plots were made and no problems were found.

The track plot looks ok.  The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The values ranged from 80% to 115%, with most between 85% and 100%; there is no obvious temporal or spatial pattern in the distribution of values.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. The following comments were entered:

Analysis methods:

-----------------

Oxygen samples are analyzed on an automated Winkler titration system

following the procedures of Carpenter (1965). Samples were analyzed at

sea using an automated titration system consisting of a Brinkman Dosimat

(model 665) and a PC 900 Colorimeter.

For details on precision see file 2008-07-DO-dups.xls.

Salinity samples were collected in glass bottles and analyzed on a

Guildline model 8400B Autosal. Salinometers are standardized with

IAPSO standard  seawater.

Nutrient samples were collected in plastic tubes and analyzed fresh at sea

using a Technicon AAII autoanalyzer following methods described in

Barwell-Clarke and Whitney (1996). For details on nutrient duplicates and

precision see file QF 2008-07nuts.xls.

References:

1. Carpenter, J.H. 1965. The Chesapeake Bay Institute Technique for the

Winkler Dissolved Oxygen Method. Limmnol. & Oceanogr., 10: 141-143.

2. J. Barwell-Clarke and F. Whitney. 1996. Institute of Ocean Sciences

Nutrient Methods and Analysis. Canadian Technical Report of Hydrography

and Ocean Sciences, No. 182, 43 pp.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### 24. Thermosalinograph Data

a.) Checking calibrations

There were 2 identical files containing TSG data. All parameters were entered correctly. A copy was saved as 2008-07-TSG.con.

The history of the T/S sensors was obtained; 2008-01 is the only cruise processed so far using that equipment since it was last recalibrated.

There was no secondary thermistor at the inlet and no note of loop sampling. There was no flow meter.

There is a note in the log that the TSG shut down sometime on the 27th of March, though it was hoped that only the display was shut down.

b.) Converting to IOS Headers, adding position headers and time channels, preliminary checks

The data were converted to CNV files using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Salinity:T0:C0 and Time Julian and then converted to IOS HEADER format (*.IOS). The Latitude and Longitude were unavailable. There is a 1-day gap between the two files, so the problem noted in the log did lead to loss of data.

4 text files containing ship positions and times were obtained. These were opened in EXCEL, unneeded lines were removed, and positions were calculated in decimal format; longitude signs were changed. The first 3 files were combined since they correspond to the first of the TSG files and that was saved as 2008-07-0001.csv, while the last text file was saved as 2008-07-0002.csv. Those 2 files were then converted to IOS HEADER files (ios1). The times in the text files are 1 hour later than the CTD log entries for positions that match. The time was changed and Julian time was calculated (ios2). The added channel name was changed to Time:Julian using Header Edit (*.hdr2). The various IOS and HDR2 files were combined into files (2008-07-0001.cmb and 2008-07-0002.cmb, respectively. Those files were then merged on the Time:Julian channel (2008-07-*.mrg.) There are many records with no positions, but a track plot shows a reasonable amount of data available for all areas.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add time and date channels in IOS SHELL format and the output files were named *.ATC. Track plots were made and look reasonable.

Time-series plots were produced.

CTDEDIT was used to clean spikes in temperature, salinity and conductivity for 4 records: 3783, 3785, 5049, 5050.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within .3db of 4db and exported to a spreadsheet 2008-07-ctd-surf4.xls. The TSG files were opened in EXCEL, median and standard deviations (over 2minutes) were calculated for temperature and salinity, and the file was then reduced to the times when CTDs were run. Those files were combined in a spreadsheet (2008-07-ctd-tsg-comp.xls). Data were removed where there were no TSG data available.

This spreadsheet will also be used in step (e) to compare temperature and salinity.

d.) Alignment check

Recent uses of this equipment showed no alignment problems. There are variations in alignment, but they are not systematic. This step was skipped.

e.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 There was no intake thermistor.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity and temperature channels for the CTD and the TSG. There were 50 casts that could be used. Graphs were prepared comparing the TSG temperature and salinity with those of the CTD.

The temperature differences were plotted against standard deviations in the TSG temperature; the variability was extremely high and did not show much relationship to the standard deviation. (There is little overall variability in temperature through this cruise from either the CTD or TSG.)  Early in the cruise the differences were on the order 0.2 to 0.3, but then the differences grow, but not in a regular way. The TSG was high by 0.46C˚ on average and the median difference was 0.38C˚.

For salinity there is even more variability and no useful pattern emerges from the plot of differences versus standard deviation in the TSG data. The TSG salinity shows a lot more variability than the CTD. The TSG salinity was lower than the CTD by 0.67 units and the median was lower by 0.43 units.

When casts in Queen Charlotte Sound were studied there were smaller differences in temperature (an average of 0.42C˚ using 5 stations and 0.28C˚ if one is rejected). However, the salinity differences were very large, ranging from 1 to 4.6psu, with an average of ~2psu. So well-mixed waters still show a large difference in salinity. (See 2008-07-CTD-tsg-comp.xls)

- Loop Bottle Comparisons There were no loop bottles.

- History This instrument was used on the TULLY during 2008-01. At that time a correction of -0.19C˚ was applied to most files and a temperature-dependent correction to others, with a result that would be about 0.16C˚ at 7˚C. The salinity was corrected by adding 0.055. When this instrument was used on the VECTOR in April 2006 a correction of -0.20C˚ was applied. The salinity correction is not relevant since the conductivity sensor has been recalibrated twice since then.

Discussion

The temperature error is at least close to what we expect, but the salinity error is very large. We don’t have flow rate information so it is possible that there may have been flow problems. There have been problems in the past with bubbles in the system. That might affect salinity more than temperature. Part of the problem may be that the match of positions is not perfect. We found an error of 1 hour, but it is possible that the error is slightly higher or lower – a few minutes difference could lead to larger errors in areas of great variability. However, the errors in salinity in Queen Charlotte Sound do not fit that explanation since the waters are quite well-mixed.

The only calibration information available are history and the comparison to the CTD. The latter data are too confusing to justify recalibration. The temperature correction for the VECTOR in April 2006 was slightly higher than the correction for the TULLY in January 2008. A correction of -0.20C˚ to temperature and +0.055 to the salinity will be applied. The following note will be entered in the headers:

NMEA positions were not recorded in the thermosalinograph record.

Positions were obtained from separate files with times and positions, but

were not available for all records.

Calibration was based on the results of 2008-01 when this instrument was

used on the TULLY and 2006-11 when it was used on the VECTOR.

Comparisons with CTD data suggest that there are much larger errors for

this data, but the results were highly variable and could not be used for

recalibration. There was no loop samples, intake thermistor or flow meter.

g.) Recalibration

CALIBRATE was used to apply equation to apply an offset of -0.2 to the Temperature:Primary channel and  +0.055 to the Salinity:T0:C0 channel.

h.) Preparing Final Files

REMOVE was used to remove channels Scan_Number, Conductivity:Primary and Flag.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Jan08 | Factory “ |  |  |
| Conductivity | 2399 | 15jan08 | “ |  |  |
| Secondary Temp. | 4700 | 16Jan08 | “ |  |  |
| Secondary Cond. | 3184 | 15Jan08 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS | 5MAR08 | IOS |
| SBE 43 DO sensor | 1176 | 14Feb2007 | Factory |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/2487       Cruise ID#:	2008-07

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 01/12/07 | Factory |  |  |
| Conductivity | 2487 | 01/12/07 | “ |  |  |
| Flow Meter | ? | ? |  |  |  |