## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2008-09

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Aquaculture

Area: Broughton Archipelago

Chief Scientist: Galbraith M.

Platform: Princeton

Date: 12 February 2008 – 15 February 2008

Processed by: Germaine Gatien

Date of Processing: 15 May 2008 – 19 May 2008

Number of original CTD casts: 7

Number of casts processed: 4 (no useful data in 3 files.)

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#4132) was used.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available. The times in the log book are in Pacific Standard Time. To match the headers in the archived files add 8 hours to the log times.

The CTD soaked for 1 to 1.5 minutes for these casts; study of this CTD during 2007-11 suggests that a longer soak (at least 2 minutes) may lead to better salinity data.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors can be as high as 0.05units in areas of high gradients. However, the temperature gradients are fairly low for this cruise and the descent rate was quite steady, so such errors are likely to be smaller.

No calibration sampling was available for this cruise and no information was available for any other cruise since the instrument was last calibrated.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard. The file names are consecutive, so a match was made to the CTD casts in the log sheets, assuming they are in the same order. According to the log book there are no data for events #1, 9 and 13.

The dates in the log book are in Pacific Standard Time.

Configuration file 4129.con was saved as 2008-09.ctd.con. The parameters were confirmed by Bon van Hardenberg.

- Conversion of Raw Data

- The raw data were converted using conversion file 2008-09-ctd.con. There are 7 files with consecutive names and the times in the headers confirm that they correspond to events #1, 4, 5, 8, 9, 12 and 13. The casts were checked and there is no useful data for events #1, 9 and 13, which agrees with notes on the log sheets. The file names for the other files were changed to match the log records and time. Plots were made of those 4 casts and the data looks ok, though 2 of the upcasts stop prematurely.

- FILTER

Tests were run to ensure that the recommended settings worked well for this data.

The temperature channel was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the conductivity. The alignment of T with C is improved by this step. The pressure was filtered with a time constant of 1 second to increase the pressure resolution. The pressure looks slightly smoother after this step, but was quite smooth already.

- CELLTM

The manufacturer’s recommendation is that CELLTM be run for this instrument. They suggest values like (α = 0.04, 1/β = 1). Tests were run using a variety of settings and any changes worsened the correspondence of up and downcasts on a T-S surface. The temperature ranges were very small for this data, so it is likely that there is no error to correct. This step will be skipped.

- DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets and in the SeaBird headers, as well as a list of corrected positions and station names prepared by Moira Galbraith. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to insert details.

8. Checking Headers

A track plot, header check and cross-reference listing were produced and no errors were found.

Track plots were added to the end of this report.

The surface check was run. All 4 surface values were 0 with near-zero conductivity values. Two upcasts have records with 0 pressure values on the upcast and the conductivity is very close to zero for those two. There is no evidence that the pressure needs adjusting.

9. Test Plots and Fix Time

ADD TIME CHANNEL was used to add 8 hours to the start times so they are in UTC.

10. SHIFT

Initial checks show some noise in the salinity, and some evidence that the conductivity is leading the temperature very slightly. For other CTD19+s a shift of +2.25 records proved effective in removing salinity spikes and instabilities in T-S plots. That setting proved effective for these data as well.

Temperature was shifted by +2.25 records followed by a recalculation of salinity for all casts.

11. DELETE

Before running DELETE, CLIP was run to remove the first 200 records to simplify editing, since those records were in the soak period.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

12. CTDEDIT

CTDEDIT was used to remove a few bad records and to smooth salinity for most casts. All casts needed a little editing. Comments were added to the headers indicating what editing was done.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and no problems noted. No further editing will be done.

14. Intercomparisons

COMPARE – The only salinity sampling was from the surface, where the gradient is too large to do a reasonable comparison with the CTD.

Previous Use of CTD – The instrument has not been used since its latest recalibration.

Repeat Casts – There were no repeat casts.

Historic ranges – No local climatology is available.

15. CALIBRATION

No recalibration was applied due to lack of information.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors can be as high as 0.05 units in areas of high gradients.

For this cruise the temperature gradients were small and the descent rate

fairly steady, so the errors should not be that large.

No calibration sampling was available for this cruise and no information was available for any     other cruise since the instrument was last calibrated.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2008-09

Dates: 	Start:	12 February 2008	    End:	15 February 2008

Location: Broughton Archipelago                 Vessel:    Princeton

Party Chief: Galbraith M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SeacatPlus | 4132 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/SEACATPLUS v1.6b / 4132

Cruise ID#:	2008-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4132 | 5Jan08 | Factory |  |  |
| Conductivity | 4132 | 5Jan08 | Factory |  |  |
| Pressure Sensor | 4132 | 5Jan08 | Factory |  |  |