## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-Sep-2017 | Corrected MISSION metadata field in header. R.H. |

## PROCESSING NOTES

Cruise: 2008-10

Agency: OSD

Location: WCVI

Project: LaPerouse

Party Chief: Yelland D.

Platform: John P. Tully

Date: August 31, 2008 – September 8, 2008

Processed by: Germaine Gatien

Date of Processing: 8 February 2009 – 25 February 2009

Number of original CTD casts:  68	Number of CTD casts processed:  68

Number of bottle casts: 		61	Number of bottle casts processed: 61

Number of original TSG files: 	 4	Number of TSG files processed:     4

# INSTRUMENT SUMMARY

A Sea-Bird Model SBE 911+ CTDs (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), SBE 43 DO sensor (#1176 - on the primary pump), a Seapoint Fluorometer (#2228) with a 10X cable (on the secondary pump), a Biospherical QSP-400 PAR sensor (#4656) and an altimeter (#1252). The deck unit was a model 911+ (#0424) and the logging computer was an HP Compaq. Seasave v7.16 was used.

A thermosalinograph (Sea-Bird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P) and a flow meter; temperature sensor #4652 was mounted at the intake.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order and there was a post-cruise report from the chief scientist. Some data were missing from the set originally provided. Eventually all of the files were found, but there are clearly problems in how the data are downloaded from the acquisition system. This is a time-consuming and recurring problem.

Niskin Bottle #7 failed to close after cast #1. According to the log the solenoid was bad and no spare pylon could be found.

Study of one cast during which data acquisition started before the initial lowering to 10db indicates that the new deployment method did not corrupt the data from the top section of the full cast. However, this cast was in open water where ship drift is likely to mean different waters were being sampled, and there was a wait of more than 15s at the top before beginning the full cast, which is probably enough for equilibration to occur. Until evidence is found to suggest otherwise, it is recommended that there be a wait of at least 20s at the top before the full cast in conditions where there may not be significant ship drift.

The scatter in the fit of extracted chlorophyll against the SBE Fluorescence is higher than usual. No evidence could be found of problems with the chlorophyll analysis. An unusual fit was also noted with this fluorometer for a few casts during 2008-61, although the particulars were quite different. One suggestion is that the CTD may have been mounted incorrectly in the rosette leading to reduced pump efficiency.

No recalibration was applied to salinity due to doubts about the salinometer. The analyst noted some problems with the instrument at about the time these analyses were run and the results for other cruises analyzed at that time have given results that are not trusted. Further, the study of duplicates showed more variability than usual. When these sensors are next calibrated at the factory the results will be studied and recalibration may be applied later.

February 25, 2008 NOTE: A post-cruise factory report shows the primary salinity was low by no more than 0.0012 due to conductivity and temperature sensor drifts. No recalibration is necessary.

A new algorithm was used for the conversion of the SBE Dissolved Oxygen sensor data. This is expected to produce more reliable data below 1000db. The results look encouraging from the first few cruises for which this method was used.

The dissolved oxygen data in the CTD files should be considered:

•	±0.6 ml/l from 0 – 125db

•	±0.2 ml/l from 125 – 400db

•	±0.1 ml/l below 400db

There was sampling of hypoxic waters in Saanich Inlet, so the effect on the dissolved oxygen sensor was studied. The evidence is not clear, but it seems that the sensor was not affected by the sampling.

There are 4 thermosalinograph files. The two large ones lacked positions, but a file was available with the ship’s positions and times; this was used to patch positions into the TSG files but there are some gaps in those channels. The intake temperature looked bad with values much higher than the lab temperature throughout the first file and part of the second one. For those two files, a proxy intake temperature was created by recalibrating the lab temperature to remove the effect of heating in the loop. The salinity was poor in the 2nd file and the comparison with bottles and CTD was greatly hampered by that. There were few points of comparison in the other files. What little data were available fit reasonably well with the history of the sensor, but when it is next recalibrated at the factory these results should be revisited.

A TSG flow test was done and the results show that increasing the flow from a setting of 0.6 to 0.9 results in a significantly shorter passage time in the loop. It was too short to conclude if data quality was affected.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The data for 7 casts were missing when this job was started, but were later found. There were no bottle files for those casts, so this does not affect the comparisons done before they were found.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as a post-cruise report from the Chief Scientist. A few problems were noted in rosette bottle firing and one instance of a large difference between temperature and salinity channel pairs.

Extracted chlorophyll, salinity, nutrients and NH4 bottle data were obtained in spreadsheet format.

The titrated dissolved oxygen files were provided in individual files without flag channels or comments, though there are comments on the rosette sheet.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The configuration files were all the same and were all correct except that the parameters were missing for the new algorithm for the SBE Dissolved Oxygen sensor. Those were entered, picking the same value for E (3.85*e-02) as was determined in studies of data from cruise 2008-27. The edited file was saved as 2008-10-ctd.con.

Some data files were missing from the set originally provided to the processor. All were eventually found by the chief scientist, but he is frustrated by the shipboard problems preventing good back ups of data from the acquisition system. As well as the danger of losing data, much time is taken up in searching for missing data.

Data were converted using the configuration file 2008-10-ctd.con. The file names for casts 101 to 107 had an extra zero in them, so those were fixed. Also the cast identified as #107 should be #106 according to the log, so the name was changed. The file named #11 should be #13, so was also renamed.

A few casts were examined and all expected channels are present and the data look ok. As usual the two temperature channels are in reasonable agreement on the downcast. The upcast data are much noisier so there are significant differences. Conductivity is similar.

The new deployment method with a soak at the 10db, a rise to the surface and then a complete full cast, was used for this cruise. There is some concern that if there is not a short wait at the top before the full cast, the data might be of poorer quality in quiet waters where the ship would not drift into “unmixed” waters. In most cases data acquisition does not start before the full cast, but for cast #39 it was on throughout the soak period. It was hoped this might be useful to study the effect of the new approach. However, there was a 20s wait which is probably enough to let things settle down, and the cast was not in a quiet area. Nonetheless, the data were studied – see section 10 for results.

Rosette files were converted using a start time of -5s and duration of 10s. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT. The file names for casts 101 to 107 were corrected in the same ways as noted above for the CNV files. All BOT files were plotted. The only significant outliers were in cast #22. CTDEDIT was used to clean a few primary salinity points in the file and to enter a comment about the editing in the header of the file. The edited file was copied to BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on a few casts using settings (α = 0.01, β=7), (0.02, 7), (0.03, 7), (0.02, 9), (0.03, 9) and (0.0245, 9.5) to see what settings looked best for this cruise. The best choice overall was (0.0245, 9.5) for the primary channel and (0.03, 9) for the secondary. CELLTM was run on all casts using those values.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A few casts were plotted to check for agreement between the pairs of T and C sensors. The casts are mostly quite shallow and the deep ones had a very noisy descent rate, so the plots were really too noisy to make useful estimates. However, the differences do not look large overall, falling reasonably close to the results of 2008-27 which preceded it and had the same equipment. The salinity differences are slightly pressure dependent. The pressure dependence in the temperature differences were not seen in this cruise, but that was only seen below 2000db during 2008-27, and none of these casts are that deep.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and checked. Random checks of times, positions and station names were made using the cross-reference listing. There were a few errors in station names; those were corrected. The files named 72 and 107 should be 73 and 106 according to the log book. Those were corrected.

The cruise track was plotted. No problems were found.

The average surface pressure is 1.6db with values as low as 0.1; this is low for the Tully and a few other recent cruises using this pressure sensor have shown similar results, so it looks like it is time to increase the pressure offset a little. At the recalibration stage an offset of 0.2db will be applied and in future a setting of +6.7db will be used for this sensor.

The altimeter readings from the headers of the CLN and BOT files were exported to spreadsheets and all casts were checked. Plots were made and the log book was checked. The algorithm worked well where the CTD got close to the bottom, except for casts #2 and 82. It recorded erroneous low values for cast #82, which does not appear to have gotten within 15m of the bottom, so the header was removed from both CTD and bottle files for that cast. For cast #2 the reading was noisy at the bottom and the algorithm result looks a little too low, so a text editor was used to change the value from 1.9db to 5db based on profiles and log book information, and the header note was changed to explain how it was derived.

##### Cast #39 study

##### There has been some concern that the habit of lowering the CTD to 10db, returning to the surface and then immediately lowering the CTD for the full cast, might lead to poor data in the top 10db. It is unlikely this will be a problem in open waters since ship drift would generally ensure that different waters are sampled than were stirred up by the CTD coming up to the surface. But in quiet areas there could be a problem and there has been some suggestion of trouble in CTD data from such areas, with T-S data sometimes looking unstable, though that problem could be from ship effects such as prop wash. Data acquisition does not usually start until the beginning of the full cast, but during this cruise the whole process was accidentally recorded during cast #39 (and a few other casts). It was thought that this might provide an opportunity to study whether there is a significant problem. Unfortunately, this was an open-water cast so it is not likely to be useful, and the wait between the upcast and downcast was about 15s which may be long enough for the water to settle to ambient conditions.

Plots were made of the 3 sections (down, up, down) between 2 and 12db to see if anything could be learned. The first downcast had very noisy data. At the top that is probably because the pump had just come on and the descent rate was low; the pump status is 1 throughout, but it is likely that it had just been turned on. At 8db there was a severe slowing of the descent and a clear shed wake signal is seen. During the upcast there is some suggestion that the CTD was going through the wake from the downcast, a possibility since there was no wait at the bottom. Upcasts do tend to be noisy anyway due to passing through the wake of the rosette package, so this is not clear. The 2nd downcast looks quite smooth with just a little noise around 5db. The upcast section is saltier and colder than both downcasts presumably because it drags water with it from below.

This study does not show any clear effects from the new deployment approach, but whether that is because there was a short wait at the top, or because the ship drifted into different water or because there is never a problem is unknown. It would be useful to repeat this experiment in calm conditions and with little drift.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The BOT files were then bin-averaged.

SALINITY

Salinity was analyzed at IOS on an Autosal including 19 duplicates. The Autosal data were delivered in spreadsheet format. The rosette samples were separated and saved as 2008-10-sal.csv. Loop samples were copied to 2008-10-sal-loop.csv and then removed from the main file.

The spreadsheet was simplified, and duplicates were found and copied to a separate file, 2008-10-sal-duplicates.xls. The averages and differences were found. There were 3 duplicates with differences >0.01, and only 9 had differences <0.0025. These differences seem quite large. Plots of differences against sampling depth show some reduction in deeper waters, but there are some large differences even at ~2000m. This does not lend great confidence in the Autosal results. For now, the averages will be used for the creation of SAL files, but the duplicate values will be entered as header comments for the 5 cases with differences >0.003.

Using all 19 duplicates Sp = 0.0040 and excluding 6 outliers with differences >0.003, Sp=0.0012 where

Sp = Square Root (sum of squares of differences / 2*number of pairs)

For 2008-26 Sp=0.0012 using all 7 pairs and 0.0005 when 1 pair were excluded.

For 2008-27 Sp=0.0027 using all 10 pairs and 0.0008 if 3 out of 10 duplicates were excluded. For that cruise 3 of the 4 largest differences, the “a” and “b” samples were analyzed in different sessions.

Both those cruises were run under similar conditions to this one, and the salinity from 2008-27 was analyzed just before this data. Two other cruises were analyzed right after this one and for 2008-51 the differences between CTD and Autosal results were surprisingly large, raising doubts about the analysis. There were no duplicates for that cruise. Another piece of evidence is that when multiple bottles were fired at 2000db during 2008-27, there was a lot of variability in the results, with a standard deviation of 0.005. While this could indicate a problem with sampling procedures, those are not known to have changed recently, whereas there have been concerns about the Autosal. There have been some repairs since these data were run.

There is not a long history of taking salinity duplicates, so this may just be illustrating the limits of the Autosal. But it is hoped that the large differences for this cruise are a reflection of problems with the instrument and that results will improve. (See 2008-10-sal-duplicates.xls.)

The spreadsheet was edited to enter the average values for the duplicates, add “f” flag to those, and for cases of differences >0.003 a “c” flag and comment with the 2 values. Event #s were added to the file and it was then converted to individual SAL files.

DISSOLVED OXGYEN

Dissolved oxygen files (*.oxy) were provided. They lacked flag and comment channels. The ADD CHANNEL routine was run to add those, and then the rosette sheets were reviewed to determine what samples needed flags and comments. The OXY file for cast #31 contained only 1 value, whereas the rosette sheet lists 7 bottles sampled, one of them with duplicates. The data from the rosette sheet and the O2 file were used to recreate a proper ADD file, though the flask volumes, blank and standard values were not known, so were left blank. Those fields are not required for the bottle comparison.

File #25 contained errors in sample numbers; the rosette sheet indicates what needs to be corrected. The ADD file was corrected. The second reading for sample #63 looks like 2.272 on the rosette sheet, but is entered into the file as 2.772 which actually makes much more sense. It is assumed that the file value is correct since the two original values in the O2 file are quite close. No flag was entered but the discrepancy was mentioned in a header comment.

There were duplicates but no averaging had been applied. A spreadsheet was prepared, 2008-10-DO-duplicates.xls, in which the individual values were recorded, averages determined and differences calculated to see if there are outliers. Four outliers were identified and the values were checked against the rosette sheets and oxy files and no incorrect entries were found. There is an explanation for the outlier of cast #5 since there is a note about a bubble in the 1st sample, so the second reading was entered in the ADD file. For all other duplicates the average values were entered into the ADD files with flag “f”, as well as a “c” flag for those that are outliers. For the latter the original values were entered in a comment.

Using all duplicates Sp = 0.075 and excluding 4 outliers Sp=0.022 where

Sp = Square Root (sum of squares of differences / 2*number of pairs)

For 2008-26 and 2008-27 the results were Sp=0.090 and 0.085 when all data were included, so these results look good.

All files were examined since many have records out of sample # order and there were a few extraneous lines that were removed.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2008-10nuts.xls which was simplified and reordered on sample number and saved as 2008-10nuts.csv. Extraneous headers were removed, header names were changed to standard format and the file was then converted to individual NUT files. Sample #30 has a comment “sample unknown sample”. Comments about 3 bottles that did not fire were removed since this is already in the CHL files as a general comment.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file QF2008-10CHL.xls which included a report on precision. The file was edited to remove extraneous lines and columns, header names were changed to standard format, the file was sorted on sample number, and saved as QF2008-10chl.csv. That file was converted to individual CHL files. There were loop samples which were saved separately as QF2008-10 CHL-loop.csv.

NH4

NH4 data were obtained in file QF2008-10NH4.xls which included a report on precisions. The file was edited to remove extraneous lines and columns, header names were changed to standard format, the file was sorted on sample number, and saved as QF2008-10NH4.csv. That file was converted into individual NH4 files. Comments about 3 bottles that did not fire were removed from the NH4 files, since this is already in the CHL files as a general comment so will appear in the headers.

The SAL, CHL, ADD, NUT and NH4 files were merged with CST files in five steps. (Output: MRG1, MRG2, MRG3, MRG4 and MRG5), MRG5 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

Data were exported to a spreadsheet and compared to the rosette sheets to ensure all expected data are present. No problems were found, other than the fact the Niskin #7 did not close on any but the first cast; this problem was noted in the log.

##### Compare

Salinity

COMPARE was run. The primary salinity is low by from 0.0065 to 0.008 depending on what data are included and the secondary is low by from 0.007 to 0.01. All the samples from cast #24 and many from cast #1 are minor outliers. The following outliers were investigated:

- Cast #1, Samples #5-11, especially #9 – Fairly large difference for sample #9 given the depth and the CTD salinity data were not noisy. Most of the differences from this cast are higher than expected but for some samples there are possible explanations like noisy CTD data or being close to the surface. However, the analyst reported problems with sample #11 and the duplicates differed a lot for sample #5. So given definite problems with 5, 9 and 11 and doubts about others, it was decided to assign “c” flags to samples 5-11.

- Cast #24, Sample #58 – Fairly large difference (0.085) and there is no indication on the rosette sheet that there was a sample taken and sample label indicated it was from station LG05, not LC01. There is no sample for LC01, cast #58, so I think this is a confusion of sample number and station number. This salinity value was moved to cast #58, with a “c” flag and note explaining it.

- Cast #24, Sample #61 – Moderate outlier, but CTD data are unusually noisy despite long enough wait and little motion in CTD. This is a 5db sample and probably in active mixing zone or affected by ship effects. No flag was assigned.

- Cast #33, Sample #89- This was a severe outlier from the surface bottle of cast #33. The CTD salinity is much lower than the bottle value, which looks like that seen around 30db. There were spikes in the secondary temperature at about 27db, which might be evidence of a misfire. The nutrients and chlorophyll values also look way out of line. Sample #33 was flagged “d” for all analyses.

- All other outliers were fairly minor and most were at the surface or the CTD salinity had a fairly high standard deviation.

When those corrections were made it was found that both salinity channels were low, the primary by about ~0.006 and the secondary by ~0.007 with a standard deviation in the fit of about 0.009 for each. When only data below 200db are used and one point is excluded, the fits are quite flat and the primary is low by 0.007 and the secondary by 0.009 with standard deviations of 0.002 and 0.003, respectively. There is a lot of variability, not unexpected around 50db and from 100 to 150 in haloclines. When the 5db samples are excluded the fit of differences against salinity is quite flat suggesting that non-linearity in the Autosal is not an issue. There is no obvious time-dependence. However, the differences are larger than expected; together with the duplicate study this raises some doubts about the reliability of the results. These data were analyzed in October and the analyst noted problems with the Autosal at about that time. Repairs were done after these data were analyzed. So there are reasons to doubt the results. Hence recalibration is not recommended at this time. (See 2008-10-sal-comp1.xls.)

Dissolved Oxygen

COMPARE was run using pressure as the reference channel. The best fit was against differences between the sensor and bottles. When outliers were removed based on residuals, the fit for all data was:

CTD-BOT = 1.0722 DOX-CTD + 0.0106

Many of the outliers are associated with large standard deviations in the CTD data including many from around 10db where the sensor generally does not do well at matching bottles.

The CTD sampled hypoxic waters in cast #1 (first 3 bottles have near-zero values). In the past sampling hypoxic waters was found to change the calibration of the sensor, which then took about 4 casts to return to the expected calibration. Sea-Bird have made some changes to the sensors to deal with that problem and there has been no recent evidence of a problem, but there has also been little anoxic sampling. This cruise provides an opportunity to study the effect in some detail.

The fit for cast #1 is difficult to assess with one major outlier and 3 anoxic samples and 1 very low one, so we don’t expect the sensor to do very well. The slope of the fit is 1.04 if we include all but the major outlier at 10db, or 1.06 if we force it through the origin and 1.05 if we include only 3 of the top 4 bottles with the offset left free (the one not included is a 10db bottle that is an obvious outlier). Cast #2 has only a few points and the slope is ~1.05. For casts #3 and 4 the slopes are ~1.078 and 1.07. Are these signs of hypoxia affecting the sensor? That is not clear, but if it is a factor the effect appears to wear off quickly.

When the first 2 casts are excluded the fit is not very different:

CTD-BOT = 1.0727 DOX-CTD + 0.0095

Further when all bottles are plotted together and those from casts 1 and 2 excluded from the fit so they show up in red on the plot, it becomes apparent that the scatter is not systematic. For cast #2 the sample from Niskin #2 stands out slightly from the pack but the other 3 do not stand out at all. For cast #1 there is the one clear outlier at 10db. The bottle at 125db is within the pack and the other bottles are minor outliers falling on either side of the trendline with no discernable pattern. This suggests that the trendline for cast #1 is different only because the sensor is not as precise at very low values and possibly because there are no data with values >4.1mL/L. That cast was in Saanich Inlet whereas all others were from the WCVI region.

So there seems no justification for using a different calibration for the first cast. The following note was placed in the headers of the CTD and bottle file to indicate that DO values may not be as reliable as for other casts:

CTD Dissolved Oxygen Data

The CTD sampled anoxic waters during this cast which may possibly have caused a shift in

calibration for the DO sensor. This is hard to judge because of the limited range of DO

sampled during this cast. If there was any effect it was short-lived, as the next cast

does not stand out in the general comparison of CTD sensor and bottle DO.

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity. The only outliers were from bottles that had already been flagged.

Fluorescence

COMPARE was run using the CTD CHL and the Titrated Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD CHL. There was more variability than usual. Casts #33-55 showed a good correspondence between CHL and FL, but before and after that the fluorescence is only about 40% of the CHL. Even when only data with fluorescence <2mg/l are included, the slope is very low. Choosing data from 5db, 10db and 20db only showed the same general picture.

Data were separated into days on which CHL samples were analyzed. Plots were made of CHL and FL in the order that the analyses were run. This shows that the lowest ratios of CHL/FL were from data analyzed on Sept. 4, especially the first half of the data. However, this is not proof of problems with the chlorophyll, since the fluorescence is also very low early in the session, rising later. Nonetheless, we usually get a pretty good match when the CHL is low. This may mean the CHL values are too high. From the Sept. 5 analyses, the ratio of FL/CHL rises through the session whereas there does not seem to be any systematic variation in CHL and FL. For Sept. 6 there is a sharp rise, but there is a systematic variation as CHL falls. For Sept 7 there is a lot of variation, though the average is probably ok. The last day, Sept. 8 has a fairly tight FL/CHL ratio despite a lot of scatter in both FL and CHL. There is no clear evidence of a bad day or two. When a plot was made of differences against file pair number and points with CHL > 5 were excluded, then the analysis of Sept. 6 and 7 stand out as having differences averaging close to zero, but they also stand out as having little high CHL; so it is likely that the values included in the fit are also lower, on average, for those two days.

Comparisons were done to 2008-27 and 2007-42 but they are all so different with different ranges of CHL and different sampling patterns. There is no clear evidence except that the scatter is larger than usual. Either there are some CHL values that are too large, or the fluorometer was not working well and giving low values. Cruise 2008-61 using the same fluorometer has a section in the northern part of the Strait of Georgia where the fluorometer gave very low values. While that may partly be due to a very high gradient in CHL so that the sensor was in low CHL water a metre below that Niskin bottle, that explanation does not account for all the low values. (See CHL-Study by day of analysis.xls. and 2008-10-chl-fluor-comp.xls.)

One other possible explanation for the poor fit is if the CTD was mounted incorrectly in the rosette, as was discovered to be the case in early January 2009. If that affected the pump efficiency or otherwise affected flow to the fluorometer, it might explain these results. It is not known what cruises used the CTD in that configuration.

##### 13. Shift

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset.  This is a rough estimate as the upcast data are noisy. The usual shift of +24 records (1s) was found to be appropriate. SHIFT was run applying a shift of +24 records to the fluorescence channel.

Conductivity

Tests were run on 3 casts using shifts between -1s and +1s and T-S plots were prepared to compare the results. A setting of -0.3s worked best overall for the primary sensor; +0.2s was best for the secondary sensor.

All casts were put through two runs of SHIFT using those settings. (Output: *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. There were few distinctive features to aid this judgment. Values of +40 to +80 were tried and +50 looked best; in the past the best choice for this sensor was usually +110 or +120. Perhaps using the new algorithm affects this step.

SHIFT was run using +50 records for all casts.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings

##### DETAILED EDITING

It is not obvious which sensor pair to archive. The primary salinity is a little closer to the bottles, but the comparison is not trusted. The primary sensor has been used for all 3 previous uses of this equipment, including 2008-27 which preceded this one and used all the same equipment. A few files were examined in detail and no significant differences in noise level or spikes were seen. So the primary sensors were selected.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. All casts required editing.

In the course of editing it was found that the DEL files for casts #8, 39, 45, 47, 51, 56, 68-74, 79, 97, 101 & 102 included data from the initial soak period, so the data from the before the full cast were removed; All files were then put through DELETE again.

##### Initial Recalibration

No recalibration will be applied to salinity at this point. When the sensor is next recalibrated at the factory or if comparisons from another cruise make the picture clearer, this can be revisited.

File 2008-10-ctd.ccf was prepared to add 0.2db to the pressure and to apply the following equation to the CTD Dissolved Oxygen channel:

CTD-BOT = 1.0727 DOX-CTD +0.0095

This was applied to SAM and MRGCLN2 bottle files.

COMPARE was then rerun. This showed that the DO correction was effective. (See 2008-10-sal-comp2 and 2008-10-dox-comp2.xls.)

The same DO and salinity calibration was applied to the edited downcast files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction can be applied to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Files were bin-averaged to 0.5m-bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When all data are included except for a few outliers, the differences vary from +0.03ml/l at DO=0 to +0.05ml/l at DO=7 and the fit against pressure is

DOX_BOT = DOX_CTD +0.00005 * Pressure - 0.0819
(See 2008-10-dox-comp3.xls.)

The thinned files were recalibrated using the above equation and the comparison was rerun. That showed that the results were good. (See 2008-10-comp4.xls.)

Recalibration using file 2008-10-recal2.ccf was applied to the downcast files only. (Output:COR2)

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. Two casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: Both sensors were used for 2008-26, 2008-41 and 2008-27. The value of the bottle comparisons were limited by either salinometer problems or very few bottles. The primary looked low by ~0.003 for 2 of them and was within 0.001 for another. The results for the secondary were highly variable, being high by 0.002 for 2 and low by 0.002 for another.

2. Dissolved Oxygen – Since the sensor was recalibrated in Feb. 2007 it has been used for 9 other cruises that have been processed. Of these 1 had very short bottle stops, 1 was in a shallow inland area with a limited DO range and another gave very confusing results. The results of the past 4 uses look quite close.

3. Pressure – This sensor is older and prone to drift. An offset of +6.5db has been used since March 2008, but for cruise 2008-61 in November, it was set to +6.7db.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were only a few minor excursions from those ranges with temperature slightly high for casts #70 (~30db) and 102 (60-65db), and salinity slightly low in parts of the top 25db of casts #41, 42 and 71. None of these excursions suggest systematic problems with the instrument.

Repeat Casts – There were no repeat casts.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts #1, 39-45, 60-78, 82-93 & 108 because the instrument was not mounted on the CTD for those casts.

A second SBE DO channel (with umol/kg units) was added.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

•	±0.6 ml/l from   0 –  125db

•	±0.2 ml/l from 125 –  400db

•	±0.1 ml/l below 400db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found.

The final files were named CTD.

The header check was rerun and one error was found and corrected; the routine was rerun and no further problems were found.

Profile plots were made and no problems were found.

The track plot looks ok.

The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface ranged from ~80% to 140% with most values between 100 and 110%. The lowest values were near the coast and mostly in the south, with the highest values fairly randomly distributed. These data are noisy but the values look reasonable overall.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

The PAR channel was removed from casts #1, 39-45, 60-78 & 82-93 because the instrument was not mounted on the CTD for those casts.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

Header Check was rerun and a problem was found with cast #73 – samples had not been merged because of confused cast #s. This was fixed and the routine was rerun and no further problems were found.

Checks were made against rosette log sheets and all expected samples appear to be present.

A cross-reference list was produced and no problems were found

##### 24. Thermosalinograph Data

The TSG data were provided in 4 HEX files.

a.) Checking calibrations

All calibration information was correct in the con file.

b.) The data were converted to CNV files using 2008-10-0001.con.

Those files were then converted to IOS SHELL format files *.ios.

The positions are missing from files 1 and 2, but are present for files 3 and 4. There is a file Combined RMC.xls which contains times, dates and positions for the ship’s track, so that file was simplified and saved as RMC-track.csv. The positions were changed to decimal format and Julian Time calculated. Sections were removed that contained no data and Julian time was calculated. The spreadsheet was converted to IOS SHELL format 2008-10-0001.pos and also saved as 2008-10-0002.pos. Those files were then merged with the *.ios files for files 1 and 2 with output named *.mrg. There are many gaps in positions in the output file. Rather than interpolate positions, the gaps were left.

ADD TIME CHANNEL was used to derive separate date and time channels from the Time:Julian.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Time-series plots were examined on-screen and confirmed that the secondary temperature looks very bad early in the track; part-way through the second file it starts to look better at 21:17 on Sept. 5th. The flow rate is ~0 just before the data improve, and there is a note in the log “increased flow thru TSG remote temp”. There is no change in the flow rate recorded on the TSG except for the short stop, so it is presumed the change was something that affected just the flow to the secondary thermistor, not the whole system.

A flow rate test was run at ~20:00 on Sept. 7th. The log notes that two settings were checked: 0.6 and 0.9.

The track plot looks fine; there are gaps as there were gaps in the file used to add positions.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within 0.3db of 4db and exported to a spreadsheet which was saved as 2008-10-ctd-tsg-comp.xls.

The TSG files were opened in EXCEL, median and standard deviations (over 2minutes) were calculated for temperature, salinity and fluorescence, and the file was then reduced to the times when CTDs were run. Those files were added to the CTD data in file 2008-10-ctd-tsg-comp.xls. The positions were compared and were very close with average differences <0.0001º and no difference was larger than 0.0002º. For many there was no position available in the TSG file.

This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T0   A study of differences between the intake and lab temperatures show that the intake thermistor was not giving good data through File 1 and most of File 2. Towards the end of File 2 the flow stops for a few scans and after that the differences look reasonable, though noisy. A flow test was done at the very end of the file. The flow rate had been ~1.1 early in the file but went up to ~1.2 at about 0800 on Sept. 6. Just before the test it was decreased to ~1.0, then up to 1.7, then down to about 1.1. When the flow rate was increased the differences were very noisy at first, but soon settled down. When the flow rate was reduced the differences were reduced and fairly quiet, but that section is too short to conclude much. During the time with highest flow rate it appears that the offset between the intake and lab temperatures was lower at about 8 scans and there is less heating in the loop. Standard deviations were used to determine the best offset. The following table summarizes the test results:

| Intake Temperature - Lab Temperature |  |  |  |
| --- | --- | --- | --- |
| Flow ~1 |  |  |  |
|  | no offset | offset 8 | offset 12 |
| avg | -0.1702 | -0.1873 | -0.1984 |
| std dev | 0.24568 | 0.15177 | 0.07431 |
| Flow ~1.6 |  |  |  |
| avg | -0.2293 | -0.1816 | -0.1510 |
| std dev | -0.22877 | 0.06269 | 0.15734 |

A study of file #4 when the flow rate was ~1.1 shows little noise when a 12scan offset is applied.                                The average difference was -0.236. For file #3 the data were noisier and the average difference is -0.213 without an offset and -0.218 with an offset, and the standard deviations are much smaller with the offset, and a quick comparison of peaks in the two temperature channels suggests that is a reasonable offset. No systematic study was done to test offsets for that file. So the temperature differences found were -0.20, -0.22 and -0.24C˚ for those three legs. While this might look like a steady drift, these are short sections of data. Moreover, the mean temperature will affect the warming in the loop. The most reliable section of data is probably the 4th file, so a fit of differences against intake temperature was examined:

T1-T0(offset) = 0.0056 * T1 – 0.3011

The average for (T1-T0(offset) for the good section of file 2, and files 3 and 4 are 13 ˚C , 14 ˚C and 12 ˚C. The fit above would predict differences of about 0.23, 0.22 and 0.24 C˚ respectively. That fits the observations for files #3 and #4, but is a little high for file #2. It is interesting that the short end section of File #2 (after the test) had an average of -0.23C˚ which does fit the pattern and that is the section with the lowest flow rate. In the section of this file where we can compare the two temperature channels the flow rate was mostly ~1.2, though it was lower earlier.  So it appears that the amount of heating in the loop is both flow-dependent and temperature dependent. For casts #3 and 4 this is not a problem as we have both temperature channels working, but for casts 1 and 2 we need some way to estimate the heating to create a proxy intake temperature. File #1 has an average lab temperature of 12.4˚C but the range is from 9.5 to 14.6˚C. The flow rate is 1.1 on average. The first file has slightly lower temperatures than the second and the flow is slightly lower, so we might anticipate slightly more heating in the loop, but there is not enough information to make an estimate of that. Probably the best we can do is to base an estimate on file #2, so about 0.20C˚ heating in the loop.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There were 56 casts that could be used.

If all data are included the TSG lab temperature is found to be high by 0.28 Cº and the intake temperature by 1.7Cº. The differences are very noisy and often very high until cast #52 when they change abruptly to believable values, though the last 4 casts look out of line as well. There is no note in the log about the TSG from that time. From cast #52 - #98 the lab temperature is higher than the CTD by an average of 0.185C˚. The intake temperature looks bad until cast #79; from #79 to #98 the intake temperature is lower than the CTD by an average of 0.013C˚ and when only the 5 with the lowest standard deviation in the temperature are used, the difference is only 0.005C˚. This implies heating in the loop by ~0.2C˚.

Using a similar approach for salinity, the TSG salinity is low by 0.377 using all data and by 0.35 when casts with lowest standard deviation in salinity are selected. The median value shows it low by 0.34.

The ratio of TSG fluorescence to CTD fluorescence ranges from 1.8 to 4.3 with an average of 2.4 and a median of 2.1. Excluding outliers based on standard deviation leads to a lower ratio ~2.1. Plotting TSG against CTD FL gives a trendline with slope ~2.2, or ~2.3 if it is forced through the origin. During 2007-26 that slope was ~2.5 and 1.6 for 2008-01 (for that cruise it was noted that the TSG fluorometer was quite unresponsive, possibly due to a film build-up). So it appears the fluorometer did perform as usual for this cruise. (See 2008-10-ctd-tsg-comp.xls.)

- Loop Bottle Comparisons  There were 10 loop bottles. The results of all analyses were combined in a single file 2008-10-loop-combined.csv. That was used to prepare file 2008-10-TSG-loop-comp.xls. For the 4 loop samples taken during CTD casts, the rosette bottle salinity and CHL were also included. The times and positions of loop samples were found in the log. The corresponding TSG fluorescence and salinity values (using a median over a 2-minute window) were found in the TSG files.

The TSG salinity was lower than the loop bottle salinity by an average of 0.44 if all samples are used and by 0.40 if the flow test sample and one other outlier are excluded. Early in the cruise, roughly the time of File #1 the salinity is low by 0.25, but later during File #2 it is low by 0.42. For the 5 CTD casts the loop samples and downcast CTD values agreed, on average, but the scatter was high. The downcast CTD salinity and loop samples are closer and the loop samples and rosette samples are extremely close.

The loop CHL samples are very close to the rosette CHL samples. The upcast CTD fluorescence data from the stop at about 5db are about 75% of the loop CHL, and about 82% if two CHL samples with duplicate differences >10% are excluded. The TSG fluorescence is about 1.5 times the loop CHL. (See 2008-10-TSG-loop-comp.xls.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in December 2007. They have been used on 3 other cruises that had calibration sampling.

  - During 2008-01 there was no intake thermistor and the lab temperature was estimated to be high by 0.19Cº. The salinity data were thought to be low by 0.055.

  - During 2008-32 the lab temperature was believed to be high by 0.228 Cº and the salinity data were bad.

  - During 2008-26 the lab temperature was considered high by 0.15 and salinity low by 0.125.

  - During 2008-26 the lab temperature was considered high by 0.17 and the salinity low by 0.17.

Conclusions

If we use only the data from cast #52 onwards and exclude a few outliers, the lab temperature is higher than that of the CTD by an average of 0.185Cº and the intake temperature is low by about 0.013. That implies that the lab temperature is high by about 0.20Cº. That is very close to the average difference found for file #2 for the period when the intake temperature was working properly. This is the best method we have for adjusting lab temperature to produce a proxy intake temperature for files #1 and #2. The estimate should be considered ±0.04Cº.

For salinity the TSG is lower than the loop salinity by ~0.44 using all data. This is a surprisingly large difference given the history of this instrument. However, examination of the data itself, suggests that there is some odd noise in the 2nd file (abrupt changes by as much as 0.8) and that is the source of most of the calibration information. While the TSG is lower than the CTD by 0.38 over the whole cruise, if we divide the data into early (file #1) and late (file #2), it is low by 0.25 and 0.42. Many of the loop samples also overlap with File #2 but the 3 from the time of File #1 show the salinity to be low by 0.735, 0.225 and 0.230. If we reject the first one which does seem way out of line with all the comparisons, then the TSG is low by an average of 0.228. The history of the instrument shows some drift and a difference of 0.23 looks quite reasonable. A difference of 0.4 does not. The salinity will be recalibrated by adding 0.023, but this should be revisited after a report arrives from the factory after the next recalibration of the instrument.

Whatever was affecting the salinity does not seem to have affected the temperature and fluorescence; there was no unusual noise and no jumps were seen in those channels. This probably fits the idea that fresh water is getting into the system intermittently. While this might also affect temperature and fluorescence, it might not be as dramatic and, at least for temperature it might not be systematic.

The TSG fluorescence is about 1.5 times the loop CHL and about 2.4 times that of the CTD fluorescence, which is fairly close to the results for 2008-26.

f.) Editing

The time-series plots were examined in a graphical editor. No editing was applied as Files #1, 3 and 4 showed no problems, and the noise in File #2 was not editable. While there are obvious blocks of data that look odd, it is unclear whether the data are suddenly resuming normal values after a period of drifting to poor values, or whether it is suddenly shifting to bad values. A header note will be entered to warn users that the data are of dubious quality. Only 2 decimal places will be included for the salinity in file #2 in the archive so that users may notice and read the header.

g.) Recalibration

As the intake temperature is not useful for files #1 and 2, a new channel will be derived. First ADD CHANNEL was used to add channel TEMPERATURE:LAB. That was then put through CALIBRATE using file 2008-10-recal1.ccf to set TEMPERATURE:LAB = Temperature:Primary. Then file 2008-10-recal2.ccf was used to recalibrate Temperature:Primary by subtracting 0.20 Cº  and Salinity by adding 0.23 for all files.

Thus we have:

- Temperature:Lab as the uncorrected temperature recorded in the lab.

- Temperature:Primary as the lab temperature adjusted to remove the effects of ship heating; this is a proxy for the intake temperature.

For files #3 and 4 HEADEDIT was used to rename Temperature:Primary as Temperature:Lab and Temperature:Secondary was renamed Temperature:Primary. CALIBRATE was used to apply file 2008-10-recal2.ccf. REORDER was used to ensure that the intake temperature will be selected preferentially by programs accessing the archive.

h.) Preparing Final Files

REMOVE was used to remove channels Temperature:Intake from all files.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. The comment includes notes to indicate that for files #1 and 2 the Temperature:Primary was recalibrated to correct for heating in the loop based on the historic observations and comparison with CTD data and that salinity recalibration has more uncertainties than usual. For File #2 the salinity was given only 2 decimal places and a note was added to indicate the quality is lower than usual.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

The sensor history was updated.

Particulars from the Daily Log Book:

2. PAR on

5. Large difference between sensor pairs.

9. No trip bottle #7 on this and many other casts.

31. No communication at bottom – run again.

39. PAR off.

39. Logged all data including the soak period – maybe can test effect of not stopping at surface.

42. Smoke on deck from stack during sampling.

47. PAR on.

56. Smoke on deck from stack during sampling.

60. PAR off

63. CTD upcast came up to 1900 and was then lowered to get bottles.

66. “with cups on”

79. PAR on

94. PAR on

108. At about this time a TSG flow test was done.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEA-BIRD | 911+ | 443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Jan08 | Factory “ |  |  |
| Conductivity | 1766 | 07May08 | “ |  |  |
| Secondary Temp. | 4700 | 16Jan08 | “ |  |  |
| Secondary Cond. | 2173 | 07May08 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 14Feb2007 | Factory |  |  |
| PAR | 4656 | 11Feb2003 | IOS |  |  |
| Fluorometer | 2228 | ? | IOS |  |  |
| Pressure Sensor | 63507 | 25/Oct/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG

Make/Model/Serial#:	SEA-BIRD/21/2487       Cruise ID#:	2008-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 01/12/07 | Factory |  |  |
| Conductivity | 2487 | 01/12/07 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 8/01/01 | “ |  |  |
| Temperature 2 | 4652 | 22/Dec/06 |  |  |  |
| Flow Meter | ? | ? |  |  |  |