## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2008-13

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Aquaculture

Area: Broughton Archipelago

Chief Scientist: Galbraith M.

Platform: Princeton

Date: 25 February 2008 – 29 February 2008

Processed by: Germaine Gatien

Date of Processing: 17 March 2008 – 23 May 2007

Number of original CTD casts: 39

Number of casts processed: 39

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19+ SEACAT CTD (S/N#5299) was used. Salinity samples were gathered using a Niskin bottle that was mounted about 1m above the CTD.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard. There was a cast named event 31b; since such a file name cannot be used in IOS SHELL, it was renamed 32. That event # was assigned to a NET cast, so it is assumed this will not cause confusion.

Sheets from the CTD Daily Log were available. The times in the log book are in Pacific Standard Time. To match the headers in the archived files add 8 hours to the log times.

The CTD was allowed to soak for at least 2.5 minutes at a depth of 1 to 1.5m for all casts checked; study of another SBE19 CTD during 2007-11 suggests that such long soaks (2-3 minutes) lead to better salinity data. This new instrument seemed to equilibrate within 1.5 minutes so that a shorter soak may be sufficient, but it is probably wise to be cautious and wait at least 2 minutes.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard so were adjusted. The file names are consecutive, so a match was made to the CTD casts in the log sheets, assuming they are in the same order. This will be examined later.

The dates in the log book are in Pacific Standard Time.

This is a new CTD. Configuration file 5299.con is based on the factory calibrations.

- Conversion of Raw Data

- The raw data were converted using conversion file 5299.con. A few casts were checked. The upcast conductivity is bad for cast #2, and the pumps do not appear to have come on until about 10m of the downcast of cast #3. After that the data looks ok. The upcast and downcast data are offset somewhat vertically. Descent rates look good, steady and generally >0.5m/s. It looks like the pumps came on above 1db with the exception of cast #2. For that cast the upcast should be investigated.

- FILTER

Tests were run to ensure that the recommended settings worked well for this data.

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. A few casts were checked and the conductivity is slightly smoother.

The pressure was filtered with a time constant of 1 second to increase the pressure resolution. The pressure looks slightly smoother after this step.

- CELLTM

The manufacturer’s recommendation is that CELLTM be run for this instrument. They suggest values like (α = 0.04, 1/β = 1). Tests were run using a variety of settings and any changes worsened the correspondence of up and downcasts on a T-S surface.The temperature ranges were very small for this data, so it is likely that there is no error to correct. This step will be skipped.

- DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets and in the SeaBird headers. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to add details. Where bottom depth was not available from the log that line was removed. Where the log recorded bottom depth in feet, it was converted and entered in metres.

8. Checking Headers

A track plot was produced and showed that many of the positions look slightly wrong. Moira Galbraith prepared corrections for all files. A text editor was used to replace the original positions with the new ones in the *.HDR files and a few station names were changed.

Track plots and header checks turned up no further problems. Track plots were added to the end of this report.

The surface check was run. The average surface pressure is +0.1db; there are 2 negative values at the surface (-0.1 and -0.2db) but the corresponding conductivity is ~0 and the upcast surface is 0db for each. The conductivity values are reasonable for very near the surface. The pressure offset applied looks appropriate and no further correction is needed.

9. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. The only problem noted was with the upcast of cast #2 and possibly the top of the downcast of #3, as noted before. The downcast data looks ok for #2. ADD TIME CHANNEL was used to add 8 hours to the start times so they are in UTC.

The soak period was about 2.5 minutes for all casts checked. The conductivity suggests that 1.5 minutes is probably enough. There were a few casts where there was a lot of variability beyond that stage, but this is likely reflecting actual conditions. Cast #3 looks very odd as noted above.

10. SHIFT

This was a new instrument and was used for this cruise and for 2008-18 and -38. Initial checks show no spiking in salinity, though there is some evidence that the conductivity is leading the temperature very slightly. Tests done on cast #12 from 80 to 90db suggest that a shift of +1.5 records (~+0.3s) to the temperature makes T and S line up better with respect to pressure and the T-S plot looks a little more believable. For cast #17 a shift of +1.5 looks good. For cast #52 a shift of +1 is slightly better but +1.5 looks good. The gradients are not sharp and the differences slight and using larger shifts still looks ok.

During 2008-18 using this instrument, the salinity was noisier and a setting of +2 or +2.5 records produced the best results. During 2007-36, -37, -38 and -39 using a similar instrument a setting of +2.25 records looked best.

Temperature was shifted by +2.25 records followed by a recalculation of salinity for all casts.

11. DELETE

To reduce the amount of editing needed CLIP was run to remove the first 200 records; the data removed will all come from the early part of the soak period, so no useful data will be removed.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

12. CTDEDIT

CTDEDIT was used to remove bad data and to clean salinity as needed.

The following casts required no editing: 18, 39-41, 49-52, 55, 66.

All other casts required some editing. Notes about editing were entered in the headers.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined. Cast #3 looks a little odd at the surface, but is stable, so no further editing was applied.

14. Intercomparisons

Previous Use of CTD – The instrument is new. It was used for 2 other cruises in the same area in early 2008. From 2008-18 the CTD appears to be low by from 0.005 to 0.01 with an average of ~0.0075. There was only 1 bottle during 2008-38 and the CTD was lower than that bottle by 0.014.

Repeat Casts – There were no repeat casts.

Historic ranges – No local climatology is available.

15. CALIBRATION

There were 4 salinity samples gathered about 1m above the bottom. When compared to CTD data after editing and bin-averaging the differences of (CTD – Bottle) are -0.0096, -0.0073 and +0.0014. The fourth sample was from an area with noisy CTD data that was removed in editing. Before editing the difference was -0.0157. Of the 3 values available after editing, the one with the CTD higher than the bottle was the shallowest sample and was in an area of fairly high temperature and salinity gradients. So the best two available bottles show the CTD to be low by about 0.01 and 0.007.  (See 2008-13-sal-comp.xls.)

The results of 2008-13, -18 and -38 all suggest that the CTD is low by from 0.005 to 0.014. There are some concerns that the Autosal is giving salinity values that are too high, particularly for salinity values <34. The errors from an initial test of Autosal linearity indicated errors on the order of 0.005 to 0.009 for the range of salinity sampled in Broughton. (See file “Linearity test-Autosal.xls”.)

No recalibration will be applied to these 3 cruises.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Particulars

31a. The cast with event #31a in the log book was renamed as #32 to enable standard naming. The event named #32 in the log book was a NET cast, so it is hoped this will not cause confusion. There was a surface salinity sample during the NET cast.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2008-13

Dates: 	Start:	25 February 2008	    End:	29 March 2008

Location: Broughton Archipelago                 Vessel:    Princeton

Party Chief: Galbraith M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SeacatPlus | 5299 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/SEACATPLUS v1.6b / 5299

Cruise ID#:	2008-13

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 20Dec07 | Factory |  |  |
| Conductivity | 5299 | 20Dec07 | Factory |  |  |
| Pressure Sensor | 5299 | 18Dec07 | Factory |  |  |