## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected the Salinity:Bottle precision lost during HPLC addition. S.H. |
| 4 Jan 2021 | Added HPLC Data. S.H. |

## PROCESSING NOTES

Cruise: 2008-19

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG-JdF Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: April 14, 2008 – April 20, 2008

Processed by: Germaine Gatien

Date of Processing: 20 June 2008 – 29 September 2008

Number of original CTD casts: 	74		Number of CTD casts processed: 74

Number of bottle casts: 	22 (There was 1 other bottle file with no sampling)

Number of bottle casts processed: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1176), a Seapoint Fluorometer (#2228) with a 10X cable, a PAR sensor (#4656), a Surface PAR sensor (#16504) and an altimeter (#1252). The deck unit and the logging computer serial numbers were not recorded in the log.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order though there was no list of equipment used. There were a few errors concerning rosette sampling: one came from confusing station name and event number and another was due to an error on the rosette log sheet. Having sample numbers recorded in both log and rosette sheets helped resolve the errors.

There were problems with the date entered in the files for casts between 0600 and 1200 hours UTC. A similar problem occurred during cruise 2008-07 which preceded this one and used mostly the same equipment. This could be a problem with acquisition software.

A recent change to the method for deployment of the CTD may be leading to noisier data in the top 10db; it is recommended that the CTD be kept at the surface for at least 15s before starting the full cast to ensure that the water stirred in raising the CTD from the 10db soak level has had time to settle. This is most likely to be significant in quiet waters with little ship drift.

The dissolved oxygen data in the CTD files should be considered

- ±0.6  ml/l from     0 –100db

- ±0.25ml/l from 100 – 300db

- ±0.1  ml/l below 300db

Due to doubts about the reliability of the salinometer analyses, recalibration was not applied to the secondary salinity (primary salinity was used for 1 cast, so was recalibrated only to make it match the secondary); the calibration should be revisited when the sensor is next recalibrated at the factory.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets.

Salinity data were available in spreadsheet format.

Nutrients data were provided in spreadsheet format.

Dissolved oxygen data were available in individual files.

CHL data were available in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The calibration constants were checked for all instruments and no errors were found.

All CTD con files were the same; one was saved as 2008-19-ctd.con.

##### Conversion of Raw Data

Data were converted using configuration file 2008-19-ctd.con. The pressure offset was set to 6.5db which proved reasonable for the last use of this equipment, so it was left unchanged.

A few casts were examined and all expected channels are present and look reasonable when plotted. As usual the upcast temperature and conductivity traces are noisier and further apart than in the downcasts. For the deepest casts, the fluorescence is about 0.12ug/l to 0.15ug/l near the bottom.

The transmissivity has small spikes at depth, but not large ones.

The dissolved oxygen shows the usual offset, but the response time does not look too bad.

The descent rate was quite steady early in the cruise, but sometimes very noisy and fairly low in the 2nd half.

The altimetry looks reasonable near the bottom.

Rosette files were converted using a start time of -2s and duration of 5s. While 23 files were created there are only rosette sheets for 22. Cast #66 shows 5 bottles were fired in rapid succession, but there is no indication that there was any sampling for that cast. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and the only problems noted were in cast #66 for which bottles were fired without stopping. It is believed that there was no sampling from that cast. No editing was done.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on a few casts to determine the best choice of parameters for CELLTM. Settings tried were (α = 0.01, β=7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9), (0.04, 7), (0.04, 9) and (0.0245, 9.5). The differences among them were slight and varied from feature to feature. This may be due to small-scale noise in the data. The choice of (0.0245, 9.5) looked reasonable for all cases, so was applied to all casts for both channels since it was found best during 2008-07 when all but the secondary conductivity sensor were also used.

CELLTM was run using (α = 0.0245, β = 9.5).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are very noisy, especially in the shallow temperature, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 8 | 200 300 | -0.0001 XN -0.0003 XN | -0.00025 -0.0003 XN | -0.0027 XN -0.003 XN | Very Steady ~.8 |
| 25 | 200 300 400 | -0.0001 XN +0.0003 XN -0.0001 | -0.0003 -0.0003 XN -0.0003VN | -0.0031 XN -0.0031 -0.0029 | Very Steady ~.8 |
| 55 | 200 300 | -0.0001 XN Too noisy | -0.0004 XN -0.0003 XN | Too noisy Too noisy | Very steady ~.7 |

The differences are noisier than usual even for shallow data, but they do show that the primary salinity is higher than the secondary by about 0.003 and the temperature and conductivity differences are reasonably small, on average, but with great variability around those averages.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and a cross-reference listing were produced. Errors were found and corrected in the station names for casts #59 and 64. Times and positions look fine. The NMEA date does not agree with the system upload time for casts #16-22, 52-54 and 64-67. Notes in the log book indicate that it was realized there was a problem at sea. The dates in the log book agree with the system upload time except for the first set for which there is a note added that the entries were wrong. All the dates that are in error are from 0600 to 1200 hours UTC. A similar problem occurred during 2008-07. It is possible it is due to a bug in a new acquisition system. The dates were changed in the headers to agree with the log book dates.

The cruise track was plotted and no problems found.

The average surface pressure is 2.9db, which is fairly high for the Vector. It is possible that the offset is slightly too high, but since the drift in pressure sensors always leads to increasingly large offsets, this is not likely to be a significant source of error. It is more likely that the CTD was just started deeper than usual. For cast #74 during the upcast when the pressure was 1.2db the salinity was 31.95 which looks reasonable, so the pressure is probably not too high.

The altimeter readings from the headers of CLN files were exported to a spreadsheet and the entries were checked against plots of altimetry against pressure. All values looked fine up to cast #50. As noted in the log the chains were put on for cast #51 due to rough weather and the altimeter began picking up the weight rather than the bottom. All values from #51 to the end of the cruise are bad except for cast #68 when the chains probably were taken off. The chains were put on again for cast #69. The altimetry headers were removed from both CTD and BOTTLE files from casts #51-67 and #69-75.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Note that samples #84 and 85 were repeated for casts 15 and 16, so those were renamed 984 and 985 for cast #16. Cast #66 was removed because no sample numbers were assigned. There were some errors in the chemistry files from analysts: the rosette cast at station 72 was named event #68 on the rosette log and should have been #69; the nutrients from event #50 were named as #46 (46 was the station name).

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (SAMAVG.)

SALINITY

The salinity data were delivered in spreadsheet format 2008-19sal.xls. The sample numbers were combined with the station names in a single entry. Those were separated into 2 columns. There was no flag column so one was added. Headers were changed to standard formats and the spreadsheet was simplified and saved as2008-19-sal.csv. This was converted into individual SAL files.

There were no duplicate salinity samples.

DISSOLVED OXGYEN

Dissolved oxygen files (*.add) were provided with a flag channel but the format was wrong for the quality flag. The original files were renamed ADD1 and then put through Header Edit to add a column. The output files were named ADD. One sample number had a typo and one file needed to be reordered on sample number. One cast had 2 entries for 1 bottle, a common problem and easily fixed using the rosette log sheets for confirmation of the right value. No flags or comments had been added. There were a few comments on the rosette sheets that might warrant flags, so “c” flags were entered plus comments, but those should be reviewed when the comparison to the CTD DO is done.

- sample #12, cast #1 had a note “Bad End Point”

- sample #196, cast #52, note: “murky sample”

- sample #244, cast #61, note: “surface sample taken with a bucket”

- sample #269, cast #68, note: “surface sample taken with a bucket

File 2008-19-0006.ADD was renamed 2008-19-0007.ADD. There was no DO sampling during event #6. This was a case of accidentally using the Station Name rather than Event Number in naming the file.

File 2008-19-0068.ADD was renamed 2008-19-0069 as there is no record of rosette sampling for cast #68 and there is no ADD file for cast #69 where there was rosette sampling.

There were no duplicate samples.

NUTRIENTS

The nutrient spreadsheet was simplified and saved as 2008-19nuts.csv. Extraneous columns were removed, header names were changed to standard format and lines were removed for which there was no nutrient sampling. Data were sorted on sample number. File 2008-19nuts.csv was then converted to individual NUT files. (After COMPARE was run on salinity “d” flags were added to all bottle analyses for sample #218; W. Richardson agreed that was appropriate.)

CHL

The CHL spreadsheet was edited to change headers to standard names and remove extraneous lines; it was then saved as 2008-19chl.csv. Data were sorted on sample number. File 2008-19chl.csv was then converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

COMPARE was run using pressure as the reference channel. There was one extreme outlier: sample #218, cast #56 looks like the bottle closed well above 50db. Since CHL, DO and Nutrients are consistent with a shallower closing all bottle values were flagged “d”. There were 3 other outliers differing by >0.04, but none of those were associated with suspicious DO, CHL or Nutrient values. Two of those 3 have fairly high standard deviation in the CTD salinity.

There was a lot of scatter in the differences; when outliers were excluded there were 11 points left in the fit. The average of differences shows the primary salinity to be low by 0.0010 and the secondary low by 0.0037.

There are some doubts about the Autosal analyses done before July 2008 so plots were also made of differences versus salinity. There is some salinity-dependence in the fits and the average of the 3 deepest, highest salinity samples shows the primary to be high by 0.0004 and the secondary low by 0.003, whereas the 3 lowest salinity samples show the primary to be low by 0.0031 and the secondary low by 0.0057. The differences between the salinity channels are consistent with the observations in section 7.

A non-linearity test was run on the Autosal in April 2008 on standards 10, 30, 35 and 37 salinity units. The Autosal-Standard difference found was 0 at 35psu, +0.009 at 30psu and +0.001 at 10psu. Comparisons of bottle data from other cruises compared to post-cruise factory calibrations support the results at 35 and 30. Interpolating between those would indicate that the Autosal would read high by ~0.0025 in the 33-34 range into which fell the 3 deepest bottles. The average of those 3 bottles indicate that the primary CTD was higher than the bottles by 0.0004 and the secondary CTD was lower by 0.003. If we assume that the Autosal is too high by 0.0025, then the primary would actually be high by 0.0029 and the secondary low by 0.0005. This is a very rough estimate, so recalibration should not be applied until the sensors are next sent to the factory for recalibration. Using the secondary sensors looks like the best choice.

CTD Fluorescence

COMPARE was run using extracted chlorophyll and CTD fluorescence. The fits show fluorescence to be about 56% of extracted chlorophyll values, with the ratio FL/CHL having lowest values in the central part of the Strait of Georgia, and highest in Saanich Inlet, the northern end of the Strait of Georgia and most of Juan de Fuca Strait. The ratio is slightly lower for low values of CHL.

Dissolved Oxygen

COMPARE was run using pressure as the reference channel. There was only one severe outlier and that came from a bottle that had already been flagged by the analyst; the flag was changed from “c” to “d”.  When a few other outliers were identified by residuals the fit was:

CTD-BOT = 1.039 DOX-CTD + 0.0022

During 2008-07 in late March in Broughton Archipelago the fit was:

CTD-BOT = 1.0254 DOX-CTD + 0.0081

and during 2007-43 in June 2007 in the Juan de Fuca Strait and the Strait of Georgia, it was:

CTD-BOT = 1.0351 DOX-CTD + 0.0098

For 2008-28 a new approach was used because the fit looked unrealistic for this instrument at low DO values. When bottle values are very low the CTD DO sensor does not reach zero and certainly never <0. To force a more realistic value at low DO, the offset was set so that CTD DO = 0.07 when Bottle DO = 0, and the resulting fit was:

CTD-BOT = 1.0408 DOX-CTD - 0.07

For this data the same approach was used as for 2008-28 but a slightly lower offset seemed appropriate, so it was set to 0.01 and the fit was then:

CTD-BOT = 1.0405 DOX-CTD - 0.01

This is remarkably close to the results for 2008-28 which was a cruise in mainland B.C. inlets.

The fit against file pair number shows that cast #1 is slightly out of line, but it was the only one that sampled low oxygen water. There was also a lot of scatter in the fit and it was the first cast of the cruise so the CTD may not have performed well. Excluding that cast from the fit increased the slope only slightly to 1.0416. A few individual casts were examined to see if there is any obvious effect of sampling low DO water on the calibration for cast #1 and those that followed:

| Event # | Slope of fit | Range of DO values | Site |
| --- | --- | --- | --- |
| 1 | 1.0228 | 0-10 | Saanich Inlet |
| 4 | 1.0507 | 2-8 | West of Texada I. |
| 7 | 1.0419 | 4-8 | West of Texada I. |
| 9 | 1.0367 | 3-9 | West of Texada I. |
| 32 | 1.05 | 2-9 | Central St. of Georgia |
| 37 | 1.0314 | 3-6 | Haro Strait |
| 58 | 1.0396 | 2-6 | Near Victoria in JdeF |
| 74 | 1.0532 | 1-6 | Mouth of Juan de Fuca |

There is too much variability to conclude anything, but cast #1 certainly has the lowest slope. However, it also has the largest range and the sharpest DO gradients near the surface and the bottom. The only other cast with a surface gradient that is similar is cast #9 but it does not have much gradient below 20m. We would not really expect to see much effect from hypoxic sampling on cast #4, as the instrument was used on cast #2 and would have largely reverted to the old calibration if there had been any change. In other cruises from the past few years which sampled in anoxic waters, changes in calibration have not been noted. (See 2008-19-dox-comp1.xls.)

Plots of Bottle DO and CTD DO versus Salinity suggested that another bottle value be flagged – sample #175 for cast #43 was also a moderate outlier in COMPARE. So it was flagged “c”. The flags assigned earlier were reviewed; sample #12 from cast #1 was changed from a “c” flag to a “d” flag.

##### 13. Shift

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The usual shift of +24 records (1s) was found to be appropriate and was applied. (Output: SHFFL)

Conductivity

Tests were run on 6 casts using shifts between -1s and +1s and T-S plots were prepared to compare the results. The range of values was small, so the results were not clear, but a setting of -0.5s looked best for the primary sensor and -0.3s for the secondary. All casts were put through two runs of SHIFT using those settings. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on 3 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. SHIFTS of from +100 to +130 records were tested and +110 looked best overall, though the differences were not large.

SHIFT was run using +110 records for all casts.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

An examination of a few casts shows no significant difference in the noise level in the two salinity channels. COMPARE shows the primary closer to the bottles, but it is likely that the Autosal results are high so that the secondary is likely more accurate. The secondary conductivity sensor has not been used on any cruises that have been processed, so might be expected to have not drifted much. So the secondary Temperature and Salinity were selected for further processing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. It is believed that for many casts the CTD had been lowered to 10db then returned to the surface and lowered again after a short stop. There is a lot of noise in the top 10db and some of it is probably due to the stirring by the initial lowering. In areas like Haro Strait where instabilities are not unexpected, editing was not applied unless data were clearly corrupted. The descent rate varied from extremely steady to extremely noisy.

Cast #28 had sections of bad data in the secondary channels, but the primary looks fine, so primary channels were edited for that cast.

All casts required editing except for cast #40. Note was made of the editing details in the headers of the relevant files.

##### Initial Recalibration

The secondary salinity is probably very close to the bottle values, though the bottle calibration shows it to be low by ~0.003; no recalibration of secondary salinity will be applied at this point, though information from the next factory check might lead to a future recalibration. Because primary salinity is to be archived for cast #28, it is necessary to recalibrate that channel to match the calibration of the secondary salinity. (The primary was found to be higher than the secondary salinity by 0.0027.)

File 2008-19-ctd.ccf was prepared to apply the following equation to the CTD Dissolved Oxygen channel in the SAM and MRG files (done on temp files, redo later):

CTD-BOT = 1.0405 DOX-CTD - 0.01

and to subtract 0.0027 from Salinity:T0:C0.

COMPARE was then rerun to check that the results were as expected and they were. (See 2008-19-dox-comp2.xls and 2008-19-sal-comp2.xls.)

The same DO calibration was applied to the edited CTD files.

##### Special Fluorometer Processing

The COR1 files were clipped to 15+0db and processed separately for A. Peña. (Output: CLIP)

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: The primary conductivity sensor was used for 2008-07 when the salinity was found to be low by 0.0007. No data from the secondary have been processed since it was last recalibrated.

2. Dissolved Oxygen – This sensor was used for 2007-46, 2007-43, 2008-07 and 2008-28 in May and June 2007 and 2 in March 2008. The 1st of these produced confusing results; the next two had slopes of 1.0310 and 1.0254, but the offset had not been forced as was done for 2008-19. 2008-28 had a slope of 1.0484 and the offset had been forced to 0.07. For all 4 a second recalibration was applied to downcasts.

3. Pressure – This sensor has been drifting significantly in recent years. When used for 2007-61 and 2008-07 in October and March, offsets of +6.3db and +6.5db were used.

Historic ranges –The salinity was slightly above the historic maximum around 85db for cast #54 but the range itself looks odd. Cast #59 also has a few sections of high salinity near the bottom. Cast #73 has temperatures below the historic minima at ~20db and salinity above the maxima in the top 70db. This cast was the southernmost at the mouth of Juan de Fuca Strait and likely sampled waters not represented in the climatology. These excursions do not suggest a problem with the calibration of T and S, but are more likely reflective of limitations in the climatology which has not been updated since 1997.

Repeat Casts – There were no repeat casts.

Comparison of Nearby Casts – For lines across Juan de Fuca Strait there is an obvious gradient with the southernmost sites having colder and saltier water except at the bottom; this is seen as far east as Victoria.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. A further correction will be applied to at least partly correct for response time. To do this we compare downcast data to bottle data from the same pressure.

Files were bin-averaged to 0.25m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. There is a lot of scatter in the comparison especially near the surface. The fit shows some pressure dependence with higher differences near the surface. The fit against DO is fairly flat at DO values <6.5ml/l; above that there is a lot of noise and the fit depends strongly on what criterion is chosen for eliminating outliers. A simple offset looks appropriate. Using large differences and large standard deviations in the CTD data to eliminate outliers leads to an estimate of -0.07ml/l for the offset to be applied. Returning to the fit of differences against pressure with an offset of -0.07 in mind, it is apparent that it looks a little low between 20 and 50db but reasonable elsewhere. (See 2008-19-dox-comp3.xls.)

The thinned files were recalibrated by subtracting 0.07ml/l and the comparison was rerun. That showed that the recalibration was applied properly. (See 2008-19-comp4.xls.)

Recalibration using file 2008-19-recal2.ccf was applied to the downcast files only. (AVG and CLIP) to subtract 0.07ml/l. (Output: COR2 and CLIPCOR)

The clipped, recalibrated files were then bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts except #28 REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

For cast #28 REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added; REORDER was run to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered

• ±0.6  ml/l from     0 –100db

• ±0.25ml/l from 100 – 300db

•   ±0.1  ml/l below 300db

The Standards Check routine was run and no problems were found. The final files were named CTD.

A header check turned up no problems.

Profile plots were made and no problems were found.

The track plot looks ok.  The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The values ranged from 70% to 160%, with the highest value in Saanich Inlet and 130-150% in the northern section of the Strait of Georgia except for the 2 most northerly sites which were well mixed. The lowest values were in Juan de Fuca Strait. For the most extreme cases, bottle values were checked against CTD and look ok.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts the following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars from log:

3. Cast aborted – too rough

19-22. Date may be wrong – should be 16th, not 17th. For cast #22 the time was 0435PDT, April 16.

52. NMEA has day as 19th, not 18th.

61. Surface samples taken with a bucket over the side because bottle did not fire.

66. 5 bottles fired, but there is no record of any sampling for that cast.

67. Display says day is 20th, should be 19th.

70. Surface samples taken with a bucket over the side because bottle did not fire.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Jan08 | Factory “ |  |  |
| Conductivity | 2399 | 15jan08 | “ |  |  |
| Secondary Temp. | 4700 | 16Jan08 | “ |  |  |
| Secondary Cond. | 3396 | 21Dec08 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS | 5MAR08 | IOS |
| SBE 43 DO sensor | 1176 | 14Feb2007 | Factory |  |  |
| Fluorometer | 2228 |  | IOS |  |  |
| PAR | 4656 | 11Feb2003 |  |  |  |
| Surface PAR | 16504 | 2Jan2004 |  |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |