## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2008-21

Agency: PBS, Salmon and Freshwater Ecosystems, Sidney, B.C.

Location: Broughton Archipelago	Project: Sea Lice Monitoring

Party Chief: Hargreaves B.		Platform: Franciscan NO 1

Date: 23 April 2008 –30 April 2008

Processed by: Germaine Gatien

Date of Processing: 25 July 2008 – 1 August 2008

Number of original CTD casts: 100 + test

Number of CTD casts processed: 99 (cast #45 empty, cast 100 incomplete)

# INSTRUMENT SUMMARY

A SeaBird Model SBE-25 CTD (#334) was used.

# SUMMARY OF QUALITY AND CONCERNS

The pressure sensor for all other casts is considered ±1db.

The times in the log are believed to be in PDT; the times in the headers of the archived files are in UTC.

There were problems with the correspondence of file names to event numbers for the first day of the cruise. The files were renamed in the order of collection. Details are given in section 3.

There was no salinity calibration sampling and no history for the sensors since they were last recalibrated, so no comment can be made on the salinity accuracy.

The CTD precision is considered ±0.001psu when stopped. In motion the salinity is considered ±0.005psu. (In areas of rapid temperature change the SBE25 does not perform this well, but for this data set the gradients were generally fairly low.)

# PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension HEX.

2. Preliminary Steps

The Daily Log was obtained. There was no note of any problems encountered.

All file names were non-standard.

The cruise summary sheet was completed.

## 3. Conversion of Raw Data

All file names were changed to standard format. The consecutive number was included in the name but there were problems with the first day’s data:

- There was a cast given consecutive number 0 which won’t work in IOS numbering system.

- There was no cast 10 in the log, but a file was given that name that corresponds in time to what is called #12 in the log book.

- Cast 11 is listed in the log, but there is no file with data from the time given. The file listed above is from the same site but 10 minutes later – log book list #12.

- The file named #13 corresponds to one called 15 in the log book, but then scratched out.

- The file named #14 corresponds to one called 16 in the log book, but then scratched out.

- There is a test cast mentioned and given consecutive #15 in the log, but the file is just called test.

There are 16 data files from the first day, but one is a test cast. There are 17 log entries, but no data corresponding to one of them and one is the test cast. So there are 15 “real” data files and 15 log entries to which they correspond. So a renumbering was done from 1 to 15, taking the casts in time order. A note was put on the copy of the log sheets giving the new event numbers on which the new file names will be based. The new event number is equal to the number in the original file name plus 1. After cast #15 the file names correspond to the log entries.

The conversion file 0334.con was found; it had an old calibration for the temperature sensor. A new configuration file was prepared with the April 2007 calibration for temperature. This file was saved as 2008-21-ctd.con which was then used to convert all files.

File 2008-21-0045.hex contained no data so could not be converted.

Data were plotted from a few casts. Cast #100 could not be plotted due to an error in header time. Fixing the time interval and date and time in the header allowed plotting. It is clearly not a complete file with just downcast data to 17db. The log indicates the cast was to 40db. The surface temperature is also a little odd. The data will be processed, but should be checked later to see if it is suitable for archiving. Presumably whatever caused the data from casts 101-109 to be lost, also affected this data.

The test cast plot looks fine, but it is unknown where it was run, so it will not be processed further.

All casts examined had a soak period of at least 1.5 minutes near the surface.

4. WILDEDIT

This step is usually skipped and since there is no evidence of spikes it was not run.

5. WFILTER

WFILTER was run using cosine filter, size 5 on the temperature, conductivity and salinity channels. Plots were made before and after this step. Reversals disappear from the pressure after this step but the trace is not overly smooth. When salinity was derived on filtered files it was much smoother than the unfiltered version.

6. CELLTM

SeaBird recommend the use of (α, 1/β) = (0.04, 8) for CELLTM for the SBE 25. For 2008-20 a variety of settings were tested, but as for this cruise, the casts were all shallow and variable, so tests were not easy to interpret. A choice of (0.04, 9) was selected as best overall and that was also chosen when this sensor was used for 2007-02. A setting of (0.04, 9) was applied to this data set as well. Plots of a few casts show the setting to be reasonable.

7. DERIVE

Program DERIVE was run to calculate salinity.

8.  Conversion to IOS Headers

The IOSSHELL routine was used to convert the CNV files to IOS Headers.

9. Checking Headers

The times in the headers agrees with the log book times. The latter are said to be UTC times, but given that work goes on between 07:00 and 17:00 it would seem likely they are really in local time, PDT. This was confirmed by Brent Hargreaves.

HEADEDIT was used to add lines for latitude, longitude, station name and water depth.

A text editor was then used to enter the relevant details in all the files.

A cross-reference listing was produced and details checked against the log records. No errors were found.

ADD TIME CHANNEL was used to add 7 hours to the times to put them into UTC.

HEADER CHECK was run, and no problems were found.

The track plot looked fine.

For 2008-20 pressure was recalibrated by adding 0.7db to the pressure as was done for cruises 2006-01 and 2006-02 when this sensor was used. During 2007-02 no offset was applied. The surface check gave values from +0.5db and -0.6db for surface pressure with the exception of the first cast when the surface was at -0.8db. The salinity had clear “in-water” values for most casts with a positive surface value. The salinity for cast 78 was 30.5 at pressure 0.5db. So applying an offset of +0.7db looks reasonable.

##### 10. Test plots

Plots were made and show spikes at the surface and bottom, but no other obvious problems.

11. Salinity bottles

The only salinity samples taken were surface bucket samples not suitable for calibration checks. They had not been analyzed at the time of processing.

12. SHIFT

Conductivity

Tests were run on several casts of 2008-20 (with the same sensors and similar conditions), using values from -0.6 to +0.6 and the best choice proved to be -0.25s. During 2007-13 and -14 when this sensor was also used a setting of -0.6s looked best. A few checks were made on casts from this cruise and a setting of -0.25s looked reasonable for most data. All casts were put through SHIFT with a setting of -0.25s.

13. DELETE

DELETE was run on all casts using the following parameters:

Surface Record Removal: Last Press Min.	Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted from 10db to 10db above the maximum pressure.

COMMENTS ON WARNINGS: The only warning pertained to upcast data, so is not of concern. The scan numbers increase as expected, but pressure shifts suddenly between consecutive scan #s, as though the CTD failed to record data for awhile.

- DETAILED EDITING

Based on results during the processing of 2008-20, and a few checks of this data set, CLIP was run to remove data from 0 to 0.7db. Then the files were metre-averaged and T-S plots were prepared; casts with unstable features or other suspicious features were examined in CTDEDIT to see if the oddities were likely due to instrumental problems so that editing was justified. For most it was not.

The following files were edited lightly: 8-10, 29, 38, 43, 45, 48-51, 56, 61, 62, 64-66, 74, 77 and 92.

Notes of editing details were made in the headers.

At this point cast #100 which is missing data below 18.6db was examined on a T-S surface and found to be very different from cast 99 which was nearby. Casts from the same two stations during cruise 2008-20 were examined and were much closer. However, when the files were clipped to show only data from the top 18.6db the similarities are not so strong.  The two stations are separated by an island and local variability is likely so that fresher, cooler conditions at cast #100 are not clearly out of line. A note was added to the header to indicate that data are missing.

- Inter-comparisons

Sensor History – The sensors were used for cruise 2008-20 when there was no salinity calibration sampling. There is no history of use prior to that since the last recalibration. The pressure sensor was used for two cruises in 2006 when an offset of +0.7db was applied, and for one cast in 2007 when a 0db offset was found appropriate. During 2008-20 an offset of +0.7db was used.

Comparison of repeat casts –There are no repeat casts, and while some casts are close there is a lot of small-scale variability. Casts #19 and 20 are about 2.7km apart; temperature and salinity are virtually indistinguishable along lines of constant σt in the bottom 3db for those casts.

Historic Ranges – Historic ranges were not available.

- Recalibration, Bin Average and REMOVE

All files were recalibrated by adding 0.7db to the pressure.

The files were bin averaged using 1db bins.

REMOVE was run to remove Scan_Number, Conductivity:Primary, Descent Rate and Flag channels.

- HEADER EDIT and final checks

Header Edit was used to fix formats and channel names and to add the following note to the headers:

The pressure sensor is considered ±1db

The CTD salinity precision is considered ±0.001psu when stopped. In motion

the salinity is considered ±0.005psu except in areas of rapid temperature

change where the salinity has been edited heavily and should be considered

±0.1psu.

There were no salinity calibration samples and there is no history for these

sensors since their most recent recalibration, so no comment can be made

on the accuracy.

The final files were named CTD.

A cross-reference listing was produced.

A header check was run on the CTD files and no errors found.

A track plot was produced and looked fine.

Particulars

45. Empty hex file

100. File incomplete

101-109 – data lost

## Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2008-21

Dates: Start: 28 March 2008 End: 2 April 2008

Location: Broughton Archipelago

Vessel: Franciscan No 1

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 25 | 0334 | No | Yes |

CTD CALIBRATION INFORMATION

Make/Model/Serial#:	SEABIRD/SBE25/0334	Cruise ID#:		2008-21

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 25 Apr 07 | Factory |  |  |
| Conductivity | 1729 | 21 Aug 07 | Factory |  |  |
| Pressure | 0464 | 23Mar06 | Factory |  |  |