## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2008-23

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Aquaculture			Area: Broughton Archipelago

Chief Scientist: Galbraith M.		Platform: FRANCISCAN NO 1

Date: 19 June 2008 – 27 June 2008

Processed by: Germaine Gatien

Date of Processing: 25 July 2008 – 4 August 2008

Number of original CTD casts: 112

Number of casts processed: 111 (1 file contained only surface data)

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#5299) was used.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available.

The times in the log book are in Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

The soak period for the CTD was at least 2 minutes which was sufficient to reach equilibrium.

There was no salinity calibration sampling.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. However, this cruise had large gradients only at the surface and the descent rate was generally very steady, so the errors should be small.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard so were adjusted. 
The dates in the log book are believed to be in Pacific Daylight Time. Computer times were in the same time zone.

Configuration file 2008-23-ctd.con is based on the factory calibrations.

- Conversion of Raw Data

- The raw data were converted using conversion file 2008-23-ctd.con.

- Plots were made of a few casts and T and C look reasonable. The CTD was lowered to a level between 0.6db and 0.9db and soaked for about 500 or 600 scans, so >2 minutes. The temperature and conductivity data look as though they had equilibrated before the cast continued in most cases. Where there was variability after 2 minutes that is likely due to real variability rather than sensor problems.

- FILTER

The same parameters were used as for 2008-13, 2008-18 and 2008-22.

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 1 seconds to increase the pressure resolution.

- CELLTM

Tests for 2008-13 and 2008-18 showed no improvement in data using CELLTM and SeaBird don’t recommend it. This step will be skipped.

- DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets and in the SeaBird headers. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to add details. In a few cases the longitude was obviously wrong in the log.

8. Checking Headers

A cross-reference list was produced and checked against the log sheets. A few errors were found and fixed.

A header check was run and no errors were found there.

The track plot looked fine and was added to the end of this report.

The surface check was run. The average surface pressure is +0.4db with low surface conductivity values; there are no negative values. The pressure calibration looks appropriate and no correction is indicated.

9. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. There is lots of variability but reasonable correspondence. Cast #6 was found to contain only a little surface data; the HEX file was very small. This cast will not be processed further.

ADD TIME CHANNEL was used to add 7 hours to the start times so they are in UTC.

10. SHIFT

During 2007-26, -27, -28 and -29 using a different SBE19+ a setting of +2.25 records was found best for adjusting temperature to align it with conductivity, and the same setting was found appropriate for this particular instrument when it was used during 2008-18, 2008-38 and 2008-22.

Temperature was shifted by +2.25 records followed by a recalculation of salinity for all casts. A few casts were checked and temperature and conductivity lined up better after this step than before. There were no salinity spikes in either case.

11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

12. CTDEDIT

Before editing CLIP was run to remove records from the first 1db since the CTD was generally parked around 0.7 to 0.9db for soaking, so there is little useful data in the top 1db; eliminating that will make editing much quicker.

T-S plots of the clipped files were made to check which files needed editing. For those that have unstable features in T-S space, the files were opened in CTDEDIT.  For some casts bad records were removed from the top or bottom, and for a few casts light smoothing of salinity was applied where instabilities appeared to be due to a mismatch of sensors. Most casts required no editing.

The following casts required light editing: 5, 7, 20, 37, 63, 72, 76, 81, 91, 103-105 and 111.

Notes were entered in the headers about editing applied.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and no problems noted. No further editing will be done.

14. Intercomparisons

Previous Use of CTD – This CTD was used during 2008-13 when 2 bottles indicated that the CTD was low by 0.007 and 0.010 and during 2008-18 when 4 bottles showed the CTD to be low by from 0.005 to 0.01. During 2008-38 there was only one calibration sample that indicated the CTD was low by 0.014. There were some concerns that Autosal results in early 2008 were too high by as much as 0.01 at salinity ~30.

Repeat Casts – There are no repeat casts, and while some casts are close there is a lot of small-scale variability. Nonetheless, near the bottom, temperatures and salinity for casts #52 and 53 (about 1.7km apart at Marilyn Falls and Brown Point) are within 0.005C˚ and 0.001 salinity units along lines of constant σt. So there is no evidence of instrumental problems.

Historic ranges – No local climatology is available.

15. CALIBRATION

There were no calibration salinity samples. The results of other recent uses of this instrument suggest that the CTD salinity is low by from 0.005 to 0.014. There are some concerns that the Autosal was giving salinity values that were too high, particularly for salinity values <34. The errors from an initial test of Autosal linearity indicated errors on the order of 0.005 to 0.009 for the range of salinity sampled in Broughton.

No recalibration will be applied, though this should be revisited when the conductivity sensor is next recalibrated.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

There was no salinity calibration sampling and information available from previous uses was considered unreliable, so no recalibration has been applied to the salinity data.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients. However, this cruise had large gradients only at the surface and the descent rate was generally very steady, so the errors should be small.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2008-23

Dates: 	Start:	19 June 2008	    End:	27 June 2008

Location: Broughton Archipelago                 Vessel:    Franciscan No 1

Party Chief: Hargreaves B.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SeacatPlus | 5299 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/SEACATPLUS v1.6b / 5299

Cruise ID#:	2008-23

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 20Dec07 | Factory |  |  |
| Conductivity | 5299 | 20Dec07 | Factory |  |  |
| Pressure Sensor | 5299 | 18Dec07 | Factory |  |  |