## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2008-33

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Aquaculture

Area: Broughton Archipelago

Chief Scientist: Galbraith M.

Platform: Princeton

Date: 7 April 2008 – 10 April 2008

Processed by: Germaine Gatien

Date of Processing: 24 June 2008 – 28 June 2008

Number of original CTD casts: 65

Number of casts processed: 65 downcast files plus 1 upcast file

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#5299) was used.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available. The times in the log book are believed to be in Pacific Daylight Time. To match the headers in the archived files add 7 hours to the log times.

The soak period for the CTD varied from 45s to almost 2 minutes. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though 1.5 minutes may be sufficient.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients.

There was no data in the top 20db of the downcast for cast #80. It is likely that the data was actually acquired, but lost in the separation of files. For the time a partial downcast file (2008-33-0080.ctd) was prepared and the upcast data was used to create a full cast file (2008-33-9080.ctd), but if the complete downcast file is retrieved, it should be processed and substituted for the downcast in the archive and the upcast should be removed from the archive.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The dates in the log book are believed to be in Pacific Daylight Time.

Computer times were in Pacific Standard Time

- Conversion of Raw Data

- The raw data were converted using conversion file 2008-18-ctd.con. The same sensors were used for that cruise as well as 2008-13 and 2008-38.

- After conversion the file names were changed to standard format.

- Plots were made of a few casts and T and C look reasonable. There was a very long soak for the 1st cast which is an excellent idea. After that the soak time was shorter than recommended, ranging from 45s to 2.5min, with most well under 1.5 minutes. While the pumps come on within 30 seconds, a longer soak ensures that the sensors are fully equilibrated. Nonetheless, there is no evidence of significant problems with the downcast temperature and salinity for these data.

- FILTER

The same parameters were used as for 2008-13, 2008-18 and 2008-38.

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 1 seconds to increase the pressure resolution.

- CELLTM

Tests for 2008-13 and 2008-18 showed no improvement in data using CELLTM. This step will be skipped.

- DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The bottom depths and positions are missing from the headers. That information is available from the log sheets and in the SeaBird headers. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to add details. The log recorded bottom depth in feet, so it was converted and entered in metres.

8. Checking Headers

A cross-reference list was produced and checked against the log sheets. One position error was found and corrected.

One time differed (allowing for the 1 hour difference noted above), but the header entry looks right and the log time does not make sense. The time for the first cast differs significantly, but the log time may include set up time which was likely longer than for other casts.

Header Check was run and several positions were found to be wrong in the log book. Moira Galbraith confirms that the positions should be as follows:

- Cast 73 station BSLd- latitude 50 40.47 should probably be 50 50.465.

- Cast 98 station BG1 -  latitude 50 40.53 should probably be 50 48.533.

After those errors were fixed the track plots looked fine and were added to the end of this report.

The surface check was run. The average surface pressure is +0.2db with very low surface conductivity values; there are no negative values. The pressure calibration looks appropriate and no correction is indicated.

For cast #80 the surface pressure is 20.5db. The headers show that there should be 1405 records, but there are only 775, so there must have been an error in separating the original hex file into individual files. Darren Tuele was contacted to see if he can retrieve the full downcast data. He was unavailable at the time of processing, so upcast data was processed for this cast. If the full file becomes available, it should be processed and used to replace file 2008-33-0080.ctd.

ADD TIME CHANNEL was used to add 8 hours to the start times so they are in UTC.

9. SHIFT

When this instrument was used during 2008-03, -18 and -38, the temperature channel was shifted by +2.25 records followed by a recalculation of salinity for all casts. Tests were done on one cast and that setting looked best overall, though the effect on particular features varied presumably reflecting minor variations in alignment of the sensors.

For cast #80 an upcast file was created by putting the SHF file through REVERSE. The file was named #9080.

10. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. The upcast generally shows evidence of the package carrying deeper water up with it. Near the surface the traces are offset by up to 6m. For a few casts the downcast data in the top 10db looks a little suspicious. This is probably due to shed wakes. After editing this should be revisited in case upcast data looks more reliable. Apparent instabilities can be due to drifting through the cast, and there may be some small real instability.

11. DELETE

CLIP was used to remove the first 200 records as they do not contain useful data.

Then DELETE was run on the CLIP files using the following parameters:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

The only warning in the Delete log was for cast #80 – this was because CLIP had removed most of the data since the surface data was not present. DELETE was rerun on this cast using the SHF file and that worked properly.

Cast #9080 (upcast data) was also put through DELETE.

12. CTDEDIT

CTDEDIT was used to remove bad records and to smooth salinity as required. The range of T and S was small for most casts, so that small variations tend to look like major instabilities in the T-S plots. Editing was applied only where there was evidence of problems such as surface variability during equilibration, shed wake contamination or salinity variations that appear to be due to conductivity/temperature mismatch.

All casts required some editing. After the initial editing plots were made and further editing applied to some casts.

Notes were entered in the headers about editing applied.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and while there are some small instabilities. Further editing was not considered appropriate.

Cast 80 and 9080 were compared, and they are close in T-S space though the upcast data is slightly cooler and fresher along σt lines. A few other casts were checked to see if this was a consistent feature and it was not. So the small differences may reflect real changes during the cast.

14. Intercomparisons

Previous Use of CTD – This CTD was used during 2008-13 when 2 bottles indicated that the CTD was low by 0.008, during 2008-18 when 4 bottles showed the CTD to be low by from 0.005 to 0.01 and during 2008-38 when the CTD was found to be low by 0.014.

Repeat Casts – There were no repeat casts.

Historic ranges – No local climatology is available.

15. CALIBRATION

There were 5 salinity samples gathered about 1m above the bottom. When compared to edited CTD data before bin-averaging the differences showed the CTD to be low by 0.041 when all samples were used and 0.018 when one outlier was excluded.

The results of the comparisons from all the cruises that used this CTD so far in 2008 show it to read low by from 0.005 to 0.018. There are some concerns that the Autosal is giving salinity values that are too high, particularly for salinity values <34. The errors from an initial test of Autosal linearity indicated errors on the order of 0.005 to 0.009 for the range of salinity sampled in Broughton. Moreover, there is a tendency in the region for temperatures to increase quite sharply with a consequent tendency to produce erroneous low salinity spikes. There are too many doubts about the comparison to justify recalibration.

No recalibration will be applied to these cruises.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Particulars

13. Possible bottom creep.

90/91. repeat casts at TRC3.0.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2008-33

Dates: 	Start:	7 April 2008	    End:	10 April 2008

Location: Broughton Archipelago                 Vessel:    Princeton

Party Chief: Galbraith M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SeacatPlus | 5299 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/SEACATPLUS v1.6b / 5299

Cruise ID#:	2008-33

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 20Dec07 | Factory |  |  |
| Conductivity | 5299 | 20Dec07 | Factory |  |  |
| Pressure Sensor | 5299 | 18Dec07 | Factory |  |  |