## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruises: 2008-38

Agency: Ocean Sciences Division, Sidney, B.C.

Project: Aquaculture

Area: Broughton Archipelago

Chief Scientist: Galbraith M.

Platform: Princeton

Date: 21 April 2008 – 22 April 2008

Processed by: Germaine Gatien

Date of Processing: 16 May 2008 – 23 May 2008

Number of original CTD casts: 13

Number of casts processed: 13

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#5299) was used.

SUMMARY OF QUALITY AND CONCERNS

All file names were non-standard.

Sheets from the CTD Daily Log were available.

The times in the log book are in Pacific Standard Time. To match the headers in the archived files add 8 hours to the log times.

The soak period for the CTD varied from 1 minute to more than 2 minutes. It is recommended that the SBE19+ be soaked for at least 2 minutes to improve salinity quality, though 1.5 minutes may be sufficient.

Salinity for this cruise has been edited but large errors likely remain. SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times, especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. SeaCat salinity errors are expected to be as high as 0.05units in areas of high gradients.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

CTD Daily Log sheets were obtained.

The original file names were non-standard so were adjusted. 
The dates in the log book are in Pacific Daylight Time.

Computer times were in Pacific Standard Time

This is a new CTD. Configuration file 2008-38-ctd.con is based on the factory calibrations.

- Conversion of Raw Data

- The raw data were converted using conversion file 2008-38-ctd.con.

- Plots were made of a few casts and T and C look reasonable. The soak time was a little shorter for some casts than seen during 2008-13. For some casts with less than 2 minutes of soaking the salinity does not appear to have fully equilibrated before the downcast began.

- FILTER

The same parameters were used as for 2008-13 and 2008-18.

The temperature and conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 1 seconds to increase the pressure resolution.

- CELLTM

Tests for 2008-13 and 2008-18 showed no improvement in data using CELLTM. This step will be skipped.

- DERIVE

Program DERIVE was run to calculate salinity.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers. The station names, bottom depths and positions are missing from the headers. That information is available from the log sheets and in the SeaBird headers. The EDIT HEADERS routine was used to add the lines for that information and a text editor was used to add details. Where bottom depth was not available from the log that line was removed. Where the log recorded bottom depth in feet, it was converted and entered in metres.

8. Checking Headers

Track plots look fine and were added to the end of this report.

The surface check was run. The average surface pressure is +0.09db with very low surface conductivity values; there are no negative values. The pressure calibration looks appropriate and no correction is indicated.

A cross-reference list was produced and checked against the log sheets. No errors were found.

9. Test Plots and Fix Time

Profiles were plotted for all casts and up and down traces compared. There is lots of variability but reasonable correspondence.

ADD TIME CHANNEL was used to add 8 hours to the start times so they are in UTC.

10. SHIFT

There is no history for this instrument, but it was used for 2 other cruises in Feb/March of 2008.

Initial checks show no significant spiking in salinity. During 2007-26, -27, -28 and -29 using a different SBE19+ a setting of +2.25 records was found best, and the same setting was found appropriate for this particular instrument when it was used during 2008-18.

Temperature was shifted by +2.25 records followed by a recalculation of salinity for all casts.

11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min    Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0	     Pressure not filtered

Swells deleted. Warning message if pressure difference of 2.00

Minimum Drop Rate 0.2m/s over 5 records between 10db and 10db above maximum pressure

There were no warnings in the Delete log.

12. CTDEDIT

CTDEDIT was used to remove bad records and to smooth salinity as required.

All casts required some editing. Notes were entered in the headers about editing applied.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value = .000

Average value will be used.	   Interpolated values are NOT used for empty bins

Page plots were examined and no problems noted. No further editing will be done.

14. Intercomparisons

Previous Use of CTD – This CTD was used during 2008-13 when 2 bottles indicated that the CTD was low by 0.007 and 0.010 and during 2008-18 when 4 bottles showed the CTD to be low by from 0.005 to 0.01.

Repeat Casts – There were no repeat casts.

Historic ranges – No local climatology is available.

15. CALIBRATION

There was 1 salinity sample gathered about 1m above the bottom. When compared to edited CTD data before bin-averaging the differences showed the CTD to be low by 0.014.

The results of 2008-13, -18 and -38 all suggest that the CTD is low by from 0.005 to 0.014. There are some concerns that the Autosal is giving salinity values that are too high, particularly for salinity values <34. The errors from an initial test of Autosal linearity indicated errors on the order of 0.005 to 0.009 for the range of salinity sampled in Broughton.

No recalibration will be applied to these 3 cruises.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity, Descent_Rate and Flag.

The HEADEDIT routine was used to add the following comment to the headers:

SeaCat salinity is prone to error due to a mismatch of conductivity and

temperature response times, especially when the descent rate is non-uniform.

Such errors are likely to be larger than calibration errors. SeaCat salinity

errors are expected to be as high as 0.05 units in areas of high gradients.

The standards check routine was run and no further problems found.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

c.) The sensor histories were updated.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2008-38

Dates: 	Start:	21 April 2008	    End:	22 April 2008

Location: Broughton Archipelago                 Vessel:    Princeton

Party Chief: Galbraith M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SeacatPlus | 5299 | No | Yes |

CTD Calibration Information

Make/Model/Serial#: SEABIRD/SEACATPLUS v1.6b / 5299

Cruise ID#:	2008-38

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5299 | 20Dec07 | Factory |  |  |
| Conductivity | 5299 | 20Dec07 | Factory |  |  |
| Pressure Sensor | 5299 | 18Dec07 | Factory |  |  |