## Notes
Run Information
Autosal Salinometer Model 8400A
RunDate=29-Jul-2008
SalinometerSerialNumber=68572
StartZeroReading=0.000030
EndZeroReading=0.000030
StartStandbyNumber=5975.000000
EndStandbyNumber=5973.000000
Bath Temperature=240C
Lab Information
LabTemperature=22.50C
SampleTemperature=22.50C
Analyst Observations and Comments
< | the analysis of these sample ran very smoothly with no complications at all.
< | the Autosal remained stable throughout the Run Time.
< | samples were tightly closed and very minimal to no salt on the rims or in the caps
< | some of the labels were hard to read because of fading or possibly low ink in printer?
< | overall everything ran smooth resulting in confident, stable readings

## 2008-41
2008-41 Salinity Data
the analysis of these sample ran very smoothly with no complications at all.
Information found on sample bottles
written on Logsheets by analyst: Douglas Sieberg
23 | additional info | sample | station | depth/nisk# | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Comments
1 |  | 2 | S1 | bot 2 | 24C | 0.78753 | 6.889e-06 | 0 | 0.78753 | 26.7982
1 |  | 11 | S59 | bot 2 | 24C | 0.92866 | 4.2164e-06 | 0 | 0.92866 | 32.2102
1 |  | 25 | S56 | bot 2 | 24C | 0.92442 | 2.2401e-06 | 0 | 0.92442 | 32.0453
1 |  | 39 | S46 | bot 2 | 24C | 0.79168 | 5.3463e-06 | 0 | 0.79168 | 26.9552
1 |  | 51 | S42 | bot 2 | 24C | 0.79478 | 5.297e-06 | 0 | 0.79478 | 27.0726
1 |  | 51D | S42 | bot 2 | 24C | 0.79481 | 4.5659e-06 | 0 | 0.79481 | 27.0736
1 |  | 68 | S39 | bot 3 | 24C | 0.79519 | 1.1819e-05 | 0 | 0.79519 | 27.0881
1 |  | 84 | S27 | bot 2 | 24C | 0.79423 | 6.6042e-06 | 0 | 0.79423 | 27.0519
1 |  | 100 | S2 | bot 2 | 24C | 0.79268 | 4.818e-06 | 0 | 0.79268 | 26.993
1 |  | 116 | S3 | bot 3 | 24C | 0.79384 | 4.9815e-06 | 0 | 0.79384 | 27.0368
1 |  | 131 | S6 | bot 2 | 24C | 0.79026 | 4.3336e-06 | 0 | 0.79026 | 26.9014
1 |  | 144 | S9 | bot 2 | 24C | 0.78915 | 6.0358e-06 | 0 | 0.78915 | 26.8594
1 |  | 156 | S12 | bot 2 | 24C | 0.79039 | 2.2391e-06 | 0 | 0.79039 | 26.9065
1 |  | 172 | S14 | bot 2 | 24C | 0.78991 | 4.7426e-06 | 0 | 0.78991 | 26.888
1 |  | 188 | S16 | bot 2 | 24C | 0.7878 | 5.5668e-06 | 0 | 0.7878 | 26.8082
1 |  | 199 | S62 | bot 2 | 24C | 0.94765 | 5.2457e-06 | 0 | 0.94765 | 32.9493
1 |  | 205 | S62 | bot 8 | 24C | 0.78014 | 2.2391e-06 | 0 | 0.78014 | 26.5187
1 |  | 210 | S65 | bot 2 | 24C | 0.9464 | 9.7984e-06 | 0 | 0.9464 | 32.9006
1 |  | 221 | ADCP | bot 3 | 24C | 0.96556 | 2.7261e-06 | 0 | 0.96556 | 33.6488
1 |  | 230 | S69 | bot 2 | 24C | 0.97308 | 3.865e-06 | 0 | 0.97308 | 33.9431
1 |  | 242 | S72 | bot 2 | 24C | 0.97375 | 6.5625e-06 | 0 | 0.97375 | 33.9691
1 |  | 255 | S75 | bot 2 | 24C | 0.97399 | 4.9257e-06 | 0 | 0.97399 | 33.9787
1 |  | 269 | S102 | bot 2 | 24C | 0.97503 | 5.3586e-06 | 0 | 0.97503 | 34.0193

## Raw Data with Logsheet info
[Sample Run Header File]
RunID=Line P
ShipsName=John P. Tulley
CruiseNumber=2008-26
Analyst=DGS
| LabTemperature=22.5 |  |  |  | all data recorded by Data Logger Program
| RunDate=29-Jul-2008 |  |  |  | this sheet included Line P cruise 2008-26 and cruise 2008-41 data
SalinometerSerialNumber=68572
StartZeroReading=0.000030
EndZeroReading=0.000030
StartStandbyNumber=5975.000000
EndStandbyNumber=5973.000000
Information found on sample bottles
written on Logsheets by analyst: Douglas Sieberg
216 | Sample ID | Bottle Label | additional info | sample | station | depth/nisk# | DateTime | Bath Temperature | Uncorrected Ratio | Uncorrected Ratio StandDev | Correction | Adjusted Ratio | Calculated Salinity | Calculated Salinity StandDev | Comments
1 | Calibration#1 | K15 |  |  |  |  | 2008-07-29 17:32 | 24C | 0.99984 | 3.541e-06 | 0 | 0.99984
1 | Line P#1 | 13 | 2008-26 | 13 | P2 | bot 1 | 2008-07-29 17:33 | 24C | 0.96788 | 3.4659e-06 | 0 | 0.96788 | 33.7394 | 0.00013687
1 | Line P#2 | 14 |  | 14 | P2 | bot 2 | 2008-07-29 17:38 | 24C | 0.9664 | 3.5298e-06 | 0 | 0.9664 | 33.6813 | 0.00013829
1 | Line P#3 | 15 |  | 15 | P2 | bot 3 | 2008-07-29 17:40 | 24C | 0.95562 | 3.9739e-06 | 0 | 0.95562 | 33.26 | 0.00015705
1 | Line P#4 | 15D |  | 15D | P2 | bot 3 | 2008-07-29 17:42 | 24C | 0.95576 | 4.4374e-06 | 0 | 0.95576 | 33.2656 | 0.00017295
1 | Line P#5 | 16 |  | 16 | P2 | bot 4 | 2008-07-29 17:44 | 24C | 0.95562 | 3.8647e-06 | 0 | 0.95562 | 33.2603 | 0.00015042
1 | Line P#6 | 17 |  | 17 | P2 | bot 5 | 2008-07-29 17:46 | 24C | 0.9478 | 2.4753e-06 | 0 | 0.9478 | 32.9549 | 9.9543e-05
1 | Line P#7 | 18 |  | 18 | P2 | bot 6 | 2008-07-29 17:48 | 24C | 0.9276 | 2.4744e-06 | 0 | 0.9276 | 32.1691 | 9.4885e-05
1 | Line P#8 | 19 |  | 19 | P2 | bot 7 | 2008-07-29 17:50 | 24C | 0.92217 | 3.8749e-06 | 0 | 0.92217 | 31.9582 | 0.000149
1 | Line P#9 | 24 |  | 24 | P2 | bot 12 | 2008-07-29 17:52 | 24C | 0.91984 | 5.49e-06 | 0 | 0.91984 | 31.8676 | 0.00021358
1 | Line P#10 | 56 |  | 56 | P4 | bot 1 | 2008-07-29 17:54 | 24C | 0.98639 | 3.9742e-06 | 0 | 0.98639 | 34.4651 | 0.00015606
1 | Line P#11 | 57 |  | 57 | P4 | bot 2 | 2008-07-29 17:57 | 24C | 0.98609 | 1.4046e-06 | 0 | 0.98609 | 34.4532 | 5.1197e-05
1 | Line P#12 | 58 |  | 58 | P4 | bot 3 | 2008-07-29 17:58 | 24C | 0.9842 | 5.4852e-06 | 0 | 0.9842 | 34.379 | 0.00021574
1 | Line P#13 | 59 |  | 59 | P4 | bot 4 | 2008-07-29 18:00 | 24C | 0.98422 | 4.1786e-06 | 0 | 0.98422 | 34.3797 | 0.00016529
1 | Line P#14 | 59D |  | 59D | P4 | bot 4 | 2008-07-29 18:02 | 24C | 0.98423 | 4.6251e-06 | 0 | 0.98423 | 34.3802 | 0.00017853
1 | Line P#15 | 60 | discard | 60 | P4 | bot 5 | 2008-07-29 18:04 | 24C | 410.1166 | 914.8526 | 0 | 410.1166 | 494696800 | 1106175616
1 | Line P#16 | 61 | flip around | 61 | P4 | bot 6 | 2008-07-29 18:06 | 24C | 0.97943 | 3.1533e-06 | 0 | 0.97943 | 34.1919 | 0.00012256
1 | Line P#17 | 60 | flip around | 60 | P4 | bot 5 | 2008-07-29 18:09 | 24C | 0.9821 | 3.1415e-06 | 0 | 0.9821 | 34.2964 | 0.00012417
1 | Line P#18 | 62 |  | 62 | P4 | bot 7 | 2008-07-30 11:42 | 24C | 0.97612 | 2.7426e-06 | 0 | 0.97612 | 34.062 | 0.00010865 | standby 24+5976 July 30 08
1 | Line P#19 | 63 |  | 63 | P4 | bot 8 | 2008-07-30 11:45 | 24C | 0.97521 | 4.2257e-06 | 0 | 0.97521 | 34.0266 | 0.00016741
1 | Line P#20 | 64 |  | 64 | P4 | bot 9 | 2008-07-30 11:46 | 24C | 0.9744 | 5.5032e-06 | 0 | 0.9744 | 33.9947 | 0.00021503
1 | Line P#21 | 65 |  | 65 | P4 | bot 10 | 2008-07-30 11:49 | 24C | 0.97204 | 4.2256e-06 | 0 | 0.97204 | 33.9025 | 0.00016736
1 | Line P#22 | 66 |  | 66 | P4 | bot 11 | 2008-07-30 11:51 | 24C | 0.97083 | 1.7702e-06 | 0 | 0.97083 | 33.8547 | 6.8803e-05
1 | Line P#23 | 67 |  | 67 | P4 | bot 12 | 2008-07-30 11:53 | 24C | 0.9696 | 2.2391e-06 | 0 | 0.9696 | 33.8067 | 8.7013e-05
1 | Line P#24 | 68 |  | 68 | P4 | bot 13 | 2008-07-30 11:55 | 24C | 0.96534 | 1.1199e-06 | 0 | 0.96534 | 33.6401 | 4.4364e-05
1 | Line P#25 | 69 |  | 69 | P4 | bot 14 | 2008-07-30 11:57 | 24C | 0.95174 | 3.2012e-06 | 0 | 0.95174 | 33.1087 | 0.0001264
1 | Line P#26 | 70 |  | 70 | P4 | bot 15 | 2008-07-30 11:59 | 24C | 0.93696 | 3.1417e-06 | 0 | 0.93696 | 32.5329 | 0.0001197
1 | Line P#27 | 71 |  | 71 | P4 | bot 16 | 2008-07-30 12:04 | 24C | 0.93566 | 4.8362e-06 | 0 | 0.93566 | 32.4821 | 0.0001892
1 | Line P#28 | 72 |  | 72 | P4 | bot 17 | 2008-07-30 12:06 | 24C | 0.92816 | 4.1892e-06 | 0 | 0.92816 | 32.1905 | 0.00016096
1 | Line P#29 | 73 |  | 73 | P4 | bot 18 | 2008-07-30 12:08 | 24C | 0.92238 | 4.5996e-06 | 0 | 0.92238 | 31.9661 | 0.00017706
1 | Line P#30 | 74 |  | 74 | P4 | bot 19 | 2008-07-30 12:10 | 24C | 0.92242 | 3.4522e-06 | 0 | 0.92242 | 31.9678 | 0.00013116
1 | Line P#31 | 126 |  | 126 | P8 | bot 1 | 2008-07-30 12:13 | 24C | 0.98944 | 3.8599e-06 | 0 | 0.98944 | 34.5847 | 0.00015554 | station P8
1 | Line P#32 | 126D |  | 126D | P8 | bot 1 | 2008-07-30 12:16 | 24C | 0.98957 | 4.4576e-05 | 0 | 0.98957 | 34.5898 | 0.0017522
1 | Line P#33 | 127 |  | 127 | P8 | bot 2 | 2008-07-30 12:21 | 24C | 0.98814 | 0.00010823 | 0 | 0.98814 | 34.5338 | 0.0042515
1 | Line P#34 | 128 | noisy | 128 | P8 | bot 3 | 2008-07-30 12:31 | 24C | 0.98778 | 7.0318e-06 | 0 | 0.98778 | 34.5195 | 0.00027391 | noisy readings-possible side arm air blockage
1 | Line P#35 | 129 | noisy-check side arm | 129 | P8 | bot 4 | 2008-07-30 13:27 | 24C | 0.98429 | 3.4658e-06 | 0 | 0.98429 | 34.3824 | 0.00013799 | checked and found blockage in air lines for side arms-dried and clered air lines
1 | Line P#36 | 130 |  | 130 | P8 | bot 5 | 2008-07-30 14:06 | 24C | 0.98218 | 4.9162e-06 | 0 | 0.98218 | 34.2997 | 0.00019056
1 | Line P#37 | 131 |  | 131 | P8 | bot 6 | 2008-07-30 14:08 | 24C | 0.97911 | 4.1137e-06 | 0 | 0.97911 | 34.1791 | 0.00016193
1 | Line P#38 | 132 |  | 132 | P8 | bot 7 | 2008-07-30 14:11 | 24C | 0.97372 | 3.6105e-06 | 0 | 0.97372 | 33.9681 | 0.00013924
1 | Line P#39 | 133 |  | 133 | P8 | bot 8 | 2008-07-30 14:13 | 24C | 0.97287 | 2.7425e-06 | 0 | 0.97287 | 33.9347 | 0.00010447
1 | Line P#40 | 134 |  | 134 | P8 | bot 9 | 2008-07-30 15:17 | 24C | 0.97224 | 4.5996e-06 | 0 | 0.97224 | 33.91 | 0.00018084
1 | Line P#41 | 135 |  | 135 | P8 | bot 10 | 2008-07-30 15:20 | 24C | 0.97136 | 1.1529e-06 | 0 | 0.97136 | 33.8755 | 4.4364e-05
1 | Line P#42 | 136 |  | 136 | P8 | bot 11 | 2008-07-30 15:22 | 24C | 0.97043 | 3.8435e-06 | 0 | 0.97043 | 33.8394 | 0.00015042
1 | Line P#43 | 137 |  | 137 | P8 | bot 12 | 2008-07-30 15:23 | 24C | 0.9688 | 7.7116e-06 | 0 | 0.9688 | 33.7755 | 0.00030228
1 | Line P#44 | 138 |  | 138 | P8 | bot 13 | 2008-07-30 15:26 | 24C | 0.96491 | 3.0873e-06 | 0 | 0.96491 | 33.6231 | 0.00012073
1 | Line P#45 | 139 |  | 139 | P8 | bot 14 | 2008-07-30 15:28 | 24C | 0.94845 | 2.4744e-06 | 0 | 0.94845 | 32.9806 | 9.8045e-05
1 | Line P#46 | 140 |  | 140 | P8 | bot 15 | 2008-07-30 15:30 | 24C | 0.93688 | 6.1868e-06 | 0 | 0.93688 | 32.5295 | 0.00024235
1 | Line P#47 | 141 |  | 141 | P8 | bot 16 | 2008-07-30 15:32 | 24C | 0.93646 | 1.0288e-05 | 0 | 0.93646 | 32.5132 | 0.00040079
1 | Line P#48 | 142 |  | 142 | P8 | bot 17 | 2008-07-30 15:34 | 24C | 0.93599 | 5.5551e-06 | 0 | 0.93599 | 32.4952 | 0.00021636
1 | Line P#49 | 143 |  | 143 | P8 | bot 18 | 2008-07-30 15:36 | 24C | 0.9357 | 2.4908e-06 | 0 | 0.9357 | 32.4839 | 9.869e-05
1 | Line P#50 | 148 |  | 148 | P8 | bot 23 | 2008-07-30 15:40 | 24C | 0.93555 | 4.1725e-06 | 0 | 0.93555 | 32.4781 | 0.00016101
1 | Line P#51 | 199 |  | 199 | P12 | bot 1 | 2008-07-30 15:42 | 24C | 0.99108 | 3.1415e-06 | 0 | 0.99108 | 34.6493 | 0.00012312
1 | Line P#52 | 200 |  | 200 | P12 | bot 2 | 2008-07-30 15:44 | 24C | 0.99104 | 5.246e-06 | 0 | 0.99104 | 34.6476 | 0.00020619
1 | Line P#53 | 201 |  | 201 | P12 | bot 3 | 2008-07-30 15:46 | 24C | 0.99105 | 3.2144e-06 | 0 | 0.99105 | 34.6479 | 0.00012599
1 | Line P#54 | 202 |  | 202 | P12 | bot 4 | 2008-07-30 15:49 | 24C | 0.99058 | 3.4658e-06 | 0 | 0.99058 | 34.6296 | 0.00013927
1 | Line P#55 | 203 |  | 203 | P12 | bot 5 | 2008-07-30 15:51 | 24C | 0.98955 | 8.3651e-06 | 0 | 0.98955 | 34.5889 | 0.00033035
1 | Line P#56 | 204 |  | 204 | P12 | bot 6 | 2008-07-30 15:53 | 24C | 0.98745 | 3.9746e-06 | 0 | 0.98745 | 34.5067 | 0.00015697
1 | Line P#57 | 205 |  | 205 | P12 | bot 7 | 2008-07-30 15:55 | 24C | 0.98619 | 4.5003e-06 | 0 | 0.98619 | 34.4572 | 0.00017791
1 | Line P#58 | 206 |  | 206 | P12 | bot 8 | 2008-07-30 15:58 | 24C | 0.98398 | 4.4783e-06 | 0 | 0.98398 | 34.3702 | 0.00017403
1 | Line P#59 | 207 |  | 207 | P12 | bot 9 | 2008-07-30 16:00 | 24C | 0.98177 | 5.9406e-06 | 0 | 0.98177 | 34.2838 | 0.00023223
1 | Line P#60 | 208 |  | 208 | P12 | bot 10 | 2008-07-30 16:02 | 24C | 0.97846 | 2.4908e-06 | 0 | 0.97846 | 34.1538 | 9.4885e-05
1 | Line P#61 | 209 |  | 209 | P12 | bot 11 | 2008-07-30 16:04 | 24C | 0.97416 | 5.3463e-06 | 0 | 0.97416 | 33.9855 | 0.00020976
1 | Line P#62 | 210 |  | 210 | P12 | bot 12 | 2008-07-30 16:05 | 24C | 0.97214 | 3.3587e-06 | 0 | 0.97214 | 33.9063 | 0.00012951
1 | Line P#63 | 211 |  | 211 | P12 | bot 13 | 2008-07-30 16:07 | 24C | 0.97102 | 3.1309e-06 | 0 | 0.97102 | 33.8624 | 0.00012108
1 | Line P#64 | 212 |  | 212 | P12 | bot 14 | 2008-07-30 16:09 | 24C | 0.97039 | 3.4551e-06 | 0 | 0.97039 | 33.8377 | 0.00013615
1 | Line P#65 | 213 |  | 213 | P12 | bot 15 | 2008-07-30 16:11 | 24C | 0.96989 | 4.3337e-06 | 0 | 0.96989 | 33.8181 | 0.00016797
1 | Line P#66 | 214 |  | 214 | P12 | bot 16 | 2008-07-30 16:13 | 24C | 0.96774 | 3.3588e-06 | 0 | 0.96774 | 33.7338 | 0.00013084
1 | Line P#67 | 215 |  | 215 | P12 | bot 17 | 2008-07-30 16:15 | 24C | 0.9539 | 5.3307e-06 | 0 | 0.9539 | 33.1928 | 0.00020742
1 | Line P#68 | 216 |  | 216 | P12 | bot 18 | 2008-07-30 16:17 | 24C | 0.94053 | 7.6416e-06 | 0 | 0.94053 | 32.6719 | 0.00029618
1 | Line P#69 | 217 |  | 217 | P12 | bot 19 | 2008-07-30 16:19 | 24C | 0.93706 | 4.189e-06 | 0 | 0.93706 | 32.5368 | 0.00016096
1 | Line P#70 | 218 |  | 218 | P12 | bot 20 | 2008-07-30 16:20 | 24C | 0.93627 | 6.7012e-06 | 0 | 0.93627 | 32.5061 | 0.00025988
1 | Line P#71 | 218D |  | 218D | P12 | bot 20 | 2008-07-30 16:23 | 24C | 0.93641 | 3.3588e-06 | 0 | 0.93641 | 32.5113 | 0.00012945
1 | Line P#72 | 219 |  | 219 | P12 | bot 21 | 2008-07-30 16:25 | 24C | 0.93608 | 4.6101e-06 | 0 | 0.93608 | 32.4986 | 0.00017888
1 | Line P#73 | 220 |  | 220 | P12 | bot 22 | 2008-07-30 16:27 | 24C | 0.93622 | 4.6248e-06 | 0 | 0.93622 | 32.5041 | 0.00018082
1 | Line P#74 | 221 |  | 221 | P12 | bot 23 | 2008-07-30 16:32 | 24C | 0.93626 | 6.7762e-06 | 0 | 0.93626 | 32.5055 | 0.00026511
1 | Line P#75 | 243 |  | 243 | P16 | bot 1 | 2008-07-30 16:34 | 24C | 0.99173 | 3.3454e-06 | 0 | 0.99173 | 34.6749 | 0.00012966
1 | Line P#76 | 244 |  | 244 | P16 | bot 2 | 2008-07-30 16:35 | 24C | 0.99162 | 4.0093e-06 | 0 | 0.99162 | 34.6703 | 0.00015952
1 | Line P#77 | 245 |  | 245 | P16 | bot 3 | 2008-07-30 16:37 | 24C | 0.99128 | 1.1199e-06 | 0 | 0.99128 | 34.6569 | 4.5217e-05
1 | Line P#78 | 246 |  | 246 | P16 | bot 4 | 2008-07-30 16:39 | 24C | 0.9907 | 2.4915e-06 | 0 | 0.9907 | 34.6342 | 9.9578e-05
1 | Line P#79 | 247 |  | 247 | P16 | bot 5 | 2008-07-30 16:41 | 24C | 0.99072 | 2.4744e-06 | 0 | 0.99072 | 34.6351 | 9.4885e-05
1 | Line P#80 | 247D |  | 247D | P16 | bot 5 | 2008-07-30 16:46 | 24C | 0.99073 | 4.6251e-06 | 0 | 0.99073 | 34.6356 | 0.00018325
1 | Line P#81 | 248 |  | 248 | P16 | bot 6 | 2008-07-30 16:48 | 24C | 0.98946 | 2.7425e-06 | 0 | 0.98946 | 34.5856 | 0.00010761
1 | Line P#82 | 249 |  | 249 | P16 | bot 7 | 2008-07-30 16:50 | 24C | 0.98752 | 3.3439e-06 | 0 | 0.98752 | 34.5092 | 0.0001316
1 | Line P#83 | 250 | 1250 m | 250 | P16 | bot 8 | 2008-07-30 16:52 | 24C | 0.98609 | 2.4909e-06 | 0 | 0.98609 | 34.4532 | 9.9769e-05
1 | Line P#84 | 251 | 1000m | 251 | P16 | bot 9 | 2008-07-30 16:54 | 24C | 0.98408 | 4.6354e-06 | 0 | 0.98408 | 34.3744 | 0.00018396
1 | Line P#85 | 252 | 800m | 252 | P16 | bot 10 | 2008-07-30 16:56 | 24C | 0.982 | 2.4915e-06 | 0 | 0.982 | 34.2927 | 9.7029e-05
1 | Line P#86 | 253 | 600m | 253 | P16 | bot 11 | 2008-07-30 16:59 | 24C | 0.97873 | 1.1532e-06 | 0 | 0.97873 | 34.1643 | 4.4364e-05
1 | Line P#87 | 254 | 400m | 254 | P16 | bot 12 | 2008-07-30 17:01 | 24C | 0.97512 | 3.6105e-06 | 0 | 0.97512 | 34.023 | 0.00013904
1 | Line P#88 | 255 | 300m | 255 | P16 | bot 13 | 2008-07-30 17:03 | 24C | 0.97257 | 3.9739e-06 | 0 | 0.97257 | 33.9231 | 0.0001534
1 | Line P#89 | 256 | 250m | 256 | P16 | bot 14 | 2008-07-30 17:06 | 24C | 0.97129 | 2.4915e-06 | 0 | 0.97129 | 33.873 | 0.00010062
1 | Line P#90 | 257 | 200m | 257 | P16 | bot 15 | 2008-07-30 17:08 | 24C | 0.97064 | 5.5978e-06 | 0 | 0.97064 | 33.8473 | 0.00021962
1 | Line P#91 | 258 | 175m | 258 | P16 | bot 16 | 2008-07-30 17:10 | 24C | 0.97025 | 2.8792e-06 | 0 | 0.97025 | 33.8323 | 0.00011147
1 | Line P#92 | 259 | 150m | 259 | P16 | bot 17 | 2008-07-30 17:12 | 24C | 0.96889 | 2.2725e-06 | 0 | 0.96889 | 33.7788 | 8.8188e-05
1 | Line P#93 | 260 | 125m | 260 | P16 | bot 18 | 2008-07-30 17:14 | 24C | 0.9626 | 5.7911e-06 | 0 | 0.9626 | 33.5328 | 0.00022418
1 | Line P#94 | 261 | 100m | 261 | P16 | bot 19 | 2008-07-30 17:16 | 24C | 0.94248 | 5.3448e-06 | 0 | 0.94248 | 32.7476 | 0.00020743
1 | Line P#95 | 262 | 75m | 262 | P16 | bot 20 | 2008-07-30 17:18 | 24C | 0.93829 | 4.3338e-06 | 0 | 0.93829 | 32.5847 | 0.00016863
1 | Line P#96 | 263 | 50m | 263 | P16 | bot 21 | 2008-07-30 17:19 | 24C | 0.93778 | 1.3713e-06 | 0 | 0.93778 | 32.5649 | 5.4326e-05
1 | Line P#97 | 264 | 25m | 264 | P16 | bot 22 | 2008-07-30 17:21 | 24C | 0.93733 | 2.2398e-06 | 0 | 0.93733 | 32.5471 | 8.7058e-05
1 | Line P#98 | 265 | 10m | 265 | P16 | bot 23 | 2008-07-30 17:25 | 24C | 0.93679 | 5.4536e-06 | 0 | 0.93679 | 32.5261 | 0.00021083
1 | Line P#99 | 266 | 5m | 266 | P16 | bot 24 | 2008-07-30 17:28 | 24C | 0.9368 | 1.1529e-06 | 0 | 0.9368 | 32.5266 | 4.4364e-05
1 | Line P#100 | 2 | 2008-41 | 2 | 2008-41 S1 | bot 2 | 2008-07-30 17:30 | 24C | 0.78753 | 6.889e-06 | 0 | 0.78753 | 26.7982 | 0.00026176 | 2008-41 S1
1 | Line P#101 | 11 | 2008-41 | 11 | S59 | bot 2 | 2008-07-30 17:34 | 24C | 0.92866 | 4.2164e-06 | 0 | 0.92866 | 32.2102 | 0.00016236 | S59
1 | Line P#102 | 25 | 2008-41 | 25 | S56 | bot 2 | 2008-07-30 17:36 | 24C | 0.92442 | 2.2401e-06 | 0 | 0.92442 | 32.0453 | 8.7876e-05 | S56
1 | Line P#103 | 39 | 2008-41 | 39 | S46 | bot 2 | 2008-07-30 17:39 | 24C | 0.79168 | 5.3463e-06 | 0 | 0.79168 | 26.9552 | 0.00019939 | S46
1 | Line P#104 | 51 | 2008-41 | 51 | S42 | bot 2 | 2008-07-30 17:41 | 24C | 0.79478 | 5.297e-06 | 0 | 0.79478 | 27.0726 | 0.00020104 | S42
1 | Line P#105 | 51AD | 2008-41 | 51AD | S42 | bot 2 | 2008-07-30 17:48 | 24C | 0.79481 | 4.5659e-06 | 0 | 0.79481 | 27.0736 | 0.00017225 | S42
1 | Line P#106 | 68 | 2008-41 | 68 | S39 | bot 3 | 2008-07-30 17:56 | 24C | 0.79519 | 1.1819e-05 | 0 | 0.79519 | 27.0881 | 0.00044879 | S39
1 | Line P#107 | 84 | 2008-41 | 84 | S27 | bot 2 | 2008-07-30 17:58 | 24C | 0.79423 | 6.6042e-06 | 0 | 0.79423 | 27.0519 | 0.00025119 | S2
1 | Line P#108 | 100 | 2008-41 | 100 | S2 | bot 2 | 2008-07-30 18:00 | 24C | 0.79268 | 4.818e-06 | 0 | 0.79268 | 26.993 | 0.00018042 | S2
1 | Line P#109 | 116 | 2008-41 | 116 | S3 | bot 3 | 2008-07-30 18:03 | 24C | 0.79384 | 4.9815e-06 | 0 | 0.79384 | 27.0368 | 0.00018636 | S3
1 | Line P#110 | 131 | 2008-41 | 131 | S6 | bot 2 | 2008-07-30 18:05 | 24C | 0.79026 | 4.3336e-06 | 0 | 0.79026 | 26.9014 | 0.00016427 | S6
1 | Line P#111 | 144 | 2008-41 | 144 | S9 | bot 2 | 2008-07-30 18:07 | 24C | 0.78915 | 6.0358e-06 | 0 | 0.78915 | 26.8594 | 0.00022777 | S9
1 | Line P#112 | 156 | 2008-41 | 156 | S12 | bot 2 | 2008-07-30 18:10 | 24C | 0.79039 | 2.2391e-06 | 0 | 0.79039 | 26.9065 | 8.307e-05 | S12
1 | Line P#113 | 172 | 2008-41 | 172 | S14 | bot 2 | 2008-07-31 10:47 | 24C | 0.78991 | 4.7426e-06 | 0 | 0.78991 | 26.888 | 0.00018019 | S14
1 | Line P#114 | 188 | 2008-41 | 188 | S16 | bot 2 | 2008-07-31 10:52 | 24C | 0.7878 | 5.5668e-06 | 0 | 0.7878 | 26.8082 | 0.00020943 | S16
1 | Line P#115 | 199 | 2008-41 | 199 | S62 | bot 2 | 2008-07-31 10:55 | 24C | 0.94765 | 5.2457e-06 | 0 | 0.94765 | 32.9493 | 0.00020332 | S62
1 | Line P#116 | 205 | 2008-41 | 205 | S62 | bot 8 | 2008-07-31 10:57 | 24C | 0.78014 | 2.2391e-06 | 0 | 0.78014 | 26.5187 | 8.4873e-05 | S62
1 | Line P#117 | 210 | 2008-41 | 210 | S65 | bot 2 | 2008-07-31 11:00 | 24C | 0.9464 | 9.7984e-06 | 0 | 0.9464 | 32.9006 | 0.00038416 | S65
1 | Line P#118 | 221 | 2008-41 | 221 | ADCP | bot 3 | 2008-07-31 11:02 | 24C | 0.96556 | 2.7261e-06 | 0 | 0.96556 | 33.6488 | 0.00010761 | ADCP
1 | Line P#119 | 230 | 2008-41 | 230 | S69 | bot 2 | 2008-07-31 11:04 | 24C | 0.97308 | 3.865e-06 | 0 | 0.97308 | 33.9431 | 0.0001486 | S69
1 | Line P#120 | 242 | 2008-41 | 242 | S72 | bot 2 | 2008-07-31 11:06 | 24C | 0.97375 | 6.5625e-06 | 0 | 0.97375 | 33.9691 | 0.00025459 | S72
1 | Line P#121 | 255 | 2008-41 | 255 | S75 | bot 2 | 2008-07-31 11:09 | 24C | 0.97399 | 4.9257e-06 | 0 | 0.97399 | 33.9787 | 0.00019336 | S75
1 | Line P#122 | 269 | 2008-41 | 269 | S102 | bot 2 | 2008-07-31 11:12 | 24C | 0.97503 | 5.3586e-06 | 0 | 0.97503 | 34.0193 | 0.00020838 | S102
1 | Line P#123 | 354 | 2008-26 LINE P | 354 | P 20 | bot 1 | 2008-07-31 11:56 | 24C | 0.99187 | 7.8471e-06 | 0 | 0.99187 | 34.6803 | 0.00030767
1 | Line P#124 | 355 | 3500m | 355 | P 20 | bot 2 | 2008-07-31 12:01 | 24C | 0.99166 | 5.7678e-06 | 0 | 0.99166 | 34.6721 | 0.00022842
1 | Line P#125 | 356 | 3500m | 356 | P 20 | bot 3 | 2008-07-31 12:04 | 24C | 0.99166 | 3.4659e-06 | 0 | 0.99166 | 34.6721 | 0.00013203
1 | Line P#126 | 357 | 3000m | 357 | P 20 | bot 4 | 2008-07-31 12:07 | 24C | 0.99128 | 3.5942e-06 | 0 | 0.99128 | 34.6571 | 0.00014191
1 | Line P#127 | 358 | 2500m | 358 | P 20 | bot 5 | 2008-07-31 12:09 | 24C | 0.99059 | 4.2256e-06 | 0 | 0.99059 | 34.6299 | 0.00016562
1 | Line P#128 | 359 | 2000m | 359 | P 20 | bot 6 | 2008-07-31 12:11 | 24C | 0.9895 | 7.3235e-06 | 0 | 0.9895 | 34.5872 | 0.00028951
1 | Line P#129 | 360 | 1500m | 360 | P 20 | bot 7 | 2008-07-31 12:15 | 24C | 0.98721 | 4.3195e-06 | 0 | 0.98721 | 34.4972 | 0.00016889
1 | Line P#130 | 361 | 1250 m | 361 | P 20 | bot 8 | 2008-07-31 12:17 | 24C | 0.98567 | 5.0658e-06 | 0 | 0.98567 | 34.4367 | 0.00019895
1 | Line P#131 | 362 | 1000m | 362 | P 20 | bot 9 | 2008-07-31 12:20 | 24C | 0.98374 | 2.2259e-06 | 0 | 0.98374 | 34.3609 | 8.7013e-05
1 | Line P#132 | 363 | 800m | 363 | P 20 | bot 10 | 2008-07-31 12:23 | 24C | 0.98137 | 8.2147e-06 | 0 | 0.98137 | 34.2679 | 0.00032231
1 | Line P#133 | 364 | 600m | 364 | P 20 | bot 11 | 2008-07-31 13:54 | 24C | 0.97807 | 2.8543e-06 | 0 | 0.97807 | 34.1387 | 0.00011125
1 | Line P#134 | 365 | 400m | 365 | P 20 | bot 12 | 2008-07-31 13:56 | 24C | 0.97442 | 4.3336e-06 | 0 | 0.97442 | 33.9953 | 0.0001684
1 | Line P#135 | 366 | 300m | 366 | P 20 | bot 13 | 2008-07-31 13:58 | 24C | 0.97128 | 6.6287e-06 | 0 | 0.97128 | 33.8725 | 0.00025966
1 | Line P#136 | 367 | 250m | 367 | P 20 | bot 14 | 2008-07-31 14:00 | 24C | 0.97008 | 3.6099e-06 | 0 | 0.97008 | 33.8257 | 0.00014279
1 | Line P#137 | 368 | 200m | 368 | P 20 | bot 15 | 2008-07-31 14:03 | 24C | 0.96847 | 3.3593e-06 | 0 | 0.96847 | 33.7624 | 0.00013116
1 | Line P#138 | 369 | 175m | 369 | P 20 | bot 16 | 2008-07-31 14:06 | 24C | 0.96768 | 3.4516e-06 | 0 | 0.96768 | 33.7314 | 0.0001362
1 | Line P#139 | 370 | 150m | 370 | P 20 | bot 17 | 2008-07-31 14:09 | 24C | 0.95961 | 2.8904e-06 | 0 | 0.95961 | 33.4159 | 0.00011219
1 | Line P#140 | 371 | 125m | 371 | P 20 | bot 18 | 2008-07-31 14:11 | 24C | 0.93906 | 2.8765e-06 | 0 | 0.93906 | 32.6146 | 0.00011145
1 | Line P#141 | 372 | 100m | 372 | P 20 | bot 19 | 2008-07-31 14:13 | 24C | 0.93905 | 4.1892e-06 | 0 | 0.93905 | 32.614 | 0.00016328
1 | Line P#142 | 373 | 75m | 373 | P 20 | bot 20 | 2008-07-31 14:16 | 24C | 0.93831 | 2.2394e-06 | 0 | 0.93831 | 32.5853 | 8.5352e-05
1 | Line P#143 | 374 | 50m | 374 | P 20 | bot 21 | 2008-07-31 14:18 | 24C | 0.93807 | 1.1529e-06 | 0 | 0.93807 | 32.576 | 4.5227e-05
1 | Line P#144 | 375 | 25m | 375 | P 20 | bot 22 | 2008-07-31 14:20 | 24C | 0.93808 | 3.707e-06 | 0 | 0.93808 | 32.5764 | 0.00014538
1 | Line P#145 | 376 | 10m | 376 | P 20 | bot 23 | 2008-07-31 14:23 | 24C | 0.93807 | 3.2141e-06 | 0 | 0.93807 | 32.5759 | 0.00012777
1 | Line P#146 | 377 | 5m | 377 | P 20 | bot 24 | 2008-07-31 14:26 | 24C | 0.93807 | 5.6944e-06 | 0 | 0.93807 | 32.5761 | 0.00022469
1 | Line P#147 | 461 |  | 461 | -- | bot 1 | 2008-07-31 14:28 | 24C | 0.99199 | 5.443e-06 | 0 | 0.99199 | 34.685 | 0.00021533
1 | Line P#148 | 462 | 4000m | 462 | -- | bot 2 | 2008-07-31 14:30 | 24C | 0.99189 | 3.4657e-06 | 0 | 0.99189 | 34.681 | 0.00013414
1 | Line P#149 | 463 | 3500m | 463 | -- | bot 3 | 2008-07-31 14:32 | 24C | 0.9918 | 3.5319e-06 | 0 | 0.9918 | 34.6773 | 0.00013853
1 | Line P#150 | 464 | 3000m | 464 | -- | bot 4 | 2008-07-31 14:35 | 24C | 0.9913 | 3.8648e-06 | 0 | 0.9913 | 34.658 | 0.00014997
1 | Line P#151 | 465 | 2500m | 465 | -- | bot 5 | 2008-07-31 14:37 | 24C | 0.99062 | 2.2401e-06 | 0 | 0.99062 | 34.6311 | 8.9601e-05
1 | Line P#152 | 466 | 2000m | 466 | -- | bot 6 | 2008-07-31 14:39 | 24C | 0.9896 | 3.0499e-06 | 0 | 0.9896 | 34.591 | 0.00011982
1 | Line P#153 | 467 | 1500m | 467 | -- | bot 7 | 2008-07-31 14:42 | 24C | 0.98776 | 7.7123e-06 | 0 | 0.98776 | 34.519 | 0.00030494
1 | Line P#154 | 468 | 1250 m | 468 | -- | bot 8 | 2008-07-31 14:44 | 24C | 0.98633 | 2.2394e-06 | 0 | 0.98633 | 34.4628 | 8.7013e-05
1 | Line P#155 | 469 | 1000m | 469 | -- | bot 9 | 2008-07-31 14:46 | 24C | 0.98456 | 2.4912e-06 | 0 | 0.98456 | 34.3932 | 9.4885e-05
1 | Line P#156 | masaya | asked me to run a sample for him |  |  |  | 2008-07-31 14:49 | 24C | 1.0216 | 5.5968e-06 | 0 | 1.0216 | 35.8518 | 0.00022036
1 | Line P#157 | 470 | 800m | 470 | -- | bot 10 | 2008-07-31 14:54 | 24C | 0.9825 | 2.4909e-06 | 0 | 0.9825 | 34.3121 | 9.4885e-05
1 | Line P#158 | 471 | 600m | 471 | -- | bot 11 | 2008-07-31 14:56 | 24C | 0.97978 | 3.4551e-06 | 0 | 0.97978 | 34.2056 | 0.00013386
1 | Line P#159 | 472 | 400m | 472 | -- | bot 12 | 2008-07-31 14:58 | 24C | 0.97598 | 1.1397e-06 | 0 | 0.97598 | 34.0565 | 4.2703e-05
1 | Line P#160 | 473 | 300m | 473 | -- | bot 13 | 2008-07-31 15:00 | 24C | 0.97297 | 5.9204e-06 | 0 | 0.97297 | 33.9385 | 0.00023176
1 | Line P#161 | 474 | 250m | 474 | -- | bot 14 | 2008-07-31 15:02 | 24C | 0.97094 | 4.3343e-06 | 0 | 0.97094 | 33.859 | 0.00016914
1 | Line P#162 | 475 | 200m | 475 | -- | bot 15 | 2008-07-31 15:05 | 24C | 0.96938 | 4.465e-06 | 0 | 0.96938 | 33.7981 | 0.0001752
1 | Line P#163 | 476 | 175m | 476 | -- | bot 16 | 2008-07-31 15:07 | 24C | 0.96723 | 3.6105e-06 | 0 | 0.96723 | 33.7139 | 0.00014389
1 | Line P#164 | 477 | 150m | 477 | -- | bot 17 | 2008-07-31 15:09 | 24C | 0.96048 | 2.2259e-06 | 0 | 0.96048 | 33.4498 | 8.7015e-05
1 | Line P#165 | 478 | 125m | 478 | -- | bot 18 | 2008-07-31 15:11 | 24C | 0.9423 | 6.1184e-06 | 0 | 0.9423 | 32.7405 | 0.00024124
1 | Line P#166 | 479 | 100m | 479 | -- | bot 19 | 2008-07-31 15:13 | 24C | 0.94092 | 5.8913e-06 | 0 | 0.94092 | 32.6869 | 0.00022853
1 | Line P#167 | 480 | 75m | 480 | -- | bot 20 | 2008-07-31 15:15 | 24C | 0.94069 | 3.9595e-06 | 0 | 0.94069 | 32.6781 | 0.00015403
1 | Line P#168 | 481 | 50m | 481 | -- | bot 21 | 2008-07-31 15:18 | 24C | 0.94021 | 2.4909e-06 | 0 | 0.94021 | 32.6594 | 9.4885e-05
1 | Line P#169 | 482 | 25m | 482 | -- | bot 22 | 2008-07-31 15:22 | 24C | 0.94038 | 6.5469e-06 | 0 | 0.94038 | 32.6661 | 0.00025516
1 | Line P#170 | 483 | 10m | 483 | -- | bot 23 | 2008-07-31 15:24 | 24C | 0.94038 | 4.6205e-06 | 0 | 0.94038 | 32.6658 | 0.00017966
1 | Line P#171 | 484 | 5m | 484 | -- | bot 24 | 2008-07-31 15:27 | 24C | 0.94048 | 3.3587e-06 | 0 | 0.94048 | 32.6697 | 0.00012966
1 | Line P#172 | P20-360 | 1500m | 360 | P 20 | bot 7 | 2008-07-31 16:34 | 24C | 0.98719 | 4.3338e-06 | 0 | 0.98719 | 34.4965 | 0.00016798
1 | Line P#173 | P26-463 | 3500m | 463 | P 26 | bot 3 | 2008-07-31 16:36 | 24C | 0.99167 | 1.628e-05 | 0 | 0.99167 | 34.6723 | 0.00064004
1 | Line P#174 | M6-578 |  | 578 | M6 | bot 3 | 2008-07-31 16:39 | 24C | 0.78081 | 1.003e-05 | 0 | 0.78081 | 26.544 | 0.00037931
1 | Line P#175 | M5-579 |  | 579 | M5 | bot 1 | 2008-07-31 16:42 | 24C | 0.78136 | 4.6354e-06 | 0 | 0.78136 | 26.5647 | 0.00017619
1 | Line P#176 | M4-581 |  | 581 | M4 | bot 2 | 2008-07-31 16:44 | 24C | 0.9619 | 3.9742e-06 | 0 | 0.9619 | 33.5055 | 0.00015407
1 | Line P#177 | M4-582 |  | 582 | M4 | bot 3 | 2008-07-31 16:46 | 24C | 0.96088 | 5.0219e-06 | 0 | 0.96088 | 33.4657 | 0.00019476
1 | Line P#178 | M4-582-D |  | 582 D | M4 | bot 3 | 2008-07-31 16:48 | 24C | 0.9609 | 8.3649e-06 | 0 | 0.9609 | 33.4666 | 0.0003264
1 | Line P#179 | M4-583 |  | 583 | M4 | bot 4 | 2008-07-31 16:50 | 24C | 0.95975 | 1.1242e-05 | 0 | 0.95975 | 33.4216 | 0.00044045
1 | Line P#180 | M4-594 |  | 594 | M4 | bot 15 | 2008-07-31 16:52 | 24C | 0.75187 | 4.2766e-06 | 0 | 0.75187 | 25.4534 | 0.00016125
1 | Line P#181 | M3-596 |  | 596 | M 3 | bot 2 | 2008-07-31 16:55 | 24C | 0.95683 | 1.6767e-05 | 0 | 0.95683 | 33.3075 | 0.00065511
1 | Line P#182 | M3-597 |  | 597 | M3 | bot 3 | 2008-07-31 16:58 | 24C | 0.95472 | 4.6355e-06 | 0 | 0.95472 | 33.2251 | 0.00017824
1 | Line P#183 | M3-606 |  | 606 | M3 | bot 12 | 2008-07-31 17:00 | 24C | 0.77722 | 3.6104e-06 | 0 | 0.77722 | 26.4085 | 0.00014035
1 | Line P#184 | M2-608 |  | 608 | M2 | bot 2 | 2008-07-31 17:02 | 24C | 0.95278 | 3.2142e-06 | 0 | 0.95278 | 33.1492 | 0.00012684
1 | Line P#185 | M2-609 |  | 609 | M2 | bot 3 | 2008-07-31 17:04 | 24C | 0.95094 | 1.3713e-06 | 0 | 0.95094 | 33.0775 | 5.2235e-05
1 | Line P#186 | M2-615 |  | 615 | M2 | bot 9 | 2008-07-31 17:07 | 24C | 0.77372 | 3.5954e-06 | 0 | 0.77372 | 26.2762 | 0.00013468
1 | Line P#187 | RI1-617 |  | 617 | RI 1 | bot 2 | 2008-07-31 17:11 | 24C | 0.95423 | 3.8747e-06 | 0 | 0.95423 | 33.2059 | 0.00014849
1 | Line P#188 | Ri1-617 dup |  | 617 D | RI 1 | bot 2 | 2008-07-31 17:14 | 24C | 0.95422 | 4.3337e-06 | 0 | 0.95422 | 33.2053 | 0.00016839
1 | Line P#189 | Ri1-618 |  | 618 | RI 1 | bot 3 | 2008-07-31 17:16 | 24C | 0.95255 | 2.4909e-06 | 0 | 0.95255 | 33.1403 | 9.7246e-05
1 | Line P#190 | Ri1-619 |  | 619 | RI 1 | bot 4 | 2008-07-31 17:19 | 24C | 0.95191 | 1.7702e-06 | 0 | 0.95191 | 33.1152 | 6.8797e-05
1 | Line P#191 | Ri1-629 |  | 629 | RI 1 | bot 14 | 2008-07-31 17:22 | 24C | 0.70064 | 7.3649e-06 | 0 | 0.70064 | 23.5383 | 0.00027377
1 | Line P#192 | Ri2-630 |  | 630 | RI 2 | bot 1 | 2008-07-31 17:25 | 24C | 0.92279 | 2.2391e-06 | 0 | 0.92279 | 31.9822 | 8.5299e-05
1 | Line P#193 | Ri2-631 |  | 631 | RI 2 | bot 2 | 2008-07-31 17:28 | 24C | 0.91266 | 3.9746e-06 | 0 | 0.91266 | 31.5893 | 0.00015279
1 | Line P#194 | Ri2-632 |  | 632 | RI 2 | bot 3 | 2008-07-31 17:30 | 24C | 0.79469 | 4.1892e-06 | 0 | 0.79469 | 27.0692 | 0.00015913
1 | Line P#195 | Ri2-633 |  | 633 | RI 2 | bot 4 | 2008-07-31 17:33 | 24C | 0.73671 | 4.3337e-06 | 0 | 0.73671 | 24.8848 | 0.00016433
1 | Line P#196 | Ri3-634 | 20m | 634 | RI 3 | bot 1 | 2008-07-31 17:35 | 24C | 0.92291 | 2.5219e-06 | 0 | 0.92291 | 31.9867 | 9.7284e-05
1 | Line P#197 | Ri3-635 | 20m | 635 | RI 3 | bot 2 | 2008-07-31 17:38 | 24C | 0.91358 | 3.4657e-06 | 0 | 0.91358 | 31.6248 | 0.00013329
1 | Line P#198 | Ri3-636 | 5m | 636 | RI 3 | bot 3 | 2008-07-31 17:40 | 24C | 0.78483 | 4.1786e-06 | 0 | 0.78483 | 26.6961 | 0.00015891
1 | Line P#199 | check |  |  |  |  | 2008-08-05 11:04 | 24C | 0.99982 | 3.5392e-06 | 0 | 0.99982 | 34.993 | 0.00013811
1 | Calibration#2 | K15 |  |  |  |  | 2008-08-05 11:06 | 24C | 0.99984 | 6.6996e-06 | -1.9073e-06 | 0.99984
1 | Line P#200 | Ri3-637 |  | 637 | RI 3 | bot 4 | 2008-08-05 11:09 | 24C | 0.74811 | 3.974e-06 | -2.0266e-06 | 0.74811 | 25.3122 | 0.00014926
1 | Line P#201 | Ri4-639 |  | 639 | RI 4 | bot 2 | 2008-08-05 11:12 | 24C | 0.95272 | 3.8622e-06 | -1.967e-06 | 0.95272 | 33.1468 | 0.00015026
1 | Line P#202 | Ri4-640 |  | 640 | RI 4 | bot 3 | 2008-08-05 11:15 | 24C | 0.95203 | 2.5036e-06 | -2.1458e-06 | 0.95203 | 33.1201 | 9.7279e-05
1 | Line P#203 | Ri4-650 |  | 650 | RI 4 | bot 13 | 2008-08-05 11:19 | 24C | 0.75641 | 3.4665e-06 | -1.967e-06 | 0.75641 | 25.6242 | 0.00013108
1 | Line P#204 | Ri5-651 |  | 651 | RI 5 | bot 1 | 2008-08-05 11:22 | 24C | 0.56471 | 1.1196e-06 | -1.9073e-06 | 0.56471 | 18.5542 | 4.0091e-05
1 | Line P#205 | Ri6-652 |  | 652 | RI 6 | bot 1 | 2008-08-05 11:26 | 24C | 0.74443 | 1.7367e-05 | -1.967e-06 | 0.74443 | 25.1742 | 0.00065074
1 | Line P#206 | P26-563 | N4 | 563 | P26 | 2500m | 2008-08-05 11:30 | 24C | 0.9905 | 3.9988e-06 | -1.967e-06 | 0.9905 | 34.6265 | 0.0001569
1 | Line P#207 | Loop2 | 2008-26 | Loop2 | -- | -- | 2008-08-05 11:33 | 24C | 0.93828 | 2.7426e-06 | -1.967e-06 | 0.93828 | 32.584 | 0.00010865
1 | Line P#208 | Loop3 | 2008-26 | Loop3 | -- | -- | 2008-08-05 11:35 | 24C | 0.93778 | 2.2397e-06 | -1.9073e-06 | 0.93778 | 32.5647 | 8.7024e-05
1 | Line P#209 | Loop4 | 2008-26 | Loop4 | -- | -- | 2008-08-05 11:38 | 24C | 0.93808 | 3.541e-06 | -1.9073e-06 | 0.93808 | 32.5763 | 0.00013623
1 | Line P#210 | Loop5 | 2008-26 | Loop5 | -- | -- | 2008-08-05 11:40 | 24C | 0.93798 | 2.7425e-06 | -1.967e-06 | 0.93798 | 32.5723 | 0.00010447
1 | Line P#211 | Loop6 | 2008-26 | Loop6 | -- | -- | 2008-08-05 11:43 | 24C | 0.93787 | 2.2258e-06 | -1.967e-06 | 0.93787 | 32.5682 | 8.5299e-05
1 | Line P#212 | Loop7 | 2008-26 | Loop7 | -- | -- | 2008-08-05 11:46 | 24C | 0.93621 | 2.0945e-06 | -1.967e-06 | 0.9362 | 32.5034 | 8.1169e-05
1 | Line P#213 | Loop8 | 2008-26 | Loop8 | -- | -- | 2008-08-05 11:48 | 24C | 0.92703 | 2.4908e-06 | -1.9073e-06 | 0.92702 | 32.1465 | 9.6609e-05
1 | Line P#214 | Loop9 | 2008-26 | Loop9 | -- | -- | 2008-08-05 11:50 | 24C | 0.91666 | 1.8035e-06 | -1.9073e-06 | 0.91666 | 31.7444 | 6.817e-05
| 216 | Total including cal checks, 2008-26 and 2008-41 cruises