## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 2 April 2025 | Updated channel names & formats in TOB files. G.G. |
| 18 August 2010 | Transmissivity was corrected. See end of report for details. |

## PROCESSING NOTES

Cruise: 2008-50

Agency: OSD

Location: WCVI

Project: LaPerouse

Party Chief: Juhasz T.

Platform: John P. Tully

Date: Sept. 11, 2008 – September 21, 2008

Processed by: Germaine Gatien

Date of Processing: 8 February 2009 – 18 March 2009

Number of original CTD casts:  44	Number of CTD casts processed:  43 (no downcast for #5)

Number of bottle casts: 		44	Number of bottle casts processed: 42 (no samples from #21&31)

Number of original TSG files: 	 1	Number of TSG files processed:     1

# INSTRUMENT SUMMARY

A Sea-Bird Model SBE 911+ CTDs (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), SBE 43 DO sensor (#1176 – probably on the primary pump), a Seapoint Fluorometer (#2228) with a 10X cable (probably on the secondary pump), a Biospherical QSP-400 PAR sensor (#4656) and an altimeter (#1252). Based on details from the log book for 2008-10 which immediately preceded this cruise it is assumed that the deck unit was a model 911+ (#0424) and the logging computer was an HP Compaq. Seasave v7.16 was used.

A thermosalinograph (Sea-Bird 21 S/N 2487) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P) and a flow meter; temperature sensor #4652 was mounted at the intake.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good order. Especially useful were comments on the CTD deployment method and rosette sampling routines, and note was made of any deviations from those routines. While there was no list of equipment, this was because it immediately followed another cruise using the same sensors. However, this fact is not necessarily known to the person who does the processing many months later and the party chief is not always available to provide the information. So it would be helpful to at least make note in the log of why the information is missing or where it can be found, and especially to note when a TSG was present, because the TSG data sometimes go astray and the processor has no way of knowing it exists if there is no mention of it in the log.

As on the previous cruise Niskin Bottle #7 failed to close. There were also problems with Niskin #1 for the latter part of the cruise. The words “did not fire” were used about that one, yet the data do appear in the rosette files and the rosette sheet has the “confirmed pressure” ticked, so this is assumed to mean that the bottle did not close.

The downcast data for event #5 were not archived, but the upcast data were, so the bottle file is fine.

As during 2008-10 the comparison of CTD fluorescence and chlorophyll looks a little odd. For that cruise it was speculated that the pump might not have been working well. While that may be valid, the data from this cruise suggest that the variability is real; in sampling a wide variety of waters a wide range was encountered in the ratio of CTD fluorescence to extracted chlorophyll and geographic patterns are seen. For this cruise the bottle stops were relatively long and quiet, so the comparison may better reflect how the sensor behaved for downcasts.

No recalibration was applied to salinity despite large differences between bottles and CTD. There are doubts about the Autosal salinometer, and a post-cruise calibration showed little drift in the primary temperature and conductivity sensors since the previous calibration in May 2009.

As for 2008-10 a new algorithm was used for the conversion of the SBE Dissolved Oxygen sensor data. This is expected to produce more reliable data below 1000db. The results look encouraging from the first few cruises for which this method was used.

The dissolved oxygen data in the CTD files should be considered:

•	±1 ml/l from 0 – 75db

•	±0.4 ml/l from 75 – 200db

•	±0.2 ml/l below 200db

This is not as good as usual (especially near the surface) because there were many near-shore casts with large temperature gradients which affect the quality of the SBE DO data.

Comparison with bottles suggests that the CTD dissolved oxygen sensor performed well in the low oxygen waters of Effingham Inlet.

There is 1 thermosalinograph file with both intake temperature and lab temperature. The salinity has some odd jumps in values suggestive of fresh water getting into the system occasionally. This was noted during the time when the ship was steaming to and from Bowie Seamount.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. The log book lacked a list of equipment, but given that the cruise immediately followed 2008-10 and had the same sensors listed in the configuration files it was assumed that the CTD was configured in the same way. A few problems noted in the log book were with the altimeter, noise in the fluorometer for one cast and a few bottles not closing and cast #5 downcast data were not archived. The CTDs were lowered to 10m at which point the pump was started, the CTD was then returned to the surface and lowered after a 30s wait. This information is very useful in processing. The technique should ensure no problems with the surface data, but if any casts look odd at the surface at the CTDEDIT stage, then the pre-DELETE data should be examined to see if there would be any benefit in removing the leading upcast data.

Salinity, extracted chlorophyll and nutrients data were obtained in spreadsheet format.

The titrated dissolved oxygen files were provided in individual files without flag channels or comments, though there are comments on the rosette sheet.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

##### Conversion of Raw Data

The configuration files were all the same and were all correct except that the parameters were missing for the new algorithm for the SBE Dissolved Oxygen sensor. Those were entered, picking the same value for E (3.85*e-02) as was determined in studies of data from cruise 2008-27. The edited file was saved as 2008-50-ctd.con.

Data were converted using the configuration file 2008-50-ctd.con and SBE Data Processing v7.18c.

A few casts were examined and all expected channels are present and the data look ok. The upcast data are much noisier than the downcast with larger differences between the temperature and conductivity channels. The altimeter data mostly look ok, but in one case it is fine until the bottom, where it suddenly goes to 0 values without moving downwards. So the altimeter header entries will have to be checked for reasonableness. The dissolved oxygen and transmissivity data look as expected.

Of special note are the fluorescence data; upcasts and downcast are notably different. One cast was examined that shows near surface upcast data to be very low compared to the downcast until the stop for the surface bottle. The fluorescence then increases to be close to the downcast value. There were problems noted in processing 2008-10 that were thought to be either in the chlorophyll analysis or a pump problem. These data seem to support the idea that the pump was not operating at peak efficiency, possibly because the CTD was mounted wrong. The fluorometer was on the secondary pump, so the secondary T/S data may also be compromised. Some fluorescence is off-scale. The bottle stops were at least 30s long and the descent rate during stops shows conditions were relatively quiet, so if the pump was not working well during upcasts, the fluorometer would have time to “catch up”; thus the comparison may better reflect how the fluorometer behaved during the downcast.

Rosette files were converted using a start time of -5s and duration of 10s. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers and the output files were named *.BOT. All BOT files were plotted. The only significant outliers were in cast #15. CTDEDIT was used to clean a few primary salinity points in the file and to enter a comment about the editing in the header of the file. The edited file was copied to BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

For the data from 2008-10 tests were run on a few casts using settings (α = 0.01, β=7), (0.02, 7), (0.03, 7), (0.02, 9), (0.03, 9) and (0.0245, 9.5) to see what settings gave the best results. The best choice overall was (0.0245, 9.5) for the primary channel and (0.03, 9) for the secondary. Since this cruise had the same equipment configuration, only a quick test was done to ensure that the choices for 2008-10 worked well for these data too. CELLTM was run on all casts using (α = 0.0245, β=9.5) for the primary conductivity and (α = 0.03, β=9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A few casts were plotted to check for agreement between the pairs of T and C sensors. The casts are mostly quite shallow. The differences are all small, but those in conductivity and salinity do have some pressure dependence, which could be another indication of one pump not operating at full efficiency. The deeper casts are all mid-cruise so there is no basis to comment on time dependence

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 28 | 500 925 | +0.0008 +0.0008 | +0.0001 +0.0003 N | +0.0003 -0.0006 | Noisy/high |
| 29 | 500 925 1450 | +0.0008 VN +0.0009 VN +0.0008 | +0.00025 N +0.00013 N +0.00008 N | +0.0019 VB +0.0006 XN +0.0001 N | Noisy/very high |
| 31 | 500 925 1050 | +0.0008 +0.0008 XN +0.0009 XN | +0.0002 +0.0001 N +0.0001 N | +0.0014 +0.0006 N +0.0003 N | Mod/High |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and checked. Corrections were made to a few station names. Random checks of times and positions turned up no errors.

The cruise track was plotted. No problems were found.

The average surface pressure is 1.4db with values as low as -0.1; this is low for the Tully and a few other recent cruises using this pressure sensor have shown similar results, so it looks like it is time to increase the pressure offset a little. At the recalibration stage an offset of 0.2db will be applied and in future a setting of +6.7db will be used for this sensor.

The altimeter readings from the headers of the CLN and BOT files were exported to spreadsheets and all casts were checked. Plots were made and the log book was checked. Despite some serious spiking near the bottom, the algorithm worked well. There were two casts in Broughton Archipelago during which the altimetry went down while stopped at the bottom, but this is assumed to be caused by drifting into shallower waters – the change was not large.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There were some related problems in assigning the numbers:

- For casts #33, 34, 36, 38 and 41 the log notes a single bottle at the surface for each but 2 bottles actually fired, both at the surface. Only one sample number was assigned. The CTD data in the ROS files show they recorded right at the surface. There is no note in the log about a problem, but on the rosette sheets the line for Niskin #1 is frequently crossed out from cast #35 onwards and occasionally there is a note “didn’t fire”. So it was assumed that for these casts, the sample came from bottle Niskin #2. A pad value was entered for Niskin #1 in these cases.

- Cast #45 – The first bottle misfired. No sample number had been assigned, so a pad value was entered in the addsamp file.

- Cast #51, Niskin #7 was assigned a sample number, that was scratched out and the number was used again for Niskin #8. Both did appear in the rosette file, so a pad value was entered for Niskin #7.

- Casts 52/53 – Sample #342 was assigned twice. In the case of cast #53 it was not sampled because of firing problems, so a pad value will be entered for that bottle.

The ADDSAMP file was edited to remove all lines with pad values when it was found there were problems with the merging process. The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The BOT files were then bin-averaged.

SALINITY

Salinity was analyzed at IOS on an Autosal including 7 duplicates. The Autosal data were delivered in spreadsheet format. The spreadsheet was simplified and saved as 2008-50-sal.csv. Duplicates were copied to spreadsheet 2008-50-sal-duplicates.xls, and in the main spreadsheet the 2 values were replaced with an average and flag “f” was entered and it was then converted to individual SAL files.

For the duplicate study averages and differences were found. Of the 7 duplicates 1 had a difference of 0.01 and another >0.67. It was discovered that the latter pair were not actually duplicates, but mislabelled samples from 2 different casts. The values were reassigned with a “c” flag and comment explaining the confusion. For the other pair a “c” flag was assigned and the value closest to the CTD data was used, but the original 2 values were placed in a comment to go into the headers.

Sp = 0.0028 using all 6 duplicates, Sp=0.0012 if 1 outlier with a difference >0.01 is excluded and Sp=0.0007 if 2 are excluded, where

Sp = Square Root (sum of squares of differences / 2*number of pairs)

For 2008-10 Sp=0.0040 using all 19 duplicates and 0.0012 when 6 outliers were excluded.

For 2008-26 Sp=0.0012 using all 7 pairs and 0.0005 when 1 pair were excluded.

For 2008-27 Sp=0.0027 using all 10 pairs and 0.0008 if 3 out of 10 duplicates were excluded. For that cruise 3 of the 4 largest differences, the “a” and “b” samples were analyzed in different sessions.

There are known to have been problems with the Autosal during the period when the samples from 2008-50, 2008-10 and 2008-27 were run.  The poorest duplicates were the last 2 pairs analyzed in the group, so stability was probably an issue. Repairs were done to the Autosal after these analyses and a loose hose was found. (See 2008-50-sal-duplicates.xls.)

DISSOLVED OXGYEN

Dissolved oxygen files (*.oxy) were provided. They lacked flag and comment channels. The ADD CHANNEL routine was run to add those, and then the rosette sheets were reviewed to determine what samples needed flags and comments. Those flags were added and a few corrections were made as indicated in notes in the rosette log:

- The sample numbers/nominal pressures in the ADD file for cast #12 were confused. They were changed to match the rosette log being careful to keep the values and flask information in line. There were duplicates for sample #99, but the analyst reported that the first was too low and there was a bubble on the lens, so the second value was used rather than an average.

- Records were out of order for cast #13.

- There were two lines for one sample for cast #16 – the first had a zero value - it was removed.

- Cast #17 has one sample # entered incorrectly – it was replaced with the rosette sheet value.

- The sample from Niskin #8, cast #22 was run as #174, but should have been called #175.

- The sample from Niskin #2, cast #43 was run as #252, but should have been called #253.

- The sample from Niskin #9, cast #51 is in the OXY file as 335, but the rosette sheet indicates it should be #334.

- There were a few samples (332, 334, 364) with rosette notes that indicate there was some confusion about flask numbers. It is unknown if the volumes were corrected, so “c” flags were assigned. (After running Compare the flags were reviewed – one value looks a little off and the others seemed ok, but it seems wise to leave the flag anyway. The header note was updated.

There were duplicates but no averaging had been applied. A spreadsheet was prepared, 2008-50-DO-duplicates.xls, in which the individual values were recorded, averages determined and differences calculated to see if there are outliers. The average values were placed in the first of the duplicate lines in the ADD files, a flag “f” was added to it and the second duplicate line was removed.

Four outliers were identified with differences >3%. For 2 of those “c” flags were assigned and one with a 43% difference was given a “d” flag and notes with the original values were placed in the headers. Another pair with a 5% difference had low values, so the differences are not large, so they will not be flagged. (After running COMPARE  the pair that were “d” flagged were replaced with the single value that looked best, and the flag was changed to “c”. ) The flagged DO values were checked against a list of flask numbers thought to have had errors in the recorded volumes during the June-Dec. 2008 period and none were on the list.

Using all duplicates Sp = 0.048 and excluding 3 outliers Sp=0.017 where

Sp = Square Root (sum of squares of differences / 2*number of pairs)

When the 3 largest outliers are excluded Sp=0.017.

For 2008-26, 2008-27 and 2008-50 the results were Sp=0.090, 0.085 and 0.075 when all data were included, so these results look very good.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2008-50nuts.xls which was simplified and reordered on sample number and saved as 2008-50nuts.csv. Extraneous headers were removed, header names were changed to standard format and the file was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were provided in spreadsheet format. This was simplified and reordered on sample number and saved as 2008-50chls.csv. Extraneous headers were removed, header names were changed to standard format and the file was then converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in 4 steps. (Output: MRG1, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only.

At this point it was discovered that PAR data had not been converted. There was PAR data available for casts 1-28, so new ROS files were created which were converted to IOS HEADER format, cleaned to add event numbers, sample numbers were added, and the files were averaged on bottle numbers creating SAMAVG2 files. Those files were merged with the original converted files which had been renamed SAMAVG1. The merged files were called SAMAVG.

File MRGCLN1 was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

There is a lot of scatter in the plots of differences between the CTD salinity and Bottle salinity. The only extreme outlier was the surface sample from cast #42, but the standard deviation in the CTD salinity is extremely high, so the bottle value is probably fine for that sample.

When samples from above 10m are excluded both CTD salinity channels are found to be low by ~0.006, and when only the 6 bottles below 200m are included they are low by ~0.01. Similar results were found for 2008-10 and 2008-61. Doubts have been raised about the Autosal performance in a single session during which all three of those cruises were analyzed.

There has been a concern about the linearity of the Autosal. Results in the first half of 2008 suggest that the Autosal was reading high at lower salinity. It is believed this problem has been resolved, but a check was made by plotting differences against CTD salinity. There were many outliers at low salinity, but those are all samples from the top 10m. When those are removed there is no evidence of the non-linearity noted in the Autosal earlier in 2008. However, there is a lot of scatter so this is not strong evidence that the linearity is good either. (See 2008-50-sal-comp1.xls.)

Dissolved Oxygen

COMPARE was run using pressure as the reference channel. There were 6 severe outliers, most are in areas of high gradient in temperature and dissolved oxygen.

- Cast #4, sample 37 – all the CTD data look odd, high temperature, low transmissivity, very high fluorescence. The high SBE DO is probably due to the high temperature – that sort of response is a limitation in the SBE sensor and something Sea-Bird are working on improving.

- Cast #16 – the bottle value for the outlier is already flagged “d”

- Cast #21 – this problem led to the discovery of a discrepancy between the log book and rosette log and hence sample numbers. Cast #21 had only 1 bottle. The cast called #21 in the rosette log had just one stop planned but no sampling was done.

- Cast #25 – already flagged “d”

- Cast #43 – already flagged “d”

- Cast #44 – already flagged “d”

The best fit was against differences between the sensor and bottles. When the severe outliers were removed and then other outliers based on residuals, the fit for all data was:

CTD-BOT = 1.0826 DOX-CTD - 0.025

The results for the previous cruise 2008-10 were:

CTD-BOT = 1.0722 DOX-CTD + 0.0106

and for cruise 2008-61 which followed it:

CTD-BOT = 1.1390 DOX-CTD - 0.011

The CTD sampled low DO waters in many casts in Effingham Inlet in the final section of the cruise. To test how well the sensor performed in low DO, a fit was done that excluded all DO<1ml/l and used the same criteria for other exclusions. The fit was very similar:

CTD-BOT = 1.082 DOX-CTD - 0.023

Sea-Bird report that the sensors no longer have a problem with hypoxic sampling, and these results are certainly very encouraging.

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity. Most of the significant outliers had either bottle values that had already been flagged, or the CTD data had a high standard deviation so looked unreliable. However, one error was discovered in cast #22, where it was discovered there were 2 values for one sample # in the ADD file. One was a low value and clearly from a bad titration; the wrong one had been merged into the bottle file. That was fixed.

Fluorescence

COMPARE shows that the fluorescence is about 63% of extracted chlorophyll values. All but one sample are from the LD and LG lines. A single sample from Effingham Inlet has a similar ratio of FL to CHL. That ratio shows little variation with pressure or with CHL or CTD Fluorescence values, but it does show variation with time/position. When the data from the LD line were plotted the offshore casts had a ratio of about 0.8 while the near-shore ones were closer to 0.4. Only cast #34 looking intermediate. On the LG line which was north of LD a similar pattern is seen except the offshore ratios are higher at about 1.1 and the inshore lower at 0.3 to 0.4.

##### 13. Shift

Because the same equipment was used for 2008-10 which immediately preceded this cruise, it was presumed that the same shifts would be suitable.

Fluorescence

SHIFT was run applying a shift of +24 records to the fluorescence channel.

Conductivity

All casts were put through two runs of SHIFT using -0.3s for the primary sensor and +0.2s for the secondary. (Output: *.SHFC0 and SHFC1).

Dissolved Oxygen

SHIFT was run using +50 records for all casts.

Afterwards, tests were run to ensure that the choices did improve the data:

- For the fluorescence the improvement is not as good as usual. However, from the comparison of chlorophyll with fluorescence for 2008-10, we have some doubts about the fluorometer values during upcasts. The values seem very low and may have been affected by poor pump performance due to the way the CTD was mounted in the rosette. The downcast data look more reliable. Applying the usual setting of +24 records is the best choice.

- The primary salinity is greatly improved. While the secondary is also improved, it looks like a better setting might emerge if more tests were done, but this channel is unlikely to be archived anyway. If it is needed for a cast, then this step can be rerun.

- The dissolved oxygen shift did a good job.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #5 which had no downcast data.

Because the log noted a variation in deployment method for cast #51, the DEL file was checked to ensure that the first downcast to 11db was not selected, and it was not.

The DELETE output files were copied to EDT.

##### DETAILED EDITING

It is not obvious which sensor pair to archive. The bottle comparison is not trusted, and the two channels look quite close. The primary sensor has been used for all 4 previous uses of this equipment, including 2008-27 and 2008-10 which preceded this one and used all the same equipment. A few files were examined in detail and no significant differences in noise level or spikes were seen. So the primary sensors were selected.

Graphical editing was done using program CTDEDIT with output EDU. On-screen plots of descent rate and pump status were also used. All casts required some editing.

The EDU files were then copied to EDT.

Two errors were discovered at this point. First, the PAR channel had not been converted. Second, the wrong configuration file had been used at the DERIVE stage, so the salinity and dissolved oxygen files were wrong. Since the editing was not affected by these errors, new CNV files were prepared and put through all steps through all steps up to the EDITING phase. They were then merged with the EDT files using sample number to ensure the right match; the output files were named NEW.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: Both sensors were used for 2008-26, 2008-41, 2008-27 and 2008-50. The value of the bottle comparisons were limited by either salinometer problems or very few bottles. The primary looked low by ~0.003 for 2 of them, was within 0.001 for another and was low by 0.007 for the most recent. The results for the secondary were highly variable, being high by 0.002 for 2, low by 0.002 for another and low by 0.009 for the most recent. Results from other cruises using different sensors has cast doubts on Autosal results from around the time that the samples from 2008-10 were analyzed. Problems were found with the Autosal after those analyses were done.

2. Dissolved Oxygen – Since the sensor was recalibrated in Feb. 2007 it has been used for 11 other cruises that have been processed. There has been a slow but fairly steady drift in calibration. This sensor has been found to give quite good detail and no obvious problem with hypoxic sampling.

3. Pressure – This sensor is older and prone to drift. An offset of +6.5db has been used since March 2008, but for cruise 2008-61 in November, it was set to +6.7db and 2008-10 was recalibrated by adding 0.2db. That looks appropriate for this cruise as well.

Post-cruise calibration – When the primary sensors were examined at the factory in January 2009 the “as received” drift was -0.0001 salinity units per month in conductivity and the temperature had drifted higher at a rate of -~0.00054C˚/yr. So for 2008-50 one could estimate that the effect on salinity from both these drifts was towards values low by either 0.0005 or 0.0011 depending on whether we assume the drift was linear with time, or that it varied with use. This was probably that last cruise on which these sensors were deployed before recalibration.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were only 2 minor excursions from those ranges with temperature slightly high around 1000db for cast #29 and 700db for cast #32. None of these excursions suggest systematic problems with the instrument.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

No recalibration will be applied to salinity since the post-cruise calibration shows little drift.

File 2008-50-ctd.ccf was prepared to add 0.2db to the pressure and to apply the following equation to the CTD Dissolved Oxygen channel:

CTD-BOT = 1.0826 DOX-CTD - 0.025

This was applied to SAM and MRGCLN2 bottle files.

COMPARE was then rerun. This showed that the DO correction was effective. (See 2008-50-sal-comp2 and 2008-50-dox-comp2.xls.)

The same DO calibration was applied to the edited downcast files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but usually a further correction is applied to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Files were bin-averaged to 0.5m-bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The differences were small with the best fit being that of differences against DOX_CTD:

DOX_BOT = 0.9148 * DOX_CTD +0.146

There was a lot of scatter in the fits and the choice of what to exclude is quite arbitrary. A variety of fits were tried and used to recalibrate the thinned files. COMPARE was rerun to test the choice. The fit given above produced the best results as seen in the final plots vs pressure or DO concentration.  (See 2008-50-dox-comp3.xls. for the initial comparison of thinned downcast data to bottles. See 2008-50-comp4.xls for the comparison after the final recalibration of the thinned downcast files.)

Recalibration using file 2008-50-recal2.ccf was applied to the downcast files only. (Output:COR2)

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. Two casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts #29-55.

A second SBE DO channel (with umol/kg units) was added and the channels were reordered to put the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

•	±1 ml/l from 0 – 75db

•	±0.4 ml/l from 75 – 200db

•	±0.2 ml/l below 200db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found.

The final files were named CTD.

The header check was rerun and profile plots made; no problems were found..

The track plot looks ok. The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface ranged from ~68%(Kn0) to ~150% (station TrC2) with moderate values in the off-shore casts, high values in Effingham Inlet and the inlets of Broughton Archipelago, but very low in Queen Charlotte Sound. These data were collected in very different environments, so some variability is expected, but it is likely that the high temperature gradients near the surface are leading to some poor CTD DO values. Looking at bottle values does suggest that the surface SBE DO data for the Broughton and Effingham areas are a little too high.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix headers, formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

Header Check was rerun and a few problems were found and resolved.

The CHE files for casts #21 and #31 were removed because there was no sampling – the Niskin bottle didn’t close.

A few checks were made against rosette log sheets and all expected samples appear to be present.

A cross-reference list was produced and no problems were found

##### 23. Thermosalinograph Data

The TSG data were provided in 1 HEX files.

a.) Checking calibrations

All calibration information was correct in the con file.

b.) The data were converted to CNV files using 2008-50-0001.con.

The output file was then converted to IOS SHELL format files *.ios.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to derive separate date and time channels from the Time:Julian.

Time-series plots were examined on-screen. The flow rate changed from ~0.7 to ~1.3 on Sept. 14. There are some odd sudden changes in salinity that do not appear to be seen in temperature. Given that the cruise visited inlets, there could be some large local variability, so this may be ok, but this should be explored later in CTDEDIT. The temperature difference is very noisy in patches, especially near the end which is probably in Effingham Inlet. There are other sections where the differences are quite flat.

c.)  Checking Time Channel

The CTD data, after editing and metre-averaging, were thinned to reduce the files to a single point at or within 0.3db of 4db and exported to a spreadsheet which was saved as 2008-50-ctd-tsg-comp.xls.

The TSG files were opened in EXCEL, median and standard deviations (over 2minutes) were calculated for the two temperature channels, salinity and fluorescence, and the file was then reduced to the times when CTDs were run. Those files were added to the CTD data in file 2008-50-ctd-tsg-comp.xls. The positions were compared and were very close with average differences <0.0001º and no difference was larger than 0.00025º.

This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T0   A study of differences between the intake and lab temperatures shows the lab temperature to be high by an average of 0.15Cº, but the standard deviation was 0.38 Cº: To see if some of the noise is caused by the time it takes for the water to move through the loop differences were calculated by offsetting the two temperatures, looking at comparing the lab temperature to that of the intake temperature 6 to 15 records (3 to 7.5 minutes) earlier. Just for comparison a shift of -3 records (1.5 minutes) was also tested though it makes no sense to compare the lab temperature to a later intake temperature.

| Offset in records | Avg Diff | Stdev |
| --- | --- | --- |
| -3 | 0.15050 | 0.44248 |
| -1 | 0.14957 | 0.40283 |
| 0 | 0.14911 | 0.38117 |
| 1 | 0.14916 | 0.36790 |
| 2 | 0.14921 | 0.35389 |
| 3 | 0.14925 | 0.33891 |
| 6 | 0.14939 | 0.28368 |
| 9 | 0.14950 | 0.20071 |
| 10 | 0.14954 | 0.17487 |
| 11 | 0.14957 | 0.16125 |
| 12 | 0.14960 | 0.16301 |
| 15 | 0.14972 | 0.24153 |

So the minimum difference is with no offset, but there is nothing to say the minimum is correct. The minimum standard deviation is probably more significant and that occurs with the 5.5 minute offset. In either case the temperature measured in the lab is higher than the intake temperature by about 0.15C˚.

Next a plot was made of temperature differences with an offset of 11 versus intake temperature. That showed considerable temperature dependence as has been noted in other cruises, as water closer to the ship temperature warms up less than cooler water.

| Intake Temp | Temp Difference |
| --- | --- |
| 8 | 0.2766 |
| 10 | 0.2122 |
| 12 | 0.1478 |
| 14 | 0.0834 |
| 16 | 0.0190 |

The relationship was T0-T1 = -0.0322 * T1 + .5342. If we wanted to recalibrate lab temperature it would be awkward to do this sort of offset, but it is not needed for this cruise since the intake temperature is available. If it is ever needed then how to do a time-dependent recalibration would have to be considered. It might be more easily done in EXCEL than in IOS SHELL.

During the first 2730 records when the flow rate was low, the average differences was 0.23 Cº with average temperatures only slightly lower than for the rest of the record.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There were 43 casts that could be used.

If all data are included the TSG intake temperature is found to be high by an average of 0.124 Cº but the standard deviation is 0.684. A plot of these differences against cast # shows that the differences in Broughton and Effingham are incredibly noisy. When only the LD and LG lines are included, then the intake temperature is high by 0.276 Cº and the standard deviation is 0.243. The salinity differences look even more unrealistic with the TSG looking low by 1.91 using all data and a standard deviation of 2.3; using only the offshore lines the TSG salinity is low by 0.22 Cº with a standard deviation of ~0.1. The lab temperature is high by 0.28 Cº for the LG-LD lines. That suggests heating in the loop of about 0.19 during offshore stops. That is more than we expect based on studying the full record from the TSG. The ratio of CTD FL to Extracted CHL does not vary so much from one area to another, though there does seem to be a steady increase with time. Using just the offshore the ratio is ~4.1 with a standard deviation of 1.4. The ratio is 4.3 using all the data. (See 2008-50-ctd-tsg-comp.xls.)

- Loop Bottle Comparisons There were no loop salinity samples.

- Calibration History

The TSG primary temperature and conductivity were recalibrated in December 2007. They have been used on 3 other cruises that had calibration sampling.

  - During 2008-01 there was no intake thermistor and the lab temperature was estimated to be high by 0.19Cº. The salinity data were thought to be low by 0.055.

  - During 2008-32 the lab temperature was believed to be high by 0.228 Cº and the salinity data were bad.

  - During 2008-26 the lab temperature was considered high by 0.15 and salinity low by 0.125.

  - During 2008-26 the lab temperature was considered high by 0.17 and the salinity low by 0.17.

  - During 2008-10 the lab temperature was found to be high by about 0.2 and the salinity low by 0.23, but there were not many points of comparison.

Conclusions

We have temperature in the lab and at the intake so there is no need to recalibrate temperature.

For salinity the TSG is lower than the CTD by 0.22 using only the offshore data. But while we have lots of CTD data for this cruise, we have no loop samples. For 2008-10 the TSG was found to be low by an average of 0.228; there were loop samples for that, though less CTD data with which to compare. The fact that the results of 2008-10 and 2008-50 are this close is encouraging. It is reasonable to apply the same recalibration as for 2008-10. The salinity will be recalibrated by adding 0.023, but this should be revisited after a report arrives from the factory after the next recalibration of the instrument.

The TSG fluorescence is about 4.1 times the CTD fluorescence, whereas it was higher by a factor of 2.4 for 2008-10 and 3.3 for 2008-27. There was some geographic variability in the CTD fluorescence vs CHL that may be a factor in the differences between this cruise and the earlier one.

f.) Editing

The time-series plots were examined in a graphical editor. There were no obvious problems in the temperature and fluorescence channels. There were unbelievable jumps in salinity values, but it is not clear which data are good and which bad, so no editing was applied. Similar blocks of suspicious data were seen during 2008-10. All of the jumps seen in 2008-50 were during the long steam to Bowie Seamount and back, so this problem may be related to high speeds. It is conceivable that high speeds might affect the depth from which the loop water is drawn, but that should affect temperature too. There has been some speculation that fresh water could be getting into the loop occasionally. Is there anything that would lead to this happening during long periods of steaming, but not at other times? Or is that just a period of low variability so that we see the shifts that might otherwise be masked?

g.) Recalibration

CALIBRATE was used to apply file 2008-50-recal1.ccf which subtracts 0.23 from the salinity.

h.) Preparing Final Files

HEADEDIT was used to rename Temperature:Primary as Temperature:Lab and Temperature:Secondary was renamed Temperature:Intake and to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header.

REORDER was used to ensure that the intake temperature will be selected preferentially by programs accessing the archive. The final file is named 2008-50-0001.tob.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

The sensor history was updated.

Particulars from the Daily Log Book:

5. Upcast only (downcast not archived)

15. Very close to bottom

16. Fluorometer noisy

17. No sample #143

21. Bottle #1 misfired, no sample. Error in station name.

31. Bottle did not trigger, no sample.

33. Altimeter not working.

34. Altimeter ok.

42. Bottle 2 did not fire

51. Problem in first 9m.

REVISION: August 18, 2010

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly but steadily until July 2009; then a sudden shift occurred, so that maximum values between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study was made of transmissivity that led to a decision to apply post-processing corrections to all cruises between March 2008 and June 2010.

Transmissivity data from this cruise were corrected by multiplying the original values by correction factor 1.106. This was based on assumptions that deep offshore transmissivity from a June 2009 cruise should be about 62%/m and that drift was linear with time between March 2008 and July 2009. The corrections produced reasonable results for all cruises in that period.

For details on how the correction factor was derived see:

OSD_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer 1005DR Corrections.doc

These data should be considered estimates.

Revisions done by: Germaine Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEA-BIRD | 911+ | 443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 16Jan08 | Factory “ |  |  |
| Conductivity | 1766 | 07May08 | “ |  |  |
| Secondary Temp. | 4700 | 16Jan08 | “ |  |  |
| Secondary Cond. | 2173 | 07May08 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 14Feb2007 | Factory |  |  |
| PAR | 4656 | 11Feb2003 | IOS |  |  |
| Fluorometer | 2228 | ? | IOS |  |  |
| Pressure Sensor | 63507 | 25/Oct/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG

Make/Model/Serial#:	SEA-BIRD/21/2487       Cruise ID#:	2008-50

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2487 | 01/12/07 | Factory |  |  |
| Conductivity | 2487 | 01/12/07 | “ |  |  |
| Wetlab/Wetstar Fluorometer | WS3S-713P | 8/01/01 | “ |  |  |
| Temperature 2 | 4652 | 22/Dec/06 |  |  |  |
| Flow Meter | ? | ? |  |  |  |