## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 December 2021 | ### Corrected Salinity:Bottle precision that was lost during HPLC addition. S.H. |
| ### 5 February 2021 | ### Added HPLC Data and corrected Bottle and CTD to pad some off scale fluorescence values. S.H. |
| 27 May 2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

NOTE: Salinity was not recalibrated because of known problems with the Autosal performance. Calibration should be revisited when the next drift report is received from the factory.

UPDATE: Post-cruise calibration reports from Feb. 10, 2009 indicate that the drift in the T and C sensors were very low and that the combined effect on salinity would lead to errors of <0.001 for this cruise. April 2009

## PROCESSING NOTES

Cruise: 2008-51

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG-JdF Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: September 8, 2008 – September 15, 2008

Processed by: Germaine Gatien

Date of Processing: 10 November 2008 –  11 December 2008

Number of original CTD casts: 73	Number of CTD casts processed: 73

Number of bottle files: 22		Number of bottle casts processed: 21 (no sampling for 1 file)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0585) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1117), a Seapoint Fluorometer (#2841) with a 10X cable, two PAR sensors (#4565 and later #4601), a Surface PAR sensor (#16504) and an altimeter (#1233 and late in cruise #1252). The deck unit and the logging computer serial numbers were not recorded in the log.

# SUMMARY OF QUALITY AND CONCERNS

Preparations for this cruise were severely hampered by the absence of a CTD technician familiar with the program and a general shortage of scientific personnel. It appears that the fluorometer was not pumped as is usually done for Pacific cruises where the descent rate can be very noisy. For most casts in the Strait of Georgia, where the descent rate was very steady, this is not likely to be a problem, but for such casts as those at the mouth of Juan de Fuca Strait, there will be spikes associated with CTD descent rate variations. Filtering will reduce the spikes somewhat, but caution is recommended in interpretation of these data. Another potential problem is that the 30X cable was used with a range of 0 to 5ug/l, so that there are some off-scale values in fluorescence.

The Daily Science Log Book was generally in good order but there was no list of equipment used. Notes in the log indicate there were problems with equipment. The rosette log sheets occasionally lack a record of what sampling was done. The dissolved oxygen data lacked quality flags and comments, but there were comments entered on the rosette sheets that were useful in assigning flags.

The Altimeter was changed at sea because of spikes in the data, but they continued after the change. Such spikes are generally seen in these data particularly if there is a weight on the rosette package. The altimeter # was entered incorrectly in the configuration file and the same one was used even after the switch in sensors; since the calibration coefficients are the same this is not a problem, and the altimeter serial number does not appear in the header, so the serial number was left uncorrected in the configuration file.

Transmissivity and fluorescence are nominal and unedited except that some records were removed in editing temperature and salinity. The fluorescence is noisy and has higher than usual values below 100db for casts #2 to 5; deep transmissivity is lower than usual for the same casts.

The new dissolved oxygen algorithm was used with the nominal value used for parameter E. While tests were run by varying E, there proved to be insufficient deep data to judge what value is best. This should not be a problem for data this shallow, but if this sensor is used for deep sampling, further tests should be run to establish the best value of E.

The dissolved oxygen data in the CTD files should be considered:

•	±0.5  ml/l from  0– 50db

•	±0.4 ml/l from 50–150db

•	±0.2 ml/l below 150 db

There is evidence that the salinity is low by 0.012 salinity units, but because there are some doubts about the salinometer performance, salinity was not recalibrated. It should be re-evaluated when the drift check is received from the factory for either of the conductivity sensors and after two other cruises are processed that used the same salinometer at about the same time, but had different sensors.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets.

Salinity data were available in spreadsheet format.

Nutrient data were provided in spreadsheet format.

Dissolved oxygen data were available in individual OXY files without flag channel or comments.

CHL data were available in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

There were a number of problems in preparing the configuration file.

- The calibrations for the two temperature sensors were not on file, but Ron Lindsay obtained copies from SeaBird.

- The secondary conductivity sensor was from the Arctic group and has apparently not been calibrated since 2002. I do not have a record of use for this sensor but the Arctic calibration spreadsheet indicates that it was used as a spare in 2006 and 2007, so it has probably not been used much. It may have been used on the Laurier in 2008.

- The altimeter was changed at cast #64 to try to get rid of spikes that were occurring. The switch did not help. The serial number in the configuration file is wrong for both sensors according to notes in the log book, but since it does not get written into the headers and the coefficients are the same, this was not corrected.

- Errors were found in the calibration constants for the PAR, pressure and dissolved oxygen sensors. The offset for the pressure sensor was set to -0.6db whereas in recent use a value of +0.4db has been found appropriate.

- The transmissometer is said to be serial number 1050DR but no such instrument is on file. There is a 1005DR but there is no calibration on file for the date in the con file, 9 June 2008. The entry was not changed.

- The parameters entered for the DO sensor were those used for the old DO algorithm (Owens-Millard equation) and they were from a pre-repair calibration in Jan. 2008, whereas a calibration from Feb. 2008 after repair is the correct choice. The coefficients for the Owens-Millard equation were corrected and those for the Murphy-Larson equation were entered, as this new algorithm is now available. The value of E may have to be adjusted, but for the full files that can be done just before running DELETE as the concentration will not be calculated until then. The bottle file conversion will have to be rerun if it is decided that this factor needs adjustment. However, given that there are no casts deeper than 500db this is unlikely to be a problem.

- The fluorometer configuration gives the gain as 30X which is not the usual. A few checks will be done to ensure this information is correct.

##### Conversion of Raw Data

Data were converted using configuration files 2008-51-ctd.con.

A few casts were examined and all expected channels are present. As usual the upcast temperature and conductivity traces are noisier and further apart than in the downcasts. There is fine-scale noise in the secondary conductivity ~±0.00015 units.

For the deepest casts, the fluorescence is about 0.08ug/l near the bottom.

For casts #2-5 the fluorescence data are unusual, with higher dark values and spikier traces than those seen elsewhere in the cruise. These casts are all in Haro Strait and the deep transmissivity is relatively low, so there is presumably some effect from deep mixing.

The dissolved oxygen voltage shows the usual offset between upcasts and downcasts.

There were pressure spikes in cast #10 but WILDEDIT fixed that.

The descent rate varies from cast to cast, but overall was steadier than usual. A few shallow casts had very low descent rate, but were mostly quite steady.

The altimetry does have a lot of spikes especially near the bottom. The header readings will have to be checked carefully.

Rosette files were converted using a start time of -5s and duration of 10s.  Four runs were done using different values for E: 3.5, 3.6, 3.7 and 3.85.

CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and the only significant outliers were found in cast 52, bottle #1. These were cleaned using CTDEDIT. The output files were then copied to BOT.

Tests were done preparing merged bottle files so that COMPARE could be run using the different values for E to see if there is a significant difference in the fits. As is described in section 11, it was difficult to say which worked best, the differences were small and given no deep data there was no way to compare sections with similar DO values but different pressures. The fits vary with the choice of E, but the scatter is similar, so the recalibration should be appropriate for all of them. Given the absence of strong evidence for any other choice, the nominal value of E=3.6 will be used.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on a few casts to determine the best choice of parameters for CELLTM. Settings tried were (α = 0.01, β=7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9), (0.04, 7), (0.04, 9) and (0.0245, 9.5). The differences among them were slight and varied from feature to feature. The choice of (0.03, 9) looked best overall for the primary conductivity and (0.03, 9) and (0.0245, 9.5) were very close for the secondary conductivity. CELLTM was run using (α = 0.03, β = 9) for both the primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences in conductivity and salinity were extremely noisy, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 44 | 250 | +0.0007 | -0.0003 | -0.004 XN | Moderate, V. Steady |
| 50 | 200 275 400 | ? XN +0.0007 +0.0007 | -0.0004 VN -0.00035 VN -0.00025 | ~-0.0044 XN ~-0.004 XN -0.003 XN | Moderate, V. Steady “ “ |
| 72 | 200 275 | XN ~0 XN | -0.00025 VN -0.003 VN | -0.0034 XN -0.0034 XN | Moderate, V. Steady “ |

The secondary conductivity has a lot of fine-scale noise.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and a cross-reference listing were produced. Times and positions look fine.

The cruise track was plotted and no problems found.

The average surface pressure is 2.9db, which is a little deep for the Vector. For cast #35 when the CTD sat at 1.1db for a few seconds, the salinity settled to about 25 which looks reasonable.

The altimeter readings from the headers of CLN files were exported to a spreadsheet; because there were many notes in the log about spiking near the bottom all entries were checked against plots of altimetry against pressure. All values looked reasonable given the algorithm averages over the bottom 2db. There was no header entry for cast #1 because the CTD did not go near the bottom. The bottom depth was entered incorrectly in the log and in the file as 100db; based on records from 2008-19, the bottom depth was changed to 177db in the CLN file.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There was a rosette file for cast #66, but there were no sample numbers assigned to that cast, so this is presumed to be an accidental firing.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (output:SAMAVG).

The altimeter header entries were checked in the SAMAVG files. The bottom depth for cast #1 was changed to 177db as was done for the CLN files. For cast #60 spiking in the altimetry led to a poor estimate. Based on examination of plots and the results for the full downcast file the header entry was changed from 12.3db to 9.6db.

SALINITY

The salinity data were delivered in spreadsheet format together with the results of several other cruises. The relevant data were extracted and saved as 2008-51-sal.xls. There was one comment but no flag. Headers were changed to standard formats and the spreadsheet was simplified and saved as 2008-51-sal.csv. This was converted into individual SAL files.

It later became apparent that sample #24 was mislabelled as being from station 59, it should have said station 102 which is cast #7. That was corrected.

There were no duplicate salinity samples.

DISSOLVED OXGYEN

Dissolved oxygen files (*.oxy) were delivered without flag channel. ADD CHANNEL was used to add a quality flag channel, with the resulting files named *.ADD. Information on the rosette sheets were used to assign flags to samples 19 (cast #3), 313 (cast 31) and many samples from cast #55. The latter was a case where the program did not work for some flask #s so the analyst entered alternate #s to enable titration. It is unknown whether the volumes were adjusted or not. For now “c” flags were assigned to the doubtful samples. (COMPARE later showed that these flagged values were minor outliers, so the flags were left in place.) There were many notes on the rosette sheets about problems, but in most cases these were cases of a repeat titration due to problems with the first run. In each case the sample number had to be adjusted for the second run, so a 0 was added. In each case the entry for the first sample was removed from the ADD file and the zero then was removed from the second sample #.

There were no duplicate samples.

NUTRIENTS

The nutrient spreadsheet was simplified and saved as 2008-51-nuts.csv. Extraneous columns were removed, header names were changed to standard format and lines were removed for which there was no nutrient sampling. Data were sorted on sample number. There were a few comments, but no flags. There were a few values with flags but no comments. File 2008-51nuts.csv was converted to individual NUT files.

CHL

The CHL spreadsheet was edited to change headers to standard names and remove extraneous lines; it was then saved as 2008-51-chl.csv. An event # column was added and filled in based on rosette log records. Data were sorted on sample number. File 2008-51-chl.csv was then converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG) and put through CLEAN to remove SeaBird headers.

##### Compare

Salinity

COMPARE was run using pressure as the reference channel. There were many outliers – the same ones for both sensors. Two extreme outliers were easy to explain since the analyst had noted a cap missing for one and bubbles present for the other. One other had 4 readings with each larger than the previous. The other outliers have no obvious cause, but the analyst did comment on problems with the Autosal at the time of the analysis. These samples were analyzed at the end of a session that included the analyses for 2008-27, 2008-10 and 2008-50. The comparison with CTD for 2008-27 suggests the Autosal results were bad from the latter part of the session. For 2008-10 the bottle comparison showed the CTD primary salinity to be low by 0.006, yet a post-cruise calibration showed very little drift. 2008-50 comparisons have not yet been done. So there is cause to doubt the bottle data.

Flat fits were achieved for both sensor pairs when 8 outliers were removed, but there were only 13 values left. The primary salinity was low by 0.012 and the secondary by 0.015. Those differences are much larger than we normally see. Other recent Autosal results suggest the linearity problem that may have affected analyses during the year previous to the summer of 2008 has been resolved, but just to check that, a fit was done against salinity. When the outliers identified for the other fits are excluded, the trendline is quite flat, and the outliers are all in the opposite direction to what would be expected from linearity problems with the exception of the one that is well explained by the missing cap.

The following samples were flagged and with explanatory comments entered in the headers:

Cast 1   – sample 1, flagged “c” (4 readings)

Cast 39 – sample 123 – flagged “c”

Cast 41 – sample 135 – flagged “d” (bubbles)

Cast 45 – sample 151 – flagged “c”

Cast 48 – sample 168 – flagged “c”

Cast 63 - sample 240, flagged “d” (no cap)

CTD Fluorescence

COMPARE was run using extracted chlorophyll and CTD fluorescence. The fits show fluorescence to be about 80% of extracted chlorophyll values, for values of CHLa <2. The 30X cable was used and it is clear that the fit flattens at about CTD_CHL = 4ug/l. When the dark value, 0.08ug/l, is subtracted from the CTD data, the offset for the fit of bottle vs CTD is ~0.08.

Dissolved Oxygen

Test runs of COMPARE were made using the 4 different E settings in the conversion files, to determine if there is a significant difference between them. It is not clear what criteria to use. For some cruises with deep data there is a DO minimum, so we can compare the correction at a particular DO value from different depths to ensure they are as close as possible. For these data there are only minor reversals, not significant enough to be useful. Setting the value of E higher produced fits with higher slopes and smaller offsets, with the R2 value slightly higher for the fit, but that does not look significant. The differences are not sufficiently large to be noticeable when plotted. It seems best to stay with the nominal value as the correction should work well for any of these values. If this sensor is ever used for deep data, this will have to be revisited.

When outliers were identified based on standard deviations in the CTD data and residuals in the differences, plus a few near-surface values with positive differences, the fit was found to be:

CTD-BOT = 1.0624 DOX-CTD + 0.0777

When the data from cast #55 are excluded (due to uncertainties about the flask volumes) the fit is:

CTD-BOT = 1.0627 DOX-CTD + 0.077

These are much higher slopes than found for the last use of this sensor, but the fit for that cruise was out of step with previous uses so it was sent to the factory where it was found that the membrane was torn. It has been repaired and recalibrated since then and the new algorithm is being used for the first time. The flag was removed from sample #19 since it looks fine in COMPARE; the other flags were left in place. Most of the flagged samples from cast #55 look ok in COMPARE, but a few are slight outliers – the flags will be left in place.

Plots of Bottle DO and CTD DO versus Salinity turned up a few outliers that had also been removed from COMPARE as outliers. The following were flagged “c”:

Cast #3, sample #12 – outlier in profile, in DO vs Sal and outlier in COMPARE.

Cast #7, sample #26 - outlier in profile, in DO vs Sal and outlier in COMPARE.

Cast #23, sample #93 - outlier in profile, in DO vs Sal and outlier in COMPARE.

Cast #41, sample #140 – outlier in profile, in DO vs Sal and outlier in COMPARE.

Cast #65, sample #270 - outlier in DO vs Sal and outlier in COMPARE.

See file 2008-51-dox-comp1.xls for the results with E=3.6. (For the other choices see 2008-51-dox-comp1-35.xls, 2008-51-dox-comp1-37.xls and 2008-51-dox-comp1-385.xls.)

The merging of the bottles (as described in the final paragraph of section 10) was rerun to ensure the newly applied flags are in present in all steps.

##### 12. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. A shift of +24 records (1s) has been found appropriate for all cruises in the past year, but for this data it appears to be much too high. Throughout the cruise the offset appears to be no more than 0.4s (~10 records). At times no shift seems necessary. The fluorescence is generally quite low and noisy, so this is not an easy estimate to make.

This is a different fluorometer than has been used on other cruises this year, so perhaps the response is different, or perhaps there was some difference in the plumbing of the pump. Since the CTD came from the Arctic group and they do not normally pump fluorescence, a few casts were investigated to see if there are signs of that being the case for this cruise. To judge this we need a cast with a noisy descent rate and reasonably high fluorescence near the surface. Cast #18 fits the bill and contains clear spikes in fluorescence at descent rate minima, with those spikes being larger than are seen in the temperature data. So they are not due to shed wake effects.

SHIFT was NOT applied to the fluorescence.

Dissolved Oxygen

Tests were run on 3 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. SHIFTS of from +30 to +120 records were tested and +60 looked best overall. This is quite different from previous uses of this equipment when +90 to +130 records were used. There are a number of possible explanations: repairs to the sensor were done in February 2008, a new algorithm was used for converting the oxygen and the oxygen concentration was derived after the initial conversion so that the window for that derivation can include time before the current record. Variations in pump and plumbing might also have occurred, since this cruise had an Arctic-group CTD and some scientific personnel accustomed to Arctic methods.

SHIFT was run using +60 records for all casts.

Conductivity

Tests were run on 3 casts using shifts between -0.5s and +0.9s and T-S plots were prepared to compare the results. Overall a setting of +0.5s looked best for both. All casts were put through two runs of SHIFT using those settings. (Output *.SHFC0 and SHFC1).

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concerned the upcasts of #6, 7 and 10 so will not affect the downcast files and none appear to be in areas of concern for upcast bottles; there were frequent pad values in several channels near the pressure level of the warnings.

##### DETAILED EDITING

The primary salinity was selected for archiving. It is a little closer to the bottles than the secondary and the secondary has a lot of fine-scale noise in it. While this should average out, it makes editing more difficult.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. There were some spikes in salinity that required cleaning. The descent rate was noisy in Juan de Fuca Strait, so editing was required to remove the effects of shed wake corruption. In the Strait of Georgia the descent rate was very steady, so shed wakes were only a problem near the bottom; however, there was considerable noise in the T-S data in the top 5 or 6db, with features that do not look like interleaving. This is probably due to the method of deployment - the CTD had been lowered to 10db to soak, and was then brought to the surface and immediately was lowered for the full cast. In such quiet sea conditions the ship may not have drifted much (except where currents were strong), so the CTD would have sampled waters that had recently been stirred by the rosette package. For cast #67 in an area where tidal currents are usually very high, there are unstable sections, but those do look like the interleaving one would expect in this area. In contrast cast #64 which had a very steady descent rate has unstable features in the top 10db that look random – that cast is from open waters.

For the most part such data were left unedited.

All casts required some editing except for #67.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: The primary conductivity sensor was used for 2007-67 when there were many severe outliers and the sensor was found to be low by 0.0067. Those bottle samples were analyzed during the period for which there is suspicion that the Autosal had errors due to non-linearity, and that would have led to apparent negative differences for low salinity. The primary was deployed on an SBE25 during 2008-12 but that data set has not been processed yet. The secondary sensor is an old one with no record of use since it was recalibrated in 2002. It may have been used in the Arctic in 2008, but even if it was not, there can be sensor drift sitting on a shelf for 6 years.

2. Dissolved Oxygen – This sensor was recalibrated in February 2008 and there is no record of its being used on any other cruises since then.

3. Pressure – This sensor has not shown significant drift, and the offset of +0.4db has been in use since mid-2004.

Historic ranges –There were many excursions from the historic ranges, most in the top 30m with temperature and salinity sometimes above the maxima and sometimes below the minima. Since the excursions are in both directions, they do not suggest a problem with the calibration of T and S, but are more likely reflective of limitations in the climatology (which has not been updated since 1997) and show how variable a region this is. The temperature excursions in the Strait of Georgia were all on the high side and were all above 30db.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

The comparison with bottles suggests that the sensors are reading low by 0.012 and 0.015. These are much larger differences than we are used to seeing, and it seems a little odd that they are both so large. There is no evidence of a non-linearity problem, but problems seen with the Autosal during these analyses reduce confidence in the COMPARE results. Data from 2008-10 analyzed at about that time suggested a sensor was low by 0.006, but the post-cruise factory report showed little drift. The history of the sensors is not very helpful, but does suggest that large differences could be possible. Being used on an SBE25 on a small ship might possibly have led to damage. The secondary has no history at all but in 6 years significant drift is quite possible. There is unlikely to be useful calibration information from 2008-12, so that is not worth waiting for. When the sensors are next sent for analysis, this issue should be revisited.

Given the unusually large differences and the analyst’s concerns about the Autosal, it would be unwise to recalibrate based only on the bottle comparison. If either sensor is sent to the factory for recalibration, the drift report may give some indication of whether these differences are due to calibration drift: Also when 2008-10and 2008-50 are processed, it will be easier to judge what errors may have occurred in the salinity analysis since they involved different conductivity sensors. So for now, no recalibration will be applied to the salinity.

File 2008-51-ctd.ccf was prepared to apply the following equation to the CTD Dissolved Oxygen channel in the SAM and MRGCLN2 files:

1.0627 DOX-CTD + 0.077

COMPARE was then rerun to check that the results were as expected and they were. (See 2008-51-dox-comp2.xls and 2008-51-sal-comp2.xls.)

The same calibration control file was then applied to the edited CTD files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. There is also an error expected due to slow response time which may be partly corrected by the conversion algorithm and/or SHIFT, but to test whether further correction is needed a comparison was done between the downcast data and bottle data from the same pressure.

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. There is a lot of scatter in the comparison especially near the surface, but when outliers are removed based on residuals, a fit is found that is flatter than that usually seen. An examination of the data suggests this is due to the fact that there is little variation in DO below 50m for many casts, so the results are not particularly sensitive to response time errors. A fit against file pair number confirms that there is almost no error for the casts in the Strait of Georgia where most of the water column was quite well-mixed in DO. For the Juan de Fuca Strait casts, there is more variability and larger differences. The fit against Do also shows almost no correction for DO values from 3 to 4ml/l, which is the range found below the surface in northern and middle part of the Strait of Georgia. Applying the fit against DO would appear to be the best choice to ensure that data in the low DO gradient zones are not affected, but errors for high and low DO are corrected.

The fit against DO when outliers are excluded based on residuals leads to a correction as follows:

DO (Corrected) = 0.9973 *DO (after 1st recalibration) + 0.0132

The corrections will be <0.005ml/l for DO values from 3.4 to 6.2ml/l. (See 2008-51-dox-comp3.xls.)

The thinned files were recalibrated by applying the above correction and the comparison was rerun. That showed that the recalibration was applied properly. (See 2008-51-comp4.xls.)

Recalibration using file 2008-51-recal2.ccf was then applied to the downcast files only since the bottle data are collected after the CTD DO sensor has had time to equilibrate. (Output: COR2)

The COR2 files were then clipped to 150db (output:CLIP) and bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

T-S plots were examined on-screen and no problems were found.

For all casts REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added; REORDER was run to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorometer appears to have been UNPUMPED for this cruise

which may result in spikes associated with descent rate variations.

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

•	±0.5  ml/l from  0– 50db

•	±0.4 ml/l from 50–150db

•	±0.2 ml/l below 150 db

The Standards Check routine was run and no problems were found. The final files were named CTD.

A header check turned up no problems.

Profile plots were made. The transmissivity at the bottom was near-zero for cast #41 and quite low for #42 and #45. The altimetry is spiky but does not suggest that the CTD hit bottom, and there is no note of such a problem in the log. There is a lot of variability from one cast to the next.

The track plot looks ok.  The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The values ranged from 60% at station 73 to 160% in Saanich Inlet. The distribution seems quite random, with some high values at one cast and a low one nearby. The casts near the mouth of Juan de Fuca had moderate values, but mid-way along the strait variability was high. The CTD DO values look the same in upcast and downcast traces for the extreme values and where available bottle files confirm the extremes are real. Because a 10X cable was chosen fluorescence goes off-scale for the plots with high DO saturation, but this does indicate biological activity was high, explaining the high values. Similarly, it is low for the casts with low saturation. Some of the variability will also be due to the minimum pressure sampled by the CTD which varies from cast to cast. There is no evidence that the dissolved oxygen sensor malfunctioned.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts the following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. (Output: CHE)

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars from log:

Many problems noted during test cast – ship returned to IOS to fix CTD problems.

7. Lower rosette ring contacted ship during recovery.

34. Altimeter spike

36. Altimeter spike

41. Altimeter spike at bottom

64. Altimeter was replaced but still spiked

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 29May08 | Factory “ |  |  |
| Conductivity | 1764 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2710 | 29May08 | “ |  |  |
| Secondary Cond. | 2676 | 31Dec02 | “ |  |  |
| Transmissometer | 1005DR | 9Jun08 | IOS | 5MAR08 | IOS |
| SBE 43 DO sensor | 1117 | 8Feb2008 | Factory |  |  |
| Fluorometer | 2841 | 1Oct2006 | IOS |  |  |
| PAR | 4601 | 19dec2003 |  |  |  |
| Surface PAR | 16504 | 2Jan2004 |  |  |  |
| Pressure Sensor | 77511 | 13/03/2000 | Factory |  |  |
| Altimeter | 1233/1252 | 14Feb2007 | ? |  |  |