## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 18-Dec-2013 | Added units to nutrient data. J.L. |
| 18 August 2010 | Transmissivity data were corrected. See details at end of report. |
| 2 July 2009 | Nitrate data replaced with updated values. Channel order changed. See end of report for details. |

## PROCESSING NOTES

Cruise: 2008-61

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia/Juan de Fuca Water Properties Survey

Party Chief: Chandler P.

Platform: Vector

Date: November 19, 2008 – November 26, 2008

Processed by: Germaine Gatien

Date of Processing: 10 November 2008 – 14 February 2008

Number of original CTD casts: 	74	Number of CTD casts processed: 73 (1 was shallow test cast)

Number of bottle files: 		21	Number of bottle casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1176), a Seapoint Fluorometer (#2228) with a 10X cable, a PAR sensors (#4656), a Surface PAR sensor (#16504) and an altimeter (#1252). The deck unit was a Seabird model 11 (#0424) and the logging computer was PAC02588.

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book and rosette sheets were in good order.

There were many date errors in the SeaBird headers. Some were noted in the log, but not all.

There were many notes about spiky altimetry, but the values near the bottom were fine.

Transmissivity and fluorescence are nominal and unedited except that some records were removed in editing temperature and salinity.

The fluorescence data in the rosette files from Northern Strait of Georgia look low; this could be due to poor pumping during the bottle stop, though it was not poor enough to have a noticeable effect on salinity. The CTD was mounted incorrectly in the rosette for this cruise which may explain this, and may also account for the unusual alignment of the dissolved oxygen data and the drift in the DO calibration.

A few recent changes in DO processing were applied to these data. The new dissolved oxygen algorithm was used, with parameter E being the one derived in the processing of 2008-27. Based on recommendations from SeaBird, a 10s window was used for rosette files centred on the bottle firing time, and the concentration of dissolved oxygen was derived later in the processing than in the past so that the time-dependent calculations have access to information before and after the record concerned.

The dissolved oxygen data in the CTD files should be considered:

- ±0.5 ml/l from  0– 200db

- ±0.2 ml/l below 200 db

The comparison between the salinity bottles and CTD was noisy, with no obvious explanation except that these casts are fairly shallow so that local salinity gradients are relatively high. The salinometer is reported to have worked well and the CTD data looks fine. Possibly this is another result of improper mounting of the CTD in the rosette. On average the salinity was found to be low by 0.0025, with the deepest bottle showing it low by 0.0015 and two bottles showing it to be high by 0.003 and 0.006. No recalibration was applied. This can be revisited when the sensors are next calibrated.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book was obtained together with rosette log sheets. There was note of a number of errors in dates in the headers.

Salinity, nutrient and CHL data were available in spreadsheet format.

Dissolved oxygen data were available in individual OXY files without flag channel or comments.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The configuration files were checked. Surface PAR parameters had been omitted from the first cast, but were added for cast #2. A copy of 2008-61-0002.con was saved as 2008-61-ctd.con and then edited to fix a few problems. The offset was missing for the PAR sensor for all casts. The dissolved oxygen parameters were correct for the Owens-Millard method, but they had not been entered for the Murphy-Larson method which will be used, so those were added. The value of E was set to 3.85 based on the results of a study during the processing of 2008-27.

##### Conversion of Raw Data

Data were converted using configuration files 2008-61-ctd.con. The dates were fixed in 5 files.

A few casts were examined and all expected channels are present. As usual the upcast temperature and conductivity traces are noisier and further apart than in the downcasts. There were some pressure spikes.

For the deepest casts, the fluorescence is about 0.13ug/l near the bottom.

There may be a problem with fluorescence late in the cruise. During cast #62 it suddenly goes down to very low values, and other than a few spikes, no high values are seen after that.

The dissolved oxygen voltage shows the usual offset between upcasts and downcasts.

There were pressure spikes in some casts.

The descent rate is noisy for some casts, but in the Strait of Georgia it was quite steady.

The altimetry looks useful near the bottom.

Rosette files were converted using a start time of -5s and duration of 10s.  Salinity and DO concentration were included.

The dates were fixed in 2 files.

The data were then converted to IOS HEADER format using the IOS SHELL routine.

CLEAN was run to add event numbers and the output files were named *.BOT.

All BOT files were plotted and a few significant outliers were found in casts #57 and 66. These were cleaned using CTDEDIT and notes were entered in the headers about the editing done. The output files were then copied to BOT.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The spikes in pressure have been removed, but there are still some spikes in other parameters for cast #31 but those are in the upcast and don’t appear to affect bottle data. It is possible that further runs will be needed, but some channels with remaining spikes (such as PAR) don’t respond well to this routine.

##### CELLTM

Tests were run on a few casts to determine the best choice of parameters for CELLTM. Settings tried were (α = 0.01, β=7), (0.01, 9), (0.02, 7), (0.02, 9), (0.03, 7), (0.03, 9), (0.04, 7), (0.04, 9) and (0.0245, 9.5). The differences among them were slight and varied from feature to feature. The choice of (0.03, 7) looked best overall for the primary conductivity and (0.02, 7) for the secondary. CELLTM was run using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences in conductivity and salinity were extremely noisy, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 44 | 300 | -.0002 | +0.0009 N | +0.009 XN | Moderate, V. Steady |
| 51 | 300 400 | -.0002 -.0003 | +0.0008 +0.00085 | +0.0085 XN +0.009 XN | Moderate, V. Steady “ |
| 62 | 300 | -.0008 VN | +0.0007 VN | +0.0085 XN | Moderate, Steady |
| 75 | 296 | -.0003 VN | +0.00085 N | +0.0085 XN |  |

Profile plots were made of the dissolved oxygen data and it was discovered that there was an error in how the new value for E was entered in the configuration file. It was also discovered that the dates were wrong in many files. The configuration file and header dates were fixed and DERIVE was rerun. The DO profiles then looked normal.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check and a cross-reference listing were produced. One more error was found a header date. The start times were fixed to match the log book and the checks were rerun until all problems were resolved.

The cruise track was plotted and no problems were found.

The average surface pressure is 1.4db, which is a little shallow for the Vector, though not unreasonable especially given the new method of soaking at depth and just coming to the surface briefly before the full cast is run. One cast was found that had negative pressures recorded during the upcast. The transmissivity records from that time suggest that the pressure is within ±0.2db. The error is small, but it looks like it will soon be time to increase the pressure offset.

The altimeter readings from the headers of CLN files were exported to a spreadsheet; because there were many notes in the log about the altimeter many entries were checked against plots of altimetry versus pressure. The only problem noted was with cast #1 for which the CTD did not get close to the bottom, but spikes were misinterpreted as real data, so the header entry was removed from both the CTD and BOT files.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The ADDSAMP file was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files (output: SAM). The BOT files were then bin-averaged (output:SAMAVG).

SALINITY

The salinity data were delivered in spreadsheet format. There were no comments or flags, but after discussion with the analyst a comment was added for one sample. Headers were changed to standard formats, a flag channel as added and the spreadsheet was simplified and saved as 2008-61-sal.csv. This was converted into individual SAL files. There were no duplicate salinity samples.

DISSOLVED OXGYEN

Dissolved oxygen files (*.oxy) were delivered without flag channel. ADD CHANNEL was used to add a quality flag channel, with the resulting files named *.ADD. There were many comments on the rosette sheets that appear to all refer to oxygen titration problems, so those were used to guide the addition of flags and comments. For now “c” flags were assigned only if there were comments and the values seem out of line. In a few cases there were values in the files, but the analyst had crossed them off or erased them from the rosette sheets. In those cases the values were noted in header comments, but were replaced with -99 in the file and “e” flags were assigned. There were many notes on the rosette sheets about problems, but in most cases these were addressed by the analyst and do not require any comments in the files.

The file for cast #53 was lost, but the data was found in the file for cast #49. It was removed from #49 and used to create a file for #53.

Profile plots of DO versus pressure were made to check for problems, and it was found that Cast #1 had two samples listed for 100db and none for 75db; the rosette sheet shows samples at 100db and 75db, so the 2nd sample pressure was changed to 75db in the file. No other problems were noted.

There were no duplicate samples.

NUTRIENTS

The nutrient spreadsheet was simplified and saved as 2008-61-nuts.csv. Extraneous columns were removed, header names were changed to standard format and lines were removed for which there was no nutrient sampling. Data were sorted on sample number. There were a few comments and flags. File 2008-61-nuts.csv was converted to individual NUT files. Duplicates were studied by the analyst and can be found in QF2008-61nuts.xls.

CHL

The CHL spreadsheet was edited to change headers to standard names, add event numbers and remove extraneous lines; it was then saved as 2008-61chlarc.csv. Data were already sorted on sample number. That file was then converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3, and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only (Output MRGCLN1.) That file was then merged with SAMAVG files (Output:MRG) and put through CLEAN to remove SeaBird headers.

##### Compare

Salinity

COMPARE was run using pressure as the reference channel. There is a lot of scatter among the 21 values, and there is only 1 bottle from any cast. One bottle is a clear outlier and was flagged “c” (sample #98, cast #26). When data are excluded from above 150m and for bottles for which the standard deviation of the CTD salinity is >0.0006 the primary salinity is low by an average of 0.0112 while the secondary is low by 0.0025. The deepest bottle (from 350m) has a low standard deviation in the CTD data and indicates that the two channels are low by 0.010 and 0.0015. Two bottles in Juan de Fuca show the primary closer to the bottles and the secondary higher than the bottles. There is a slight hint of time dependence, but given the very different salinity gradients and depth of sampling in different parts of the survey, this is not necessarily due to calibration drift. A plot of differences against salinity also shows a small slope, but this may reflect that the lower salinity samples come from higher gradient regions, so are not as reliable.

CTD Fluorescence

COMPARE was run using extracted chlorophyll and CTD fluorescence. The fits show fluorescence to be about 70% of extracted chlorophyll values, though the ratio is much lower for the two values of CHLa >2. The ratio is also low for most of the casts at the northern end of the Strait of Georgia where the CTD fluorescence values are extremely low. In section 13 there is further discussion of the low CTD fluorescence values in the bottle files; there is such a sudden drop-off in values that the difference in height between the CTD and Niskin bottle may lead to a very significant difference in chlorophyll measured. The ratio ranges from 0.7 to 1.5 in the south and central parts of the Strait of Georgia.

Dissolved Oxygen

The fit of differences against DO_CTD was quite tight. When a few data were excluded based on high standard deviation in the CTD DO and then a few more based on residuals in the differences, the fit was found to be:

CTD-BOT = 1.1390 DOX-CTD - 0.0111

The slope is much higher than seen previously for this sensor, which could mean the calibration is drifting significantly or that the membrane has been damaged or that the way the CTD was mounted has affected the readings. However, this cruise sampled a small range of DO values and only shallow casts that were frequently well-mixed. The shift found appropriate for the CTD DO was also significantly lower than for previous uses. During 2008-51 when a different DO sensor was used, it was also found that the slope was higher and the shift lower. That was thought to be due to something in the set-up such as the plumbing, or maybe due to using the new algorithm for DO. The latter does not explain the higher slope in this case, since it is higher than the slope for 2008-27 when the new algorithm was also used for the same sensor. Given the nature of this particular cruise it may be unwise to over-interpret the results, but an eye will be kept on how the sensor performs on other cruises.

There were no severe outliers, so no changes were made to the DO flags. (See file 2008-61-dox-comp1.xls.)

##### 12. SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. A shift of +24 records (1s) has been found appropriate for all cruises in the past year, and looks right though there were few casts with high enough chlorophyll to make a judgment.

SHIFT was applied to SBE fluorescence with a setting of +24 records. Comparing profiles from before and after the shift show that it was effective.

Conductivity

Tests were run on 3 casts using shifts between -1s and +1s and T-S plots were prepared to compare the results. Overall a setting of +0.2s proved best for the primary and -0.5s for the secondary. All casts were put through two runs of SHIFT using those settings. (Output *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on 3 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. SHIFTS of from +60 to +80 records were tested and +70 looked best overall. This is quite different from most previous uses of this equipment when +90 to +130 records were used, but is close to the results for 2008-51 when +60 proved best. There are a number of possible explanations: repairs to the sensor were done in February 2008, a new algorithm was used for converting the oxygen and the oxygen concentration was derived after the initial conversion so that the window for that derivation can include time before the current record.  Another explanation is a change in the plumbing: it was discovered in January 2009 that the CTD was mounted incorrectly in the rosette (backwards and upside down), and it is believed it was used in that configuration for this cruise and possibly for some other cruises. The 2008-51 cruise used a different CTD but may have been plumbed the same way.

SHIFT was run using +70 records for all casts.

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warnings concerned the upcast of #31; this will not affect the downcast data and the spikes that caused it did not occur during upcast bottle stops, so there is no need to take any action.

At this stage plots were made to determine if there was a problem with the fluorometer later in the cruise. The data have very low values and some casts with virtually no signal. However, the profiles vary from one to the next so that it does not look like a total failure to record a signal and there is often a shallow layer with values of >0.4ug/L below which the values drop abruptly to near zero. At sites without such a layer and very low values overall, the DO and temperature show the water column was well-mixed, so it is possible that the little chlorophyll available was distributed over the water column with very low concentrations. All the casts in question are from the northern section of the Strait of Georgia. For cast #70 the downcast data shows a surface layer with extracted CHL values ~0.45ug/L down to 4.8db. By 4.9db the values are ~0.07ug/L. No high values are seen by the CTD fluorometer during the upcast but the extracted chlorophyll values for that cast were 0.6ug/L and 0.3ug/l at 1db and 5db. For cast #68 a similar pattern is seen with downcast fluorescence about 0.45ug/L down to 6.0db and 0.07ug/L at 6.2db. The upcast fluorescence values are very low, but the extracted chlorophyll values from the upcast bottles at 1db and 5db are 0.44ug/L and .59ug/L. So downcast fluorescence looks ok. The CTD being a metre below the bottle could make a significant difference in a high gradient, though that does not seem to explain all of the low values. Perhaps the pump was not operating at full efficiency.

##### DETAILED EDITING

The secondary salinity was selected for archiving since it is closer to the bottles; there was no significant difference in spikiness between the two channels.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. The descent rate was noisy near the mouth of Juan de Fuca Strait and for some casts in the northern part of the Strait of Georgia, so editing was required to remove the effects of shed wake corruption. For many casts in the Strait of Georgia the descent rate was very steady, so shed wakes were only a problem near the bottom. However, there was considerable noise in the T-S data in the top 5 or 6db, with features that do not look like interleaving. This is probably due to the method of deployment - the CTD had been lowered to 10db to soak, and was then brought to the surface where acquisition began and it was immediately lowered for the full cast. In quiet sea conditions the ship may not have drifted much (except where currents were strong), so the CTD may have sampled waters that had recently been stirred by the rosette package. For the most part such data were left unedited.

All casts required some editing except for #30, 49,

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity: This is the first known use of both sensors since they were last recalibrated. Both have performed reasonably well in the past.

2. Dissolved Oxygen – This sensor was recalibrated in February 2007 and has been used for many cruises since then. It has been found to perform well, with good detail in the profiles. When used in 2008-27 the new DO algorithm was used for conversion.

3. Pressure – This sensor has shown considerable drift in the past, but an offset of +6.5db has been used since March 2008. For 2008-10 in September the offset was increased to +6.7db.

Historic ranges –All the temperature data fell within the historic ranges, but there were a few excursions in the salinity data, with low values at cast 6 between 20 and 40db and values that were slightly above the historic maxima at many of the casts in northern Strait of Georgia around 70 to 150db. Since the excursions are in both directions, they do not suggest a problem with the calibration of T and S, but are more likely reflective of limitations in the climatology (which has not been updated since 1997) and show how variable a region this is.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

The comparison with bottles suggests that the secondary salinity sensor is reading low by 0.0025, but the noise in the comparison does not make this conclusion a confident one. The deepest bottle showed it too be low by 0.0015 and several samples indicate it is actually higher than bottles. The data will not be recalibrated at this time. When the sensors are next recalibrated or if there is a more trustworthy comparison on a future cruise, this decision can be revisited.

The pressure looks like it might be a little low and the results of 2008-10 also suggest that it is time to add about 0.2db to the offset.

File 2008-61-ctd.ccf was prepared to apply the following equation to the CTD Dissolved Oxygen channel in the SAM and MRGCLN2 files:

CTD_corrected = 1.1390 DOX_CTD - 0.0111

and to add 0.2db to the pressure channel.

COMPARE was then rerun to check that the results were as expected and they were. (See 2008-61-dox-comp2.xls and 2008-61-sal-comp2.xls.)

The same calibration control file was then applied to the edited CTD files.

##### Final Calibration of DO & Pressure

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. There is also an error expected due to slow response time which may be partly corrected by the conversion algorithm and/or SHIFT, but to test whether further correction is needed a comparison was done between the downcast data and bottle data from the same pressure.

Files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. The tightest fit was of differences versus pressure value and leads to a correction:

DO (Corrected) = DO (after 1st recalibration) -0.0643 + 0.0001 * Pressure

with an average of offset of 0.05ml/L for the points included in the fit. (See 2008-61-dox-comp3.xls.)

The thinned files were recalibrated using the above correction, and the comparison was rerun. That showed that the recalibration was applied properly. (See 2008-61-comp4.xls.)

Recalibration using file 2008-61-recal2.ccf was then applied to the downcast files only since the bottle data are collected after the CTD DO sensor has had time to equilibrate. Pressure was also recalibrated by adding +.2db since this was not done earlier to these files. (Output: COR2)

The COR2 files were then clipped to 150db (output:CLIP) and bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved for Angelica Peña. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files put through REMOVE and named *.BOF and saved for the use of Angelica Peña. A readme.doc file was prepared with some notes on the preparation of those files. (Note that the bottle files used include Reference PAR.)

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

T-S plots were examined on-screen and no problems were found.

For all casts REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added; REORDER was run to put the two SBE DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the project name and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

•	±0.5  ml/l from  0– 200db

•	±0.2 ml/l below 200 db

On average the CTD salinity looked lower than bottles by about 0.0025

units, but the comparison was too noisy to justify recalibration.

The Standards Check routine was run and no problems were found. The final files were named CTD.

A header check turned up no problems.

Profile plots were made. No problems were noted, though some strikingly high temperature gradients were noted, which might explain the noise in the salinity comparison.

The track plot looks ok.  The cross-reference lists turned up no problems.

As a final check of dissolved oxygen data % saturation was calculated and plotted. The values varied from 70 to 95% with the lowest values in Haro Strait and the eastern end of Juan de Fuca Strait and the highest in the southern part of the Strait of Georgia.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

For all casts the following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units. Then the files were reordered to put the two SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods and to fix the project name.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

It was discovered at this point that the Reference PAR had not been converted for the bottle files. A new set of files (*.ros2) were converted, then converted to IOS SHELL format. They were treated in the usual way to add sample numbers, bin average on bottle numbers, and fix formats. The previously created CHE files were reordered on sample number and then merged with the new files choosing all channels from the original CHE files and only Reference PAR from the new ones. The output files were once again named CHE.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars from log:

14-16. Error in date in file headers. Fixed.

35. Altimeter spiky.

37. Error in date in file header. Fixed.

46. CTD hit bottom.

52. Alt. came on @250m.

55. Started alt at 35m from bottom.

56. Alt. started at 319.

65. Rough.

66. Error in date. Fixed.

72. Alt. started ~300m.

73. Alt started ~70m.

Note: There were many other errors in dates.

REVISION: July 2, 2009

- Nitrates were recalculated in June 2009 to correct for too short a cadmium column resulting in incomplete colour development. All results coded "c" for this reason. Previous values were     removed and new ones inserted. See “QF2008-61nuts recalculated nitrate.xls” and “AstoriatoTechniconDec15.xls”.

- The order of channels was changed at the same time to put the Reference PAR above the bottle sample channels.

REVISION: August 18, 2010

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly but steadily until July 2009; then a sudden shift occurred, so that maximum values between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study was made of transmissivity that led to a decision to apply post-processing corrections to all cruises between March 2008 and June 2010.

Transmissivity data from this cruise were corrected by multiplying the original values by correction factor 1.141. This was based on assumptions that deep offshore transmissivity from a June 2009 cruise should be about 62%/m and that drift was linear with time between March 2008 and July 2009. The corrections produced reasonable results for all cruises in that period.

For details on how the correction factor was derived see:

OSD_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer 1005DR Corrections.doc

These data should be considered estimates.

Revisions done by: Germaine Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 06May08 | Factory “ |  |  |
| Conductivity | 1729 | 13Jun08 | “ |  |  |
| Secondary Temp. | 2449 | 06May08 | “ |  |  |
| Secondary Cond. | 2424 | 07May08 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 8Feb2008 | Factory |  |  |
| Fluorometer | 2228 | ? | IOS |  |  |
| PAR | 4656 | 11Feb2003 |  |  |  |
| Surface PAR | 16504?? | 2Jan2004 |  |  |  |
| Pressure Sensor | 63507 | 25/10/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |