## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1 April 2025 | Updated channel names & formats in TOB files. GG |

## PROCESSING NOTES

Cruise: 2009-01

Agency: PBS, Salmon and Freshwater Ecosystems, Nanaimo, BC

Location: North-East Pacific

Project: High Seas Salmon

Party Chief: Morris J.

Platform: W.E. Ricker

Date: February 25, 2009 – March 22, 2009

Processed by: Germaine Gatien

Date of Processing: July 1, 2009 – July 13, 2009

Number of original CTD casts: 144	Number of CTD casts processed: 140 (3 contain only surface data, 1 cast not logged)

Number of original TSG casts: 1*		Number of TSG casts processed: 1

*Only 1 TSG file was available at the time of original processing, but more are believed to have been collected.

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with a Seapoint Fluorometer (#2229) and an SBE43 Dissolved Oxygen sensor (#0047) on the primary pump with a 10X cable. The deck unit was a model 911 (#0471). The data logging computer was PACOSAPFS03. The salinometer used was a model 8400B Autosal.

# SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard.

There was no deep salinity calibration sampling. There were many samples from ~10m. No recalibration was applied to the salinity, but this should be revisited when the conductivity sensors are next recalibrated at the factory. The primary sensors were used for casts 174 & 424; the secondary, for all others.

Dissolved Oxygen data were recalibrated based on the results of cruise 2008-01 because there was no calibration sampling during this cruise. Since it is likely that some drift had occurred between the two cruises, these data must be considered less reliable than usual.

For cast #138 the data come from the upcast and there is no data below 49db. The pumps were not turned on until the CTD reached 48.9db during the upcast.

Only one thermosalinograph file (with about 41 hours of data) was available at the time of processing, though it is believed that the TSG was running throughout the cruise. There were no loop samples and no intake thermistor. The temperature and salinity from the thermosalinograph are considered bad. Only time and positions will be archived.

# PROCESSING SUMMARY

##### 1. Seasave

This step was completed at sea; some of the raw data files have extension HEX and others DAT.

The file names were non-standard and had to be corrected before conversion could be done.

##### 2. Preliminary Steps

The Log Book was obtained. There were good notes describing problems encountered including the following that may affect processing:

- Cast #132 – new version of SEASAVE used, change from HEX files to DAT files.

- Cast #138, pumps not turned on for downcast. Part of upcast also has pumps off.

- Around time of cast #147 – problem with date in TSG files – rebooted SEASAVE – fixed problem.

- After cast #197 – Daylight Savings Time starts, but Ricker clocks did not change until later. However, UTC was used in the log and appears to be used in the data files, so this does not appear to be a problem. (The local time is noted in another column, but that looks wrong in some cases.)

- After cast #204 Ricker clocks changed.

- The configuration file for cast #259 was missing, but there was one for cast #258 which was cancelled, so that one is assumed to be what was used for #259.

- Cast #262 – cancelled due to computer crash, but there is a data file which contains almost nothing.

- Cast #265 – used a different computer but could not turn pumps on.

- Cast #268 – back to CTD computer.

The bottle salinity, nutrient and chlorophyll data were obtained in spreadsheet formats.

The cruise summary sheet was completed. The histories of the dissolved oxygen, conductivity and pressure sensors were obtained as well as that of the thermosalinograph conductivity.

One configuration file was saved as 2009-01-ctd.con. Calibration constants were checked. No errors were found in the CTD con files, but the oxygen calibration parameters are not the ones that should be chosen for the new SeaBird algorithm. Those parameters were entered; nominal values were entered for E, H1 and H3 since there has so far been no opportunity to fine-tune these parameters for this sensor.

##### 3. Conversion of Raw Data

All data were converted using the con file given above.

A few casts were examined and all expected channels are present and look reasonable.

The primary temperature looks noisy during upcasts.

The dark value of the CTD fluorescence is reasonable (<0.13mg/m3).

The files for casts #41, 129 and 262 contain only a little surface data. The log notes that #129 and #262 were cancelled, but #41 is said to be a cast to 70db.

##### 4. WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.    Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### 5. CELLTM

Tests were run on 3 casts using a variety of settings for CELLTM to determine the best choice of parameters. The best choice varied for different depths and casts. For the primary there was a lot of noise in the upcasts making it hard to pick, but overall (0.02, 7) seems best. For the secondary the choice of (0.03, 9) was clearly best.

CELLTM was run using (0.02, 7) for the primary and (0.03, 9) for the secondary conductivity for all casts.

##### 6. DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### 7. Test Plots and Channel Check

Few deep casts with steady descent rate were available to judge the differences between sensor pairs. Three casts to at least 150m were found with steady descent rates; plots were made to check for agreement between the pairs of T and C sensors. The following values are rough estimates from downcast data:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 235 | +0.0007 | -0.0001 | -0.0017 | High, steady |
| 32 | 160 | X Noisy | -0.0002 | X Noisy | High, steady |
| 259 | 240 | +0.0004 | +0.0002 | -0.0002 | High, fairly steady |
| 412 | 200 | X Noisy | -0.0002 | -0.0019 | High, steady |

The temperature differences are very noisy, and the variations in conductivity and salinity differences don’t seem related to time or pressure. The differences are not large.

##### 8. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### 9. Checking Headers

The header check was run and an error in agency name was noted – that will be fixed at the end when HEAD EDIT is run. The station name was missing for 4 casts, so those were added to the CLN files. No further errors were found.

The track plots (using event #s and station names) were produced and added to the end of this report.

The average surface pressure is ~2.6db which is a little lower than usual for the Ricker. Two casts were examined that had surface pressures <0.5db during the upcast with the pumps running, and the salinity at that level was about 3 units, so this is clearly very close to the surface. The offset looks appropriate.

Casts #129 and 262 will not be processed further since there is only data from the surface.

##### 10. SHIFT

Fluorescence

In the past it has been found appropriate to advance the Fluorescence by +24 records (1s) relative to the other channels. A few casts were examined to compare the offset between the upcast and downcast fluorescence with that of the temperature traces to ensure this setting is appropriate. The results were highly variable, complicated somewhat because the CTD was stopped at 15db during the upcast, so this is a little hard to judge, but +24s looks like a reasonable choice.

A shift of +24 records (1s) was applied. (Output: SHFFL)

Conductivity

Tests were run on the primary and secondary conductivity channels using a variety of settings. There were variations from feature to feature but the best results overall were with a choice of -0.5s for the primary and -1.1s for the secondary.

All casts were put through SHIFT using -0.5s for the primary and -1.1s for the secondary conductivity.

Dissolved Oxygen

Tests were run on a few casts to compare the offset between downcast and upcast oxygen compared to that in the temperature traces after a variety of offsets were applied to the DO data. Settings from +40 to +60 records all did a reasonable job. SHIFT was run using an offset of +50.

##### 11. REVERSE

##### All files were put through REVERSE to be used to compare with bottles. The file for cast #138 will also be used for archiving since the pumps were not on for the downcast.

##### 12. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              		Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There was a warning for cast #41 because it contained only surface data; it will not be processed further. There was also a warning about cast #258 which concerned a pressure jump at the surface. The first record in the DELETE file was removed and the rest of the data look fine.

All DEL files were copied to *.EDT.

DELETE was run on the reversed file for cast #138 and the output file was copied to EDT.

##### 13. COMPARISON WITH BOTTLES

Bottle File preparation

There were no deep bottle samples, but there were samples from ~10m for all casts. There is a note in the log that the Niskin bottle was 5m above the CTD. Examination of the CTD upcast data shows that it typically stopped around 14.5db, so the bottle data are presumed to come from 9.5db and should be compared with CTD data from that level. The stops do vary somewhat, so there is room for error.

All CTD casts were put through REVERSE, bin-averaged using a 0.5m bin size. Then data from 9.5m was extracted for comparison with the bottle files and that was exported to a spreadsheet.

The salinity spreadsheet was simplified by removing columns not needed for the comparison with CTD data. The sample # and station name were combined in a single column, so those were separated. The file was saved as 2009-01-sal.csv and it was then converted into individual SAL files.

Extracted chlorophyll data were obtained in spreadsheet format; this was simplified and saved as file 2009-42-chl.csv. All data were the average of duplicates so were flagged “f”, and those with duplicate variability >10% (as calculated by the analyst) were flagged “fc” and a comment of explanation was added. The spreadsheet was then converted to individual CHL files.

The nutrient spreadsheet was simplified and saved as 2009-01-nuts.csv and converted to individual NUT files.

MERGE was used to combine the SAL, CHL and NUT data into merged files; the data from those files were exported to a spreadsheet and then the CTD and Bottle spreadsheets were combined in file 2009-01-upcast-ctd-bottle.csv. A few manipulations were needed to ensure that the event numbers matched. This file was then simplified by removing some columns and saved as 2009-01-bottle_comp.xls. Then comparisons were made between CTD salinity and fluorescence and bottle samples.

Salinity

The differences between CTD salinity and bottle salinity are highly variable with an average using all available data showing the primary CTD to be low by ~0.02 and the secondary low by about 0.017. Looking at a plot of differences a section was found with little variability between casts #231 and 364, and the average for that section shows the two CTD salinity channels to be low by 0.005 and 0.008. Finally an average of casts for which Absolute Value of differences <0.01 showed the primary salinity to be high by 0.0002 and the secondary low by 0.0017. Given that the secondary conductivity sensor appears not to have been used since its last factory calibration, being off by 0.002 seems more likely than 0.008. It is possible the bottle sample came from a little higher than estimated so that we are comparing it to CTD data that is from too low a level.

A look at downcast CTD data shows the differences between bottles and CTD were noisier and larger, on average, presumably due to temporal variations.

Fluorescence

The CTD fluorescence at 9.5db is lower than the extracted fluorescence from the bottles with a ratio of Extracted CHL to CTD Fluorescence of ~1.5 at low values to ~2.5 for CHL>2ug/l. We do expect fluorescence to be a little low during upcasts since the pump does not seem to work as efficiently. However, a quick check against downcast data gives only a slightly lower ratio.

##### 14. DETAILED EDITING

The choice of which sensor pair to use is not obvious. The primary appears to be slightly closer to the bottles, but this comparison is rough. The primary conductivity sensor has been used a lot, but without any strong bottle comparisons. The secondary has probably not been used since factory recalibration, so should be more reliable. There are no notable spikes in either salinity channel. The secondary pair was chosen for editing. In the process of editing it was found that the primary data were better for 2 casts.

CTDEDIT was used to clean the primary temperature and salinity data for the following casts: 174, 424.

CTDEDIT was used to clean the secondary temperature and salinity data for all other casts.

Upcast data were used for file #138.

On-screen plots of descent rate and pump status were used to guide editing.

All casts required some editing.

Note was made of the editing details in the headers of the relevant files.

##### 14. Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensor was recalibrated in May 2008 and has been used many times, but there has not been a good bottle comparison. This is the first cruise processed at IOS using the secondary conductivity sensor since it was calibrated in June 2008, but it is possible that it has been used.

- The pressure sensor was recalibrated in May 2006 and there is no evidence of the calibration drifting since then.

- The dissolved oxygen sensor was repaired and recalibrated in Feb. 2007. It was used on 2008-01 and 2008-15, but there was no calibration sampling for the latter cruise.

Historic ranges – Profile plots were made of T and S with local climatology superimposed where available. Most data were within the historic ranges; the excursions, sometimes high, sometimes low, are mostly associated with odd looking historic ranges. Others are from casts close to shore; the climatology covers fairly large regions and the near-shore areas are probably not well sampled within it. There is no suggestion of a systematic error.

##### 15. Initial Recalibration

Given the rough nature of the bottle comparison, no salinity recalibration is warranted. When these sensors are next calibrated at the factory, or used for a cruise with rosette sampling, this decision can be revisited. (If that happens casts #174 and 424 will need separate treatment as they have primary channels archived.)

There is no basis for recalibration of the dissolved oxygen, except to use the results of cruise 2008-01 since it is the only other cruise with calibration sampling that used this sensor since it was last recalibrated. The use of the new DO algorithm is not expected to change the fit much, but we do expect some drift. The second recalibration will not be applied as that tends to vary from cruise to cruise and is a small correction.

File 2009-01-recal1.ccf  was used to apply the following correction to the DO channel:

Dissolved Oxygen (Corrected) = 1.0633 * Dissolved Oxygen - 0.0194

##### 16. Fluorescence Processing and special files for Angelica Peña

The EDT files were put through a median filter, size 11, applied to the fluorescence channel only.

Five sets of files were prepared for the use of Angelica Peña:

TWO DOWNCAST DATA SETS

- The EDT files were clipped to 150db, bin-averaged with 0.25db bins, unwanted channels removed and HEADEDIT used to fix headers, formats, channel names and to add comments. (Output: FCTD1)

- The EDT files were clipped to 150db, median filtered (width 11 applied to fluorescence only), bin-averaged with 0.25db bins, unwanted channels removed and HEADEDIT used to fix headers, formats, channel names and to add comments. (Output: FCTD2)

THREE UPCAST DATA SETS - Because there were no rosette files, special files were prepared with upcast data from around the depths at which chlorophyll was sampled. The REV files described in section 11 were put through CLIP. (Output: REVCLIP) In a few cases there were negative values near the surface so CLIP was set to remove data between -1db and 20db.

- The clipped and reversed files were put through DELETE, then REMOVE and HEADEDIT were run to remove unwanted channels and fix formats and add a comment. (Output: UPFCTD1)

- Another set of upcast files were created by applying a median filter of fixed width 11 to the fluorescence channel of the REVCLIP files followed by DELETE, REMOVE and HEADEDIT. (Output: UPFCTD2)

- Finally a set of upcast files were created by running DELETE, REMOVE, BIN AVERAGE and HEADEDIT. (Output: UPFCTD3) The bins were size 0.25db.

##### 17.  BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. No further editing was considered necessary.

Transmissivity and Fluorescence profiles were plotted on-screen. The only odd feature noted was a spike in fluorescence at the bottom of cast #94. The cast was shallow and near-shore, so this was left unchanged.

##### 18. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from casts #174 and 424: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag.

The following channels were removed from all other casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag.

A second CTD dissolved oxygen channel was derived with units umol/kg. REORDER was run to put the two DO channels together.

Dissolved Oxygen saturation was calculated and surface values were plotted. The values north of Vancouver Island varied from about 90% to 110% which suggests that the calibration is reasonable. Values were somewhat higher to the south with most between 100% and 130%; the higher saturations are associated with higher fluorescence values. Values were lower in areas with strong mixing.

HEADER EDIT was used to change the agency name to "PBS, Salmon And Freshwater Ecosystems, Nanaimo, B.C.", fix formats and channel names, to fix the platform name and geographic description and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

Dissolved Oxygen data were recalibrated based on the results of cruise 2008-01

because there was no calibration sampling during this cruise. Since it is likely

that some drift had occurred, these data must be considered less reliable than usual.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### 19. Thermosalinograph Data

a.) Checking calibrations

There was no configuration file provided. One was prepared based on a post-cruise calibration from shortly after this cruise and just before repairs were done to the TSG. That was named 2009-01-tsg-15apr09.con was prepared.

b.) Converting to IOS Headers and adding position headers and time channels

Only 1 hex file was available as processing began. When converted the CTD data had temperatures that look too high (~14C˚ around Texada Island in late February) and salinity too low (~26). So alternate conversions were tested using parameters from the 2008 calibration and the post-repair calibration. The differences between the pre-cruise and post-cruise data sets show that calibration drift will only account for temperature errors on the order of 0.001C˚ and 0.015 in salinity. Even the post-repair calibrations show only small differences. The post-cruise pre-repair calibration described above was used.

The data were converted to a CNV file using a SeaSoft routine. The channels converted were: Scan_Number, Temperature:Primary, Conductivity:Primary, Latitude, Longitude, Salinity:T0:C0 and Time Julian. There was no channel for Flow Rate, no secondary temperature and no fluorometer.

The files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add time and date channels and the output files were named *.ATC.

An initial track plot was produced and it looks fine.

It is believed that this data will only be used for the positions since the T/S data look unreliable. So REMOVE was used to remove the following channels: Scan_Number, Conductivity:Primary, Temperature:Primary, Salinity:T0:C0, Flag.

Since no one is available to look for the data, the processing job was finished with just the 1 TSG file. If and when more data turn up, they can be added.

HEADER EDIT was used to change the DATA TYPE to THERMOSALINOGRAPH and to add the depth of sampling to the header and the following comment about the data.

The temperature and salinity data were removed since the values are clearly bad. Only time and positions are considered reliable.

As a final check plots were made of the cruise track and data; no problems were noted.

##### 20. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the TSG and CTD sensors.

Particulars: (Notes from log book)

46. Not logged.

129. Cancelled due to SEASAVE problem.

132. New version of SEASAVE used.

138. Pumps not on for downcast – came on partway up.

147. Around this time problem with date in TSG files – rebooted SEASAVE – fixed problem.

197. Daylight Savings Time starts, but Ricker clocks did not change until later. UTC was used in the log and the local time is noted in another column for some casts. Data file times seem ok.

204. Ricker clocks changed. Does not appear to affect the data files.

259. Configuration file missing, but there was one for cast #258 which was cancelled, so that one is assumed to be what was used for #259.

262. Cancelled due to computer crash.

265. Used a different computer but could not turn pumps on.

268. Back to CTD computer.

# Institute of Ocean Sciences

# CRUISE SUMMARY

| Cruise ID#: 2009-01 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Dates: Start: 25 February 2009 End: 22 March 2007 |  |  |  |  |  |
| Location: NE Pacific |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: Morris J. |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506         Cruise ID#:		2009-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4752 | 06/03/07 | Factory |  |  |
| Conductivity | 2173 | 07/05/08 | “ |  |  |
| Secondary Temp. | 2968 | 22/08/07 | “ |  |  |
| Secondary Cond. | 2399 | 13/06/08 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Transmissometer | 723DR | 24/01/08 | IOS |  |  |
| Pressure Sensor | 69698 | 26/05/2006 | Factory |  |  |

# TSG Calibration Information

Make/Model/Serial#:	SEABIRD/21/2488       Cruise ID#:	2009-01

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 23/1/08 | Factory | 15/04/09 | Factory |
| Conductivity | 2488 | 23/1/08 | “ | 15/04/09 | Factory |