## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 26 Feb 2025 | Corrected channel name for channel Nitrate_plus_nitrite:ISUS in CTD and CHE files. G.G. |
| 10 July 2020 | Fixed Julian Time channel in TOB files. G.G. |

## PROCESSING NOTES

Cruise: 2009-06

Agency: OSD

Location: North Pacific	/ Bering Sea / Chukchi Sea

Project: C3O/IPY

Party Chief: Vagle S.

Platform: Sir Wilfrid Laurier

Date: July 10, 2009 - July 22, 2009

Processed by: Germaine Gatien

Date of Processing: 30 January 2019 – 7 March 2019

Number of original HEX files: 28

Number of CTD files:  28	Number of CHE files:	28

Number of original TSG files: 16

# INSTRUMENT SUMMARY

SeaBird Model SBE 911+ CTD (#0941) was used for this cruise. It was mounted in a 24-bottle rosette and attached were a Wetlabs CSTAR transmissometer (#993DR), a SeaPoint Fluorometer (#2336) with a 10X cable for events 1-13 and a 3X cable for events 14-28, an SBE 43 DO sensor (#1483), a Rinko III DO sensor (#005), a Rinko III Temperature sensor (#005), an ISUS (#72), a SPAR sensor (#20281) and an altimeter (#1252).

A thermosalinograph (SeaBird 21 S/N 3274) was mounted with a fluorometer (SCF2841) with either a 10X or a 30X cable and a remote temperature sensor #0271.

The data logging computer was SWL #1.

The data acquisition program was Seasave 7.18c.

The deck unit was a Seabird model 11, serial # 11P53201-0800.

The salinometer used at IOS was a Guildline model 8400B Autosal, serial # 69086. Samples were analyzed in January 2010.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was generally in good order with notes about problems encountered. Information about the TSG was not entered in the opening section of the log. Information about bottle sampling was found in a spreadsheet and rosette logs.

Salinity samples were analyzed 6 months after the end of the cruise. Both CTD salinity channels were both found to be low by 0.0008psu and 0.0013psu compared to bottles. This estimate was based on data below 75db so the effect of incomplete flushing of Niskin bottles should not be large. Evaporation and adsorption of samples may have raised the bottle values so the CTD salinity could be reading closer to correct or possibly even a little high. A post-cruise calibration shows that primary salinity was reading low by ~0.001psu and secondary reading low by ~0.0005psu after several other cruises.

The comparison of the SBE DO sensor with titrated samples produced a fit similar to that found during cruise 2010-01 after which the sensor was returned to the factory for service. The membrane was found to be torn at that time so response may be poor.

The SBE DO sensor has a fairly long response time so data accuracy is not as high when it is in motion as it is during stops for bottles. This will be especially true when vertical DO gradients are large. To get an estimate of the accuracy of the SBE DO data during downcasts (after recalibration) a rough comparison was made between downcast SBE DO and upcast titrated samples. Some of the difference will be due to problems with flushing of Niskin bottles (especially in high gradient zones) and/or analysis errors, so the following statement likely underestimates SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 35db

±0.25 mL/L from 35db to 250db

±0.08 mL/L below 250-400db

±0.015 mL/L below 400db

A SeaPoint fluorometer was mounted on the CTD for this cruise. While the CTD fluorescence data are expressed in concentration units, they do not always compare well to extracted chlorophyll samples, particularly for casts far from shore. It is recommended that users check extracted chlorophyll values where available. During this cruise SBE fluorescence generally was higher than CHL for CHL<1ug/L. For CH L≥ 1ug/L the ratio gradually dropped with FL/CHL at 0.7 when CHL = 22.34ug/L.

Data from a Rinko oxygen sensor were converted but have not been processed for the OSD Data Archive. Sarah Zimmermann did a comparison of the Rinko data with the SBE43 data and details can be found in report 2009-06 RINKO SBE oxygen sensor comparison*.docx.

The raw voltage from an ISUS sensor will be archived.

TSG file names were changed to format: 2009-006-YYYYMMDD-HHMMSS

The TSG performed well through the cruise. TSG intake temperature reads higher than the CTD at 5m by an average of 0.07Cº while salinity reads lower by about 0.006psu. The post-cruise calibration shows little drift in temperature so this error may be due to a small vertical offset in the 2 sources with the loop drawing water from a little above the CTD sampling depth. This would also lead to salinity reading a little low. The post-cruise calibration showed a drift of ~0.009psu over 9 months and at least 4 cruises, of which this is the first. Salinity was recalibrated by adding 0.003psu as a rough guess of how much drift occurred by the middle of this cruise. Fluorescence values are much lower than CTD fluorescence and extracted chlorophyll samples from the rosette when CHL >1ug/L.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

The format for cruise numbers had changed after this cruise but before processing was done, so processed files will have the new standard, 2009-006.

##### Preliminary Steps

The Daily Science Log Book and rosette log sheets were obtained and a cruise report was available. File 2009-06_SWL_Chem*.xlsx contains the chemistry data and a rosette log file and there are also paper rosette log sheets. Scans were made of the paper logs.

The pressure sensor, conductivity and DO sensors had all been recalibrated since the previous use on record so there is no sensor useful history.

There were a variety of calibration control files used.:

- Event #1 had Scan Time added turned off and apparently had the wrong altimeter serial number.

- Event #2 had Scan Time added turned off and had the right altimeter serial number.

- Event #3 had Scan Time added turned ON and it stayed on for the rest of the cruise.

- Event #14 had a change to the fluorometer gain to 10X and it stayed that way through the rest of the cruise.

The calibration constants were checked for all instruments and a few problems were found:

- The transmissivity calibration factor M needs fixing; it had been calculated incorrectly.

- The altimeter # is wrong in the first 2 files.

- The date for the Rinko calibration was incomplete.

- The serial number was wrong for the SPAR and the calibration date was wrong.

The calibration files for casts 2, 3 and 14 were corrected and saved as 2009-006-ctd1.con (casts 1-2), 2009-006-ctd2.con(casts 3-13),  and 2009-006-ctd3.con(casts 14-28). The errors noted above were corrected in each of them.

There were post-cruise calibrations from November 2009 for many sensors

The pre-cruise calibrations were chosen for processing.

All files were converted using either 2009-06ctd1.con, 2009-06ctd2.con or 2009-06ctd3.con as appropriate.

All channels were plotted for 1 cast each using the 3 configurations. The temperature and conductivity channels track fairly well on downcasts but are much noisier during upcasts. There are some spikes in the primary channels. Altimetry looks useful. The transmissivity, SPAR, fluorescence and dissolved oxygen have the right shapes. Descent rates are high.

##### Hysteresis Study

The nominal hysteresis factor was applied, but there was no sampling below 1000m, so no tests were run to fine-tune the factor.

##### Bottle File Preparation

The ROS files were converted in the same way as the full files except that bottle number and bottle position channels were included and oxygen concentration and salinity were derived.

The files were put through CLEAN to create BOT files.

Temperature and salinity were plotted for all BOT files. There were a few spikes in salinity. CTDEDIT was used to remove such spikes from casts #25 and 35. The edited files were copied to BOT.

A preliminary header check turned up errors in the 3rd configuration file. That was fixed and conversion rerun. The report also shows that station names were not entered.

The BOT files were averaged on bottle number to enable preparation of file ADDSAMP.csv. That file was edited to add sample numbers and was then used to add sample numbers to the BOT files, creating SAM files. The files were then bin-averaged on bottle number to create SAMAVG files.

The bottle sample data were presented in a single spreadsheet, 2009-06-SWL_Chem.xlsx. This was simplified to include only the columns needed for addition to the BOT files and saved as 2009-06-chem.csv.

Some changes were made to the data provided:

- For event #1 flag 2 was added to all samples to indicate that bottles were fired on the fly.

- Flags and comments had to be separated.

File 2009-006-samples.csv was then converted to individual MRG1 files. Those were put through CLEAN and SORT (on bottle #) and then merged with the SAMAVG files with output MRG. These files were put through CLEAN to remove SeaBird headers and comments and to add 0 flags where flag channels were blank and to remove channels with no data.

A file was prepared with file names and station names and that was added to the MRGCLN2 headers using routine MRG CSV file to headers.

Next, text file 2009-06-bot-hdr.txt was prepared to add an explanation of quality flags and some general comments from analysts.

##### Compare

Salinity

Compare was run with pressure as reference channel.

First cases were excluded when the standard deviation in the CTD salinity was >0.001psu were excluded plus all bottles from event #1 when bottles were fired on the fly. There was 1 other outlier that was excluded from the comparison; Bottle #6 of event #7 (sample 97) at 227db had a salinity higher than the CTD by about 0.1psu; the bottle value is out of line in the profile of bottles as well, so there is clearly something wrong with the bottle. It was flagged 4 with a comment. Other bottle samples were not out of line, so this is not a case of a misfire. There is no other level at which CTD salinity looks close to that value, so it doesn’t look like the sample was taken from the wrong bottle. There may have been evaporation of sample or an analysis problem.

With those exclusions the primary salinity was low by an average of 0.0014psu and the secondary was low by 0.0016psu with standard deviations of ~0.005psu in both. When bottles fired above 75db are excluded the fits are flatter with the primary salinity low by an average of 0.0008psu and the secondary by 0.0013psu with standard deviations of ~0.004psu in both. The difference may be due to incomplete flushing in the presence of higher vertical gradients, but this does not appear to be a serious problem.

There could be some evaporation and adsoprtion of samples given the long wait for analysis which might increase bottle salinity and thus make the CTD salinity appear to be lower than it really is. A post-cruise calibration showed the combination of temperature and conductivity drift implying little change in salinity over the 3 legs of this cruise; there is no need to recalibrate CTD salinity.

The only outlier requiring a change of flags is sample #97.

The event #1 samples were already flagged 2 to indicate they were fired on the fly.

For full details for the COMPARE run see file 2009-06-sal-comp1.xls.

Oxygen

COMPARE was run using the CTD DO sensor data and titrated DO samples.

The initial fit showed 2 major outliers:

Event 2 – sample 15 at 5m. The bottle wasn’t flagged and looks ok in profile. The CTD value looks out of line in profile and unnaturally low this close to the surface. The CTD value was padded.

Event 27 – sample 293 at 24m – this sample was flagged 2 due to an analysis problem that was not considered serious. There is a very high gradient at this level so the bottle contents could be significantly higher or lower than the CTD due to vertical offset or incomplete flushing of bottles. So there is no reason to doubt the bottle value. No flag was added.

Next the bottles from event #1 were removed since they were fired without stopping. They appear somewhat out of line, with the surface bottle especially so. More bottles were then excluded based on residuals until the following fit was found:

CTD Bottles = CTD DO * 1.0726 - 0.0003 with R² = 0.9718

This result is similar to that found during cruise 2010-01 after which the sensor was returned to the factory for service. The membrane was found to be torn at that time which may account for the high slope in the correction.

Fluorescence

COMPARE was run using the SeaPoint Fluorescence and the Extracted Chlorophyll from bottles. As usual the fit is poor for very low chlorophyll and there were many very low values in these data. CHL varied from 0.03 to 26.83ug/L and fluorescence from 0.12 to 22.34ug/L. At low values the fluorescence tends to much higher than CHL by as much as a factor of 14. For CH L≥ 1ug/L the ratio gradually drops with a ratio of 0.7 for FL/CHL at the maximum CHL value.

For more detail see file 2009-06-fl-chl-comp1.xls.

##### Conversion of Full Files from Raw Data

All files were converted using 2009-06ctd1-con, 2009-06ctd2.con or 2009-06ctd3.con as appropriate.

A few casts were examined and all expected channels are present.

The two temperature and conductivity channels are close overall during downcasts though there is quite a lot of noise in both channels near the surface. During the upcasts, traces differ more and there are odd excursions that are often seen in upcasts. This is likely something to do with how the CTD was mounted.

Fluorescence went off-scale during cast #27 and there are a few other spikes that look off-scale.

Altimetry looks useful when the CTD got near the bottom; dissolved oxygen, fluorescence, PAR and transmissivity profiles look normal.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity & temperature only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

The parameter “Keep data within this distance of the mean” was set to 0 so all spikes would be removed.

##### ALIGN DO

Tests were done on a few casts to determine the offset between the DO channel and the primary temperature. It is very hard to judge because the temperature is noisy on the upcast and there were many stops for bottles. There is always a slight shift in values between upcast and downcast so that matching features works better than just looking at the vertical offset of the trace. An advance of 2.5s looks reasonable.

ALIGNCTD was used to advance the DO Voltage by 2.5s relative to the pressure.

##### CELLTM

The upcast data are extremely noisy and there were many stops for bottles, so the usual tests for CELLTM settings were a little hard to interpret. For the deeper casts in places where the upcast data was fairly smooth, all settings tested improved the data somewhat, with a choice of (α = 0.0245, β=9.5) looking as good as any other choice; these values are default values. For the shallower casts all settings appear to make things worse, but an examination of profiles shows that the default setting has only a minor effect on the downcasts, but affects the upcasts much more. So CELLTM will be applied in the usual way.

CELLTM was run using (α = 0.0245, β=9.5) for both the primary and secondary conductivity.

##### DERIVE  -

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration. (The same con file was used for all since it is the same DO sensor in all versions.)

on  3 deep casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-06-0003 | 1000 | +0.0006 | ~0 | -0.0011 | Fairly High, Moderate |
| 2009-06-0007 | 1000 | +0.0006 | ~0 | -0.0013 | High, Noisy |
| 2009-06-0015 | 1000 | +0.0007 | ~0 | -0.0013 | High, Very Noisy |

All the differences are small. The differences between salinity channels are slightly higher than the -0.0005psu difference found in the bottle comparison, and lower than the +0.0008psu based on the post-cruise comparison. The bottle comparison was noisy and the post-cruise estimates are rough but this may imply that the drift in the primary conductivity occurred mostly after this cruise.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the pressure channel with interpolated values based on record number.

Next plots were examined to see how many records needed to be removed from the beginning of files to eliminate the soak periods which lasted from 1 to 2.5min. Cast #21 appears to have had no soak but this was a 2nd attempt at acquisition so it probably did have a soak that was not recorded.

CLIP was run to remove the soak records.

Program “MERGE CSV file to Headers” was used to add station names to the headers using file “stationName-mrg.csv”.

##### Checking Headers

The header check was run.  There are some slightly negative pressures. Speeds look reasonable.

Surface check was run and shows an average surface pressure for the cruise was 2.1db. There were 4 very low pressures (+0.04, +0.03, +0.05, -0.05db) with very low associated salinity values. Examination of a few casts suggests that the pressure is within 0.2db. At the end of upcasts the conductivity fell to very low values when pressure was between +0.1 and -0.2db. The factory calibrations for pressure look appropriate, as we would expect so close to recalibration. There is no evidence that the sensor was reading high as indicated in the post-cruise calibration.

The cross-reference check was compared with the log book and no problems were found.

The altimeter readings and bottom depths from the headers of the MRH files were exported to a spreadsheet. The water depth was missing from some files, so was added at this point and the export rerun. Note that there was no bottom depth available for event #15 due to a sounder problem. The altimeter readings look fine for most casts. For cast #24 it looks slightly high but there was a lot of variation near the bottom so the value likely reflects an average in an area of variable depth. For cast #23 the altimetry reading looks too high for the full file and likely arose from a spike in the upcast. The reading was adjusted from 4.96m to 3.1m. It looked ok for the bottle file.

The cruise track was plotted and added to the end of this report. No problems were found.

##### 13. Shift

Fluorescence

SHIFT was run on all casts to advance the SeaPoint fluorescence channel by +24 records.

Conductivity

Tests were run on the conductivity channels. The results were plotted for 2 deep casts and 1 shallow cast. For the deep casts the best results were with an advance of -0.5 records for the primary and -0.6 records for the secondary. Those settings also worked well for the shallow data in the secondary channels but no setting looked produced particularly good results for the primary.

SHIFT was run on all casts to advance the conductivity by -0.5 records for the primary and -0.6 records for the secondary.

Dissolved Oxygen

The Dissolved Oxygen voltage channel was aligned earlier.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

There were no off-scale fluorescence data after this step.

##### Other Comparisons

Sensor History

Temperature, dissolved oxygen and salinity sensors had all been recalibrated since they were last used.

Post- cruise reports

There were post-cruise calibrations from November 2009 for many sensors

The primary temperature was found to have drift high by 0.00018C°/year and the secondary high by 0.00054C°/year.

The primary conductivity was reported to have drifted low by 0.0001psu/month (-0.0008psu over the 8 months between calibrations.)

The secondary conductivity was reported to have no drift over the 8 months between calibrations.

The combined effect on salinity would lead to the primary being low by ~0.001psu and the secondary by ~0.0005psu.

Pressure was found to be high by 0.9039db.

Dissolved oxygen was found to be low by 8% in March 2011. The membrane was reported to be torn at that time.

According to the Arctic Group’s calibration spreadsheet fluorometer #2336 was written off in June 2010.

Historic ranges – Profile plots were made with 3-standard deviation climatology ranges of T and S superimposed. Those were only available for the first 2 casts but larger-scale climatology was available for 2 other casts. There were no excursions from those ranges.

Repeat Casts –

There were no repeat casts.

##### DETAILED EDITING

The bottle comparison shows that the primary salinity is closest to bottles, but it is much noisier than the primary. The alignment step did not improve the primary T-S traces in shallow water whereas it was more effective on the secondary. The secondary sensors showed less calibration drift and they were selected for cruise 2010-01. So the secondary channels were selected for editing and eventual archival.

CTDEDIT was used to remove large spikes, smaller spikes that appear to be due to instrumental problems and likely to affect the bin-averaged values and records corrupted by shed wakes including some surface records.

The most difficult editing for the shallower casts was in the near-surface high-gradient zone, with temperature drops on the order of 7C° in 4m with little change in salinity. There are many unstable features in that zone where small reversals in temperature lead to larger reversals in salinity. These unstable features occur while the descent rate looks reasonable steady. Tests were run to see if CELLTM had caused these instabilities but skipping that step did not improve the data. Salinity was cleaned in those zones to remove spikes that look instrumental in nature; while some unstable features remained they are small and most will be removed by bin-averaging.

All EDU files were copied to EDT.

All files were put through Bin Average using 0.5db bins and producing standard deviations. Plots were made to see if further editing was required. Many unstable features remain but most are fine-scale. Several casts were given a little further editing in CTDEDIT. There remain some small unstable features.

##### Initial Recalibration

The pressure looks ok.

We have insufficient information on which to base a recalibration of dissolved oxygen and it was recently calibrated.

The secondary salinity was found to be low by 0.0013psu in COMPARE and the post-cruise calibration of temperature and conductivity sensors implies the secondary salinity was low by ~0.0005psu, no recalibration is necessary for salinity.

File 2009-006-recal1.ccf was prepared to apply the following correction to channel Oxygen:Dissolved:SBE

CTD DO corrected = CTD DO * 1.0726 - 0.0003 with R² = 0.9718

The recalibration was first applied to the MRH and SAM files and COMPARE was rerun.

When outliers were removed (approximately the same number of outliers as in the original comparison) the CTD DO was found to be high by an average of 0.0005mL/L with a standard deviation of 0.029mL/L.

##### Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 0.5	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Some unstable features remain as described in section 16, chiefly in high-gradient zones in shallow water. Very heavy editing would be needed to remove them and often the unstable feature is due to a very small reversal in salinity. No further editing was applied.

##### Comparison of Rinko and SBE43 Dissolved Oxygen concentration

Sarah Zimmermann did a quick comparison of the Rinko DO and SBE43 DO data; details can be found in report 2009-06 RINKO SBE oxygen sensor comparison*.docx. The RINKO oxygen concentration values were lower than SBE43 values with differences between 0.05mL/L and 0.7mL/L but the differences were not quite linear. The RINKO data were not processed beyond this point. They will not be placed in the OSD Data Archive.

##### Final CTD File Steps (REMOVE and HEADEDIT)

REMOVE was run on the AVG files to remove the following channels (Output *.REM):

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate,UPoly1, Upoly2 and Flag

A second SBE DO channel (with umol/kg units) was added to the REM files.

% DO saturation was calculated and near-surface values were plotted. Most values were between 100 and 110%. Cases with values outside that range were examined more closely:

- Event #2 had a very low saturation at 80%. This is much lower than the casts before and after that one. Examination of the cast showed that the pumps did not come on until the CTD was at about 11db, a text editor was used to remove the pumped channels from the top 11.5db.

- Events #20, 23, 24 and 25 all had saturations between 113% and 120%. Some of these casts had fairly high fluorescence and low transmissivity in the top 10 to 20m. The bottles are in reasonable agreement with the CTD DO for these casts. The values are likely reflective of an biologically active area.

- Event #26 had a very high saturation at 155%. Fluorescence was ~25ug/L at 10db and transmissivity was about 10% from 10m down. Again this looks like an active area where the DO values are likely ok.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to fix the geographic area and to add the following comments:

Data Processing Notes:

----------------------

Transmissivity, Fluorescence:SeaPoint, Nitrate_plus_nitrite:ISUS:Voltage and

PAR:Reference are nominal and unedited except that some records were removed

in editing temperature and salinity.

The SBE DO sensor has a fairly long response time so data accuracy is not as high

when it is in motion as it is during stops for bottles. This will be especially

true when vertical DO gradients are large. To get an estimate of the accuracy of the

SBE DO data during downcasts (after recalibration)a rough comparison was made

between downcast SBE DO and upcast titrated samples. Some of the difference will be

due to problems with flushing of Niskin bottles (especially in high DO gradient

zones) and/or analysis errors, so the following statement likely underestimates

SBE DO accuracy.

Downcast (CTD files) Oxygen:Dissolved:SBE data are considered, very roughly, to be:

±0.4 mL/L from 0 to 35db

±0.25 mL/L from 35db to 250db

±0.08 mL/L below 250-400db

±0.015 mL/L below 400db

The post-cruise factory calibration indicates salinity drift was insignificant.

A Rinko III Dissolved oxygen sensor was mounted on the CTD, but the data from

that instrument have not been processed.

NOTE: While the CTD fluorescence data are expressed in concentration units, they

do not always compare well to extracted chlorophyll samples, particularly for

casts far from shore. It is recommended that users check extracted chlorophyll

values where available.

For further processing details see the processing report 2009-006_Processing_Report.doc.

A noted was added to the header of file #27 about the pumps being off.

The cross-reference list was produced and no problems were found.

The Standards Check routine was run and no problems were found.

The Header Check was run and no problems were found.

The final files were named CTD.

Profile plots were made and look ok.

The track plot looks ok.

The sensor history files were updated.

##### 22. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run on the MRGSORT files to remove the following channels (Output *.REM):

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate, and Flag,

Two runs of CHANGE UNITS were used to add DO channels in mass units for both SBE and bottle DO.

REORDER was used to get the 2 SBE DO channels together and the 2 Bottle DO channels.

HEADER EDIT was run to fix formats and units, change the channel name Bottle_Number to Bottle:Firing_Sequence and the name Bottle:Position to Bottle_Number and to add a comment about quality flags and analysis methods and notes about the CTD data.

Standards check was run on all files; there were problems that were sorted out with repeated runs of Header Edit with format corrections.

After HEADEDIT was run one header entry for file #1 was changed to indicate that the bottles were fired on the fly.

DATA DESCRIPTION    : Bottle:Rosette:Up:NoStop + CTD:Up

A header check was run on the final files and no problems were found.

The bottle data were exported to file 2009-06-bottles-final.csv and compared with rosette sheets.

A few discrepancies were noted:

- One problem found was that for cast #1 the CHL was flagged 2 for 2 levels at which no samples were taken. That was fixed.

- Cast #4 has some records for which there is no sampling entered. Water was collected.

- Cast 36 has a DO sample given but this is not indicated on the rosette log. There were 2 bottles at 15m and the values look the close, so this sample does appear to be from the Niskin indicated.

- Cast #7 has no silicate value for Niskin #1 but there is a flag 6. That clearly was the original plan, but as there are no values, the flag was changed to 0. There is no comment about this.

- Casts #9 & 12 had CHL data but no sampling is indicated on the log.

- There are notes about a  few bottles not being needed for the intended use, but HPLC was taken so the records will be kept in the files.

- There were no nutrients for sample 169 (#14) and 253 (#21). There are no notes to explain these, so no flags were added.

A cross-reference list turned up no errors.

The track plot was produced on screen and no errors were found.

##### 23. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

##### 24. TSG Processing

Date were provided in 12 hex files.

All CON files were the same. One con file was saved as 2009-006-tsg.con.

a.) Checking calibrations

The calibrations were checked and no errors were found.

b.) The files were converted to CNV files using the configuration file 2009-006-tsg.con.

The TSG files have non-standard format based on the date on which the file started but without time.

Sometimes there are 2 files with the same date, so an “a” was entered after the day of the month.

The names will be adjusted later.

Channel Voltage 0 was converted; plots show this is the fluorometer channel.

Tests were done to see if WILDEDIT improved the data. Many files had no suspicious spikes while others had many.  A test run on a file with salinity not converted followed by applying WILDEDIT to conductivity and then deriving salinity had no better effect than converting salinity then applying WILDEDIT to it.  Running WILDEDIT on both salinity and conductivity produced good results when run with the following settings:

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

There remain some sections with many spikes where WILDEDIT is not effective as too many scans are involved.

The files were then converted to IOS HEADER format.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

ADD TIME CHANNEL was used to add Time and Date channels based on the Julian time.

At this point the file names were changed to format 2009-006-YYYYMMDD-HHMMSS based on start times.

Time-series plots were examined on-screen. The  intake temperature is lower than lab temperature by roughly 0.25 to 0.35 degrees C with larger differences when temperatures were lower.

There is no evidence of instances of zero flow rate.

Fluorescence has a few instances of going off-scale (July 10th), so CLEAN was rerun on that file to replace FL>4.9ug/L with pad values. ADD TIME CHANNEL was rerun and the results were good.

The track plot looked good and was added to the end of this report.

c.) Editing

No editing appeared to be necessary. There are noisy patches but editing will not help.

d.) Bin-averaging

The files were bin-averaged over 6 scans to smooth the comparison with CTD data slightly.

e.)  Checking Time Channel

The CTD data were thinned to reduce the files to a single point from the downcast at or within 0.5db of 5db.

All files were opened in EXCEL and reduced to the times of CTD files. There were 28 matches in positions but for one case the CTD data were all padded.

To check for problems in the TSG clock or bad matches of TSG and CTD data, the differences between latitudes and longitudes were found. The average and median differences in latitude and longitude were all <0.0004°. This comparison shows that both the times and positions are reliable for both systems.

The spreadsheet was saved as 2009-006-ctd-tsg-comp.xlsx and will be used in the comparisons in the next section.

f.) Comparison of T, S and Fl from Loop and Rosette samples and TSG and CTD data

- T1 vs T2 The intake thermistor was connected throughout the cruise. The differences between the two TSG temperatures seen in this comparison ranged from 0.18Cº and 0.71Cº with a median of 0.30Cº. As expected, the heating in the loop is greatest where intake temperatures are lowest. At 4ºC  and 10 ºC the temperature rises by 0.37 Cº  and 0.27 Cº in the loop, respectively. The range is as expected.

- TSG vs CTD The spreadsheet comparing CTD and TSG files was examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There is a lot of variability in the comparisons which is not surprising this close to the surface where ship effects are possible in both data sets. These comparisons are all from times when the ship was stopped.

The TSG salinity is lower than the CTD salinity by a median of 0.0046psu and a standard deviation of 0.046psu. The median difference is as small as we could expect to achieve but there is a lot of variability. Thirteen casts that were fairly well mixed were selected to minimize the effect of slight mismatches in the depth of loop and CTD sampling. Those casts showed the TSG salinity reading lower than the CTD by a median of 0.006psu with a standard deviation of 0.015psu.

When all data were included the TSG intake temperature was higher than the CTD by a median of 0.075Cº but the standard deviation was 0.576Cº. When 4 large outliers were removed the TSG intake temperature was higher than the CTD by a median of 0.066Cº but the standard deviation was 0.283Cº. This difference is larger than expected but the scatter in the comparison is very large.

The ratio of TSG fluorescence to CTD fluorescence ranges from 0.1 to 1.25 with a median of 0.64.and standard deviation of 0.24. A comparison was also made to the extracted chlorophyll from 5m rosettes and the median ratio of Fluorescence/Chlorophyll was 2.5 for the CTD and 1.8 for the TSG. Fluorometers usually read higher than extracted CHL at low values; when CHL>1ug/L the ratios were 1.1 for the CTD fluorometer and 0.4 for the TSG fluorometer.

(See 2010-05-ctd-tsg-comp.xls.)

- Calibration History

As of February 2010 the intake temperature sensor had drifted by -0.00015 C°/year. The TSG primary temperature and conductivity were recalibrated in November 2009. At that time the temperature was said to have drifted by +0.0004 C°/year and the salinity to have drifted by -0.0015psu/month, so drift in salinity over 9 months was ~0.014psu. This was the first cruise on record after the factory check and there were 2 cruises between this one and the factory check. We might expect drift to be near zero for this cruise, but if the drift was linear with time the temperature would have been high by 0.0002C° in July 2009 and salinity low by ~0.008psu. The cruise that followed had too much scatter in comparisons to offer any guidance on calibration drift.

Conclusions

1. The TSG clock worked well.

2. As expected, the temperature in the loop increased more when intake temperatures were lower. When the intake temperature was 4ºC heating in the loop was ~0.37Cº  and for intake temperature 10 ºC the temperature rose by ~0.27Cº in the loop. These values look reasonable.

3. The TSG intake temperature was higher than the CTD temperature at 5m by a median of 0.066Cº when a few large outliers were excluded, but the standard deviation was 0.283Cº. This difference is larger than expected and much larger than the post-cruise calibration would explain. However, the scatter in the comparison is very large and slight mismatches in depth of sampling would be very significant in high vertical temperature gradients. Many of these CTD casts did occur in such gradients.

4. The TSG Salinity appears to be reading lower than the CTD by about 0.006psu which is about as close as we can expect to come in such a comparison. The standard deviation is high. The post-cruise calibration report suggests that the TSG salinity may be low by between 0psu (assuming close to factory calibration values) and 0.008psu (assuming linear calibration drift with time).

5. If we assume that the intake temperature sensor calibration is good based on the post-cruise calibrations, then the differences from the CTD must reflect either a vertical offset in sampling levels or local variability. If the sampling level difference is the explanation for the TSG temperature reading high, than we would also expect the TSG salinity to look low by a small amount. Thus we can’t trust the results of the salinity comparison for recalibration purposes. Overall it seems reasonable to apply a small correction to salinity to reflect that there was likely some drift over the course of the  cruise, so a correction of +0.003psu will be applied to TSG salinity

5. The fluorescence from the TSG reads low for CHL>1ug/L both in comparison to the CTD and chlorophyll samples from the rosette. The comparisons at low CHL are not useful as small temporal changes and small instrument errors in two low signals lead to large errors in the comparison. A warning that the values from the TSG fluorometer are low will be added to the header but since the general shape may be of interest to users the data are left in place.

g.) Recalibration

File 2007-006-recal1-tsg was prepared to add 0.003 to the TSG Salinity.

i.) Preparing Final Files

REMOVE was used to remove the following channels from all casts: Scan Number, Temperature:Difference, Flag and Position:New.

REORDER was run to change the order of channels.

HEADER EDIT was used to add a comment, change the DATA DESCRIPTION to THERMOSALINOGRAPH and add the depth of sampling to the header.

Those files were saved as TOB files.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and it looks fine.

# Particulars

1.  TSG fluorometer saturation on 30X gain so swapped out both TSG and CTD cables and replaced with 10X. CHLa level dropped shortly after but 10X cables left in place. Bottle tripped on fly this cast only.

1 &2. Scan time not added.

2. Change of altimeter. Pumps not on at first.

2. Miscast – not archiving, so repeated with same filename. Overshot on way up so cycled back down from 125m. Salinity inserts not rinsed for 11, 12, 13.

3. Front features in T, S, DO ~60m. ISUS noisy up and down and shows lots of hysteresis on upcast.

4. ISUS noisy.

5. ISUS cleaned and turned on 15min before cast 5.

6. Extra Niskins fired #15 at 35m, #18 at 15m, #21&22 at 5m – all for HPLC.

12. Power turned off before acquisition stopped.

14. Changed Seapoint gain cable to 3X for this cast and rest of cruise.

18. Jelly fish tentacles on frame.

25. ISUS battery changed after this cast.

26. Vent on Niskin bottle 1 was open when rosette brought on board – samples not taken.

# Institute of Ocean Sciences Cruise Summary

CTDs

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0941 | Yes | Yes |

| Calibration Information CTD #941 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 NEW | 21Feb2009 | Factory | 25Nov2009 |  |
| Conductivity | 3579 NEW | 18Mar2009 | Factory | 25Nov2009 | Factory |
| Secondary Temp. | 5073 NEW | 21Feb2009 | Factory | 25Nov2009 |  |
| Secondary Cond. | 3581 NEW | 17Mar2009 | Factory | 25Nov2009 | Factory |
| Transmissometer | CST0993 | 30Jun2009 | Local | 19May 2011 | Local |
| SBE 43 DO sensor | 1483 | 3Feb2009 | Factory | 10Mar2011 | Factory |
| SeaPoint Fluorometer | 2336 |  |  |  |  |
| Rinko III DO sensor | 005 | 25Jun2009 | Factory | 25Oct2011 | Factory |
| ISUS | 072 | April 2009 | Factory |  |  |
| SPAR | 20281 | 9April 2007 | Factory | 4April 2016 | Factory |
| Pressure Sensor | 941 (113140) NEW | 16Mar2009 | Factory | 7Dec2009 | Factory |
| Altimeter | 1252 |  |  |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/3274       Cruise ID#:	2009-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 3274 | 21/02/09 | Factory | 25Nov09 | Factory |
| Conductivity | 3274 | 21/02/09 | “ | 25Nov09 | “ |
| Temperature2, SBE38 | 0271 | 12/02/09 | “ |  |  |