#### 2009-10

CCGS JP Tully

#### Dimethylsulfide (DMS) & Dimethylsulfoniopropionate (DMSP) Report

#### August 19 - September 3, 2009

Prepared by Michael Arychuk

- Sample Collection

Samples were collected from all major stations (P2, P4, P12, P16, P20, P26) and for DMS, DMSPD (dissolved) & DMSPT (total).

  - DMS

Thirteen water samples from various depths (200m, 175m, 100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface) were collected at each of the stations P4, P12, P16 & P20.  At P2 there were eleven samples collected (100m, 75m, 50m, 40m, 30m, 25m, 20m, 15m, 10m, 5m, surface).  In all cases, samples were collected in 250mL ground glass stoppered bottles and  stored in the dark and removed one at a time before analysis.

  - DMSP

Six samples for both DMSPD and DMSPT were collected at each station; two at the surface (0m, 5m), one in the mixed layer (100m), one in the deep chlorophyll max (20m) and two in the salinity mix layer (175m, 200m).  The only exception to this was P2 where there were no 175m or 200m samples, hence, only 4 samples were collected.

- Analysis

2.1 DMS

A sample was loaded onto the stripper and purged with UHP Nitrogen for 10 minutes at ~100mL/min.  The DMS was extracted from the water and absorbed onto a Tenax TA trap kept at -80oC.  The trap was subsequently desorbed at 100oC (with a dewar containing boiling water) onto a Chromasorb 330 column which eluted onto a Flame Photometric Detector (FPD).  All samples were run immediately after being collected.

2.2 DMSPD

Approximately 50-75mL of seawater was allowed to flow directly from the niskin into a filtration funnel containing a 0.7m GF/F filter.  The first 3.5mL was collected in a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was stored for 24 or more hours in the dark and at 4˚C after which time 3mL was transferred to a 25mL serum bottle containing 21mL of MQ water.  An additional 1mL of a 5 Molar solution of sodium hydroxide was added to the bottle before it was crimped and sealed.  The bottle was stored in the dark and at 4˚C.’

2.3 DMSPT

3.5mL of seawater was collected directly from the niskin into a polypropylene tube (15mL) containing 50L of a 50% sulphuric acid solution.  The sample was stored for 24 hours or more in the dark and at 4˚C after which time 3mL was transferred to a 25mL serum bottle containing 21mL of MQ water.  An additional 1mL of a 5 Molar solution of sodium hydroxide was added to the bottle before it was crimped and sealed.  The bottle was stored in the dark and at 4˚C.

- Calibration

3.1 DMS

A four to six level calibration table was used for calculating the concentrations of DMS.  The standards were prepared in water and run under the same conditions, as described above, for the samples.  A calibration curve is valid for 12 hours.  If analysis exceeds 12 hours, a continuing calibration standard is run to ensure the calibration curve is still within acceptable limits.

- Quality Control

4.1 DMS

System blanks and duplicates were run approximately every 13 samples to ensure the system remained free of contamination and had acceptable reproducibility.  There was a slightly elevated blank level for P2 which was traced to the MQ water pump and as a result this value was subtracted off of the standards.  The issue was resolved and all additional blanks (for other stations) were non-detectable and overall duplicates did not differ by an average 15% (within the acceptable limits of 20%).  Stripping efficiency was evaluated at the beginning of the cruise and was proven to be acceptable at over 97%.  A performance evaluation mixture (PEM) was run at the start of every cast to further ensure method accuracy.

4.2 DMSP

Blanks and duplicates were collected at every station.  Blanks were done by simply treating MQ water as an actual sample.  For example, in the case of DMSPD it was put through a separate funnel and for DMSPT it was added directly to the polypropylene tube.

- Data & Results

  - DMS

There were no instrument problems on this cruise.

  - DMSP

Samples are to be run here at IOS within the next couple of months.

- Conclusions

6.1 DMS

Instrument and analysis performed well on this cruise.  There are some preventative maintenance issues that need to be addressed before the 2010 cruise season but none of them affected the overall quality of the data on this cruise.

  - DMSP

No problems to report.