## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 1 April 2025 | ### Updated channel names & formats in TOB files. GG |
| ### 24 Nov 2021 | ### Corrected Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 25 Nov 2020 | ### Added HLC data. S.H. |
| 24 July 2020 | Added DMSP data. G.Gatien |
| 8 July 2013 | Corrections to Nitrate and Phosphate data; see headers for details. |
| 20-Jul-2011 | Changed instrument serial number in the thermosalinograph files from 2487 to 2488. |
| 9 November 2010 | 2009-10-0028.CTD recreated – previous version was incomplete. |
| 5 October 2010 | Errors in 2009-10-0048.CHE were corrected – Nutrient and DMS data for 2 bottles had been associated with the wrong bottle #s. |
| 27May2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2009-10

Agency: OSD

Location: North-East Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: August 18, 2009 – September 4, 2009

Processed by: Germaine Gatien

Date of Processing: 15 September 2009 – ??? 2010

Number of original CTD casts: 58 (60 files*)	Number of CTD casts processed: 58

Number of bottle casts: 	57			Number of bottle casts processed: 57

Number of original TSG files:  23  		Number of TSG files processed: 17

*1 cast separated into down/up files and 1 had only surface data (NET cast)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), an SBE 43 DO sensor (#1176 on the primary pump), a Seapoint Fluorometer (#2228) with a 10X cable (believed to be on the secondary pump though the log indicates it was on the primary), a Biospherical QSP-400 PAR sensor (#4615) and an altimeter (#1252). The deck unit was a model 911+ (#619). Casts #3-38 were done with the LARS mid-ship station and casts #1/2 and 39-75 on the aft-deck with the mid-ship winch. Seasave version 7.16 was used. The salinometer used was a model 8400B Autosal, serial # 68572.

A thermosalinograph (SeaBird 21 S/N 2248) was mounted with a Wetlab/Wetstar fluorometer (WS3S-713P), remote temperature sensor #4652 and a flow meter.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order though no information was entered about the thermosalinograph and there were occasional problems with rosette sheets. It is recorded that the fluorometer was on the primary pump, but there is considerable evidence that it was on the secondary pump. The sampling notes from the Chief Scientist were very helpful in dealing with problems that occurred at sea.

Conditions were very unusual during the second occupation of station P4. Temperature and salinity had both increased markedly and dissolved oxygen concentration had decreased along lines of constant σt,, down to σt = 26; below that the changes are of the opposite sign, but not large. There was a 10X increase in the thermosalinograph fluorescence over a 6-minute period about 20 minutes before the CTD cast and the fluorescence was off-scale for over 3 hours. The Chief Scientist noted that the water was brown and there were many euphausids.

Data quality is thought to be lower for this cruise than usual, especially for some of the early casts. The new Tully mid-ship station was used for casts #3 to 38. There were many problems with kinks in the cable and re-terminations were noted in the log after casts #34 and 35. An examination of the descent rate shows that for cast #38 there were frequent CTD reversals with upward velocities that often exceeded 2m/s. Work was moved to the aft station from cast #39 onwards. Many records were removed from casts #33 to 38 because of corruption by shed wakes during such reversals, but it was hard to distinguish good data from bad. The number of records removed in editing data from cast #39 was about one third the number removed for cast #38; this was only about 6 hours after #38 and it is likely that the sea-state was similar, so it is assumed the difference is primarily due to the move to the aft deck. Some of the subsequent casts were extremely noisy and had a similar amount of data removed in editing as for #38, but in those cases it was generally much clearer which points were bad. DELETE removed most of the bad sections, leaving easily identifiable shed wake noise which could then be removed with CTDEDIT. For this reason it is believed that the quality of the data is better for casts #39 to the end.

A table-driven method of firing bottles was tested during 2 casts for this cruise. This worked well and required no special processing as long as sample numbers are assigned in the order of Firing # rather than Niskin #. For cast #3 it was easy to check that samples were associated with the right CTD data because a rosette log sheet was used that recorded both firing number and sample number and indicated what sampling had been done. The second cast was harder to check because the only numbers recorded are in the column named Niskin # but it looks like the entries are actually for Firing #. There was also no record of what sampling was done and the target pressures are mostly not entered. There are sample data for Nutrients and Extracted CHL. If we assume that sample numbers are arranged in firing order, then the Nutrients and CHL values look ok. The CHL sample labels show that the bottle numbers were changed from the original and the changes look right for a table-driven method. But to insure that bottle data get listed with the correct CTD data, it is recommended that the Firing # and Niskin # both be recorded.

The sample numbers for cast #44 were adjusted to correct for the top lanyards being off by one.

Sea-Bird have a new algorithm for calculating dissolved oxygen with several parameters that need fine-tuning for each instrument to produce the best deep data. This requires bottle samples from deep casts, which were available for this cruise. The tests were done and the results should be usable for other cruises until the next recalibration. One of the parameters had been fine-tuned before the most recent repair; it is notable that the results were different this time. So, at least when repairs are done, these tests should be redone and until there is further evidence to suggest otherwise they should be redone after recalibration. The best method for determining parameters E, H1 and H3 was found to be examining deep casts in COMPARE to see which values led to deep data (below DO minima) falling closest to shallow data in plots of differences (CTD – bottle DO) versus DO Concentration.

The salinity analyst noted that many bottles had loose lids, missing or loose caps and in several cases bottles were so full that some water had to be removed to enable effective shaking; in one case this led to cracked threads. The loose lid problem seemed especially prevalent during duplicate sampling. No such problems affected any of the loop samples.

Most of the salinity analysis was done using an Autosal at IOS, but some samples were run at sea using a Portasal. Because the at-sea salinometer broke there are not many samples, so there is little basis on which to recalibrate that data to match the Autosal. They do not stand out in COMPARE, so any errors are probably small, but all Portasal samples were flagged “c” to indicate lack of confidence in the calibration. The Portasal was very stable for the short time it was used as evidenced by very little variation among 2 sets of 12 bottles fired at a single depth.

There were problems with both pumps on different casts, sometimes for a whole cast, but most often for part of a cast. There were no obvious problems with the secondary pump for any upcasts, so secondary channels were chosen for all CHE files. The dissolved oxygen channel had to be removed from the CHE file for cast #49. For the CTD files, the secondary temperature and salinity were used for all casts except #34 for which the primary were selected. Fluorescence was removed from #34 and dissolved oxygen data were removed from the CTD files for casts #49 and 67.

The SBE dissolved oxygen data are considered of lower quality than usual. The calibration drift through the cruise is unusually large. It is possible that anoxic sampling in Saanich Inlet caused the problem but if so, the effect is more dramatic than ever seen before, and it is usually only seen when there is a problem with the membrane. No damage was found in a post-cruise calibration, nor was there an unusually large drift in calibration. Recalibration of the data was complex, using 8 different fits based on the calibration data available.

After most of the work had been done on this job it was found that the DO titrations were not reliable. There were two errors, a small one due to software picking the endpoints a little early, and a larger one of opposite sign, that was probably due to a bad batch of standard. The thiosulphate titer of 0.709 obtained during standardization was low and out of range leading to a suspected overestimate in DO value of ~2.5%. Tests were done using a post-cruise calibration of the SBE DO sensor and they also suggested an error of roughly 2.5%. The data were recalculated using the mean thiosulfate titer for the year (0.721).

All samples are flagged “d" since the data are considered to be of lower quality than usual. The SBE DO oxygen channels were recalibrated after running a comparison with the new titrated values. The changes do not explain the variable fits noted before. (See section 24 for details.)

The precision of the SBE dissolved oxygen channel is difficult to estimate because the comparison with bottles was very noisy, but roughly, the DO should be considered:

•	±0.5ml/l from       0- 300db

•	±0.15ml/l from 300- 500db

•	±0.06ml/l from 500-2000db

•	±0.1ml/l below 2000db (Below 2000db the values are mostly low by about 0.06.)

It has been noted in the past that the residual noise in the Dissolved Oxygen signal is amplified by the use of the TAU correction, particularly at depth. However, for this cruise little difference was noted so the effect is minimal with noise on the order of ±0.005ml/l below 1500db. The correction was applied to improve the response in the high-gradient zone.

The salinity is noisier than usual even after metre-averaging; this may be due to pump problems. A lot of editing was done where it was clear the spikes were due to misalignment of temperature and salinity in the presence of temperature gradients. Such misalignments are likely due to swinging of the rosette package due to variations in the descent rate of the CTD. Unstable features remain (mostly just below the mixed layer) that are likely instrumental rather than real but to remove them would require severe editing where there is no clear justification. Filtering salinity did not produce satisfactory results. When repeat casts were compared they generally both had spikes, but no systematic differences.

There were a number of problems with the thermosalinograph during the cruise. For many of the files the Julian time record was set to advance every 6 seconds but data were really only acquired every 30s; this was fixed by entering the correct interval in the headers and deriving time from start time and interval. The latitude and longitude channels had many bad values, many of which were single spikes that could be fixed by interpolation, but there were large sections in which the values got stuck, with absolutely no variation; for those files the positions were replaced with values from the ship tracking system, using a different GPS system. Where TSG and CTD casts coincide, the position data look fine except for those bad patches. Flow meter and intake temperature data are not available until file #21. The fluorescence channel had no signal until file #21. Many of the files are very short, some have so few records that it is impossible to assess whether the equipment was working properly, so they have not been archived.

There was excellent agreement between the loop salinity and the corresponding surface CTD rosette samples. This shows that there is no significant amount of freshwater getting into the loop. Furthermore, this result supports using the rosette data for recalibration of the Thermosalinograph salinity data; that is very useful since there are many more samples available from the rosette than from the loop. The extracted chlorophyll from the loop was lower than the CTD rosette CHL but there were few samples. The thermosalinograph intake temperature was very close to the CTD data from around 4.5m.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained as well as a summary from the Chief Scientist of problems and points of interest with reference to processing. There were a number of computer crashes and some unusual stops for boom pressure checks. The first cast was split into 2 files. A few casts were misnamed and positions are wrong for another. Cast #50 should be deleted since it was a NET cast. There were notes about occasional serious problems with primary channels.

The PAR sensor was used for casts 1-9, 14, 15, 30, 31, 35, 36, 48, 49, 59, 60. That channel will be converted for all casts, but will be removed at the end from casts with no signal.

Extracted chlorophyll, nutrients, Autosal salinity and Portasal salinity and DMS data were obtained in spreadsheet format.

The titrated dissolved oxygen files were provided in individual ADD files with a flag channel but no flags or comments were entered.

The cruise summary sheet was completed.

The history of the pressure sensor was obtained but the conductivity and DO sensors were recently recalibrated and have no relevant history.

The calibration constants were checked for all instruments. There were a number of problems:

- The calibration information for the pressure sensor was from an older calibration. The parameters were replaced with the most up-to-date ones. The offset entered, -1.5db, keeps getting used for this sensor, but it is wrong. We never see an offset <0 so it is hard to figure out where the number ever came from. It was changed to +0.4db.

- The transmissometer offset is wrong – it should be -1.1772, not -1.2737.

- The date was missing for the PAR calibration so that was added.

##### Tests for DO parameters

SBE DO sensor #1176 was tested during 2009-27 to see which value of E worked best at reducing hysteresis. A value of 3.85 was found to be better than the nominal values of 3.6. No tests have been done for the H1 and H3 parameters used in the TAU correction. And there has been a membrane replacement and sensor recalibration since the tests were done on E. So tests were re-done.

1. The first step was to investigate the value of E that minimizes hysteresis at depth.

1a. A deep cast, #29, was converted with TAU turned off and hysteresis correction selected. Values of E studied were 3.5, 3.6 and 3.85. Profiles were plotted to see if there is any obvious improvement in hysteresis but the differences were slight. The DO values at depth increase with increasing E value, but the significance of that is not clear since the values are known to be too low, in general. Comparing the DO and temperature vertical offsets between up and downcast (in areas where temperature is not offset much) the greatest difference is for E=3.85. The other two choices are very close, though E=3.5 may be slightly closer to the temperature offset. This method does not produce a clear answer as to which value is best.

1b. Next rosette files were converted for all deep casts (>2000db) with DO calibration samples (29, 37, 44, 61), using values of E between 3.5 and 3.85. DO concentration was included in the conversion. The files were then combined with titrated oxygen files and COMPARE was run for each set. Obvious outliers were excluded (3 for each run, one of which had been flagged by the analyst) and then a fit of differences against DO was made including R2 values. A plot was made of E vs R2 and the highest value is for E= 3.75. However, there is a lot of scatter especially in the shallow data. When values of DO>5ml/l were excluded the results were completely different with the best fits for lower E values. This method is too dependent on the choice of outliers.

1c. Finally, COMPARE plots were made in which the deep values (below the DO minimum) were excluded from the fit so they would show up as red. In this way it is possible to compare data at similar DO values above and below the DO minimum. Hysteresis would create an offset between them. For E=3.5 the deep CTD DO values fall below the shallow ones whereas with E=3.85 they are too high. The best fit is with E=3.65. This method looks most useful and should be used in future rather than the approaches of 1a and 1b.

2. Next, a few files were converted with hysteresis turned on (with E=3.65) but no TAU correction and parameters H1 and H3 were varied to achieve the best match between downcast and upcast data.

2a. H1 (-.02 to -0.05) A comparison of upcast and downcast data using different choices of H1 did not provide a simple answer – at one depth one choice looked better than at another. So COMPARE runs were done with the different choices using just cast #61. Doing runs of COMPARE and looking at the deep versus shallow (in the same way as for E) using a range of H1 values, the best results were between  -0.023 and -0.033, with the middle of that range probably best. In the past -0.028 proved best for this sensor, so will be used again.

2b. H3 (1200-2000)

One deep file (#61) was converted to see if varying the H3 parameters would reduce hysteresis. Looking at profile plots the choice made no notable difference. COMPARE plots show some small variations. The differences are so slight that it is hard to judge for this cast, but it does seem as if 1350s does a slightly better job of making deep data fall closer to shallow data in the COMPARE plots than does 1450s.

Configuration file 2009-10-ctd1.con was adjusted to make H1= -0.028 and H3= 1350s.

3. Finally, both hysteresis (with E, H1 and H3 as determined above) and the TAU correction were selected to see how well the TAU correction works and at what cost to the deep data noise level.  Tests were run on cast #61 to see if the TAU correction should be applied. The correction had good results in the high gradient zone in the top 300db, with a sharper gradient and upcast and downcast closer. At a depth of 4000db the noise level in the corrected data is <±0.005ml/l. In a previous study <±0.01ml/l was found after metre-averaging, so this looks good. The TAU correction also produces higher DO values at depth, but the differences are small and closer to the bottle values, so this does not look like a problem. The improvement in the high gradient zone looks sufficient to justify the noise at depth, particularly since scientists tend to rely on the titrated bottle data at depths below 1500db.

The history file for this sensor was edited to add this information since these values should be consistent from cruise to cruise until the next recalibration. I don’t know if recalibration will change these parameters, but repairs are likely to do so.

##### Conversion of Raw Data

Data were converted using configuration file 2009-10-ctd.con.

A few casts were examined and all expected channels are present.

- The two temperature channels are generally very close during the downcast, but there are some odd excursions in both. The upcast data are much noisier so there are significant differences. Conductivity is similar in spikiness, but the two channels are further apart than usual even during downcasts.

- The fluorescence looks ok with a dark value <0.1.

- Dissolved oxygen voltage has the usual offset between downcast and upcast.

- PAR look fine.

- There is no SPAR signal – this channel will be removed later.

- The transmissivity has no noticeable hysteresis; there are spikes at depth, but only small ones.

- The altimetry looks useful for some casts, but for most the CTD did not come near the bottom.

Rosette files were converted using a start time of -5s and duration of 10s. The TAU correction was used.

The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers, with output named *.BOT.

At this point a few corrections were made to headers and file names:

- For both the ROS and CNV files the names were changed from events #7 and 8 to #8 and 9, respectively.

- The ROS file name for cast #1/2 was changed from #2 to #1 to match the downcast CNV file. The files from the analysts were also changed to #1. This should ensure that no data are lost through mismatch of names.

- For cast #10 the NMEA time and positions were changed to match the log because NMEA was not logged for that cast.

- Cast #50 was deleted for both CNV and ROS files- there was only a little surface data in them.

Temperature and salinity were plotted for all BOT files and a few outliers were noted for casts 12 and 36. CTDEDIT was used to clean primary salinity for the first and to remove some points in primary temperature and salinity for the second. The output files were then copied to *.BOT. Editing details were added to the header comments.

There were some other casts with generally noisy data, but these were not a matter of just a few suspect records, so editing would not address the problems. Included were casts #47 and 49. There is mention in the chief scientist’s notes of problems with the primary channels for cast #49 and the full cast profile does show that the primary conductivity and oxygen channels are very bad. For cast #47 there is no obvious problem and the noise in the bottle file was only for one bottle. Clearly for cast #49 it will be necessary to use the secondary T and C. There are also problems likely with #29 and 67.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run comparing a variety of settings for CELLTM. The results were difficult to judge because the best setting varied from depth to depth and from cast to cast.

The best choice overall proved to be (α = 0.03, β=9) for both the primary and secondary. CELLTM was run on all casts using that setting.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

NOTE: The first time DERIVE was run using the wrong configuration file. This affected only the Oxygen:Dissolved:SBE channel. Later this step was repeated using the correct configuration file, followed by STRIP to remove all but the Pressure, Scan Number and Oxygen:Dissolved:SBE channels. Those files were then converted to IOS SHELL header format, and put through CLEAN and SHIFT (for dissolved oxygen only) and DELETE. The resulting files were later merged with the edited CTD data after removal of the original dissolved oxygen data. The merged files were named *.NEW.

This problem did not affect the bottle files for which the correct configuration file was used.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often extremely noisy so these are very rough estimates.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 16 | 1000 2000 | +0.001 VN +0.001 VN | -0.00005 N -0.00005 N | -0.002 XN -0.002 XN | High, V. noisy High, noisy |
| 38 | 1000 2000 | +0.001 +0.0009 | -0.0001 XN -0.00015 XN | -0.0025XN -0.0025XN | High, X Noisy High, X Noisy |
| 51 | 1000 2000 3000 3900 | +0.0013 VN +0.0010 N +0.0012 N +0.00012XN | -0.00011 N -0.00013 N -0.00012 N -0.00015 VN | -0.003 -0.003 -0.003 -0.003 | High, X Noisy High, X Noisy VHigh, Noisy VHigh, Noisy |
| 61 | 1000 2000 3000 3900 | +0.0011 N +0.0010 N +0.0012 N +0.0011 N | -0.00015 VN -0.00015 VN -0.00015 VN -0.00015 VN | -0.003 XN -0.003 XN -0.003 XN -0.003 XN | High, V Noisy High, V Noisy High, V Noisy High, V Noisy |

There is no suggestion of pressure dependence though it would be hard to see. The conductivity and salinity differences give very slight suggestions of temporal variation. While noisy, none of the differences are unusually large. The secondary conductivity has a lot of fine-scale noise.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. Cast #44 was too large to convert, so two channels were removed first: Descent Rate and Pump Status.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. No problems were found.

The cross-reference check was compared with the log book, and no errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run. The average surface pressure was 3.3db which is fairly typical for the Tully.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet and all casts with readings were checked. Plots were made and the log book was checked. The algorithm worked well where the CTD got close to the bottom, but it recorded erroneous low values for cast #41 when it did not get close to the bottom according to the log. The altimetry header was removed for that cast.

The water depths were also checked; it was missing for cast #41, so that was added based on the log book entry. A few others either differed significantly from the log book entry or were less than the maximum sampling depth. They were changed to match the log entry unless there was evidence from altimetry and maximum sampling depth to provide a reliable estimate. Casts affected were: 29, 37, 44, 55, 56 and 66. The same changes were made to the bottle files at the MRGCLN2 stage.

A few casts were examined at this point due to concerns about the primary pump:

- Cast #29 – the downcast looks ok. The upcast data are noisy towards the end; both sensor pairs are noisy though the primary may be a little worse.

- Cast #49 – Dissolved Oxygen, Primary Salinity are bad, both down and up, but the Fluorescence seems ok.

- Cast #67 – DO looks bad at least on way down. Seems ok up.  Primary salinity bad down, ok up. Fluorescence seems ok.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. A few problems were encountered:

- One bottle was fired during cast #48 for which no sample number was assigned. Niskin #13 was fired at 15db instead of 10db, so at the end of the cast the CTD was lowered to 10db to get that level. Sample #417 is assigned to Niskin #13 and Niskin #19 had no sample number assigned. For now, the sample from Niskin #19 will be labelled as #9417 and later the CTD data from the Niskin #19 line in the SAMAVG file will be copied to the line with chemistry samples #417 in the bottle file. The samples were taken from the real 10db bottle, so that will be the proper match. It seems less error-prone to fix it later.

- During cast #66 4 bottles were fired, so sample #573 was used for this cast but labels had already been prepared to use it for cast #67. The instance in cast #66 was renamed as #9573. There were no samples taken from that bottle.

After those corrections were made to the addsamp.csv file, it was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

Salinity analysis was done at IOS using Guildline Autosal #Model 8400B, serial # 68572 and at sea using Portasal #68627. Only cast #11 has data from both salinometers and cast #10 is the only one entirely analyzed on the Portasal. The data will later be combined – probably with flags on the Portasal.

AUTOSAL DATA

The Autosal data were delivered in spreadsheet format and included duplicates and loop samples in file “2009-10 Line P last updated160909.xls” together with the analysis log sheets. Loop data were copied to spreadsheet 2009-10-sal-loop.csv. Data for which there are both surface rosette samples and loop samples were copied to file 2009-10-sal-comp-ctd-loop.csv. The sample numbers were rearranged for cast #44 based on the chief scientist’s notes and the nutrient analyst’s assignment of sample numbers. See the OXYGEN section for details. The data make much more sense after this adjustment.

The duplicates were copied to a separate spreadsheet, 2009-10-salinity-duplicates.csv. The averages were calculated.

The full spreadsheet was saved as 2009-10-sal-auto.csv. The loop sample data were removed and the duplicates were replaced with the averages and “f” flags were inserted for those samples. Comments were found on the Autosal Analysis Log sheets and were added to the spreadsheet along with flags based on the comments:

- The duplicate samples #86 differed by 0.0044. Both were marked “loose lids” and the second had 4 readings. The first was much closer to the sample from the other Niskin bottle fired at the same level. The first reading was selected instead of the average, but was flagged “c”.

- There were many comments about loose lids and/or caps and those seemed especially prevalent among the duplicates. Others had no cap, just a lid. Flag “c” was attached for most of those comments, though for sample #140 since the second reading was very close to the first, it was felt that the loose lid on the first was probably not significant, so no flag was added.

- A number of samples were added to the spreadsheet since they were intended to be sampled but the bottles, while labelled, were empty. Flag “i” was attached. However, for the missing sample #65, there was a sample analyzed on the Portasal that may be added later if the quality of the Portasal data is deemed high enough.

The file was then converted to individual SAL files.

PORTASAL DATA

The Portasal data analyzed at sea were received in spreadsheet 2009-10P4 SAL.csv. This was saved as “2009-10-SAL-Port.csv”. Columns were added for event number and station name and those details were added based on a column that combined the two. Flag and comment channels were added. The flags were adjusted so that “d’ was entered for cases of readings differing by >0.001 and “c” for >0.0005. The data were averaged and unnecessary columns were removed.

The data were then converted to individual SALP files.

Since there are no overlapping data with the AUTOSAL and very little data from the PORTASAL, it does not seem likely that the PORTASAL data can be recalibrated to match the AUTOSAL. For now all the PORTASAL data have been flagged “c” with an explanatory note in the header. The file for cast #10 was then renamed as *.SAL. For event #11 most bottles were analyzed on the AUTOSAL but there was one analysed on the Portasal, so it was transferred to the AUTOSAL file, again with a note in the header and a “c” flag.

Salinity Studies

(i) One set of duplicates appeared to have either been mislabelled or mis-entered in the analyst’s spreadsheet. There is one sample named as #338 and two as #339. The rosette sheet shows the duplicate as coming from #339, but because of the general mix-up with cast #44 when lanyards were connected wrong this is probably sample #338 that has duplicates. There were 8 pairs of duplicates that were analyzed on the Autosal. Using all samples Sp=0.0019 where Sp is defined as:

Sp = Square Root (sum of squares of differences / 2*number of pairs)

When 2 outliers (differences >0.004) were removed Sp=0.0010. This shows reasonably good repeatability especially given that one of the outliers came from the confused cast #44. The standard deviation is 0.003 using all data and 0.0015 without the two outliers. For one of the outliers, sample #86, there is another sample from the same depth, but not the same Niskin bottle. It suggests that one of the duplicates is better than the other. Both samples had loose lids, and the 2nd appeared to be slightly unstable with 4 readings. (The 1st reading is closest to the CTD salinity when corrections are made based on COMPARE.)

(ii) There were no duplicate samples analyzed on the Portasal.

(iii) There were no cases for which there was a sample analyzed on each of the Autosal and the Portasal.

(iv) From cast #46 there were 12 bottles fired at 2000db and 12 at 500db, for cast #47 there were 12 at 1000db and during #55 there were 23 bottles fired at 2000db. When bottles are fired this rapidly using the pressure as reference means the CTD values are not independent of each other. So a separate run of COMPARE was done using just those casts and with sample number chosen as the reference channel. Standard deviations were determined for both the CTD and Bottle salinity for each set of repeat firings. (See 2009-10-sal-comp1-samp.xls.)

| Cast # | # of bottles | Pressure | Std Dev CTD SAL0 | Std Dev CTD SAL1 | Std Dev SAL:Bottle |
| --- | --- | --- | --- | --- | --- |
| Autosal |  |  |  |  |  |
| 46 | 12 | 2000 | .0003 | .0011 | .0024 (.0007*) |
| 46 | 12 | 500 | .0004 | .0012 | .0018 |
| 47 | 12 | 1000 | .0003 | .0011 | .0008 |
| 55 | 23 | 2000 | .0002 | .0013 | .0014 |
| Portasal |  |  |  |  |  |
| 10 | 10 | 1300 | .0004 | .0012 | .0008 |
| 10 | 12 | 500 | .0003 | .0012 | .0007 |

*Excluding 1 slight outlier for the 2000m data of cast #46 made a big improvement.

Excluding other bottles flagged “c” did not improve the standard deviations.

The standard deviation is low in the primary CTD salinity channels; it is higher in the secondary reflecting fine-scale noise in that channel. The bottle data show more variability, but when the one outlier is excluded for the first set, all are <0.002. It is noteworthy that the variability is lower from the Portasal used at sea than from the Autosal used at IOS.

(v) From cast #10 there were 12 salinity samples from 1300 and 12 from 500db which were analyzed on the Portasal. The results show the instrument was surprisingly stable during this period.

Conclusions

The repeatability looks very good from the CTD at depth, so the deep repeat bottles can be used to estimate that the errors from sampling and analysis are ~±0.002 for the Autosal and ~±0.001 for the Portasal. The duplicate study (Autosal only) suggests similar sampling/analysis errors.

DISSOLVED OXGYEN

Dissolved oxygen files (*.add) were provided, but there many problems with them.

- There were no flags or comments entered although the rosette sheet indicates some problems.

- Duplicates were not averaged.

- Some samples are missing that are found in the OXY files.

The OXY files were exported to a spreadsheet 2009-10-oxy.csv. There were problems with formatting in the OXY files that led to many data missing from the spreadsheet. The individual files were opened in EXCEL and the missing data copied into the s/s. Duplicates and repeat bottles at a single level were copied to file “2009-10_oxy_duplicates.xls”. The duplicates were averaged in the main spreadsheet and flag “f” attached. None of the duplicates differed by >10%.

The pooled standard deviation of pairs of samples (Sp) was calculated as

Sp = SQRT{sum (d*d)/2k}

where k = no. of pairs and d = difference between pairs. When all 7 duplicates were used Sp = 0.056 and when 1 pair were excluded Sp = 0.025. These results are not as good as those for 2009-09, but compare reasonably well with other cruises.

There were 10 cases of oxygen samples from each of 2 bottles fired at the same depth. When those were analyzed in the same way as the duplicates, a value of 0.061 was found for Sp. When the only near-surface pair were excluded, Sp=0.032 which does show good repeatability.

The main spreadsheet was edited by adding flag and comment columns. Details were added to those based on the rosette sheet comments. Some samples were added that were noted on the rosette sheet as having been empty, or where the Niskin had failed to close. For those pad values and “i” flags were assigned. Since the rosette sheet did show that sampling was intended, it is necessary to offer an explanation about what went wrong, or people may just think the samples were mislaid. There were a few problems:

- For cast 1/2 the data were not in the spreadsheet – probably because some event numbers had non-numeric elements. The event number was changed to 1 and the values were added based on the OXY file and the rosette sheet information.

- For cast 3, the file indicated cast #2, so that was fixed. Event number was changed to 1 to be consistent with other samples and CTD data files.

- For cast #20 the duplicate is said to be for sample #146 in the OXY file, but it is noted on the rosette sheet that it should really be #156. The value makes sense for #156, so the sample number was changed to #156 in the S/S.

- Cast #44, sample #337 – the DO value looks way out of line but it is noted in the nutrient file that all samples are off by one place due to lanyards being misarranged. So the sample identified as 337 is really the surface sample #360, and samples 338, 339 etc should be 337, 338 etc. The nutrients have already been rearranged in this way. So the DO data were also rearranged. SALINITY was adjusted in the same way. There was no CHL sampling for that cast.

The file was then converted to individual ADD files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2009-10nuts.xls which included a report on precisions. The file was simplified and saved as 2009-10-nuts.csv. Extraneous columns were removed and header names were changed to standard format. Data were sorted on sample number. File 2009-10-nuts.csv was then converted to individual NUT files. There were loop samples in the original file; those were saved separately as 2009-10nuts-loop.csv. Sample #146, cast #20 will be checked again later when files are merged because the nutrients analyst suggested there might be a mis-trip or mis-sample. The file was then converted to individual NUT files. (It was later discovered that the surface samples from casts #53, 54, 56, 57, 58, 64, 65 and 66 were missing from the file used for this step. They were in the loop file, so the information was added to 2009-10nuts-loop.csv and conversion to individual files was rerun.)

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file QF2009-10CHL.xls which included flags and comments and a report on precision; the duplicates had been averaged. Loop data were copied to 2009-10-chl-loop.csv. The file was edited to remove extraneous lines and columns, header names were changed to standard format, the file was sorted on sample number, and saved as 2009-10-chl.csv. There were loop samples in the main file; they were saved separately as 2009-10-CHL-loop.csv. The file was then converted to individual CHL files. (It was later discovered that the surface samples from casts #53, 54, 56, 57, 58, 64, 65 and 66 were missing from the file used for this step. They were in the loop file, so the information was added to 2009-10nuts-loop.csv and conversion to individual files was rerun.)

DMS

DMS data were obtained in file DMS 2009-10 summary.xls. The file was saved as 2009-10-dms.csv and edited. There were flag and comment columns, but no entries were made; the report 2009-10-DMS-report.doc. indicates that none were needed. All entries “<” were replaced with “0”; a note in the header will explain that the minimum detectable level is 0.1. Where there were duplicates they were averaged to get a single value for each bottle and an “f” flag was added. Headers were changed to standard format and unnecessary columns were removed. The file was then converted to individual DMS files.

The SAL, CHL, ADD, NUT and DMS files were merged with CST files in five steps. After the 5th step the files were put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output: MRG).

The altimeter readings from the headers of the BOT files were exported to a spreadsheet and all casts with readings were checked. Plots were made and the log book was checked. The algorithm worked well where the CTD got close to the bottom, but it recorded erroneous low values for many casts when it did not get close to the bottom according to the log. The altimetry headers were removed from the BOT files for casts #16, 54, 57 and 65 because the algorithm didn’t work well.

The water depths were corrected in the same way as for the full cast files as described in section 10.

##### Compare

Salinity

Compare was run early using all data. When the samples analyzed on the Portasal at sea plus differences >0.02 were excluded, the average difference was 0.0029 for the primary and 0.0003 for the secondary. When data from the top 200db are also excluded the differences are 0.0035 and 0.0011 with the CTD reading high in each case. Large outliers (differences >0.01) were:

- Sample #146 had already been flagged “c” due to a loose lid and the sign of the difference does suggest that the sample might have been affected by evaporation. The flag was changed to “d”.

- Two samples from near the surface of cast #29. The salinity gradient is low for the downcast, but the upcast CTD salinity is very noisy near the surface, so there is no evidence that the bottle value is bad. No flags will be added.

- Sample #7 from cast #1 is also in an area of noisy CTD data, so no flag will be added.

- Sample #40 from cast #9 is associated with extremely noisy CTD data, so no flag will be added.

The secondary differences are even smaller at depth (11 samples below 3000db differ by an average of 0.0001), so if this channel is chosen for archiving, no recalibration will be applied.

There is a slight suggestion of time dependence in the primary channel but the fit is very noisy; the secondary looks flat with time, but again there is so much noise in the fit that this is not very significant.

Next the samples analyzed on the Portasal were studied. The samples are all below 500db. The average differences for those samples are 0.003 for the primary and 0.001 for the secondary, so that is not significantly different from the Autosal, though the evidence is not strong. The flags will be left as “c” on the Portasal data. (For more detail see 2009-10-sal-comp1.xls.)

Dissolved Oxygen –

COMPARE was run for Dissolved Oxygen. The results are very odd. First, some severe outliers were examined:

- Cast #1 – samples from 25 and 49db are outliers, not severe, but in the opposite direction from most outliers. No flags were added at this point since it is likely the CTD that is odd.

- Cast #3 – 4 samples between 11 and 51db are outliers, again not severe but in the opposite direction from most. No flags were added at this point since it is likely the CTD that is odd.

- Cast #20, Sample 144, ~800db – Significant outlier. Bottle already flagged “d”. Add comment about COMPARE.

- Cast #29 –Two surface bottles were severe outliers. The CTD data look bad. No flags justified. All channels using the primary pump were replaced with pad values in the SAMAVG file. (DO concentration, DO Voltage, Fluorescence, Primary Conductivity, Primary Salinity)

- Cast #44, Sample #346, ~800db – Significant outlier. Bottle already flagged “c”, change to “d” and add comment.

There is a clear temporal drift. The first 2 casts are very odd. The next contains too much scatter among its few points for a fit to be estimated. Thereafter trendlines were found in the fit of differences between bottles and CTD versus CTD DO concentration, and the % correction implied by those fits was calculated.

| Cast # | Slope of fit | Max Press | % Correction |
| --- | --- | --- | --- |
| 12 | 1.0405 | 1300 | 4.4 |
| 20 | 1.0421 | 2000 | 4.8 |
| 29 | 1.0482 | 3275 | 5.3 |
| 37 | 1.0622 | 3500 | 6.5 |
| 44 | 1.0715 | 4315 | 7.5 |
| 61 | 1.0712 | 4000 | 8.0 |
| 71 | 1.0807 | 2000 | 8.4 |

So clearly the fit was changing with time. However, the fits do not all look reasonable without an offset and identifying outliers is tricky.

It is notable that the first two casts looked so odd. This could be a result of anoxic sampling, but that has not been a problem in other recent uses of this type of sensor and even in the past the effects wore off after a few casts with no anoxic sampling. If that is the explanation then it suggests there is something wrong with the sensor despite its having been repaired recently. Another possibility is that there was a pump problem, but that seems unlikely to cause a steady drift and the salinity data does not support that theory. The sensor was used for 2009-53 before this cruise, but those data had not been processed yet. (Note that the SBE DO sensor was returned to the factory after this study was done, and no damage was found.)

A fit was done excluding points below 1500db as check for hysteresis. The red points are below the DO minimum and the green above. Only one red point looks out of line and even that is not far off. When we exclude the outliers noted above, the results look satisfactory. Another plot was done with points from P>3000db in red and they cluster either side of green points, so there is no need to exclude that data from the fit.

A number of tests were done in COMPARE to develop a scheme for recalibration the DO data. No simple method emerged. The variations in slope did not seem to have a simple dependence on event number, sequential CTD deployment # nor cumulative pressure sampled.

The first approach was to do a fit for each of the casts with titrated DO samples. This is not as easy as it sounds since the trendlines often do not seem representative of data from the whole range. There is a bias in sampling to surface samples and the CTD DO sensor does not work particularly well close to the surface. By selection of outliers and forcing of the offset, the best trendline possible was found for each of the casts as follows where DO corrected = Slope * DO original + offset:

| Event # | Slope | Offset | Comment |
| --- | --- | --- | --- |
| 1 | 1 | -0.5 | Very little data; huge scatter |
| 3 | 0.953 | -0.05 | Very little data; huge scatter |
| 6 | 1.042 | -0.01 | Very little data; lot of scatter |
| 12 | 1.0475 | +0.01 |  |
| 20 | 1.051 | +0.01 |  |
| 29 | 1.0532 | +0.02 |  |
| 37 | 1.0658 | +0.01 | Not as linear a fit as many others |
| 44 | 1.0727 | +0.01 |  |
| 61 | 1.0803 | -0.01 |  |
| 71 | 1.0792 | -0.01 |  |

The only reasonable thing to do with this data is to apply a variable correction with cast #. The results were applied to a band around each of the calibration casts – interpolation seemed impossible given the variable offsets.

A number of attempts were made at recalibration using constant offsets and gradually increasing slopes. The results showed some time dependence in the fits, even when the input data were restricted to a variety of levels of DO ranges. This makes it clear that the time-dependence is not due to deeper mixed layers in some areas or different DO ranges. There is a calibration drift. The complications may well be pressure-related if there is a membrane problem. There is limit to what can be accomplished with this approach. (See 2009-10-dox-comp1.xls and DO-cal-study.xls.)

COMPARE was rerun after recalibrating the SAM files using the scheme developed and the results are in file 2009-10-dox-comp2.xls. The results are as good as can be expected from this method with little time-dependence and an average difference of 0.002 when outliers are removed.

Fluorescence

COMPARE was run using the CTD Fluorescence and the Titrated Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD FL, the FL/CHL ratio versus event # and differences versus Fluorescence and pressure. There is a distinct difference between the near-shore casts and those offshore. The range of chlorophyll is very low offshore, with a maximum of 0.44ug/l west of station P5 and most values much lower than that; for those casts the CTD Fluorescence averages about 2.4 times the extracted chlorophyll. For casts #1 -16 the fluorescence and extracted chlorophyll are quite close in value. (See 2009-10-chl-fluor-comp.xls.)

After COMPARE final checks were then made of files that need special attention. Among the checks were plots of DO against salinity and, in one case, nutrients against salinity as well.

- 1 / 2 – 2009-10-0001.mrg contains upcast CTD plus all sampling. Looks fine.

- 3 – Everything looks fine – this was a test of table driven firing.

- 15 – Table-driven cast. The data look ok, but it is impossible to be sure since the rosette log sheets have no record of Niskin #, target pressures or sampling done. The files from the Nutrient and CHL analysts use Niskin Number in the header, but this appears to actually be the firing number.

- 20 – The analyst suggests a possible mis-fire for sample #146. This is suggested by both the salinity and the silicate samples, but the others do not really look out of line. However, when plots were made of each of the DO, Nitrate, Silicate and Phosphate against Salinity:CTD and Salinity:Bottle, all look better plotted against the Salinity:Bottle than against Salinity:CTD. This shows that all the samples probably came from the same depth, but not from 400db; so it was indeed a misfire. There is one DO sample that looks bad in both plots, indicating that while it is a bad bottle, there is a different reason than closing at the wrong depth.

- 20 - Sample #144 was changed from “d” to “e” after examination of plots against salinity – severe outlier.

- 29 – As expected DO and primary salinity look bad at end of cast – The affected channels were removed from the SAMAVG file, but at this point a note was put into the header to explain why. It was first assumed that fluorescence would be bad too, but it seems fine, so was left in. Later more evidence was found to suggest that the fluorometer was on the secondary pump, not the primary.

- 44 – The data appear to be in the right order, with the top and bottom pressures associated with appropriate salinity, DO and nutrients.

- 44 - Sample #346 flag was changed to “e” since severe outlier in COMPARE and in plots against salinity.

- 48 - At this point cast #48 was rearranged as per notes from the chief scientist, so that the CTD data from 10m are substituted for those at 15m and the final line is then removed from the file. (Niskin #19 CTD data being entered where Niskin #13 CTD data were originally.)

- 66 – Sample numbers are as required.

All MRG files were put through CLEAN to remove Sea-Bird headers and comments from the secondary files.

Data were exported to spreadsheet 2009-10-bottles.xls and compared to the rosette sheets to ensure all expected data are present. A few problems were found:

- Some surface nutrients and chlorophyll had not been copied to the main page of the spreadsheet. The data were found in the loop page and added to the simplified spreadsheets. The analyst was notified.

- A few nutrient and chlorophyll samples were not shown on the rosette sheet, but checks of the s/s from analyst confirms they are correctly labelled. This included a few extra duplicate samples.

- Some salinity samples noted on the rosette sheet were not found. One was for cast #3 – a second rosette log sheet for that cast does not show salinity sampling, so probably no samples were taken. Another was a UBC cast and probably due to Portasal failure they were not gathered or not analyzed.

- Some NH4 sampling was done, but no data was found. It is assumed this was UBC sampling and not available.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The usual shift of 1s was found to be appropriate though the noisy descent rate makes the judgment a very rough one.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the primary conductivity for cast #34 and a setting of +1.0s looks best.

Tests were run on the secondary conductivity sensors for 3 casts using a variety of shifts. The best choice varies from feature to feature and from cast to cast. A setting of +0.5s worked best overall.

SHIFT was run on the primary conductivity for cast #34 only using +1.0s.

SHIFT was run on the secondary conductivity using +0.5s for all casts.

Dissolved Oxygen

Tests were run on a few casts for each sensor to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. A value of +90 seemed best.

SHIFT was run using +90 records for all casts.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #2 which was an upcast only, so will not be processed further.

##### DETAILED EDITING

The secondary temperature and salinity channels were selected for editing. They are closer to the calibration bottles and there were problems with the primary pump for some channels. There is fine-scale noise in the secondary conductivity, but the effect of that on salinity should be minimized in metre-averaging. However, there is also some larger-scale noise in salinity in both channels which is presumed to be due to variations in alignment of conductivity and temperature sensors. Tests at filtering were not satisfactory, so editing of salinity will be more extensive than usual. This noise is most troublesome at the base of the mixed layer in the high temperature gradient where alignment is more critical. It seems somewhat less after the change to the aft deck presumably because the descent rate became somewhat steadier.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. Editing was used to remove spikes where they are systematic in direction and/or likely to affect the metre-averaged results. Records were removed that were corrupted by shed wakes or near the surface before the pumps were turned on. For many casts the descent rate was extremely noisy so many records were removed. The CTD was frequently seen to go upwards by up to 2m/s (and this is a metre- averaged speed) during cast #38.

All casts required some editing. A few particular problems were noted:

- Cast #1 had pumps off until 11db, so data were removed from the top 11.3db.

- Cast #34 – secondary temperature, conductivity and salinity data are bad from about 50-350db. Comparisons were made with nearby casts to see if the primary channels were affected. This is complicated by the fact that cast #33 sampled an eddy, but the primary data from #34 looks similar to cast #35 and 36 and the data fell within the historic ranges, so it is assumed that the primary channels are ok. For this cast, primary channels were selected for archiving. This will not affect the bottle file since the secondary channels look fine for the upcast.

- Casts #34-38 – The CTD descent rate was extremely noisy, with #38 being especially so. See the plot at the end of the report to get an idea of how noisy. The descent rate exceeded -2m/s often; that is a value averaged over 2m.

At this point the EDT files were merged with the DEL2 files prepared with just scan number, pressure and revised Oxygen:Dissolved:SBE channels. The scan number was used as the merging channel. For the EDT files all channels were selected except: Oxygen:Dissolved:SBE, Descent Rate, Status:Pump, Altimeter and Flag. The output files were named *.NEW.

##### Initial Recalibration

The SAM files were recalibrated using file 2009-10-recal.ccf which includes a range of DO corrections for different casts as discussed earlier. No correction was applied to salinity except for cast #47 because the secondary salinity is very close to the bottles and does not require correction. For cast #47 the primary salinity was recalibrated by subtracting 0.0035.

COMPARE was run using the recalibrated dissolved oxygen data and again showed the recalibration was applied properly. (See 2009-10-dox-comp2.xls.)

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction can be applied to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When a few extreme outliers are excluded based on residuals, the average difference is +0.05ml/l with a range of 0.01 to 0.07. In most recent use of this sensor type a 2nd DO recalibration has not been needed when the new DO algorithm is used, but for this data it looks worth applying, another hint of trouble with this sensor. (See 2009-10-dox-comp3.xls.)

File 2009-10-recal2.ccf was prepared to apply the following correction to the SBE DO for all downcast data:

Corrected CTD DO = 0.09896 * CTD DO – 0.0169

First this was applied to the thinned files and COMPARE was rerun. The result proved satisfactory, so it was then run on all the downcast files.

Dissolved oxygen concentration data was examined at depth to see how big an effect the TAU correction has on the noise level. The noise level was ±0.005mL/L at 1800db and 3900db. This is only slightly higher than found when the TAU correction was turned off. This may indicate that the noise in the instrument itself was higher for this cruise, or perhaps the parameters were better chosen than in the past.

##### Special Fluorometer Processing

An examination of the fluorescence channel shows a dark value of ~0.095mg/m3 .

The COR1 files were clipped to 150db and processed separately for A. Peña. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective. However, there do seem to be some problems with the fluorescence in some casts – possibly related to the pump problems. All these files will be examined later.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing appeared to be necessary.

On-screen plots were checked to see if there were any problems and the following were noted:

Cast #24 – The fluorescence has patches of zero values that look unbelievable. No other data look unusual.

Cast #26 – Like #24 - fluorescence needs editing.

Cast #34 – Fluorescence unbelievable and smooth – same cast had problems with the secondary conductivity. But according to the log book the fluorometer was on the primary pump. Should remove the FL channel.

The fluorescence data look bad for the cast with problems with the secondary pump, while they look fine for casts #49 and #67 that had problems with the primary pump. So it looks as thought the fluorometer was on the secondary pump, not as indicated in the log book.

T-S plots were examined and it was decided to do a little more editing to casts #5, 58, 59 and 67. The NEW files were edited and the recalibration and bin averaging steps repeated.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

The conductivity sensors were used in July for 2009-53 but that has not been processed yet. For 2009-10 which immediately preceded this cruise the primary salinity was low by ~0.02 and the secondary by ~0.025. There was a lot of scatter and great doubt about the results.

2. Dissolved Oxygen – The sensor was used for 2009-53 which has not been processed yet.

3. Pressure – The sensor is an older one but the calibration is not drifting. An offset of +0.4db has been used since 2004.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. All salinity data fell within the local climatology. Temperatures were above the historic maxima between 900 and 1100db for the 3 deep casts at P4. Deep temperatures were very close to the historic maximum for stations P3 and P5 – P7. This is presumed to reflect real conditions rather than evidence of problems with the CTD.

Repeat Casts – There were many repeat casts. A few cases were examined in detail:

- The T-S plots for 3 casts at P4 showed a large change between the first two on the outward run and the cast done at the end of the cruise. Temperature, salinity and fluorescence had all increased markedly (fluorescence went off scale), and dissolved oxygen concentration had decreased along lines of constant σt, down to σt = 26.Below that level the changes in temperature and salinity are of the opposite sign, but not large, and there is no significant fluorescence signal. There are hints in the salinity of interleaving below 26. The temperature profile does not have the local minimum seen earlier. The chief scientist observed that the water was brown and full of euphausids during cast #75, so this is assumed to be a real change, and not a symptom of CTD trouble. The Thermosalinograph fluorescence went off-scale during that time.

- There were 2 casts on the way out at P16 and another 2 on the return trip. The main thermocline was deeper for the later pair, and all showed very noisy salinity below the mixed layer though the earlier pair was spikiest. The temperature is smoother on the return, probably because the two earlier casts were run from the mid-ship station during very rough conditions. This confirms that the salinity problems were not caused by the use of the mid-ship station, but they were exaggerated by the noisy temperature data from those earlier casts. Between the surface and 600m the salinity is lower on the return trip. The mixed layer is shallower for the later casts. When examined in T-S space the differences in temperature and salinity between the two casts at depth are on the order of 0.003C˚ and 0.0003 along lines of constant σt,, better than for most other cruises. However, we normally see a consistent direction in the differences whereas the traces for these two casts intertwine and are noisier than usual

- There were many casts at P26 with 4 that went below 500db. At about 650db they show slightly higher temperature variations than seen in the above (deeper) comparison and there are quite large spikes in salinity associated with small temperature steps. The depth of isotherms varies by about 50db among the casts.

- Because an eddy was noted at P8 casts 20 and 74 were compared to each other and to casts at P7 and P9 in profile and T-S plots. The two P8 casts are much closer to each other than either is to P7 or P9. In the top 100db the P8 casts lie between the P7 and P9 casts on a T-S plot. Below 100m the P8 casts look notably warmer and saltier than at the other two stations.

- The eddy noted at P14 was examined in the same way. Near the surface the P14 casts lay between the P13 and P15 casts on a T-S surface, but below 100db they were warmer and fresher than at those two stations.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

Some of the channels listed below were actually removed earlier when the EDT files were merged with the NEW files, but they are listed here so there is a complete list.

The following channels were removed from all casts except cast #34: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR:Reference, Status:Pump, Descent_Rate and Flag.

The following channels were removed from cast #34: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR:Reference, Status:Pump, Descent_Rate and Flag and Fluorescence:Seapoint.

The PAR channel was removed from all casts except #1-10, 30, 31, 39, 40, 42, 55, and 56 because the instrument was not mounted on the CTD for those casts.

The Oxygen:Dissolved:SBE was removed from casts #49 and #67.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

There were many problems with CTD pumps and extreme variations in the descent

rate of the CTD. Dissolved oxygen and fluorescence channels were removed from

some casts due to pump problems and more data than usual were lost due to

corruption from shed wakes. Editing of salinity was more extensive than usual

from the base of the mixed layer to about 150db. There remain some unstable

T-S features and the quality of the data are considered lower than usual.

The dissolved oxygen calibration drifted through the cruise so that the comparison

of CTD and bottles did not lead to a simple recalibration scheme. Recalibration

of the data was done in a very subjective fashion, using 10 different fits based

on the calibration data available. Thus the quality of the data is not as high

as usual.

The precision of the SBE dissolved oxygen channel is difficult to estimate

because the comparison with bottles was very noisy, but roughly, the DO should

be considered:

•	±0.5ml/l from       0- 300db

•	±0.15ml/l from 300- 500db

•	±0.06ml/l from 500-2000db

•	±0.1ml/l below 2000db (Below 2000db DO data are mostly low by about 0.06mL/L.)

For details on the processing see processing report: 2009-10-proc.doc

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were all between 90% and 110% except for casts #9 at P3 and #75 at P4 which were at ~125%. There were no bottle values for either of the casts with higher saturation. For cast #9 the fluorescence was much higher than at nearby stations. Conditions were reported as being very unusual for cast #75 with brown water and many euphausids observed and transmissivity values ~0%/m and fluorescence off-scale at 15ug/l near the surface. While dissolved oxygen is high at the surface, it is lower between 20 and 100db than for earlier casts at P4 and then higher again below that.

At this stage it was discovered that there was a problem in the Dissolved Oxygen titrations, so that the recalibrated SBE DO data are high by about 2.5% because of one problem and a little low because of another. See section 24 for details on how the SBE data were corrected.

##### Final Bottle Files – part 1

The MRGCOR1 files were put through SORT to order on increasing pressure.

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR:Reference, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from all casts except #1-10, 30, 31, 39, 40, 42, 55, and 56 because the instrument was not mounted on the CTD for those casts.

The Oxygen:Dissolved:SBE was removed from casts #49.

For cast #29 the channels on the primary pump were replaced earlier with pad values for the last two bottles.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Finally cast #48 was edited. As discussed earlier the sample labelled as #417 did not come from Niskin #13 as indicated on the rosette log sheet. It really came from Niskin #19. The firing of Niskin #13 was at the wrong level, so a final bottle was fired to gather 10m samples. The CTD data from Niskin #19 were then substituted for the CTD data of Niskin #13. The samples and CTD data come from the same       firing. The line for Niskin #19 was then removed to avoid confusion. The edited file was put through CLEAN to fix the headers.

The bottle data were exported to a spreadsheet to check that all data are present. The only problems noted were already discussed in section 13.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

At this point it was discovered that there was an error in the DO titration results. See the next section for details.

##### Final Bottle Files – part 2

There is evidence from this cruise of problems with both the CTD DO and titrated DO:

1. Nina Nemcek discovered a small software error in determining the endpoint, but also had great concern about the standardization values for this cruise. The thiosulfate titer (the 0.709ml) value was low and would generally be considered out of range. Nina had no way of knowing if this was a real result due to a real change in the thiosulfate used on this cruise or if the standard or standardization was bad. Nina was unable to find the actual blanks or standards files on the computers used to run AutoOxy for this cruise. This means that either blanks or standards were not run, in which case she doesn't know where these values came from, or the files were deleted. The analyst is sure they were run, so we are assuming that the files were accidentally erased, leaving a bad standard as the likely explanation for the low value.  Nina estimates that if the standard was bad, then the DO values may be all overestimated by up to ~0.15ml/l (with higher values more affected than lower ones.)

2. The original comparison of CTD and titrated DO was most unusual with the first two casts having a positive slope in the fit of differences against DO and huge scatter in the fit. The slopes of fits for individual casts showed a steady increase. This suggested a serious problem with the sensor membrane.

3. The sensor was sent to the factory and they indicate no problem was found and the calibration drift was not unusually high.

Tests Done

1. The bottle values were adjusted by multiplying by 0.975 to 0.98 and COMPARE was rerun. Using 0.975 * Titrated DO led to an estimate of drift that is similar to drift estimate made by checking the plot sent from the factory. That is a very rough estimate, but is consistent with an error in titrated DO of 0.15ml/l at DO=6. (see 2009-10-dox-comp-correction975.xls and 2009-10-dox-comp-correction980.xls.)

2. While the first test might suggest that the problems in COMPARE were all due to the bottle titrations, it does not fix the time-dependent drift. Further tests were done in COMPARE to try to isolate where the time-dependence comes from and it appears to be entirely due to data in the top 250m. The gradients are highly variable and we know the sensor is limited in how well it can cope with large T and DO gradients. Some casts had a deep DO maximum (50db or more) and others had no subsurface maximum. The deeper casts have a deep minimum. Conditions were more variable than noted in the June Line P data and conditions at P4 were very different when it was visited on the return trip. So the variable fits are probably just a reflection of the CTD DO sensor having difficulty coping with changes of slope in the DO gradients offshore, so that the calibration fits look different. The original recalibration is probably reasonable.

Conclusion

It is not justified to just enter the result we expect for the standard without further information. However, the post-cruise calibration does support the conclusion that the titrated values should be rerun assuming a titer of 0.721. The titrated values will all be flagged “d” because this is an estimate. An explanation was entered into the headers in the Analysis Methods section. The endpoint corrections were made before the recalculation. These corrections do not explain the problems with temporal variation, so the earlier conclusion is probably correct: this is largely due to unusual conditions and possibly some problems with equipment early in the cruise.

Correction of DO data

The titrated DO values were fixed in the CHE files using a text editor; the new values were exported to a spreadsheet and added to the analyst’s spreadsheet to check all were entered correctly. A few errors were found and the process was repeated until all values were correct.

##### 24. Final Calibration of SBE DO Data in CHE and CTD files

Finally the calibration of the SBE DO data had to be revisited because not only was the bottle data adjusted to correct for a bad standard, but also there were corrections to fix the endpoints which were misjudged slightly because of a software problem.

We have a rough estimate that the CTD data are too low by about 2.5% due to the bad standard. They endpoints were also corrected, so it was decided to rerun COMPARE using the new bottle values and the recalibrated SBE DO data from the rosette files. Because the samples from anoxic waters now have zero values from the SBE DO, the fit was forced through the origin. The fit found that the bottles = 0.9853 * SBE DO. This is slightly lower than 2.5% correction that was estimated for the bad standard, but that makes sense given that it also includes an error of the opposite sign in judging endpoints and that the estimate was rough.  CALIBRATE was used to multiply both SBE DO channels by 0.9853 and COMPARE was run again. The results look good, with an average difference of ~0.0001 ml/l There is a lot of scatter and the highest values (>6.5 ml/l) are mostly too high, middle values are a little low, and the low values look excellent. (See 2009-10-dox-comp-after-fix.xls and 2009-10-dox-comp-final.xls)

The same correction was applied to the SBE DO channels in the CTD files. A note was added to the headers to explain why.

##### 25. Thermosalinograph Data

Data were provided in 23 hex files and 8 Excel spreadsheets. There were 4 loop bottles in Juan de Fuca and 4 during Line P (during CTD casts #22, 54, 56 and 65).

Files of nutrients, chlorophyll and salinity loop samples were prepared (subsets of the files from analysts.) These were combined in file 2009-10-comp-ctd-loop.xls in which the extracted chlorophyll and salinity samples from the loop were compared to the surface samples from the CTD rosette.

There was a problem with the time channel for the first group of files. While the system was acquiring data every 30s, the configuration was set so the time advanced by only 6 seconds. This must be adjusted.

a.) Checking calibrations

The calibrations were checked and the only problems were for the fluorometer which had the wrong date and small errors in the parameters. Those were fixed and the file was saved as 2009-10-tsg1.con. At cast #20 the secondary temperature sensor was changed, so after corrections to the fluorometer parameters cast 2009-10-0020.con was saved as 2009-10-tsg2.con. And then the flow meter was added for cast #21 so that was saved as 2009-10-tsg3.con, again after corrections.

b.) The files were converted to CNV files using the configuration files mentioned above. They were then converted to IOS HEADER format.

It was noted by the chief scientist the time interval for the first 9 files was wrong. She prepared XLS files with the original Time:Julian channel, but corrected Time/Date channel. Saving this as a CSV channel enables conversion to IOS files, but there are problems in converting the Time/Date combined channel to IOS format. A different approach is to fix the time interval in the headers of the CNV files, so that the start time plus interval can be used to derive the date and time. The interval is wrong in files 1 to 19.

After fixing those, all the files were put through ADD TIME CHANNEL to add the date and time. Since the Time:Julian is wrong it should be removed for casts 1-19.

CLEAN was run to add End times and Longitude and Latitude minima and maxima to the headers.

Time-series plots were produced. There is no fluorescence signal for casts 1-19 and neither flow rate or intake temperature channel for casts 1-20. There are a few distinct spikes in salinity and fluorescence. The secondary temperature data look bad at the beginning of file 21, but soon switches to believable values. The flow rate has a few drop-outs, the longest lasting 4.5 minutes, and a change in rate at one point, but there does not seem to be much effect on the record even allowing for the 4 to 5 minutes that water takes to get from the intake to the lab. It is possible that the flow was ok and just acquisition of the flow channel was lost. No editing will be applied based on the drop-outs.

The following files have fewer than 10 records each. This is insufficient for assessing the quality given that the flow may not have been properly established, and given acquisition was stopped so quickly something may well have gone wrong. These files will not be processed further: 8, 10-12, 14 and 20. Files #4, 6-7, 13-15 and 18 are will probably be archived if they look ok in the editing checks.

The first attempt at a track plot shows there must be problems with latitude and longitude. When individual plots were produced the problems were found to be in files 1 to 4 plus 9 and 13. Bad positions appear randomly distributed and are associated with spikes in the fluorescence channel, which otherwise has no signal. The spikes do not appear to affect the temperature and salinity channels. For file #9 there was a large section in which the latitude and longitude were stuck followed by sudden large changes, so those values were replaced with pad values. CTDEDIT was used to replace all other bad positions in the IOS files with interpolated values. The edited files were copied to the IOS files. The ADD TIME CHANNEL and CLEAN routines were rerun and the track plot was ok. There are gaps, as expected. The plot was added to the end of this report. (Note that it was later found that positions were also bad for parts of files #17, 18, 21 and 22 so editing was applied to those –see the next section.)

c.)  Checking Time Channel

The CTD data, after editing, but before metre-averaging, were thinned to reduce the files to a single point at or within 0.3db of 4.5db and exported to a spreadsheet which was saved as 2009-10-ctd-tsg-comp.xls. There were many casts with no data that shallow, so the routine was rerun using a wider tolerance. Since the surface layer is quite well-mixed, the deeper values may be useful.

The 8 TSG files that coincide with CTD casts were opened in EXCEL, median and standard deviations (over 5 records) were calculated for temperature and salinity (and fluorescence where available), and the file was then reduced to the times when CTDs were run. In some cases there were no TSG data available and, for a few, the positions were not available, though the other data were.

Those files were added to the CTD data in file 2009-10-ctd-tsg-comp.xls. There were 51 matches including the one with no position available. The positions were compared. Usually when this test is done the CTD and TSG positions are very close, and for this cruise that was true for 34 matches which had average differences for latitude and longitude <0.0001º and all differences <0.0004º. So the TSG clock is working well most of the time. However, for 16 matches the differences are large and it soon became obvious why. There are patches of constant values for latitude and longitude like the ones seen earlier in file #9, but these all arise at the end of files. While differences might be small during a stop, they would never be exactly the same for so long. All files were examined and this problem was found in files #17, 18, 21 and 22. Some of those files had no matches in this comparison. CTDEDIT was used to replace those positions values with pad values, and ADD TIME CHANNEL and CLEAN were rerun.

Since the comparison is so good for the files that have reasonable positions, it is assumed that the matches are ok for comparing other variables even when the positions are not. Since the matches are only poor when the latitude and longitude are clearly bad, it is clear that the clock was working well.

This spreadsheet will also be used in step (e) to compare temperature, salinity and fluorescence.

d.) Comparison of T, S and Fl from TSG and CTD data

- T1 vs T2 The intake thermistor was connected for the last 3 files. The average difference was 0.152 for file #21 and 0.131 for file #22. A fit of differences against intake temperature for file #21 (excluding some records from the beginning and the end due to data that look bad or very noisy) leads to the following fit:

Intake temperature = 1.0119 * Lab Temp -0.3084

For file #22 there was a lot of noise in the fit with the following results if some noisy sections are excluded:

Intake temperature = 1.0114 * Lab Temp -0.3002

Doing fits including quiet parts from both those casts gives the result:

Intake temperature = 1.0119 * Lab Temp -0.3038

Even with this subset of data the scatter is huge. The fits are rough, but since there is no intake temperature available for files #1-19, it is good to have some method for calculation of a proxy for that variable. File #23 was quite short and extremely noisy.

These fits imply ship heating by ~0.16 Cº at 12º C and ~0.11 Cº at 16º C. While extrapolation of these does not seem wise, doing so does suggest there would be zero heating when the intake temperature is ~26ºC, which might be considered a very crude estimate of the average ambient  temperature of the loop. Complicating any analysis is the fact that the flow rate does vary and is not available for most of the files.

- TSG vs CTD The spreadsheets comparing CTD and TSG files were then examined to find the differences between the salinity, fluorescence and temperature channels for the CTD and the TSG. There were 51 casts that could be used.

The intake temperature looks bad when it first comes on in file #21, and that includes the comparison with cast #48, so that was excluded from the comparison spreadsheet leaving 21 points of comparison. The average difference of those points shows the TSG intake temperature to be high by 0.01 Cº. However, there are two obvious outliers and they both come from casts with no sampling between 4 and 4.5db – one is from 5db and the other from 6.1db. When those are excluded the average shows the TSG intake to be high by an average of 0.004 Cº. This is as close as we could ever expect. Using the same 19 points the TSG Lab Temperature is higher than the intake temperature by an average of 0.147 Cº but there is obvious temperature dependence in the differences with higher differences for lower temperatures. When a few outliers are excluded the fit is

Temp (CTD) = 1.0082 * Lab TEMP (TSG) -0.2725

This suggests heating by 0.17 Cº at 12 ºC and 0.14 Cº at 16 ºC. This is higher by 0.02 Cº than found by studying the TSG differences, but 0.01 Cº is accounted for above, so this is not a large difference.

The TSG Lab Temperature is higher than the CTD by 0.157 Cº when all data except 4 outliers are excluded. Those outliers are from the first 3 casts (5, 8 and 9) and the last one (75) which is not surprising given that the near-surface gradients are fairly high for those casts, so a poor match in depth would have bad results and again there is some temperature dependence.

The results of the previous section should enable recalibration of the lab temperature to produce a proxy for intake temperature for those casts on which the thermistor data were not recorded.

The TSG salinity is low by an average and a median of 0.020 when 5 outliers are excluded. There is no obvious dependence on salinity, nor on the lab temperature but both comparisons are too noisy to rule anything out. Since the temperature error is temperature dependent it is likely that salinity error is too. A rough estimate suggests an error of 0.10 to 0.15 due to the effect of using the lab temperature, leaving calibration drift to explain 0.05 to 0.10. A better correction might be possible by first recalibrating temperature and then re-deriving salinity. However, this would be quite time consuming and given the noisiness of the data, such a step does not seem justified.

The ratio of TSG fluorescence to CTD fluorescence ranges from 2.2 to 6.0, with an average of 3.3 and a median of 2.8 when one obvious bad TSG value is excluded plus another when the CTD fluorometer went off-scale. (See 2009-10-ctd-tsg-comp.xls.)

- Loop Bottle - TSG Comparisons The loop spreadsheet was edited and saved as 2009-10-loop-CTD-TSG-comp.xls. Columns were added for the TSG salinity, rosette bottle salinity and TSG fluorescence and Rosette CHL. The TSG salinity was lower than the 8 loop salinity samples by an average of 0.024 and by 0.025 if the median was used; the standard deviation was 0.009. If one value were excluded the average was 0.026. The TSG fluorescence was higher than the Loop CHL by a factor of 9 to 12 but there were only 4 points to compare.

- Loop Bottle - Rosette Comparisons The loop and rosette samples were compared as a check of the quality of both. When all available measurements were used the loop salinity was lower than the rosette by an average 0.001, which is as close as we could hope to achieve given many possible sources of error. When one outlier was excluded the average difference was <<0.001. The closeness of the two lends confidence to the use of the CTD to recalibrate the TSG salinity.

The loop CHL was lower than the rosette CHL by an average of 0.027ug/L, and by 0.048ug/L when one outlier (flagged by the analyst) was excluded. All loop values were lower than corresponding rosette sample values. (See 2009-10-loop-CTD-TSG-comp.xls.)

- Calibration History

The TSG primary temperature and conductivity were recalibrated in April 2009 and there is no history of any other uses since that time, though it is likely that it has been used but the data not yet processed. The estimate for ship heating for the Tully at this time of year is from ~0.14Cº to ~0.16Cº; that would vary with ambient temperatures and flow rate in the loop.

Conclusions

1. The TSG clock appears to be working well.

2. Latitude and Longitude were subject to discrete spikes and large sections where values were stuck.

3. Where the intake temperature is not available a proxy can be determined using the following equation:

Temp Intake = 1.0119 * Temp Lab – 0.3038

The CTD comparison would lead to slightly lower intake temperatures, but the number of points is lower in that comparison and the differences are not large for such a crude estimate.

4. Salinity is low by 0.20 based on CTD comparison and 0.24 and 0.25 based on loop samples. The former looks more reliable, so 0.020 will be subtracted from all TSG salinity values. At least half of that error is due to using lab temperature in deriving salinity; calibration errors are likely on the order of 0.05 to 0.10.

5. The fluorescence is higher than the CTD fluorescence by 2.2 to 6 times which is similar to other cruises, but the ratio is not linear with CTD FL. For FL>1 it runs between 2 and 3 times. The TSG fluorescence is higher than the loop samples by 9 to 11 times. The loop CHL is lower than the rosette bottle samples by 0.027 or 0.048ug/L depending on what data are included. All the loop samples are <0.25ug/L, so small errors might account for the very large ratio of TSG fluorescence to Loop CHL. There are too few points of comparison to speculate on which is more likely. There were more loop samples but the TSG fluorometer was not running at the time.

6. The CTD rosette salinity and loop salinity are within 0.001 which is a good agreement. This shows that calibration of the TSG salinity based on CTD data is reasonable. Loop samples are important to confirm this, but there need not be a lot of them.

7. The intake temperature is very close to the CTD temperature from about 4.5m.

8. The habit of taking loop samples at the same time as CTD sampling is very useful so that TSG, loop and CTD can be inter-compared. It is also useful to have the loops at other times to check for any evidence that data collected while moving are different from while stopped. For this cruise none of the latter were available due to the TSG fluorometer not working when moving samples were taken.

f.) Editing

The time-series plots were examined and editing was applied as follows:

File 17: A few spikes in salinity were cleaned where temperature data were fairly smooth. Late in the file there is a very noisy section. This occurred just after completion of the late CTD cast at P16 when conditions were quite rough. Positions were not available for that part of the record.

File 18-19 and 22: Salinity was cleaned lightly.

File 21: Salinity was cleaned lightly and the first 38 points were removed from the intake temperature.

The edited files were copied to *.CLN.

g.) Recalibration

As the intake temperature is unavailable for casts 1-19, an intake proxy channel was derived. First ADD CHANNEL was used to add channel TEMPERATURE:LAB. That was then put through CALIBRATE using file 2009-10-recal1.ccf to set Temperature:LAB equal to Temperature:Primary. Then file 2009-10-recal2.ccf was used to recalibrate Temperature:Primary by applying the equation

Temp Intake = 1.0119 * Temp Lab – 0.3038

and Salinity:T0:C0 by adding 0.020.

Thus we have:

- Temperature:Lab as the uncorrected temperature recorded in the lab.

- Temperature:Primary as the lab temperature adjusted to remove the effects of ship heating; this is a proxy for the intake temperature.

For casts 21-23 the two temperature channels are available, so the only calibration needed is for the salinity. File 2009-10-recal3.ccf was used to subtract 0.020 from the salinity.

h.) Preparing Final Files

REMOVE was used to remove the following channels from casts 1-19: Record #, Scan Number, Time:Julian, Conductivity:Primary, Fluorescence:URU:Seapoint and Flag.

REMOVE was used to remove the following channels from casts 21-23: Record #, Scan Number, Conductivity:Primary, Uploy0 and Flag.

HEADER EDIT was used to add a comment, change the DATA TYPE to THERMOSALINOGRAPH and add the depth of sampling to the header. For casts 1-19 the comment includes notes to indicate that the Temperature:Primary was recalibrated to correct for heating in the loop based on the historic observations and comparison with CTD data. Those files were saved as TOB files.

At this point files from the SCS system were acquired to see if positions could be determined for the parts of files 17, 18, 21 and 22 that have pad values. It is much easier to replace all the positions rather than just those with pad values. The SCS files have more data, sampling every 10s, but by opening them in EXCEL and doing a modulo calculation on the record number it is easy to remove data so that the remaining data are separated by 30s. A few checks were made and while the latitudes match those in TOB files, the longitudes are off by an amount equivalent to ~0.16km. The TOB files match the CTD header positions very well, but that is presumably because they acquired their positions from the same GPS. The bridge has a different system and that is presumably what is in the SCS files. It seems better to have data that may be a little different from that in other files than no data at all in large sections of the files.

Each TOB file was copied to file TOB2. ADD TIME CHANNEL was used to add a record number.

The SCS files were opened and the section that overlaps with a TOB2 file was saved with a format like 2009-10-0017-pos.csv. The excess data were removed. Latitude and longitude positions were converted to decimal format. A check was made that the number of records was the same in the TOB3 and spreadsheet files.

The *pos.csv files were converted to IOS SHELL files named *.pos.

The TOB3 files were then merged with the POS files with record # as common channel, choosing positions from the POS and everything else from the TOB3. The output files were named TOB.

A note was added to the headers of these 4 new TOB files to explain the different method used and warn that the positions are slightly different from those in the other files.

Time-series and plots were produced and no errors were found; a new track plot was placed in the report and it fills in many of the previous gaps.

The TSG sensor history was updated.

As a final check plots were made of the cruise track and data; no problems were noted.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

The following step will be done later:

The final loop file 2009-10 loop.xls was prepared by the chief scientist including data from the final CTD files and salinity and chlorophyll samples from the loop or from 5m bottles. That spreadsheet was simplified, date calculated in DD/MM/YYYY format, header names and formats were adjusted and unneeded channels were removed and saved as a CSV file.  It was converted to IOS format, put through CLEAN and HEADEDIT to get start and stop times and positions, and to add general comments and specific comments for flagged values. The final file was named 2009-10-surface.loop.

Particulars – including notes from log, rosette sheets:

1/2. Downcast/upcast split into two files. The analysts sometimes used 1, sometimes 1/2. The DO data are in 1. Renamed ROS file and all sample files as #1 to match downcast file. Pumps not on until 11db downcast.

2. Only 10 Niskins in file instead of 11 – Niskin 7 at 75 didn’t close. Rename as cast #1.

3. 8 Niskins closed in Table Driven order, then all closed to check if they are ok.

2. ADD file really from cast 3.

7. Original file 7 should have been 8. RENAMED

8. Last Niskin (11) didn’t get closed.

8. Original file 8 should have been 9. There will then be both a ROS cast and a net tow for 9. RENAMED.

10. No NMEA header entry so positions are from previous station and time/date wrong. The positions were fixed and the NMEA time was removed, so header reading will come from the system upload time which agrees with the log entry.- CHECK

16. Long wait at bottom to test heave comp. ~4 minutes

17. Stops at 500, 1000m and 2000m to test boom pressure.

24. LARS pressure readings down and up

29. Primary and secondary salinity odd at end of upcast.

34. Secondary temperature, conductivity and salinity bad from 50-350db of downcast; fluorescence bad too. Upcast secondary looks fine.

39. Switch to aft deck.

44. Top lanyards were off by one, so Niskin 24 closed Niskin 19 at 10 at end of cast. Could we just replace data from N13 by data from N19?

44. Winch spooling problems led to yo-yoing to try to fix.

49. Primary sensor odd, salinity and oxygen bad both down and up; fluorescence ok.

53. More winch problems – yo-yo.

66. 4 Niskins tripped instead of 3 so we used sample numbers 570-573, but labels were already made for cast 67, so sample 573 used twice. (Only UBC water in both cases.)

67. Primary sensors odd, salinity bad and oxygen bad on way down but ok on way up. Fluorescence ok.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 24Dec08 | Factory “ |  |  |
| Conductivity | 3321 | 16Jan09 | “ |  |  |
| Secondary Temp. | 4700 | 24Dec08 | “ |  |  |
| Secondary Cond. | 1766 | 16Jan09 | “ |  |  |
| Transmissometer | 983DR | 27June06 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 10Mar2009 | Factory |  |  |
| PAR | 4615 | 15Dec2000 | IOS |  |  |
| Fluorometer | 2228 | ? | IOS |  |  |
| Pressure Sensor | 77511 | 13/Mar/2000 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |

# TSG

Make/Model/Serial#:	SEABIRD/21/3312       Cruise ID#:	2009-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2488 | 24/04/09 | Factory |  |  |
| Conductivity | 2488 | 24/04/09 | “ |  |  |
| Wetlab/Wetstar FL | WS3S-713P | 18/01/01 | “ |  |  |

Descent Rate of CTD (averaged over 2m) during downcast of Event #38