## File List
Number of Files= |  | 10
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2009-10-0001.mrg | 2009-10-0001.sam | 11 | 20 | 2009-10-0001.mrg & 2009-10-0001.sam
2 | 2009-10-0003.mrg | 2009-10-0003.sam | 22 | 29 | 2009-10-0003.mrg & 2009-10-0003.sam
3 | 2009-10-0006.mrg | 2009-10-0006.sam | 31 | 37 | 2009-10-0006.mrg & 2009-10-0006.sam
4 | 2009-10-0012.mrg | 2009-10-0012.sam | 39 | 56 | 2009-10-0012.mrg & 2009-10-0012.sam
5 | 2009-10-0020.mrg | 2009-10-0020.sam | 58 | 77 | 2009-10-0020.mrg & 2009-10-0020.sam
6 | 2009-10-0029.mrg | 2009-10-0029.sam | 79 | 101 | 2009-10-0029.mrg & 2009-10-0029.sam
7 | 2009-10-0037.mrg | 2009-10-0037.sam | 103 | 125 | 2009-10-0037.mrg & 2009-10-0037.sam
8 | 2009-10-0044.mrg | 2009-10-0044.sam | 127 | 150 | 2009-10-0044.mrg & 2009-10-0044.sam
9 | 2009-10-0061.mrg | 2009-10-0061.sam | 152 | 175 | 2009-10-0061.mrg & 2009-10-0061.sam
10 | 2009-10-0071.mrg | 2009-10-0071.sam | 177 | 199 | 2009-10-0071.mrg & 2009-10-0071.sam

## (DOX_CTD - DOX_BOT)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.93168 |  | X | CALC Y
m= | -0.066955 | 0.0015663 |  | Std Dev of Resid |  | 0.0388 |  | 0.044527 | -0.0021587
b= | 0.00082263 | 0.0051262 |  | Resid Sum of Squares |  | 0.20173 |  | 6.593 | -0.44061
# of Values Extracted = |  |  | 180
# of Values Used in Fit = |  |  | 136
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 220.267 | 0.050766 | 0.001127 | 0.050766 | -0.0025764 | -0.053343 |  | 0.050766 | 0.050766
1 | YES | 200.729 | 0.047012 | 0.0015261 | 0.047012 | -0.0023251 | -0.049337 |  | 0.047012 | 0.047012
1 | YES | 176.236 | 0.044834 | 0.00091597 | 0.044834 | -0.0021792 | -0.047013 |  | 0.044834 | 0.044834
1 | YES | 151.113 | 0.044527 | 0.00097484 | 0.044527 | -0.0021587 | -0.046686 |  | 0.044527 | 0.044527
1 | NO | 125.529 | 0.21221 | 0.0015754 | 0.069208 | -0.013386 | -0.082593 |  | 0.063886 | 0.054886
1 | NO | 100.564 | 0.42818 | 0.018345 | 0.10618 | -0.027846 | -0.13403 |  | 0.063886 | 0.063886
1 | NO | 49.5756 | 2.8843 | 0.0062989 | 0.21328 | -0.19229 | -0.40557 |  | 0.51929 | 0.0052876
1 | NO | 25.5956 | 3.6486 | 0.0090746 | 0.13656 | -0.24347 | -0.38002 |  | 0.51929 | -0.030712
1 | NO | 10.6371 | 5.3874 | 0.048179 | -0.015605 | -0.35989 | -0.34428 |  | 2.128 | -0.064981
1 | NO | 5.9689 | 6.5613 | 0.022672 | -0.029714 | -0.43849 | -0.40877 |  | 2.128 | -0.064981
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.054886 | -0.0034548 | -0.058341 |  | 2.6019 | -0.094104
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.063886 | -0.0034548 | -0.067341 |  | 4.2925 | -0.32851
2 | YES | 98.3678 | 0.51929 | 0.0053849 | 0.0052876 | -0.033946 | -0.039234 |  | 0.49522 | -0.054784
2 | YES | 98.3678 | 0.51929 | 0.0053849 | -0.030712 | -0.033946 | -0.0032337 |  | 0.41302 | -0.032976
2 | NO | 51.3838 | 2.8127 | 0.075542 | 0.20366 | -0.1875 | -0.39116 |  | 0.41302 | -0.026976
2 | NO | 51.3838 | 2.8127 | 0.075542 | 0.17966 | -0.1875 | -0.36716 |  | 0.23775 | -0.013245
2 | NO | 11.1691 | 6.1357 | 0.027563 | 0.22272 | -0.40999 | -0.63272 |  | 0.21223 | -0.024768
2 | NO | 11.1691 | 6.1357 | 0.027563 | 0.46072 | -0.40999 | -0.87071 |  | 0.42334 | -0.040664
3 | NO | 106.372 | 1.5767 | 0.011722 | 0.025656 | -0.10474 | -0.1304 |  | 0.99714 | -0.072863
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.14166 | -0.076677 |  | 1.2875 | -0.075485
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.14166 | -0.076677 |  | 1.6274 | -0.1076
3 | YES | 51.4045 | 2.6019 | 0.0058068 | -0.094104 | -0.17339 | -0.079283 |  | 1.8759 | -0.076058
3 | YES | 26.3033 | 4.2925 | 0.070084 | -0.32851 | -0.28658 | 0.04193 |  | 2.2244 | -0.14162
3 | NO | 11.1766 | 5.5508 | 0.030052 | -0.2702 | -0.37083 | -0.10063 |  | 2.7552 | -0.13478
3 | NO | 5.5253 | 5.7243 | 0.011576 | -0.22071 | -0.38245 | -0.16173 |  | 3.3938 | -0.14018
4 | YES | 1300.62 | 0.49522 | 0.0021859 | -0.054784 | -0.032334 | 0.02245 |  | 4.6783 | -0.22371
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.032976 | -0.026831 | 0.006145 |  | 6.2779 | -0.32107
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.026976 | -0.026831 | 0.00014495 |  | 0.61915 | -0.06185
4 | YES | 1000.54 | 0.23775 | 0.0018914 | -0.013245 | -0.015096 | -0.0018514 |  | 0.61915 | -0.06285
4 | YES | 801.075 | 0.21223 | 0.0020806 | -0.024768 | -0.013387 | 0.01138 |  | 0.33626 | -0.018739
4 | YES | 600.918 | 0.42334 | 0.0028851 | -0.040664 | -0.027522 | 0.013142 |  | 0.21613 | -0.020871
4 | YES | 400.248 | 0.99714 | 0.001509 | -0.072863 | -0.06594 | 0.0069228 |  | 0.37322 | -0.054776
4 | YES | 301.823 | 1.2875 | 0.002412 | -0.075485 | -0.085382 | -0.0098966 |  | 1.0396 | -0.060402
4 | YES | 201.201 | 1.6274 | 0.00573 | -0.1076 | -0.10814 | -0.00054236 |  | 1.8098 | -0.10022
4 | YES | 174.556 | 1.8759 | 0.0068785 | -0.076058 | -0.12478 | -0.048722 |  | 1.9105 | -0.12346
4 | YES | 150.455 | 2.2244 | 0.011745 | -0.14162 | -0.14811 | -0.0064879 |  | 1.9582 | -0.088759
4 | YES | 125.295 | 2.7552 | 0.0045848 | -0.13478 | -0.18365 | -0.048868 |  | 2.0371 | -0.10989
4 | YES | 100.512 | 3.3938 | 0.0073132 | -0.14018 | -0.22641 | -0.086227 |  | 2.1758 | -0.15724
4 | YES | 75.7457 | 4.6783 | 0.018784 | -0.22371 | -0.31241 | -0.088701 |  | 2.6831 | -0.1639
4 | NO | 51.446 | 6.1723 | 0.010765 | -0.20073 | -0.41244 | -0.2117 |  | 3.998 | -0.21401
4 | YES | 24.7249 | 6.2779 | 0.017082 | -0.32107 | -0.41951 | -0.09844 |  | 5.7981 | -0.29987
4 | NO | 10.3947 | 6.4598 | 0.044508 | -0.26224 | -0.43169 | -0.16945 |  | 5.7126 | -0.3694
4 | NO | 5.239 | 6.5395 | 0.016358 | -0.32054 | -0.43702 | -0.11648 |  | 2.1742 | -0.16977
5 | NO | 2003.65 | 1.2821 | 0.003906 | -0.1689 | -0.08502 | 0.08388 |  | 2.1445 | -0.14653
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06185 | -0.040632 | 0.021218 |  | 1.7573 | -0.10471
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06285 | -0.040632 | 0.022218 |  | 1.2232 | -0.074834
5 | YES | 1250.17 | 0.33626 | 0.0025662 | -0.018739 | -0.021692 | -0.002953 |  | 0.60927 | -0.01373
5 | YES | 1001.62 | 0.21613 | 0.0024958 | -0.020871 | -0.013648 | 0.0072232 |  | 0.38709 | -0.057913
5 | NO | 800.329 | 0.21353 | 0.0022177 | -0.79147 | -0.013474 | 0.778 |  | 0.22771 | -0.011286
5 | YES | 599.855 | 0.37322 | 0.0022839 | -0.054776 | -0.024166 | 0.030609 |  | 0.26251 | -0.04849
5 | YES | 401.219 | 1.0396 | 0.0021928 | -0.060402 | -0.068783 | -0.0083809 |  | 0.51929 | -0.064705
5 | YES | 299.827 | 1.8098 | 0.0070129 | -0.10022 | -0.12035 | -0.020131 |  | 1.3473 | -0.13473
5 | YES | 249.318 | 1.9105 | 0.0063301 | -0.12346 | -0.1271 | -0.0036379 |  | 2.5586 | -0.17243
5 | YES | 200.212 | 1.9582 | 0.0035649 | -0.088759 | -0.13029 | -0.041531 |  | 3.273 | -0.22396
5 | YES | 173.88 | 2.0371 | 0.010025 | -0.10989 | -0.13557 | -0.025677 |  | 4.0577 | -0.23326
5 | YES | 150.707 | 2.1758 | 0.0089689 | -0.15724 | -0.14485 | 0.012386 |  | 4.3958 | -0.24818
5 | YES | 124.044 | 2.6831 | 0.012874 | -0.1639 | -0.17882 | -0.014919 |  | 4.6352 | -0.25283
5 | YES | 101.474 | 3.998 | 0.011501 | -0.21401 | -0.26686 | -0.052853 |  | 5.0843 | -0.26169
5 | YES | 75.2867 | 5.7981 | 0.037336 | -0.29987 | -0.38739 | -0.087523 |  | 6.4835 | -0.3485
5 | NO | 49.8081 | 6.2679 | 0.011322 | -0.21206 | -0.41885 | -0.20678 |  | 5.9998 | -0.38818
5 | YES | 25.2931 | 5.7126 | 0.015859 | -0.3694 | -0.38166 | -0.01226 |  | 2.5276 | -0.22339
5 | NO | 10.3099 | 5.7041 | 0.026635 | -0.26693 | -0.38109 | -0.11416 |  | 2.2661 | -0.19986
5 | NO | 5.3529 | 5.7156 | 0.02113 | -0.25144 | -0.38186 | -0.13042 |  | 1.7816 | -0.12635
6 | YES | 3276.19 | 2.1742 | 0.0052468 | -0.16977 | -0.14475 | 0.025016 |  | 1.7818 | -0.12515
6 | YES | 3000.48 | 2.1445 | 0.0049557 | -0.14653 | -0.14276 | 0.0037674 |  | 1.175 | -0.044012
6 | NO | 2501.59 | 1.7573 | 0.0045462 | -0.22471 | -0.11684 | 0.10788 |  | 0.47704 | -0.029963
6 | YES | 2501.59 | 1.7573 | 0.0045462 | -0.10471 | -0.11684 | -0.012122 |  | 0.30083 | -0.010174
6 | YES | 2000.27 | 1.2232 | 0.0036682 | -0.074834 | -0.081074 | -0.0062403 |  | 0.25587 | -0.029133
6 | YES | 1501.16 | 0.60927 | 0.0032809 | -0.01373 | -0.039971 | -0.026241 |  | 0.36527 | -0.022726
6 | YES | 1250.57 | 0.38709 | 0.0031696 | -0.057913 | -0.025095 | 0.032818 |  | 0.57156 | -0.047439
6 | YES | 1001.93 | 0.22771 | 0.0019075 | -0.011286 | -0.014424 | -0.0031376 |  | 1.396 | -0.11402
6 | YES | 801.143 | 0.26251 | 0.0024751 | -0.04849 | -0.016754 | 0.031736 |  | 2.3084 | -0.22357
6 | YES | 600.801 | 0.51929 | 0.0018098 | -0.064705 | -0.033947 | 0.030759 |  | 2.9682 | -0.21684
6 | YES | 400.283 | 1.3473 | 0.0028976 | -0.13473 | -0.089383 | 0.045351 |  | 4.1863 | -0.33474
6 | YES | 301.716 | 2.5586 | 0.0042794 | -0.17243 | -0.17049 | 0.0019455 |  | 4.8064 | -0.34155
6 | YES | 249.814 | 3.273 | 0.0030968 | -0.22396 | -0.21832 | 0.0056354 |  | 5.4656 | -0.35639
6 | YES | 200.271 | 4.0577 | 0.0056607 | -0.23326 | -0.27086 | -0.037601 |  | 6.2031 | -0.40091
6 | YES | 176.597 | 4.3958 | 0.005897 | -0.24818 | -0.2935 | -0.045315 |  | 6.4238 | -0.38922
6 | YES | 150.549 | 4.6352 | 0.0091057 | -0.25283 | -0.30952 | -0.05669 |  | 6.5239 | -0.48108
6 | YES | 124.879 | 5.0843 | 0.0094488 | -0.26169 | -0.3396 | -0.077903 |  | 6.593 | -0.36099
6 | NO | 100.671 | 5.8131 | 0.0097689 | -0.26988 | -0.38839 | -0.11851 |  | 5.6457 | -0.37828
6 | YES | 75.8548 | 6.4835 | 0.015863 | -0.3485 | -0.43328 | -0.084776 |  | 5.6245 | -0.34247
6 | NO | 50.9557 | 6.7227 | 0.016748 | -0.3433 | -0.44929 | -0.10599 |  | 5.6415 | -0.33653
6 | YES | 25.6344 | 5.9998 | 0.01886 | -0.38818 | -0.40089 | -0.01271 |  | 3.0101 | -0.21489
6 | NO | 9.8456 | 2.3884 | 0.025795 | -3.5306 | -0.15909 | 3.3715 |  | 2.9411 | -0.25786
6 | NO | 5.8409 | 2.106 | 0.0078225 | -3.82 | -0.14018 | 3.6798 |  | 2.7317 | -0.22328
7 | YES | 3500.19 | 2.5276 | 0.0043321 | -0.22339 | -0.16841 | 0.054977 |  | 2.3371 | -0.20088
7 | YES | 2999.58 | 2.2661 | 0.0053773 | -0.19986 | -0.15091 | 0.048952 |  | 1.7898 | -0.16321
7 | YES | 2497.23 | 1.7816 | 0.0045839 | -0.12635 | -0.11847 | 0.0078859 |  | 1.1834 | -0.11965
7 | YES | 2496.9 | 1.7818 | 0.0046856 | -0.12515 | -0.11848 | 0.0066685 |  | 0.49161 | -0.00739
7 | YES | 2001.58 | 1.175 | 0.0032203 | -0.044012 | -0.077848 | -0.033836 |  | 0.32903 | -0.012967
7 | YES | 1501.9 | 0.47704 | 0.0031055 | -0.029963 | -0.031117 | -0.0011545 |  | 0.35476 | -0.023241
7 | YES | 1251.29 | 0.30083 | 0.0028858 | -0.010174 | -0.019319 | -0.0091448 |  | 0.56649 | -0.060506
7 | YES | 1000.7 | 0.25587 | 0.0025214 | -0.029133 | -0.016309 | 0.012824 |  | 1.0657 | -0.11028
7 | YES | 801.449 | 0.36527 | 0.0024324 | -0.022726 | -0.023634 | -0.00090806 |  | 1.7515 | -0.14347
7 | YES | 599.478 | 0.57156 | 0.002331 | -0.047439 | -0.037446 | 0.009993 |  | 2.5384 | -0.21861
7 | YES | 400.235 | 1.396 | 0.0041091 | -0.11402 | -0.092645 | 0.02137 |  | 3.4453 | -0.29367
7 | YES | 300.436 | 2.3084 | 0.0038714 | -0.22357 | -0.15374 | 0.069831 |  | 4.1073 | -0.30468
7 | YES | 250.34 | 2.9682 | 0.0091679 | -0.21684 | -0.19791 | 0.018933 |  | 5.0372 | -0.36982
7 | YES | 200.034 | 4.1863 | 0.02518 | -0.33474 | -0.27947 | 0.05527 |  | 6.4907 | -0.43729
7 | YES | 176.444 | 4.8064 | 0.0094576 | -0.34155 | -0.32099 | 0.020554 |  | 6.4701 | -0.43592
7 | YES | 149.804 | 5.4656 | 0.010897 | -0.35639 | -0.36513 | -0.0087393 |  | 6.4631 | -0.44993
7 | YES | 126.568 | 6.2031 | 0.01604 | -0.40091 | -0.4145 | -0.013598 |  | 5.8293 | -0.41173
7 | YES | 101.691 | 6.4238 | 0.0076469 | -0.38922 | -0.42928 | -0.040055 |  | 5.8095 | -0.40253
7 | YES | 74.5608 | 6.5239 | 0.024834 | -0.48108 | -0.43598 | 0.045099 |  | 2.6327 | -0.20433
7 | YES | 49.9787 | 6.593 | 0.018445 | -0.36099 | -0.44061 | -0.079619 |  | 2.3213 | -0.16272
7 | YES | 25.418 | 5.6457 | 0.022196 | -0.37828 | -0.37718 | 0.0010935 |  | 1.794 | -0.15396
7 | YES | 10.5048 | 5.6245 | 0.016742 | -0.34247 | -0.37577 | -0.033301 |  | 1.2011 | -0.077852
7 | YES | 6.1822 | 5.6415 | 0.017912 | -0.33653 | -0.3769 | -0.040364 |  | 0.508 | -0.036
8 | YES | 4315.72 | 3.0101 | 0.0067191 | -0.21489 | -0.20072 | 0.014174 |  | 0.33658 | -0.0034205
8 | YES | 3998.96 | 2.9411 | 0.007228 | -0.25786 | -0.1961 | 0.061763 |  | 0.31054 | -0.010465
8 | YES | 3501.9 | 2.7317 | 0.0053451 | -0.22328 | -0.18208 | 0.041203 |  | 0.42635 | -0.044647
8 | YES | 3000.3 | 2.3371 | 0.0058576 | -0.20088 | -0.15566 | 0.045226 |  | 0.5665 | -0.033502
8 | YES | 2500.89 | 1.7898 | 0.004994 | -0.16321 | -0.11901 | 0.044195 |  | 1.1597 | -0.13529
8 | YES | 2003.55 | 1.1834 | 0.0039803 | -0.11965 | -0.078409 | 0.041236 |  | 2.4496 | -0.21636
8 | YES | 1501.95 | 0.49161 | 0.0036682 | -0.00739 | -0.032093 | -0.024703 |  | 4.0755 | -0.31946
8 | YES | 1250.92 | 0.32903 | 0.0029901 | -0.012967 | -0.021208 | -0.0082409 |  | 6.2541 | -0.47289
8 | YES | 999.715 | 0.35476 | 0.0024273 | -0.023241 | -0.02293 | 0.00031056 |  | 6.3995 | -0.44349
8 | NO | 802.475 | 0.4331 | 0.0019361 | -0.6209 | -0.028175 | 0.59272 |  | 6.4517 | -0.48427
8 | YES | 599.257 | 0.56649 | 0.0024068 | -0.060506 | -0.037107 | 0.023399 |  | 1.1924 | -0.097585
8 | YES | 399.823 | 1.0657 | 0.0028522 | -0.11028 | -0.070532 | 0.039746 |  | 0.54876 | -0.014236
8 | YES | 297.593 | 1.7515 | 0.0051562 | -0.14347 | -0.11645 | 0.027022 |  | 0.32417 | 0.013174
8 | YES | 247.972 | 2.5384 | 0.0075519 | -0.21861 | -0.16913 | 0.049472 |  | 0.27827 | 0.024274
8 | YES | 200.803 | 3.4453 | 0.013674 | -0.29367 | -0.22986 | 0.06381 |  | 0.35488 | -0.033116
8 | YES | 174.613 | 4.1073 | 0.0079549 | -0.30468 | -0.27418 | 0.030501 |  | 0.54376 | -0.023237
8 | YES | 149.88 | 5.0372 | 0.018226 | -0.36982 | -0.33644 | 0.033382 |  | 1.0239 | -0.07805
8 | NO | 124.694 | 6.0571 | 0.013726 | -0.46995 | -0.40472 | 0.065229 |  | 1.4423 | -0.10473
8 | YES | 100.56 | 6.4907 | 0.014136 | -0.43729 | -0.43376 | 0.0035335 |  | 1.6911 | -0.12986
8 | YES | 73.8024 | 6.4701 | 0.01843 | -0.43592 | -0.43238 | 0.0035377 |  | 2.0769 | -0.15214
8 | YES | 50.6964 | 6.4631 | 0.0187 | -0.44993 | -0.43191 | 0.01802 |  | 2.3036 | -0.15136
8 | YES | 23.8606 | 5.8293 | 0.013286 | -0.41173 | -0.38947 | 0.022256 |  | 2.5619 | -0.20505
8 | YES | 11.107 | 5.8095 | 0.018044 | -0.40253 | -0.38815 | 0.014379 |  | 5.319 | -0.39605
8 | NO | 5.5023 | 5.8033 | 0.018351 | -0.57167 | -0.38774 | 0.18393 |  | 6.1191 | -0.46193
9 | NO | 4024.09 | 2.7851 | 0.0073442 | -0.39191 | -0.18565 | 0.20625 |  | 6.5152 | -0.48879
9 | NO | 3498.86 | 2.6327 | 0.0069541 | -0.24833 | -0.17545 | 0.072882 |  | 5.5937 | -0.41628
9 | YES | 3498.86 | 2.6327 | 0.0069541 | -0.20433 | -0.17545 | 0.028882 |  | 5.5706 | -0.42942
9 | YES | 2998.76 | 2.3213 | 0.0052467 | -0.16272 | -0.1546 | 0.0081202
9 | YES | 2501.18 | 1.794 | 0.0040669 | -0.15396 | -0.1193 | 0.034661
9 | YES | 1999.9 | 1.2011 | 0.004018 | -0.077852 | -0.0796 | -0.0017479
9 | YES | 1501.12 | 0.508 | 0.0028358 | -0.036 | -0.03319 | 0.0028097
9 | YES | 1252.06 | 0.33658 | 0.0031631 | -0.0034205 | -0.021713 | -0.018292
9 | YES | 1001.45 | 0.31054 | 0.0026972 | -0.010465 | -0.019969 | -0.0095044
9 | YES | 801.223 | 0.42635 | 0.0021842 | -0.044647 | -0.027724 | 0.016924
9 | YES | 601.026 | 0.5665 | 0.002174 | -0.033502 | -0.037107 | -0.0036049
9 | YES | 399.824 | 1.1597 | 0.0042208 | -0.13529 | -0.076825 | 0.058465
9 | NO | 300.177 | 1.884 | 0.0040253 | -0.20196 | -0.12532 | 0.07664
9 | YES | 250.475 | 2.4496 | 0.0055907 | -0.21636 | -0.16319 | 0.053165
9 | NO | 198.379 | 3.6022 | 0.0062745 | -0.31876 | -0.24036 | 0.078391
9 | YES | 175.41 | 4.0755 | 0.013949 | -0.31946 | -0.27205 | 0.047406
9 | NO | 150.732 | 4.9887 | 0.043861 | -0.44525 | -0.3332 | 0.11205
9 | NO | 124.125 | 5.7935 | 0.028387 | -0.49652 | -0.38708 | 0.10944
9 | YES | 99.4924 | 6.2541 | 0.018472 | -0.47289 | -0.41792 | 0.054973
9 | YES | 74.7791 | 6.3995 | 0.023285 | -0.44349 | -0.42765 | 0.015836
9 | YES | 49.3064 | 6.4517 | 0.029448 | -0.48427 | -0.43115 | 0.053114
9 | NO | 24.7034 | 5.6485 | 0.011864 | -0.45955 | -0.37737 | 0.082185
9 | NO | 9.7331 | 5.6612 | 0.021845 | -0.44281 | -0.37822 | 0.064593
9 | NO | 5.1962 | 5.6572 | 0.024126 | -0.45378 | -0.37795 | 0.075825
10 | YES | 1998.83 | 1.1924 | 0.0047216 | -0.097585 | -0.079015 | 0.01857
10 | YES | 1518.6 | 0.54876 | 0.0029639 | -0.014236 | -0.03592 | -0.021683
10 | YES | 1249.67 | 0.32417 | 0.0025906 | 0.013174 | -0.020882 | -0.034057
10 | YES | 999.208 | 0.27827 | 0.0030413 | 0.024274 | -0.017809 | -0.042083
10 | YES | 799.271 | 0.35488 | 0.0025987 | -0.033116 | -0.022939 | 0.010178
10 | YES | 599.385 | 0.54376 | 0.0024845 | -0.023237 | -0.035585 | -0.012348
10 | YES | 402.41 | 1.0239 | 0.0073432 | -0.07805 | -0.067736 | 0.010314
10 | YES | 301.932 | 1.4423 | 0.0027367 | -0.10473 | -0.095744 | 0.008986
10 | YES | 250.688 | 1.6911 | 0.006699 | -0.12986 | -0.11241 | 0.017456
10 | YES | 199.64 | 2.0769 | 0.0036755 | -0.15214 | -0.13823 | 0.013907
10 | YES | 175.483 | 2.3036 | 0.0074563 | -0.15136 | -0.15342 | -0.0020598
10 | YES | 150.246 | 2.5619 | 0.0096489 | -0.20505 | -0.17071 | 0.034334
10 | NO | 124.167 | 3.7627 | 0.032554 | -0.40133 | -0.25111 | 0.15022
10 | YES | 99.5251 | 5.319 | 0.016428 | -0.39605 | -0.35531 | 0.040748
10 | YES | 74.3439 | 6.1191 | 0.011027 | -0.46193 | -0.40888 | 0.053056
10 | NO | 49.3314 | 6.4523 | 0.025608 | -0.49372 | -0.43119 | 0.062535
10 | YES | 40.9976 | 6.5152 | 0.035805 | -0.48879 | -0.4354 | 0.053392
10 | NO | 29.6041 | 5.6367 | 0.0091063 | -0.44334 | -0.37658 | 0.066758
10 | NO | 22.1111 | 5.612 | 0.025038 | -0.45695 | -0.37493 | 0.082018
10 | NO | 22.1111 | 5.612 | 0.025038 | -0.45495 | -0.37493 | 0.080018
10 | YES | 10.6242 | 5.5937 | 0.019454 | -0.41628 | -0.3737 | 0.042575
10 | YES | 4.5235 | 5.5706 | 0.032389 | -0.42942 | -0.37215 | 0.057262
10 | NO | 3.0114 | 5.5692 | 0.03293 | -0.45884 | -0.37206 | 0.086784
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
NO
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
NO
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
NO
NO
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
NO
NO
NO
YES
YES
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
3
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=File Pair #
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.044502 |  | X | CALC Y
m= | -0.033455 | 0.011619 |  | Std Dev of Resid |  | 0.40316 |  | 1 | -0.040666
b= | -0.0072107 | 0.080445 |  | Resid Sum of Squares |  | 28.9317 |  | 10 | -0.34176
# of Values Extracted = |  |  | 180
# of Values Used in Fit = |  |  | 180
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | File Pair #
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 220.267 | 0.050766 | 0.001127 | 0.050766 | -0.040666 | -0.091432 |  | 1 | 0.050766
1 | YES | 200.729 | 0.047012 | 0.0015261 | 0.047012 | -0.040666 | -0.087678 |  | 1 | 0.047012
1 | YES | 176.236 | 0.044834 | 0.00091597 | 0.044834 | -0.040666 | -0.0855 |  | 1 | 0.044834
1 | YES | 151.113 | 0.044527 | 0.00097484 | 0.044527 | -0.040666 | -0.085193 |  | 1 | 0.044527
1 | YES | 125.529 | 0.21221 | 0.0015754 | 0.069208 | -0.040666 | -0.10987 |  | 1 | 0.069208
1 | YES | 100.564 | 0.42818 | 0.018345 | 0.10618 | -0.040666 | -0.14685 |  | 1 | 0.10618
1 | YES | 49.5756 | 2.8843 | 0.0062989 | 0.21328 | -0.040666 | -0.25394 |  | 1 | 0.21328
1 | YES | 25.5956 | 3.6486 | 0.0090746 | 0.13656 | -0.040666 | -0.17722 |  | 1 | 0.13656
1 | YES | 10.6371 | 5.3874 | 0.048179 | -0.015605 | -0.040666 | -0.025061 |  | 1 | -0.015605
1 | YES | 5.9689 | 6.5613 | 0.022672 | -0.029714 | -0.040666 | -0.010952 |  | 1 | -0.029714
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.054886 | -0.074121 | -0.12901 |  | 2 | 0.054886
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.063886 | -0.074121 | -0.13801 |  | 2 | 0.063886
2 | YES | 98.3678 | 0.51929 | 0.0053849 | 0.0052876 | -0.074121 | -0.079409 |  | 2 | 0.0052876
2 | YES | 98.3678 | 0.51929 | 0.0053849 | -0.030712 | -0.074121 | -0.043409 |  | 2 | -0.030712
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.20366 | -0.074121 | -0.27778 |  | 2 | 0.20366
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.17966 | -0.074121 | -0.25378 |  | 2 | 0.17966
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.22272 | -0.074121 | -0.29684 |  | 2 | 0.22272
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.46072 | -0.074121 | -0.53484 |  | 2 | 0.46072
3 | YES | 106.372 | 1.5767 | 0.011722 | 0.025656 | -0.10758 | -0.13323 |  | 3 | 0.025656
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.10758 | -0.042596 |  | 3 | -0.064981
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.10758 | -0.042596 |  | 3 | -0.064981
3 | YES | 51.4045 | 2.6019 | 0.0058068 | -0.094104 | -0.10758 | -0.013473 |  | 3 | -0.094104
3 | YES | 26.3033 | 4.2925 | 0.070084 | -0.32851 | -0.10758 | 0.22093 |  | 3 | -0.32851
3 | YES | 11.1766 | 5.5508 | 0.030052 | -0.2702 | -0.10758 | 0.16263 |  | 3 | -0.2702
3 | YES | 5.5253 | 5.7243 | 0.011576 | -0.22071 | -0.10758 | 0.11314 |  | 3 | -0.22071
4 | YES | 1300.62 | 0.49522 | 0.0021859 | -0.054784 | -0.14103 | -0.086248 |  | 4 | -0.054784
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.032976 | -0.14103 | -0.10806 |  | 4 | -0.032976
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.026976 | -0.14103 | -0.11406 |  | 4 | -0.026976
4 | YES | 1000.54 | 0.23775 | 0.0018914 | -0.013245 | -0.14103 | -0.12779 |  | 4 | -0.013245
4 | YES | 801.075 | 0.21223 | 0.0020806 | -0.024768 | -0.14103 | -0.11626 |  | 4 | -0.024768
4 | YES | 600.918 | 0.42334 | 0.0028851 | -0.040664 | -0.14103 | -0.10037 |  | 4 | -0.040664
4 | YES | 400.248 | 0.99714 | 0.001509 | -0.072863 | -0.14103 | -0.068169 |  | 4 | -0.072863
4 | YES | 301.823 | 1.2875 | 0.002412 | -0.075485 | -0.14103 | -0.065547 |  | 4 | -0.075485
4 | YES | 201.201 | 1.6274 | 0.00573 | -0.1076 | -0.14103 | -0.033435 |  | 4 | -0.1076
4 | YES | 174.556 | 1.8759 | 0.0068785 | -0.076058 | -0.14103 | -0.064974 |  | 4 | -0.076058
4 | YES | 150.455 | 2.2244 | 0.011745 | -0.14162 | -0.14103 | 0.00058975 |  | 4 | -0.14162
4 | YES | 125.295 | 2.7552 | 0.0045848 | -0.13478 | -0.14103 | -0.0062482 |  | 4 | -0.13478
4 | YES | 100.512 | 3.3938 | 0.0073132 | -0.14018 | -0.14103 | -0.00085025 |  | 4 | -0.14018
4 | YES | 75.7457 | 4.6783 | 0.018784 | -0.22371 | -0.14103 | 0.082678 |  | 4 | -0.22371
4 | YES | 51.446 | 6.1723 | 0.010765 | -0.20073 | -0.14103 | 0.059703 |  | 4 | -0.20073
4 | YES | 24.7249 | 6.2779 | 0.017082 | -0.32107 | -0.14103 | 0.18004 |  | 4 | -0.32107
4 | YES | 10.3947 | 6.4598 | 0.044508 | -0.26224 | -0.14103 | 0.1212 |  | 4 | -0.26224
4 | YES | 5.239 | 6.5395 | 0.016358 | -0.32054 | -0.14103 | 0.17951 |  | 4 | -0.32054
5 | YES | 2003.65 | 1.2821 | 0.003906 | -0.1689 | -0.17449 | -0.0055876 |  | 5 | -0.1689
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06185 | -0.17449 | -0.11264 |  | 5 | -0.06185
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06285 | -0.17449 | -0.11164 |  | 5 | -0.06285
5 | YES | 1250.17 | 0.33626 | 0.0025662 | -0.018739 | -0.17449 | -0.15575 |  | 5 | -0.018739
5 | YES | 1001.62 | 0.21613 | 0.0024958 | -0.020871 | -0.17449 | -0.15362 |  | 5 | -0.020871
5 | YES | 800.329 | 0.21353 | 0.0022177 | -0.79147 | -0.17449 | 0.61699 |  | 5 | -0.79147
5 | YES | 599.855 | 0.37322 | 0.0022839 | -0.054776 | -0.17449 | -0.11971 |  | 5 | -0.054776
5 | YES | 401.219 | 1.0396 | 0.0021928 | -0.060402 | -0.17449 | -0.11409 |  | 5 | -0.060402
5 | YES | 299.827 | 1.8098 | 0.0070129 | -0.10022 | -0.17449 | -0.074268 |  | 5 | -0.10022
5 | YES | 249.318 | 1.9105 | 0.0063301 | -0.12346 | -0.17449 | -0.051029 |  | 5 | -0.12346
5 | YES | 200.212 | 1.9582 | 0.0035649 | -0.088759 | -0.17449 | -0.085728 |  | 5 | -0.088759
5 | YES | 173.88 | 2.0371 | 0.010025 | -0.10989 | -0.17449 | -0.064594 |  | 5 | -0.10989
5 | YES | 150.707 | 2.1758 | 0.0089689 | -0.15724 | -0.17449 | -0.017247 |  | 5 | -0.15724
5 | YES | 124.044 | 2.6831 | 0.012874 | -0.1639 | -0.17449 | -0.010584 |  | 5 | -0.1639
5 | YES | 101.474 | 3.998 | 0.011501 | -0.21401 | -0.17449 | 0.03952 |  | 5 | -0.21401
5 | YES | 75.2867 | 5.7981 | 0.037336 | -0.29987 | -0.17449 | 0.12538 |  | 5 | -0.29987
5 | YES | 49.8081 | 6.2679 | 0.011322 | -0.21206 | -0.17449 | 0.037574 |  | 5 | -0.21206
5 | YES | 25.2931 | 5.7126 | 0.015859 | -0.3694 | -0.17449 | 0.19491 |  | 5 | -0.3694
5 | YES | 10.3099 | 5.7041 | 0.026635 | -0.26693 | -0.17449 | 0.092442 |  | 5 | -0.26693
5 | YES | 5.3529 | 5.7156 | 0.02113 | -0.25144 | -0.17449 | 0.076952 |  | 5 | -0.25144
6 | YES | 3276.19 | 2.1742 | 0.0052468 | -0.16977 | -0.20794 | -0.038175 |  | 6 | -0.16977
6 | YES | 3000.48 | 2.1445 | 0.0049557 | -0.14653 | -0.20794 | -0.061416 |  | 6 | -0.14653
6 | YES | 2501.59 | 1.7573 | 0.0045462 | -0.22471 | -0.20794 | 0.016771 |  | 6 | -0.22471
6 | YES | 2501.59 | 1.7573 | 0.0045462 | -0.10471 | -0.20794 | -0.10323 |  | 6 | -0.10471
6 | YES | 2000.27 | 1.2232 | 0.0036682 | -0.074834 | -0.20794 | -0.13311 |  | 6 | -0.074834
6 | YES | 1501.16 | 0.60927 | 0.0032809 | -0.01373 | -0.20794 | -0.19421 |  | 6 | -0.01373
6 | YES | 1250.57 | 0.38709 | 0.0031696 | -0.057913 | -0.20794 | -0.15003 |  | 6 | -0.057913
6 | YES | 1001.93 | 0.22771 | 0.0019075 | -0.011286 | -0.20794 | -0.19666 |  | 6 | -0.011286
6 | YES | 801.143 | 0.26251 | 0.0024751 | -0.04849 | -0.20794 | -0.15945 |  | 6 | -0.04849
6 | YES | 600.801 | 0.51929 | 0.0018098 | -0.064705 | -0.20794 | -0.14324 |  | 6 | -0.064705
6 | YES | 400.283 | 1.3473 | 0.0028976 | -0.13473 | -0.20794 | -0.073209 |  | 6 | -0.13473
6 | YES | 301.716 | 2.5586 | 0.0042794 | -0.17243 | -0.20794 | -0.035512 |  | 6 | -0.17243
6 | YES | 249.814 | 3.273 | 0.0030968 | -0.22396 | -0.20794 | 0.016015 |  | 6 | -0.22396
6 | YES | 200.271 | 4.0577 | 0.0056607 | -0.23326 | -0.20794 | 0.025318 |  | 6 | -0.23326
6 | YES | 176.597 | 4.3958 | 0.005897 | -0.24818 | -0.20794 | 0.04024 |  | 6 | -0.24818
6 | YES | 150.549 | 4.6352 | 0.0091057 | -0.25283 | -0.20794 | 0.044891 |  | 6 | -0.25283
6 | YES | 124.879 | 5.0843 | 0.0094488 | -0.26169 | -0.20794 | 0.05375 |  | 6 | -0.26169
6 | YES | 100.671 | 5.8131 | 0.0097689 | -0.26988 | -0.20794 | 0.061937 |  | 6 | -0.26988
6 | YES | 75.8548 | 6.4835 | 0.015863 | -0.3485 | -0.20794 | 0.14056 |  | 6 | -0.3485
6 | YES | 50.9557 | 6.7227 | 0.016748 | -0.3433 | -0.20794 | 0.13536 |  | 6 | -0.3433
6 | YES | 25.6344 | 5.9998 | 0.01886 | -0.38818 | -0.20794 | 0.18024 |  | 6 | -0.38818
6 | YES | 9.8456 | 2.3884 | 0.025795 | -3.5306 | -0.20794 | 3.3226 |  | 6 | -3.5306
6 | YES | 5.8409 | 2.106 | 0.0078225 | -3.82 | -0.20794 | 3.612 |  | 6 | -3.82
7 | YES | 3500.19 | 2.5276 | 0.0043321 | -0.22339 | -0.2414 | -0.018008 |  | 7 | -0.22339
7 | YES | 2999.58 | 2.2661 | 0.0053773 | -0.19986 | -0.2414 | -0.04154 |  | 7 | -0.19986
7 | YES | 2497.23 | 1.7816 | 0.0045839 | -0.12635 | -0.2414 | -0.11505 |  | 7 | -0.12635
7 | YES | 2496.9 | 1.7818 | 0.0046856 | -0.12515 | -0.2414 | -0.11625 |  | 7 | -0.12515
7 | YES | 2001.58 | 1.175 | 0.0032203 | -0.044012 | -0.2414 | -0.19739 |  | 7 | -0.044012
7 | YES | 1501.9 | 0.47704 | 0.0031055 | -0.029963 | -0.2414 | -0.21144 |  | 7 | -0.029963
7 | YES | 1251.29 | 0.30083 | 0.0028858 | -0.010174 | -0.2414 | -0.23122 |  | 7 | -0.010174
7 | YES | 1000.7 | 0.25587 | 0.0025214 | -0.029133 | -0.2414 | -0.21227 |  | 7 | -0.029133
7 | YES | 801.449 | 0.36527 | 0.0024324 | -0.022726 | -0.2414 | -0.21867 |  | 7 | -0.022726
7 | YES | 599.478 | 0.57156 | 0.002331 | -0.047439 | -0.2414 | -0.19396 |  | 7 | -0.047439
7 | YES | 400.235 | 1.396 | 0.0041091 | -0.11402 | -0.2414 | -0.12738 |  | 7 | -0.11402
7 | YES | 300.436 | 2.3084 | 0.0038714 | -0.22357 | -0.2414 | -0.01783 |  | 7 | -0.22357
7 | YES | 250.34 | 2.9682 | 0.0091679 | -0.21684 | -0.2414 | -0.024556 |  | 7 | -0.21684
7 | YES | 200.034 | 4.1863 | 0.02518 | -0.33474 | -0.2414 | 0.093339 |  | 7 | -0.33474
7 | YES | 176.444 | 4.8064 | 0.0094576 | -0.34155 | -0.2414 | 0.10015 |  | 7 | -0.34155
7 | YES | 149.804 | 5.4656 | 0.010897 | -0.35639 | -0.2414 | 0.11499 |  | 7 | -0.35639
7 | YES | 126.568 | 6.2031 | 0.01604 | -0.40091 | -0.2414 | 0.15951 |  | 7 | -0.40091
7 | YES | 101.691 | 6.4238 | 0.0076469 | -0.38922 | -0.2414 | 0.14783 |  | 7 | -0.38922
7 | YES | 74.5608 | 6.5239 | 0.024834 | -0.48108 | -0.2414 | 0.23968 |  | 7 | -0.48108
7 | YES | 49.9787 | 6.593 | 0.018445 | -0.36099 | -0.2414 | 0.11959 |  | 7 | -0.36099
7 | YES | 25.418 | 5.6457 | 0.022196 | -0.37828 | -0.2414 | 0.13688 |  | 7 | -0.37828
7 | YES | 10.5048 | 5.6245 | 0.016742 | -0.34247 | -0.2414 | 0.10107 |  | 7 | -0.34247
7 | YES | 6.1822 | 5.6415 | 0.017912 | -0.33653 | -0.2414 | 0.095137 |  | 7 | -0.33653
8 | YES | 4315.72 | 3.0101 | 0.0067191 | -0.21489 | -0.27485 | -0.059962 |  | 8 | -0.21489
8 | YES | 3998.96 | 2.9411 | 0.007228 | -0.25786 | -0.27485 | -0.016991 |  | 8 | -0.25786
8 | YES | 3501.9 | 2.7317 | 0.0053451 | -0.22328 | -0.27485 | -0.051572 |  | 8 | -0.22328
8 | YES | 3000.3 | 2.3371 | 0.0058576 | -0.20088 | -0.27485 | -0.07397 |  | 8 | -0.20088
8 | YES | 2500.89 | 1.7898 | 0.004994 | -0.16321 | -0.27485 | -0.11165 |  | 8 | -0.16321
8 | YES | 2003.55 | 1.1834 | 0.0039803 | -0.11965 | -0.27485 | -0.15521 |  | 8 | -0.11965
8 | YES | 1501.95 | 0.49161 | 0.0036682 | -0.00739 | -0.27485 | -0.26746 |  | 8 | -0.00739
8 | YES | 1250.92 | 0.32903 | 0.0029901 | -0.012967 | -0.27485 | -0.26189 |  | 8 | -0.012967
8 | YES | 999.715 | 0.35476 | 0.0024273 | -0.023241 | -0.27485 | -0.25161 |  | 8 | -0.023241
8 | YES | 802.475 | 0.4331 | 0.0019361 | -0.6209 | -0.27485 | 0.34605 |  | 8 | -0.6209
8 | YES | 599.257 | 0.56649 | 0.0024068 | -0.060506 | -0.27485 | -0.21435 |  | 8 | -0.060506
8 | YES | 399.823 | 1.0657 | 0.0028522 | -0.11028 | -0.27485 | -0.16458 |  | 8 | -0.11028
8 | YES | 297.593 | 1.7515 | 0.0051562 | -0.14347 | -0.27485 | -0.13138 |  | 8 | -0.14347
8 | YES | 247.972 | 2.5384 | 0.0075519 | -0.21861 | -0.27485 | -0.056248 |  | 8 | -0.21861
8 | YES | 200.803 | 3.4453 | 0.013674 | -0.29367 | -0.27485 | 0.018814 |  | 8 | -0.29367
8 | YES | 174.613 | 4.1073 | 0.0079549 | -0.30468 | -0.27485 | 0.029829 |  | 8 | -0.30468
8 | YES | 149.88 | 5.0372 | 0.018226 | -0.36982 | -0.27485 | 0.094968 |  | 8 | -0.36982
8 | YES | 124.694 | 6.0571 | 0.013726 | -0.46995 | -0.27485 | 0.1951 |  | 8 | -0.46995
8 | YES | 100.56 | 6.4907 | 0.014136 | -0.43729 | -0.27485 | 0.16244 |  | 8 | -0.43729
8 | YES | 73.8024 | 6.4701 | 0.01843 | -0.43592 | -0.27485 | 0.16106 |  | 8 | -0.43592
8 | YES | 50.6964 | 6.4631 | 0.0187 | -0.44993 | -0.27485 | 0.17508 |  | 8 | -0.44993
8 | YES | 23.8606 | 5.8293 | 0.013286 | -0.41173 | -0.27485 | 0.13688 |  | 8 | -0.41173
8 | YES | 11.107 | 5.8095 | 0.018044 | -0.40253 | -0.27485 | 0.12767 |  | 8 | -0.40253
8 | YES | 5.5023 | 5.8033 | 0.018351 | -0.57167 | -0.27485 | 0.29682 |  | 8 | -0.57167
9 | YES | 4024.09 | 2.7851 | 0.0073442 | -0.39191 | -0.30831 | 0.083596 |  | 9 | -0.39191
9 | YES | 3498.86 | 2.6327 | 0.0069541 | -0.24833 | -0.30831 | -0.05998 |  | 9 | -0.24833
9 | YES | 3498.86 | 2.6327 | 0.0069541 | -0.20433 | -0.30831 | -0.10398 |  | 9 | -0.20433
9 | YES | 2998.76 | 2.3213 | 0.0052467 | -0.16272 | -0.30831 | -0.14559 |  | 9 | -0.16272
9 | YES | 2501.18 | 1.794 | 0.0040669 | -0.15396 | -0.30831 | -0.15435 |  | 9 | -0.15396
9 | YES | 1999.9 | 1.2011 | 0.004018 | -0.077852 | -0.30831 | -0.23046 |  | 9 | -0.077852
9 | YES | 1501.12 | 0.508 | 0.0028358 | -0.036 | -0.30831 | -0.27231 |  | 9 | -0.036
9 | YES | 1252.06 | 0.33658 | 0.0031631 | -0.0034205 | -0.30831 | -0.30489 |  | 9 | -0.0034205
9 | YES | 1001.45 | 0.31054 | 0.0026972 | -0.010465 | -0.30831 | -0.29784 |  | 9 | -0.010465
9 | YES | 801.223 | 0.42635 | 0.0021842 | -0.044647 | -0.30831 | -0.26366 |  | 9 | -0.044647
9 | YES | 601.026 | 0.5665 | 0.002174 | -0.033502 | -0.30831 | -0.27481 |  | 9 | -0.033502
9 | YES | 399.824 | 1.1597 | 0.0042208 | -0.13529 | -0.30831 | -0.17302 |  | 9 | -0.13529
9 | YES | 300.177 | 1.884 | 0.0040253 | -0.20196 | -0.30831 | -0.10635 |  | 9 | -0.20196
9 | YES | 250.475 | 2.4496 | 0.0055907 | -0.21636 | -0.30831 | -0.091952 |  | 9 | -0.21636
9 | YES | 198.379 | 3.6022 | 0.0062745 | -0.31876 | -0.30831 | 0.010446 |  | 9 | -0.31876
9 | YES | 175.41 | 4.0755 | 0.013949 | -0.31946 | -0.30831 | 0.011151 |  | 9 | -0.31946
9 | YES | 150.732 | 4.9887 | 0.043861 | -0.44525 | -0.30831 | 0.13694 |  | 9 | -0.44525
9 | YES | 124.125 | 5.7935 | 0.028387 | -0.49652 | -0.30831 | 0.18821 |  | 9 | -0.49652
9 | YES | 99.4924 | 6.2541 | 0.018472 | -0.47289 | -0.30831 | 0.16458 |  | 9 | -0.47289
9 | YES | 74.7791 | 6.3995 | 0.023285 | -0.44349 | -0.30831 | 0.13518 |  | 9 | -0.44349
9 | YES | 49.3064 | 6.4517 | 0.029448 | -0.48427 | -0.30831 | 0.17596 |  | 9 | -0.48427
9 | YES | 24.7034 | 5.6485 | 0.011864 | -0.45955 | -0.30831 | 0.15124 |  | 9 | -0.45955
9 | YES | 9.7331 | 5.6612 | 0.021845 | -0.44281 | -0.30831 | 0.1345 |  | 9 | -0.44281
9 | YES | 5.1962 | 5.6572 | 0.024126 | -0.45378 | -0.30831 | 0.14547 |  | 9 | -0.45378
10 | YES | 1998.83 | 1.1924 | 0.0047216 | -0.097585 | -0.34176 | -0.24418 |  | 10 | -0.097585
10 | YES | 1518.6 | 0.54876 | 0.0029639 | -0.014236 | -0.34176 | -0.32753 |  | 10 | -0.014236
10 | YES | 1249.67 | 0.32417 | 0.0025906 | 0.013174 | -0.34176 | -0.35494 |  | 10 | 0.013174
10 | YES | 999.208 | 0.27827 | 0.0030413 | 0.024274 | -0.34176 | -0.36604 |  | 10 | 0.024274
10 | YES | 799.271 | 0.35488 | 0.0025987 | -0.033116 | -0.34176 | -0.30865 |  | 10 | -0.033116
10 | YES | 599.385 | 0.54376 | 0.0024845 | -0.023237 | -0.34176 | -0.31853 |  | 10 | -0.023237
10 | YES | 402.41 | 1.0239 | 0.0073432 | -0.07805 | -0.34176 | -0.26371 |  | 10 | -0.07805
10 | YES | 301.932 | 1.4423 | 0.0027367 | -0.10473 | -0.34176 | -0.23703 |  | 10 | -0.10473
10 | YES | 250.688 | 1.6911 | 0.006699 | -0.12986 | -0.34176 | -0.2119 |  | 10 | -0.12986
10 | YES | 199.64 | 2.0769 | 0.0036755 | -0.15214 | -0.34176 | -0.18962 |  | 10 | -0.15214
10 | YES | 175.483 | 2.3036 | 0.0074563 | -0.15136 | -0.34176 | -0.19041 |  | 10 | -0.15136
10 | YES | 150.246 | 2.5619 | 0.0096489 | -0.20505 | -0.34176 | -0.13672 |  | 10 | -0.20505
10 | YES | 124.167 | 3.7627 | 0.032554 | -0.40133 | -0.34176 | 0.059563 |  | 10 | -0.40133
10 | YES | 99.5251 | 5.319 | 0.016428 | -0.39605 | -0.34176 | 0.054289 |  | 10 | -0.39605
10 | YES | 74.3439 | 6.1191 | 0.011027 | -0.46193 | -0.34176 | 0.12017 |  | 10 | -0.46193
10 | YES | 49.3314 | 6.4523 | 0.025608 | -0.49372 | -0.34176 | 0.15196 |  | 10 | -0.49372
10 | YES | 40.9976 | 6.5152 | 0.035805 | -0.48879 | -0.34176 | 0.14703 |  | 10 | -0.48879
10 | YES | 29.6041 | 5.6367 | 0.0091063 | -0.44334 | -0.34176 | 0.10157 |  | 10 | -0.44334
10 | YES | 22.1111 | 5.612 | 0.025038 | -0.45695 | -0.34176 | 0.11518 |  | 10 | -0.45695
10 | YES | 22.1111 | 5.612 | 0.025038 | -0.45495 | -0.34176 | 0.11318 |  | 10 | -0.45495
10 | YES | 10.6242 | 5.5937 | 0.019454 | -0.41628 | -0.34176 | 0.074513 |  | 10 | -0.41628
10 | YES | 4.5235 | 5.5706 | 0.032389 | -0.42942 | -0.34176 | 0.087651 |  | 10 | -0.42942
10 | YES | 3.0114 | 5.5692 | 0.03293 | -0.45884 | -0.34176 | 0.11708 |  | 10 | -0.45884
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (3)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.20063 |  | X | CALC Y
m= | 0.021908 | 0.010932 |  | Std Dev of Resid |  | 0.11138 |  | 0.044527 | 0.055782
b= | 0.054806 | 0.035105 |  | Resid Sum of Squares |  | 0.1985 |  | 6.5613 | 0.19855
# of Values Extracted = |  |  | 18
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 220.267 | 0.050766 | 0.001127 | 0.050766 | 0.055918 | 0.0051522 |  | 0.050766 | 0.050766
1 | YES | 200.729 | 0.047012 | 0.0015261 | 0.047012 | 0.055836 | 0.0088238 |  | 0.047012 | 0.047012
1 | YES | 176.236 | 0.044834 | 0.00091597 | 0.044834 | 0.055788 | 0.010954 |  | 0.044834 | 0.044834
1 | YES | 151.113 | 0.044527 | 0.00097484 | 0.044527 | 0.055782 | 0.011255 |  | 0.044527 | 0.044527
1 | YES | 125.529 | 0.21221 | 0.0015754 | 0.069208 | 0.059455 | -0.0097523 |  | 0.21221 | 0.069208
1 | YES | 100.564 | 0.42818 | 0.018345 | 0.10618 | 0.064187 | -0.041996 |  | 0.42818 | 0.10618
1 | YES | 49.5756 | 2.8843 | 0.0062989 | 0.21328 | 0.11799 | -0.095284 |  | 2.8843 | 0.21328
1 | YES | 25.5956 | 3.6486 | 0.0090746 | 0.13656 | 0.13474 | -0.0018179 |  | 3.6486 | 0.13656
1 | YES | 10.6371 | 5.3874 | 0.048179 | -0.015605 | 0.17283 | 0.18844 |  | 5.3874 | -0.015605
1 | YES | 5.9689 | 6.5613 | 0.022672 | -0.029714 | 0.19855 | 0.22826 |  | 6.5613 | -0.029714
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.054886 | 0.056206 | 0.0013199 |  | 0.063886 | 0.054886
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.063886 | 0.056206 | -0.0076801 |  | 0.063886 | 0.063886
2 | YES | 98.3678 | 0.51929 | 0.0053849 | 0.0052876 | 0.066183 | 0.060895 |  | 0.51929 | 0.0052876
2 | YES | 98.3678 | 0.51929 | 0.0053849 | -0.030712 | 0.066183 | 0.096895 |  | 0.51929 | -0.030712
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.20366 | 0.11643 | -0.087237 |  | 2.8127 | 0.20366
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.17966 | 0.11643 | -0.063237 |  | 2.8127 | 0.17966
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.22272 | 0.18923 | -0.033497 |  | 6.1357 | 0.22272
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.46072 | 0.18923 | -0.2715 |  | 6.1357 | 0.46072
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (4)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.97977 |  | X | CALC Y
m= | -0.068879 | 0.0024003 |  | Std Dev of Resid |  | 0.0206 |  | 0.32903 | -0.049035
b= | -0.026372 | 0.00865 |  | Resid Sum of Squares |  | 0.0072138 |  | 6.4631 | -0.47154
# of Values Extracted = |  |  | 24
# of Values Used in Fit = |  |  | 19
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
8 | YES | 4315.72 | 3.0101 | 0.0067191 | -0.21489 | -0.23371 | -0.018815 |  | 3.0101 | -0.21489
8 | YES | 3998.96 | 2.9411 | 0.007228 | -0.25786 | -0.22896 | 0.028907 |  | 2.9411 | -0.25786
8 | YES | 3501.9 | 2.7317 | 0.0053451 | -0.22328 | -0.21453 | 0.0087506 |  | 2.7317 | -0.22328
8 | YES | 3000.3 | 2.3371 | 0.0058576 | -0.20088 | -0.18735 | 0.013532 |  | 2.3371 | -0.20088
8 | YES | 2500.89 | 1.7898 | 0.004994 | -0.16321 | -0.14965 | 0.013555 |  | 1.7898 | -0.16321
8 | YES | 2003.55 | 1.1834 | 0.0039803 | -0.11965 | -0.10788 | 0.011764 |  | 1.1834 | -0.11965
8 | NO | 1501.95 | 0.49161 | 0.0036682 | -0.00739 | -0.060234 | -0.052844 |  | 0.32903 | -0.012967
8 | YES | 1250.92 | 0.32903 | 0.0029901 | -0.012967 | -0.049035 | -0.036069 |  | 0.35476 | -0.023241
8 | YES | 999.715 | 0.35476 | 0.0024273 | -0.023241 | -0.050807 | -0.027567 |  | 0.56649 | -0.060506
8 | NO | 802.475 | 0.4331 | 0.0019361 | -0.6209 | -0.056204 | 0.5647 |  | 1.0657 | -0.11028
8 | YES | 599.257 | 0.56649 | 0.0024068 | -0.060506 | -0.065392 | -0.0048855 |  | 1.7515 | -0.14347
8 | YES | 399.823 | 1.0657 | 0.0028522 | -0.11028 | -0.099778 | 0.0105 |  | 2.5384 | -0.21861
8 | YES | 297.593 | 1.7515 | 0.0051562 | -0.14347 | -0.14702 | -0.0035444 |  | 3.4453 | -0.29367
8 | YES | 247.972 | 2.5384 | 0.0075519 | -0.21861 | -0.20121 | 0.017391 |  | 4.1073 | -0.30468
8 | YES | 200.803 | 3.4453 | 0.013674 | -0.29367 | -0.26368 | 0.029984 |  | 5.0372 | -0.36982
8 | YES | 174.613 | 4.1073 | 0.0079549 | -0.30468 | -0.30928 | -0.004599 |  | 6.0571 | -0.46995
8 | YES | 149.88 | 5.0372 | 0.018226 | -0.36982 | -0.37333 | -0.0035083 |  | 6.4631 | -0.44993
8 | YES | 124.694 | 6.0571 | 0.013726 | -0.46995 | -0.44358 | 0.026376 |  | 5.8293 | -0.41173
8 | NO | 100.56 | 6.4907 | 0.014136 | -0.43729 | -0.47345 | -0.036155 |  | 5.8095 | -0.40253
8 | NO | 73.8024 | 6.4701 | 0.01843 | -0.43592 | -0.47203 | -0.036111
8 | YES | 50.6964 | 6.4631 | 0.0187 | -0.44993 | -0.47154 | -0.021615
8 | YES | 23.8606 | 5.8293 | 0.013286 | -0.41173 | -0.42789 | -0.016159
8 | YES | 11.107 | 5.8095 | 0.018044 | -0.40253 | -0.42653 | -0.023998
8 | NO | 5.5023 | 5.8033 | 0.018351 | -0.57167 | -0.4261 | 0.14557
YES
YES
YES
YES
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (5)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.9788 |  | X | CALC Y
m= | -0.079831 | 0.0026269 |  | Std Dev of Resid |  | 0.026318 |  | 0.31054 | -0.030465
b= | -0.0056741 | 0.0097087 |  | Resid Sum of Squares |  | 0.013852 |  | 6.4517 | -0.52073
# of Values Extracted = |  |  | 24
# of Values Used in Fit = |  |  | 22
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
9 | NO | 4024.09 | 2.7851 | 0.0073442 | -0.39191 | -0.22801 | 0.16389 |  | 2.6327 | -0.24833
9 | YES | 3498.86 | 2.6327 | 0.0069541 | -0.24833 | -0.21584 | 0.032485 |  | 2.6327 | -0.20433
9 | YES | 3498.86 | 2.6327 | 0.0069541 | -0.20433 | -0.21584 | -0.011515 |  | 2.3213 | -0.16272
9 | YES | 2998.76 | 2.3213 | 0.0052467 | -0.16272 | -0.19099 | -0.028267 |  | 1.794 | -0.15396
9 | YES | 2501.18 | 1.794 | 0.0040669 | -0.15396 | -0.14889 | 0.0050631 |  | 1.2011 | -0.077852
9 | YES | 1999.9 | 1.2011 | 0.004018 | -0.077852 | -0.10156 | -0.023712 |  | 0.508 | -0.036
9 | YES | 1501.12 | 0.508 | 0.0028358 | -0.036 | -0.046228 | -0.010228 |  | 0.33658 | -0.0034205
9 | YES | 1252.06 | 0.33658 | 0.0031631 | -0.0034205 | -0.032544 | -0.029123 |  | 0.31054 | -0.010465
9 | YES | 1001.45 | 0.31054 | 0.0026972 | -0.010465 | -0.030465 | -0.02 |  | 0.42635 | -0.044647
9 | YES | 801.223 | 0.42635 | 0.0021842 | -0.044647 | -0.03971 | 0.0049368 |  | 0.5665 | -0.033502
9 | YES | 601.026 | 0.5665 | 0.002174 | -0.033502 | -0.050898 | -0.017396 |  | 1.1597 | -0.13529
9 | YES | 399.824 | 1.1597 | 0.0042208 | -0.13529 | -0.098255 | 0.037035 |  | 1.884 | -0.20196
9 | YES | 300.177 | 1.884 | 0.0040253 | -0.20196 | -0.15608 | 0.045883 |  | 2.4496 | -0.21636
9 | YES | 250.475 | 2.4496 | 0.0055907 | -0.21636 | -0.20123 | 0.015125 |  | 3.6022 | -0.31876
9 | YES | 198.379 | 3.6022 | 0.0062745 | -0.31876 | -0.29325 | 0.025509 |  | 4.0755 | -0.31946
9 | YES | 175.41 | 4.0755 | 0.013949 | -0.31946 | -0.33103 | -0.01157 |  | 4.9887 | -0.44525
9 | YES | 150.732 | 4.9887 | 0.043861 | -0.44525 | -0.40393 | 0.041316 |  | 5.7935 | -0.49652
9 | YES | 124.125 | 5.7935 | 0.028387 | -0.49652 | -0.46818 | 0.028343 |  | 6.2541 | -0.47289
9 | YES | 99.4924 | 6.2541 | 0.018472 | -0.47289 | -0.50495 | -0.032057 |  | 6.4517 | -0.48427
9 | NO | 74.7791 | 6.3995 | 0.023285 | -0.44349 | -0.51656 | -0.073066 |  | 5.6485 | -0.45955
9 | YES | 49.3064 | 6.4517 | 0.029448 | -0.48427 | -0.52073 | -0.03646 |  | 5.6612 | -0.44281
9 | YES | 24.7034 | 5.6485 | 0.011864 | -0.45955 | -0.4566 | 0.002954 |  | 5.6572 | -0.45378
9 | YES | 9.7331 | 5.6612 | 0.021845 | -0.44281 | -0.45762 | -0.014802
9 | YES | 5.1962 | 5.6572 | 0.024126 | -0.45378 | -0.4573 | -0.0035182
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (6)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.99347 |  | X | CALC Y
m= | -0.079063 | 0.0015106 |  | Std Dev of Resid |  | 0.015472 |  | 0.35488 | -0.018943
b= | 0.0091149 | 0.0064289 |  | Resid Sum of Squares |  | 0.0043091 |  | 6.5152 | -0.506
# of Values Extracted = |  |  | 23
# of Values Used in Fit = |  |  | 20
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
10 | YES | 1998.83 | 1.1924 | 0.0047216 | -0.097585 | -0.085161 | 0.012424 |  | 1.1924 | -0.097585
10 | YES | 1518.6 | 0.54876 | 0.0029639 | -0.014236 | -0.034272 | -0.020036 |  | 0.54876 | -0.014236
10 | NO | 1249.67 | 0.32417 | 0.0025906 | 0.013174 | -0.016515 | -0.02969 |  | 0.35488 | -0.033116
10 | NO | 999.208 | 0.27827 | 0.0030413 | 0.024274 | -0.012886 | -0.03716 |  | 0.54376 | -0.023237
10 | YES | 799.271 | 0.35488 | 0.0025987 | -0.033116 | -0.018943 | 0.014173 |  | 1.0239 | -0.07805
10 | YES | 599.385 | 0.54376 | 0.0024845 | -0.023237 | -0.033877 | -0.01064 |  | 1.4423 | -0.10473
10 | YES | 402.41 | 1.0239 | 0.0073432 | -0.07805 | -0.071842 | 0.0062082 |  | 1.6911 | -0.12986
10 | YES | 301.932 | 1.4423 | 0.0027367 | -0.10473 | -0.10492 | -0.00018528 |  | 2.0769 | -0.15214
10 | YES | 250.688 | 1.6911 | 0.006699 | -0.12986 | -0.12459 | 0.0052713 |  | 2.3036 | -0.15136
10 | YES | 199.64 | 2.0769 | 0.0036755 | -0.15214 | -0.15509 | -0.0029479 |  | 2.5619 | -0.20505
10 | YES | 175.483 | 2.3036 | 0.0074563 | -0.15136 | -0.17302 | -0.021661 |  | 5.319 | -0.39605
10 | YES | 150.246 | 2.5619 | 0.0096489 | -0.20505 | -0.19344 | 0.011605 |  | 6.1191 | -0.46193
10 | NO | 124.167 | 3.7627 | 0.032554 | -0.40133 | -0.28837 | 0.11295 |  | 6.4523 | -0.49372
10 | YES | 99.5251 | 5.319 | 0.016428 | -0.39605 | -0.41142 | -0.015363 |  | 6.5152 | -0.48879
10 | YES | 74.3439 | 6.1191 | 0.011027 | -0.46193 | -0.47468 | -0.012743 |  | 5.6367 | -0.44334
10 | YES | 49.3314 | 6.4523 | 0.025608 | -0.49372 | -0.50102 | -0.0072997 |  | 5.612 | -0.45695
10 | YES | 40.9976 | 6.5152 | 0.035805 | -0.48879 | -0.506 | -0.017204 |  | 5.612 | -0.45495
10 | YES | 29.6041 | 5.6367 | 0.0091063 | -0.44334 | -0.43654 | 0.0067997 |  | 5.5937 | -0.41628
10 | YES | 22.1111 | 5.612 | 0.025038 | -0.45695 | -0.43459 | 0.022357 |  | 5.5706 | -0.42942
10 | YES | 22.1111 | 5.612 | 0.025038 | -0.45495 | -0.43459 | 0.020357 |  | 5.5692 | -0.45884
10 | YES | 10.6242 | 5.5937 | 0.019454 | -0.41628 | -0.43314 | -0.016863
10 | YES | 4.5235 | 5.5706 | 0.032389 | -0.42942 | -0.43131 | -0.0018958
10 | YES | 3.0114 | 5.5692 | 0.03293 | -0.45884 | -0.4312 | 0.027642
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (7)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.95959 |  | X | CALC Y
m= | -0.063433 | 0.0029863 |  | Std Dev of Resid |  | 0.030723 |  | 0.25587 | -0.037501
b= | -0.02127 | 0.011711 |  | Resid Sum of Squares |  | 0.017935 |  | 6.5239 | -0.4351
# of Values Extracted = |  |  | 23
# of Values Used in Fit = |  |  | 21
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
7 | YES | 3500.19 | 2.5276 | 0.0043321 | -0.22339 | -0.1816 | 0.041785 |  | 2.5276 | -0.22339
7 | YES | 2999.58 | 2.2661 | 0.0053773 | -0.19986 | -0.16502 | 0.034839 |  | 2.2661 | -0.19986
7 | YES | 2497.23 | 1.7816 | 0.0045839 | -0.12635 | -0.13429 | -0.0079332 |  | 1.7816 | -0.12635
7 | YES | 2496.9 | 1.7818 | 0.0046856 | -0.12515 | -0.1343 | -0.0091499 |  | 1.7818 | -0.12515
7 | NO | 2001.58 | 1.175 | 0.0032203 | -0.044012 | -0.095804 | -0.051791 |  | 0.47704 | -0.029963
7 | YES | 1501.9 | 0.47704 | 0.0031055 | -0.029963 | -0.05153 | -0.021568 |  | 0.30083 | -0.010174
7 | YES | 1251.29 | 0.30083 | 0.0028858 | -0.010174 | -0.040353 | -0.030178 |  | 0.25587 | -0.029133
7 | YES | 1000.7 | 0.25587 | 0.0025214 | -0.029133 | -0.037501 | -0.008368 |  | 0.36527 | -0.022726
7 | YES | 801.449 | 0.36527 | 0.0024324 | -0.022726 | -0.044441 | -0.021715 |  | 0.57156 | -0.047439
7 | YES | 599.478 | 0.57156 | 0.002331 | -0.047439 | -0.057526 | -0.010087 |  | 1.396 | -0.11402
7 | YES | 400.235 | 1.396 | 0.0041091 | -0.11402 | -0.10982 | 0.0041925 |  | 2.3084 | -0.22357
7 | YES | 300.436 | 2.3084 | 0.0038714 | -0.22357 | -0.1677 | 0.055866 |  | 2.9682 | -0.21684
7 | YES | 250.34 | 2.9682 | 0.0091679 | -0.21684 | -0.20955 | 0.0072916 |  | 4.1863 | -0.33474
7 | YES | 200.034 | 4.1863 | 0.02518 | -0.33474 | -0.28682 | 0.047919 |  | 4.8064 | -0.34155
7 | YES | 176.444 | 4.8064 | 0.0094576 | -0.34155 | -0.32616 | 0.015387 |  | 5.4656 | -0.35639
7 | YES | 149.804 | 5.4656 | 0.010897 | -0.35639 | -0.36797 | -0.011586 |  | 6.2031 | -0.40091
7 | YES | 126.568 | 6.2031 | 0.01604 | -0.40091 | -0.41475 | -0.013848 |  | 6.4238 | -0.38922
7 | YES | 101.691 | 6.4238 | 0.0076469 | -0.38922 | -0.42875 | -0.039528 |  | 6.5239 | -0.48108
7 | YES | 74.5608 | 6.5239 | 0.024834 | -0.48108 | -0.4351 | 0.045979 |  | 5.6457 | -0.37828
7 | NO | 49.9787 | 6.593 | 0.018445 | -0.36099 | -0.43949 | -0.078495 |  | 5.6245 | -0.34247
7 | YES | 25.418 | 5.6457 | 0.022196 | -0.37828 | -0.3794 | -0.0011188 |  | 5.6415 | -0.33653
7 | YES | 10.5048 | 5.6245 | 0.016742 | -0.34247 | -0.37805 | -0.035588
7 | YES | 6.1822 | 5.6415 | 0.017912 | -0.33653 | -0.37913 | -0.042592
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (8)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.95672 |  | X | CALC Y
m= | -0.049835 | 0.0026498 |  | Std Dev of Resid |  | 0.024042 |  | 0.22771 | -0.046427
b= | -0.035079 | 0.0096731 |  | Resid Sum of Squares |  | 0.0092483 |  | 6.7227 | -0.3701
# of Values Extracted = |  |  | 23
# of Values Used in Fit = |  |  | 18
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
6 | YES | 3276.19 | 2.1742 | 0.0052468 | -0.16977 | -0.14343 | 0.026336 |  | 2.1742 | -0.16977
6 | YES | 3000.48 | 2.1445 | 0.0049557 | -0.14653 | -0.14195 | 0.0045785 |  | 2.1445 | -0.14653
6 | NO | 2501.59 | 1.7573 | 0.0045462 | -0.22471 | -0.12265 | 0.10206 |  | 1.7573 | -0.10471
6 | YES | 2501.59 | 1.7573 | 0.0045462 | -0.10471 | -0.12265 | -0.01794 |  | 1.2232 | -0.074834
6 | YES | 2000.27 | 1.2232 | 0.0036682 | -0.074834 | -0.096036 | -0.021202 |  | 0.38709 | -0.057913
6 | NO | 1501.16 | 0.60927 | 0.0032809 | -0.01373 | -0.065442 | -0.051712 |  | 0.22771 | -0.011286
6 | YES | 1250.57 | 0.38709 | 0.0031696 | -0.057913 | -0.05437 | 0.0035432 |  | 0.26251 | -0.04849
6 | YES | 1001.93 | 0.22771 | 0.0019075 | -0.011286 | -0.046427 | -0.035141 |  | 0.51929 | -0.064705
6 | YES | 801.143 | 0.26251 | 0.0024751 | -0.04849 | -0.048161 | 0.00032813 |  | 1.3473 | -0.13473
6 | YES | 600.801 | 0.51929 | 0.0018098 | -0.064705 | -0.060958 | 0.0037471 |  | 2.5586 | -0.17243
6 | YES | 400.283 | 1.3473 | 0.0028976 | -0.13473 | -0.10222 | 0.032514 |  | 3.273 | -0.22396
6 | YES | 301.716 | 2.5586 | 0.0042794 | -0.17243 | -0.16259 | 0.0098459 |  | 4.0577 | -0.23326
6 | YES | 249.814 | 3.273 | 0.0030968 | -0.22396 | -0.19819 | 0.025767 |  | 4.3958 | -0.24818
6 | YES | 200.271 | 4.0577 | 0.0056607 | -0.23326 | -0.2373 | -0.0040349 |  | 4.6352 | -0.25283
6 | YES | 176.597 | 4.3958 | 0.005897 | -0.24818 | -0.25414 | -0.0059611 |  | 5.0843 | -0.26169
6 | YES | 150.549 | 4.6352 | 0.0091057 | -0.25283 | -0.26607 | -0.013238 |  | 6.4835 | -0.3485
6 | YES | 124.879 | 5.0843 | 0.0094488 | -0.26169 | -0.28845 | -0.026762 |  | 6.7227 | -0.3433
6 | NO | 100.671 | 5.8131 | 0.0097689 | -0.26988 | -0.32477 | -0.054895 |  | 5.9998 | -0.38818
6 | YES | 75.8548 | 6.4835 | 0.015863 | -0.3485 | -0.35818 | -0.0096812
6 | YES | 50.9557 | 6.7227 | 0.016748 | -0.3433 | -0.3701 | -0.026805
6 | YES | 25.6344 | 5.9998 | 0.01886 | -0.38818 | -0.33408 | 0.054104
6 | NO | 9.8456 | 2.3884 | 0.025795 | -3.5306 | -0.15411 | 3.3765
6 | NO | 5.8409 | 2.106 | 0.0078225 | -3.82 | -0.14003 | 3.68
YES
YES
NO
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (9)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.93395 |  | X | CALC Y
m= | -0.048323 | 0.0033181 |  | Std Dev of Resid |  | 0.027592 |  | 0.21613 | -0.031673
b= | -0.021229 | 0.010689 |  | Resid Sum of Squares |  | 0.011419 |  | 5.7981 | -0.30141
# of Values Extracted = |  |  | 20
# of Values Used in Fit = |  |  | 17
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
5 | NO | 2003.65 | 1.2821 | 0.003906 | -0.1689 | -0.083184 | 0.085716 |  | 0.61915 | -0.06185
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06185 | -0.051148 | 0.010702 |  | 0.61915 | -0.06285
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06285 | -0.051148 | 0.011702 |  | 0.33626 | -0.018739
5 | YES | 1250.17 | 0.33626 | 0.0025662 | -0.018739 | -0.037478 | -0.01874 |  | 0.21613 | -0.020871
5 | YES | 1001.62 | 0.21613 | 0.0024958 | -0.020871 | -0.031673 | -0.010802 |  | 0.37322 | -0.054776
5 | NO | 800.329 | 0.21353 | 0.0022177 | -0.79147 | -0.031547 | 0.75993 |  | 1.0396 | -0.060402
5 | YES | 599.855 | 0.37322 | 0.0022839 | -0.054776 | -0.039264 | 0.015511 |  | 1.8098 | -0.10022
5 | YES | 401.219 | 1.0396 | 0.0021928 | -0.060402 | -0.071466 | -0.011064 |  | 1.9105 | -0.12346
5 | YES | 299.827 | 1.8098 | 0.0070129 | -0.10022 | -0.10868 | -0.0084638 |  | 1.9582 | -0.088759
5 | YES | 249.318 | 1.9105 | 0.0063301 | -0.12346 | -0.11355 | 0.0099062 |  | 2.0371 | -0.10989
5 | YES | 200.212 | 1.9582 | 0.0035649 | -0.088759 | -0.11586 | -0.027099 |  | 2.1758 | -0.15724
5 | YES | 173.88 | 2.0371 | 0.010025 | -0.10989 | -0.11967 | -0.0097751 |  | 2.6831 | -0.1639
5 | YES | 150.707 | 2.1758 | 0.0089689 | -0.15724 | -0.12637 | 0.030872 |  | 3.998 | -0.21401
5 | YES | 124.044 | 2.6831 | 0.012874 | -0.1639 | -0.15089 | 0.013018 |  | 5.7981 | -0.29987
5 | YES | 101.474 | 3.998 | 0.011501 | -0.21401 | -0.21443 | -0.00041753 |  | 5.7126 | -0.3694
5 | YES | 75.2867 | 5.7981 | 0.037336 | -0.29987 | -0.30141 | -0.0015484 |  | 5.7041 | -0.26693
5 | NO | 49.8081 | 6.2679 | 0.011322 | -0.21206 | -0.32412 | -0.11206 |  | 5.7156 | -0.25144
5 | YES | 25.2931 | 5.7126 | 0.015859 | -0.3694 | -0.29728 | 0.072121
5 | YES | 10.3099 | 5.7041 | 0.026635 | -0.26693 | -0.29687 | -0.029939
5 | YES | 5.3529 | 5.7156 | 0.02113 | -0.25144 | -0.29742 | -0.045984
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (10)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.97532 |  | X | CALC Y
m= | -0.046262 | 0.0019667 |  | Std Dev of Resid |  | 0.016 |  | 0.21223 | -0.024899
b= | -0.015081 | 0.0057718 |  | Resid Sum of Squares |  | 0.003584 |  | 6.5395 | -0.31761
# of Values Extracted = |  |  | 18
# of Values Used in Fit = |  |  | 16
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
4 | YES | 1300.62 | 0.49522 | 0.0021859 | -0.054784 | -0.037991 | 0.016794 |  | 0.49522 | -0.054784
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.032976 | -0.034188 | -0.001212 |  | 0.41302 | -0.032976
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.026976 | -0.034188 | -0.007212 |  | 0.41302 | -0.026976
4 | YES | 1000.54 | 0.23775 | 0.0018914 | -0.013245 | -0.02608 | -0.012835 |  | 0.23775 | -0.013245
4 | YES | 801.075 | 0.21223 | 0.0020806 | -0.024768 | -0.024899 | -0.00013157 |  | 0.21223 | -0.024768
4 | YES | 600.918 | 0.42334 | 0.0028851 | -0.040664 | -0.034665 | 0.0059987 |  | 0.42334 | -0.040664
4 | YES | 400.248 | 0.99714 | 0.001509 | -0.072863 | -0.06121 | 0.011653 |  | 0.99714 | -0.072863
4 | YES | 301.823 | 1.2875 | 0.002412 | -0.075485 | -0.074643 | 0.00084209 |  | 1.2875 | -0.075485
4 | YES | 201.201 | 1.6274 | 0.00573 | -0.1076 | -0.090367 | 0.01723 |  | 1.6274 | -0.1076
4 | YES | 174.556 | 1.8759 | 0.0068785 | -0.076058 | -0.10187 | -0.025807 |  | 1.8759 | -0.076058
4 | YES | 150.455 | 2.2244 | 0.011745 | -0.14162 | -0.11798 | 0.023637 |  | 2.2244 | -0.14162
4 | YES | 125.295 | 2.7552 | 0.0045848 | -0.13478 | -0.14254 | -0.0077582 |  | 2.7552 | -0.13478
4 | YES | 100.512 | 3.3938 | 0.0073132 | -0.14018 | -0.17208 | -0.031903 |  | 3.3938 | -0.14018
4 | YES | 75.7457 | 4.6783 | 0.018784 | -0.22371 | -0.23151 | -0.0077968 |  | 4.6783 | -0.22371
4 | NO | 51.446 | 6.1723 | 0.010765 | -0.20073 | -0.30062 | -0.099886 |  | 6.2779 | -0.32107
4 | YES | 24.7249 | 6.2779 | 0.017082 | -0.32107 | -0.30551 | 0.015565 |  | 6.5395 | -0.32054
4 | NO | 10.3947 | 6.4598 | 0.044508 | -0.26224 | -0.31392 | -0.051684
4 | YES | 5.239 | 6.5395 | 0.016358 | -0.32054 | -0.31761 | 0.0029362
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (11)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.77133 |  | X | CALC Y
m= | 0.054101 | 0.012026 |  | Std Dev of Resid |  | 0.081774 |  | 0.063886 | 0.019554
b= | 0.016098 | 0.040707 |  | Resid Sum of Squares |  | 0.040122 |  | 6.1357 | 0.34805
# of Values Extracted = |  |  | 25
# of Values Used in Fit = |  |  | 8
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | NO | 220.267 | 0.050766 | 0.001127 | 0.050766 | 0.018844 | -0.031922 |  | 0.063886 | 0.054886
1 | NO | 200.729 | 0.047012 | 0.0015261 | 0.047012 | 0.018641 | -0.028371 |  | 0.063886 | 0.063886
1 | NO | 176.236 | 0.044834 | 0.00091597 | 0.044834 | 0.018523 | -0.026311 |  | 0.51929 | 0.0052876
1 | NO | 151.113 | 0.044527 | 0.00097484 | 0.044527 | 0.018507 | -0.02602 |  | 0.51929 | -0.030712
1 | NO | 125.529 | 0.21221 | 0.0015754 | 0.069208 | 0.027579 | -0.041629 |  | 2.8127 | 0.20366
1 | NO | 100.564 | 0.42818 | 0.018345 | 0.10618 | 0.039263 | -0.06692 |  | 2.8127 | 0.17966
1 | NO | 49.5756 | 2.8843 | 0.0062989 | 0.21328 | 0.17214 | -0.041138 |  | 6.1357 | 0.22272
1 | NO | 25.5956 | 3.6486 | 0.0090746 | 0.13656 | 0.21349 | 0.076933 |  | 6.1357 | 0.46072
1 | NO | 10.6371 | 5.3874 | 0.048179 | -0.015605 | 0.30756 | 0.32317
1 | NO | 5.9689 | 6.5613 | 0.022672 | -0.029714 | 0.37107 | 0.40078
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.054886 | 0.019554 | -0.035332
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.063886 | 0.019554 | -0.044332
2 | YES | 98.3678 | 0.51929 | 0.0053849 | 0.0052876 | 0.044192 | 0.038904
2 | YES | 98.3678 | 0.51929 | 0.0053849 | -0.030712 | 0.044192 | 0.074904
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.20366 | 0.16827 | -0.035396
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.17966 | 0.16827 | -0.011396
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.22272 | 0.34805 | 0.12532
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.46072 | 0.34805 | -0.11268
3 | NO | 106.372 | 1.5767 | 0.011722 | 0.025656 | 0.1014 | 0.075741
3 | NO | 76.1436 | 2.128 | 0.0067447 | -0.064981 | 0.13123 | 0.19621
3 | NO | 76.1436 | 2.128 | 0.0067447 | -0.064981 | 0.13123 | 0.19621
3 | NO | 51.4045 | 2.6019 | 0.0058068 | -0.094104 | 0.15686 | 0.25097
3 | NO | 26.3033 | 4.2925 | 0.070084 | -0.32851 | 0.24833 | 0.57684
3 | NO | 11.1766 | 5.5508 | 0.030052 | -0.2702 | 0.3164 | 0.58661
3 | NO | 5.5253 | 5.7243 | 0.011576 | -0.22071 | 0.32579 | 0.5465
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (12)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.92569 |  | X | CALC Y
m= | -0.057146 | 0.0080957 |  | Std Dev of Resid |  | 0.033523 |  | 1.5767 | -0.017267
b= | 0.072832 | 0.029909 |  | Resid Sum of Squares |  | 0.004495 |  | 5.7243 | -0.25429
# of Values Extracted = |  |  | 7
# of Values Used in Fit = |  |  | 6
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
3 | YES | 106.372 | 1.5767 | 0.011722 | 0.025656 | -0.017267 | -0.042923 |  | 1.5767 | 0.025656
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.048775 | 0.016206 |  | 2.128 | -0.064981
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.048775 | 0.016206 |  | 2.128 | -0.064981
3 | YES | 51.4045 | 2.6019 | 0.0058068 | -0.094104 | -0.075855 | 0.018249 |  | 2.6019 | -0.094104
3 | NO | 26.3033 | 4.2925 | 0.070084 | -0.32851 | -0.17246 | 0.15605 |  | 5.5508 | -0.2702
3 | YES | 11.1766 | 5.5508 | 0.030052 | -0.2702 | -0.24437 | 0.025833 |  | 5.7243 | -0.22071
3 | YES | 5.5253 | 5.7243 | 0.011576 | -0.22071 | -0.25429 | -0.033572
YES
YES
YES
YES
NO
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (13)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.059233 |  | X | CALC Y
m= | -0.046064 | 0.015141 |  | Std Dev of Resid |  | 0.43666 |  | 0.044527 | -0.093677
b= | -0.091626 | 0.060238 |  | Resid Sum of Squares |  | 28.0286 |  | 6.7227 | -0.4013
# of Values Extracted = |  |  | 180
# of Values Used in Fit = |  |  | 149
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 220.267 | 0.050766 | 0.001127 | 0.050766 | -0.093964 | -0.14473 |  | 0.050766 | 0.050766
1 | YES | 200.729 | 0.047012 | 0.0015261 | 0.047012 | -0.093791 | -0.1408 |  | 0.047012 | 0.047012
1 | YES | 176.236 | 0.044834 | 0.00091597 | 0.044834 | -0.093691 | -0.13852 |  | 0.044834 | 0.044834
1 | YES | 151.113 | 0.044527 | 0.00097484 | 0.044527 | -0.093677 | -0.1382 |  | 0.044527 | 0.044527
1 | YES | 125.529 | 0.21221 | 0.0015754 | 0.069208 | -0.1014 | -0.17061 |  | 0.21221 | 0.069208
1 | YES | 100.564 | 0.42818 | 0.018345 | 0.10618 | -0.11135 | -0.21753 |  | 0.42818 | 0.10618
1 | YES | 49.5756 | 2.8843 | 0.0062989 | 0.21328 | -0.22449 | -0.43777 |  | 2.8843 | 0.21328
1 | YES | 25.5956 | 3.6486 | 0.0090746 | 0.13656 | -0.25969 | -0.39625 |  | 3.6486 | 0.13656
1 | YES | 10.6371 | 5.3874 | 0.048179 | -0.015605 | -0.33979 | -0.32419 |  | 5.3874 | -0.015605
1 | YES | 5.9689 | 6.5613 | 0.022672 | -0.029714 | -0.39387 | -0.36415 |  | 6.5613 | -0.029714
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.054886 | -0.094569 | -0.14945 |  | 0.063886 | 0.054886
2 | YES | 200.987 | 0.063886 | 0.0017874 | 0.063886 | -0.094569 | -0.15845 |  | 0.063886 | 0.063886
2 | YES | 98.3678 | 0.51929 | 0.0053849 | 0.0052876 | -0.11555 | -0.12083 |  | 0.51929 | 0.0052876
2 | YES | 98.3678 | 0.51929 | 0.0053849 | -0.030712 | -0.11555 | -0.084834 |  | 0.51929 | -0.030712
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.20366 | -0.22119 | -0.42485 |  | 2.8127 | 0.20366
2 | YES | 51.3838 | 2.8127 | 0.075542 | 0.17966 | -0.22119 | -0.40085 |  | 2.8127 | 0.17966
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.22272 | -0.37426 | -0.59699 |  | 6.1357 | 0.22272
2 | YES | 11.1691 | 6.1357 | 0.027563 | 0.46072 | -0.37426 | -0.83498 |  | 6.1357 | 0.46072
3 | YES | 106.372 | 1.5767 | 0.011722 | 0.025656 | -0.16425 | -0.18991 |  | 1.5767 | 0.025656
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.18965 | -0.12467 |  | 2.128 | -0.064981
3 | YES | 76.1436 | 2.128 | 0.0067447 | -0.064981 | -0.18965 | -0.12467 |  | 2.128 | -0.064981
3 | YES | 51.4045 | 2.6019 | 0.0058068 | -0.094104 | -0.21148 | -0.11738 |  | 2.6019 | -0.094104
3 | YES | 26.3033 | 4.2925 | 0.070084 | -0.32851 | -0.28936 | 0.039154 |  | 4.2925 | -0.32851
3 | YES | 11.1766 | 5.5508 | 0.030052 | -0.2702 | -0.34732 | -0.077115 |  | 5.5508 | -0.2702
3 | YES | 5.5253 | 5.7243 | 0.011576 | -0.22071 | -0.35531 | -0.1346 |  | 5.7243 | -0.22071
4 | YES | 1300.62 | 0.49522 | 0.0021859 | -0.054784 | -0.11444 | -0.059653 |  | 0.49522 | -0.054784
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.032976 | -0.11065 | -0.077675 |  | 0.41302 | -0.032976
4 | YES | 1250.05 | 0.41302 | 0.0030083 | -0.026976 | -0.11065 | -0.083675 |  | 0.41302 | -0.026976
4 | YES | 1000.54 | 0.23775 | 0.0018914 | -0.013245 | -0.10258 | -0.089333 |  | 0.23775 | -0.013245
4 | YES | 801.075 | 0.21223 | 0.0020806 | -0.024768 | -0.1014 | -0.076634 |  | 0.21223 | -0.024768
4 | YES | 600.918 | 0.42334 | 0.0028851 | -0.040664 | -0.11113 | -0.070462 |  | 0.42334 | -0.040664
4 | YES | 400.248 | 0.99714 | 0.001509 | -0.072863 | -0.13756 | -0.064695 |  | 0.99714 | -0.072863
4 | YES | 301.823 | 1.2875 | 0.002412 | -0.075485 | -0.15093 | -0.075448 |  | 1.2875 | -0.075485
4 | YES | 201.201 | 1.6274 | 0.00573 | -0.1076 | -0.16659 | -0.058994 |  | 1.6274 | -0.1076
4 | YES | 174.556 | 1.8759 | 0.0068785 | -0.076058 | -0.17804 | -0.10198 |  | 1.8759 | -0.076058
4 | YES | 150.455 | 2.2244 | 0.011745 | -0.14162 | -0.19409 | -0.052468 |  | 2.2244 | -0.14162
4 | YES | 125.295 | 2.7552 | 0.0045848 | -0.13478 | -0.21854 | -0.083759 |  | 2.7552 | -0.13478
4 | YES | 100.512 | 3.3938 | 0.0073132 | -0.14018 | -0.24796 | -0.10778 |  | 3.3938 | -0.14018
4 | YES | 75.7457 | 4.6783 | 0.018784 | -0.22371 | -0.30713 | -0.083418 |  | 4.6783 | -0.22371
4 | YES | 51.446 | 6.1723 | 0.010765 | -0.20073 | -0.37595 | -0.17521 |  | 6.1723 | -0.20073
4 | YES | 24.7249 | 6.2779 | 0.017082 | -0.32107 | -0.38081 | -0.05974 |  | 6.2779 | -0.32107
4 | YES | 10.3947 | 6.4598 | 0.044508 | -0.26224 | -0.38919 | -0.12695 |  | 6.4598 | -0.26224
4 | YES | 5.239 | 6.5395 | 0.016358 | -0.32054 | -0.39286 | -0.072317 |  | 6.5395 | -0.32054
5 | NO | 2003.65 | 1.2821 | 0.003906 | -0.1689 | -0.15068 | 0.018215 |  | 0.33626 | -0.018739
5 | NO | 1500.62 | 0.61915 | 0.0025496 | -0.06185 | -0.12015 | -0.058296 |  | 0.21613 | -0.020871
5 | NO | 1500.62 | 0.61915 | 0.0025496 | -0.06285 | -0.12015 | -0.057296 |  | 0.21353 | -0.79147
5 | YES | 1250.17 | 0.33626 | 0.0025662 | -0.018739 | -0.10712 | -0.088377 |  | 0.37322 | -0.054776
5 | YES | 1001.62 | 0.21613 | 0.0024958 | -0.020871 | -0.10158 | -0.08071 |  | 1.0396 | -0.060402
5 | YES | 800.329 | 0.21353 | 0.0022177 | -0.79147 | -0.10146 | 0.69001 |  | 1.8098 | -0.10022
5 | YES | 599.855 | 0.37322 | 0.0022839 | -0.054776 | -0.10882 | -0.054042 |  | 1.9105 | -0.12346
5 | YES | 401.219 | 1.0396 | 0.0021928 | -0.060402 | -0.13951 | -0.079112 |  | 1.9582 | -0.088759
5 | YES | 299.827 | 1.8098 | 0.0070129 | -0.10022 | -0.17499 | -0.074772 |  | 2.0371 | -0.10989
5 | YES | 249.318 | 1.9105 | 0.0063301 | -0.12346 | -0.17963 | -0.056174 |  | 2.1758 | -0.15724
5 | YES | 200.212 | 1.9582 | 0.0035649 | -0.088759 | -0.18183 | -0.093071 |  | 2.6831 | -0.1639
5 | YES | 173.88 | 2.0371 | 0.010025 | -0.10989 | -0.18546 | -0.07557 |  | 3.998 | -0.21401
5 | YES | 150.707 | 2.1758 | 0.0089689 | -0.15724 | -0.19185 | -0.034609 |  | 5.7981 | -0.29987
5 | YES | 124.044 | 2.6831 | 0.012874 | -0.1639 | -0.21522 | -0.051317 |  | 6.2679 | -0.21206
5 | YES | 101.474 | 3.998 | 0.011501 | -0.21401 | -0.27579 | -0.061782 |  | 5.7126 | -0.3694
5 | YES | 75.2867 | 5.7981 | 0.037336 | -0.29987 | -0.35871 | -0.058846 |  | 5.7041 | -0.26693
5 | YES | 49.8081 | 6.2679 | 0.011322 | -0.21206 | -0.38035 | -0.16829 |  | 5.7156 | -0.25144
5 | YES | 25.2931 | 5.7126 | 0.015859 | -0.3694 | -0.35477 | 0.01463 |  | 0.38709 | -0.057913
5 | YES | 10.3099 | 5.7041 | 0.026635 | -0.26693 | -0.35438 | -0.087449 |  | 0.22771 | -0.011286
5 | YES | 5.3529 | 5.7156 | 0.02113 | -0.25144 | -0.35491 | -0.10347 |  | 0.26251 | -0.04849
6 | NO | 3276.19 | 2.1742 | 0.0052468 | -0.16977 | -0.19178 | -0.022012 |  | 0.51929 | -0.064705
6 | NO | 3000.48 | 2.1445 | 0.0049557 | -0.14653 | -0.19041 | -0.043882 |  | 1.3473 | -0.13473
6 | NO | 2501.59 | 1.7573 | 0.0045462 | -0.22471 | -0.17257 | 0.05214 |  | 2.5586 | -0.17243
6 | NO | 2501.59 | 1.7573 | 0.0045462 | -0.10471 | -0.17257 | -0.06786 |  | 3.273 | -0.22396
6 | NO | 2000.27 | 1.2232 | 0.0036682 | -0.074834 | -0.14797 | -0.073136 |  | 4.0577 | -0.23326
6 | NO | 1501.16 | 0.60927 | 0.0032809 | -0.01373 | -0.11969 | -0.10596 |  | 4.3958 | -0.24818
6 | YES | 1250.57 | 0.38709 | 0.0031696 | -0.057913 | -0.10946 | -0.051544 |  | 4.6352 | -0.25283
6 | YES | 1001.93 | 0.22771 | 0.0019075 | -0.011286 | -0.10212 | -0.090829 |  | 5.0843 | -0.26169
6 | YES | 801.143 | 0.26251 | 0.0024751 | -0.04849 | -0.10372 | -0.055228 |  | 5.8131 | -0.26988
6 | YES | 600.801 | 0.51929 | 0.0018098 | -0.064705 | -0.11555 | -0.050841 |  | 6.4835 | -0.3485
6 | YES | 400.283 | 1.3473 | 0.0028976 | -0.13473 | -0.15369 | -0.018953 |  | 6.7227 | -0.3433
6 | YES | 301.716 | 2.5586 | 0.0042794 | -0.17243 | -0.20948 | -0.037053 |  | 5.9998 | -0.38818
6 | YES | 249.814 | 3.273 | 0.0030968 | -0.22396 | -0.2424 | -0.018438 |  | 2.3884 | -3.5306
6 | YES | 200.271 | 4.0577 | 0.0056607 | -0.23326 | -0.27854 | -0.045281 |  | 2.106 | -3.82
6 | YES | 176.597 | 4.3958 | 0.005897 | -0.24818 | -0.29412 | -0.045933 |  | 0.30083 | -0.010174
6 | YES | 150.549 | 4.6352 | 0.0091057 | -0.25283 | -0.30514 | -0.052307 |  | 0.25587 | -0.029133
6 | YES | 124.879 | 5.0843 | 0.0094488 | -0.26169 | -0.32583 | -0.064138 |  | 0.36527 | -0.022726
6 | YES | 100.671 | 5.8131 | 0.0097689 | -0.26988 | -0.3594 | -0.089523 |  | 0.57156 | -0.047439
6 | YES | 75.8548 | 6.4835 | 0.015863 | -0.3485 | -0.39028 | -0.041781 |  | 1.396 | -0.11402
6 | YES | 50.9557 | 6.7227 | 0.016748 | -0.3433 | -0.4013 | -0.058003 |  | 2.3084 | -0.22357
6 | YES | 25.6344 | 5.9998 | 0.01886 | -0.38818 | -0.368 | 0.02018 |  | 2.9682 | -0.21684
6 | YES | 9.8456 | 2.3884 | 0.025795 | -3.5306 | -0.20165 | 3.3289 |  | 4.1863 | -0.33474
6 | YES | 5.8409 | 2.106 | 0.0078225 | -3.82 | -0.18864 | 3.6314 |  | 4.8064 | -0.34155
7 | NO | 3500.19 | 2.5276 | 0.0043321 | -0.22339 | -0.20806 | 0.015332 |  | 5.4656 | -0.35639
7 | NO | 2999.58 | 2.2661 | 0.0053773 | -0.19986 | -0.19601 | 0.0038443 |  | 6.2031 | -0.40091
7 | NO | 2497.23 | 1.7816 | 0.0045839 | -0.12635 | -0.1737 | -0.047343 |  | 6.4238 | -0.38922
7 | NO | 2496.9 | 1.7818 | 0.0046856 | -0.12515 | -0.17371 | -0.048556 |  | 6.5239 | -0.48108
7 | NO | 2001.58 | 1.175 | 0.0032203 | -0.044012 | -0.14575 | -0.10174 |  | 6.593 | -0.36099
7 | NO | 1501.9 | 0.47704 | 0.0031055 | -0.029963 | -0.1136 | -0.083637 |  | 5.6457 | -0.37828
7 | YES | 1251.29 | 0.30083 | 0.0028858 | -0.010174 | -0.10548 | -0.095309 |  | 5.6245 | -0.34247
7 | YES | 1000.7 | 0.25587 | 0.0025214 | -0.029133 | -0.10341 | -0.074279 |  | 5.6415 | -0.33653
7 | YES | 801.449 | 0.36527 | 0.0024324 | -0.022726 | -0.10845 | -0.085726 |  | 0.32903 | -0.012967
7 | YES | 599.478 | 0.57156 | 0.002331 | -0.047439 | -0.11795 | -0.070515 |  | 0.35476 | -0.023241
7 | YES | 400.235 | 1.396 | 0.0041091 | -0.11402 | -0.15593 | -0.041916 |  | 0.4331 | -0.6209
7 | YES | 300.436 | 2.3084 | 0.0038714 | -0.22357 | -0.19796 | 0.025606 |  | 0.56649 | -0.060506
7 | YES | 250.34 | 2.9682 | 0.0091679 | -0.21684 | -0.22835 | -0.01151 |  | 1.0657 | -0.11028
7 | YES | 200.034 | 4.1863 | 0.02518 | -0.33474 | -0.28446 | 0.050274 |  | 1.7515 | -0.14347
7 | YES | 176.444 | 4.8064 | 0.0094576 | -0.34155 | -0.31303 | 0.028515 |  | 2.5384 | -0.21861
7 | YES | 149.804 | 5.4656 | 0.010897 | -0.35639 | -0.34339 | 0.012991 |  | 3.4453 | -0.29367
7 | YES | 126.568 | 6.2031 | 0.01604 | -0.40091 | -0.37737 | 0.023539 |  | 4.1073 | -0.30468
7 | YES | 101.691 | 6.4238 | 0.0076469 | -0.38922 | -0.38753 | 0.0016918 |  | 5.0372 | -0.36982
7 | YES | 74.5608 | 6.5239 | 0.024834 | -0.48108 | -0.39215 | 0.088938 |  | 6.0571 | -0.46995
7 | YES | 49.9787 | 6.593 | 0.018445 | -0.36099 | -0.39533 | -0.034337 |  | 6.4907 | -0.43729
7 | YES | 25.418 | 5.6457 | 0.022196 | -0.37828 | -0.35169 | 0.026587 |  | 6.4701 | -0.43592
7 | YES | 10.5048 | 5.6245 | 0.016742 | -0.34247 | -0.35072 | -0.0082508 |  | 6.4631 | -0.44993
7 | YES | 6.1822 | 5.6415 | 0.017912 | -0.33653 | -0.3515 | -0.01496 |  | 5.8293 | -0.41173
8 | NO | 4315.72 | 3.0101 | 0.0067191 | -0.21489 | -0.23028 | -0.015392 |  | 5.8095 | -0.40253
8 | NO | 3998.96 | 2.9411 | 0.007228 | -0.25786 | -0.22711 | 0.030756 |  | 5.8033 | -0.57167
8 | NO | 3501.9 | 2.7317 | 0.0053451 | -0.22328 | -0.21746 | 0.0058217 |  | 0.33658 | -0.0034205
8 | NO | 3000.3 | 2.3371 | 0.0058576 | -0.20088 | -0.19928 | 0.0016006 |  | 0.31054 | -0.010465
8 | NO | 2500.89 | 1.7898 | 0.004994 | -0.16321 | -0.17407 | -0.010864 |  | 0.42635 | -0.044647
8 | NO | 2003.55 | 1.1834 | 0.0039803 | -0.11965 | -0.14614 | -0.026491 |  | 0.5665 | -0.033502
8 | NO | 1501.95 | 0.49161 | 0.0036682 | -0.00739 | -0.11427 | -0.10688 |  | 1.1597 | -0.13529
8 | YES | 1250.92 | 0.32903 | 0.0029901 | -0.012967 | -0.10678 | -0.093816 |  | 1.884 | -0.20196
8 | YES | 999.715 | 0.35476 | 0.0024273 | -0.023241 | -0.10797 | -0.084727 |  | 2.4496 | -0.21636
8 | YES | 802.475 | 0.4331 | 0.0019361 | -0.6209 | -0.11158 | 0.50932 |  | 3.6022 | -0.31876
8 | YES | 599.257 | 0.56649 | 0.0024068 | -0.060506 | -0.11772 | -0.057215 |  | 4.0755 | -0.31946
8 | YES | 399.823 | 1.0657 | 0.0028522 | -0.11028 | -0.14072 | -0.030439 |  | 4.9887 | -0.44525
8 | YES | 297.593 | 1.7515 | 0.0051562 | -0.14347 | -0.17231 | -0.028837 |  | 5.7935 | -0.49652
8 | YES | 247.972 | 2.5384 | 0.0075519 | -0.21861 | -0.20855 | 0.010051 |  | 6.2541 | -0.47289
8 | YES | 200.803 | 3.4453 | 0.013674 | -0.29367 | -0.25033 | 0.043336 |  | 6.3995 | -0.44349
8 | YES | 174.613 | 4.1073 | 0.0079549 | -0.30468 | -0.28083 | 0.023857 |  | 6.4517 | -0.48427
8 | YES | 149.88 | 5.0372 | 0.018226 | -0.36982 | -0.32366 | 0.046162 |  | 5.6485 | -0.45955
8 | YES | 124.694 | 6.0571 | 0.013726 | -0.46995 | -0.37064 | 0.099315 |  | 5.6612 | -0.44281
8 | YES | 100.56 | 6.4907 | 0.014136 | -0.43729 | -0.39062 | 0.046679 |  | 5.6572 | -0.45378
8 | YES | 73.8024 | 6.4701 | 0.01843 | -0.43592 | -0.38966 | 0.046252 |  | 0.32417 | 0.013174
8 | YES | 50.6964 | 6.4631 | 0.0187 | -0.44993 | -0.38934 | 0.060588 |  | 0.27827 | 0.024274
8 | YES | 23.8606 | 5.8293 | 0.013286 | -0.41173 | -0.36015 | 0.051583 |  | 0.35488 | -0.033116
8 | YES | 11.107 | 5.8095 | 0.018044 | -0.40253 | -0.35923 | 0.043293 |  | 0.54376 | -0.023237
8 | YES | 5.5023 | 5.8033 | 0.018351 | -0.57167 | -0.35895 | 0.21272 |  | 1.0239 | -0.07805
9 | NO | 4024.09 | 2.7851 | 0.0073442 | -0.39191 | -0.21992 | 0.17199 |  | 1.4423 | -0.10473
9 | NO | 3498.86 | 2.6327 | 0.0069541 | -0.24833 | -0.2129 | 0.035431 |  | 1.6911 | -0.12986
9 | NO | 3498.86 | 2.6327 | 0.0069541 | -0.20433 | -0.2129 | -0.0085686 |  | 2.0769 | -0.15214
9 | NO | 2998.76 | 2.3213 | 0.0052467 | -0.16272 | -0.19855 | -0.035836 |  | 2.3036 | -0.15136
9 | NO | 2501.18 | 1.794 | 0.0040669 | -0.15396 | -0.17427 | -0.020309 |  | 2.5619 | -0.20505
9 | NO | 1999.9 | 1.2011 | 0.004018 | -0.077852 | -0.14696 | -0.069104 |  | 3.7627 | -0.40133
9 | NO | 1501.12 | 0.508 | 0.0028358 | -0.036 | -0.11503 | -0.079026 |  | 5.319 | -0.39605
9 | YES | 1252.06 | 0.33658 | 0.0031631 | -0.0034205 | -0.10713 | -0.10371 |  | 6.1191 | -0.46193
9 | YES | 1001.45 | 0.31054 | 0.0026972 | -0.010465 | -0.10593 | -0.095466 |  | 6.4523 | -0.49372
9 | YES | 801.223 | 0.42635 | 0.0021842 | -0.044647 | -0.11127 | -0.066618 |  | 6.5152 | -0.48879
9 | YES | 601.026 | 0.5665 | 0.002174 | -0.033502 | -0.11772 | -0.084219 |  | 5.6367 | -0.44334
9 | YES | 399.824 | 1.1597 | 0.0042208 | -0.13529 | -0.14505 | -0.0097569 |  | 5.612 | -0.45695
9 | YES | 300.177 | 1.884 | 0.0040253 | -0.20196 | -0.17841 | 0.02355 |  | 5.612 | -0.45495
9 | YES | 250.475 | 2.4496 | 0.0055907 | -0.21636 | -0.20447 | 0.01189 |  | 5.5937 | -0.41628
9 | YES | 198.379 | 3.6022 | 0.0062745 | -0.31876 | -0.25756 | 0.061195 |  | 5.5706 | -0.42942
9 | YES | 175.41 | 4.0755 | 0.013949 | -0.31946 | -0.27936 | 0.040098 |  | 5.5692 | -0.45884
9 | YES | 150.732 | 4.9887 | 0.043861 | -0.44525 | -0.32143 | 0.12382
9 | YES | 124.125 | 5.7935 | 0.028387 | -0.49652 | -0.3585 | 0.13802
9 | YES | 99.4924 | 6.2541 | 0.018472 | -0.47289 | -0.37972 | 0.093176
9 | YES | 74.7791 | 6.3995 | 0.023285 | -0.44349 | -0.38641 | 0.057076
9 | YES | 49.3064 | 6.4517 | 0.029448 | -0.48427 | -0.38882 | 0.095445
9 | YES | 24.7034 | 5.6485 | 0.011864 | -0.45955 | -0.35182 | 0.10773
9 | YES | 9.7331 | 5.6612 | 0.021845 | -0.44281 | -0.3524 | 0.090409
9 | YES | 5.1962 | 5.6572 | 0.024126 | -0.45378 | -0.35222 | 0.10156
10 | NO | 1998.83 | 1.1924 | 0.0047216 | -0.097585 | -0.14655 | -0.048968
10 | NO | 1518.6 | 0.54876 | 0.0029639 | -0.014236 | -0.1169 | -0.10267
10 | YES | 1249.67 | 0.32417 | 0.0025906 | 0.013174 | -0.10656 | -0.11973
10 | YES | 999.208 | 0.27827 | 0.0030413 | 0.024274 | -0.10444 | -0.12872
10 | YES | 799.271 | 0.35488 | 0.0025987 | -0.033116 | -0.10797 | -0.074857
10 | YES | 599.385 | 0.54376 | 0.0024845 | -0.023237 | -0.11667 | -0.093437
10 | YES | 402.41 | 1.0239 | 0.0073432 | -0.07805 | -0.13879 | -0.060743
10 | YES | 301.932 | 1.4423 | 0.0027367 | -0.10473 | -0.15806 | -0.053333
10 | YES | 250.688 | 1.6911 | 0.006699 | -0.12986 | -0.16953 | -0.039664
10 | YES | 199.64 | 2.0769 | 0.0036755 | -0.15214 | -0.18729 | -0.035155
10 | YES | 175.483 | 2.3036 | 0.0074563 | -0.15136 | -0.19774 | -0.046384
10 | YES | 150.246 | 2.5619 | 0.0096489 | -0.20505 | -0.20964 | -0.004594
10 | YES | 124.167 | 3.7627 | 0.032554 | -0.40133 | -0.26495 | 0.13638
10 | YES | 99.5251 | 5.319 | 0.016428 | -0.39605 | -0.33664 | 0.059415
10 | YES | 74.3439 | 6.1191 | 0.011027 | -0.46193 | -0.3735 | 0.088438
10 | YES | 49.3314 | 6.4523 | 0.025608 | -0.49372 | -0.38885 | 0.10488
10 | YES | 40.9976 | 6.5152 | 0.035805 | -0.48879 | -0.39174 | 0.097049
10 | YES | 29.6041 | 5.6367 | 0.0091063 | -0.44334 | -0.35127 | 0.092062
10 | YES | 22.1111 | 5.612 | 0.025038 | -0.45695 | -0.35014 | 0.10681
10 | YES | 22.1111 | 5.612 | 0.025038 | -0.45495 | -0.35014 | 0.10481
10 | YES | 10.6242 | 5.5937 | 0.019454 | -0.41628 | -0.3493 | 0.066982
10 | YES | 4.5235 | 5.5706 | 0.032389 | -0.42942 | -0.34823 | 0.081186
10 | YES | 3.0114 | 5.5692 | 0.03293 | -0.45884 | -0.34816 | 0.11068
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (14)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.042773 |  | X | CALC Y
m= | -0.0058873 | 0.010527 |  | Std Dev of Resid |  | 0.077462 |  | 0.044527 | 0.074407
b= | 0.074669 | 0.033976 |  | Resid Sum of Squares |  | 0.042003 |  | 6.5613 | 0.03604
# of Values Extracted = |  |  | 10
# of Values Used in Fit = |  |  | 9
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 220.267 | 0.050766 | 0.001127 | 0.050766 | 0.07437 | 0.023604 |  | 0.050766 | 0.050766
1 | YES | 200.729 | 0.047012 | 0.0015261 | 0.047012 | 0.074392 | 0.02738 |  | 0.047012 | 0.047012
1 | YES | 176.236 | 0.044834 | 0.00091597 | 0.044834 | 0.074405 | 0.029571 |  | 0.044834 | 0.044834
1 | YES | 151.113 | 0.044527 | 0.00097484 | 0.044527 | 0.074407 | 0.02988 |  | 0.044527 | 0.044527
1 | YES | 125.529 | 0.21221 | 0.0015754 | 0.069208 | 0.07342 | 0.0042121 |  | 0.21221 | 0.069208
1 | NO | 100.564 | 0.42818 | 0.018345 | 0.10618 | 0.072148 | -0.034035 |  | 2.8843 | 0.21328
1 | YES | 49.5756 | 2.8843 | 0.0062989 | 0.21328 | 0.057688 | -0.15559 |  | 3.6486 | 0.13656
1 | YES | 25.5956 | 3.6486 | 0.0090746 | 0.13656 | 0.053189 | -0.083367 |  | 5.3874 | -0.015605
1 | YES | 10.6371 | 5.3874 | 0.048179 | -0.015605 | 0.042952 | 0.058557 |  | 6.5613 | -0.029714
1 | YES | 5.9689 | 6.5613 | 0.022672 | -0.029714 | 0.03604 | 0.065754
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (15)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.84234 |  | X | CALC Y
m= | -0.042131 | 0.0044207 |  | Std Dev of Resid |  | 0.04062 |  | 0.21613 | -0.045327
b= | -0.036221 | 0.014952 |  | Resid Sum of Squares |  | 0.028049 |  | 6.2679 | -0.3003
# of Values Extracted = |  |  | 155
# of Values Used in Fit = |  |  | 19
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
4 | NO | 1300.62 | 0.49522 | 0.0021859 | -0.054784 | -0.057085 | -0.0023012 |  | 1.2821 | -0.1689
4 | NO | 1250.05 | 0.41302 | 0.0030083 | -0.032976 | -0.053623 | -0.020646 |  | 0.61915 | -0.06185
4 | NO | 1250.05 | 0.41302 | 0.0030083 | -0.026976 | -0.053623 | -0.026646 |  | 0.61915 | -0.06285
4 | NO | 1000.54 | 0.23775 | 0.0018914 | -0.013245 | -0.046238 | -0.032993 |  | 0.33626 | -0.018739
4 | NO | 801.075 | 0.21223 | 0.0020806 | -0.024768 | -0.045163 | -0.020395 |  | 0.21613 | -0.020871
4 | NO | 600.918 | 0.42334 | 0.0028851 | -0.040664 | -0.054057 | -0.013393 |  | 0.37322 | -0.054776
4 | NO | 400.248 | 0.99714 | 0.001509 | -0.072863 | -0.078232 | -0.0053689 |  | 1.0396 | -0.060402
4 | NO | 301.823 | 1.2875 | 0.002412 | -0.075485 | -0.090466 | -0.01498 |  | 1.8098 | -0.10022
4 | NO | 201.201 | 1.6274 | 0.00573 | -0.1076 | -0.10479 | 0.0028112 |  | 1.9105 | -0.12346
4 | NO | 174.556 | 1.8759 | 0.0068785 | -0.076058 | -0.11526 | -0.039199 |  | 1.9582 | -0.088759
4 | NO | 150.455 | 2.2244 | 0.011745 | -0.14162 | -0.12994 | 0.011685 |  | 2.0371 | -0.10989
4 | NO | 125.295 | 2.7552 | 0.0045848 | -0.13478 | -0.1523 | -0.017518 |  | 2.1758 | -0.15724
4 | NO | 100.512 | 3.3938 | 0.0073132 | -0.14018 | -0.17921 | -0.039025 |  | 2.6831 | -0.1639
4 | NO | 75.7457 | 4.6783 | 0.018784 | -0.22371 | -0.23332 | -0.0096138 |  | 3.998 | -0.21401
4 | NO | 51.446 | 6.1723 | 0.010765 | -0.20073 | -0.29627 | -0.095532 |  | 5.7981 | -0.29987
4 | NO | 24.7249 | 6.2779 | 0.017082 | -0.32107 | -0.30072 | 0.020355 |  | 6.2679 | -0.21206
4 | NO | 10.3947 | 6.4598 | 0.044508 | -0.26224 | -0.30838 | -0.046142 |  | 5.7126 | -0.3694
4 | NO | 5.239 | 6.5395 | 0.016358 | -0.32054 | -0.31174 | 0.0088067 |  | 5.7041 | -0.26693
5 | YES | 2003.65 | 1.2821 | 0.003906 | -0.1689 | -0.090238 | 0.078662 |  | 5.7156 | -0.25144
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06185 | -0.062307 | -0.0004568
5 | YES | 1500.62 | 0.61915 | 0.0025496 | -0.06285 | -0.062307 | 0.0005432
5 | YES | 1250.17 | 0.33626 | 0.0025662 | -0.018739 | -0.050388 | -0.03165
5 | YES | 1001.62 | 0.21613 | 0.0024958 | -0.020871 | -0.045327 | -0.024456
5 | NO | 800.329 | 0.21353 | 0.0022177 | -0.79147 | -0.045217 | 0.74626
5 | YES | 599.855 | 0.37322 | 0.0022839 | -0.054776 | -0.051946 | 0.0028302
5 | YES | 401.219 | 1.0396 | 0.0021928 | -0.060402 | -0.080021 | -0.019619
5 | YES | 299.827 | 1.8098 | 0.0070129 | -0.10022 | -0.11247 | -0.01225
5 | YES | 249.318 | 1.9105 | 0.0063301 | -0.12346 | -0.11671 | 0.0067442
5 | YES | 200.212 | 1.9582 | 0.0035649 | -0.088759 | -0.11872 | -0.029965
5 | YES | 173.88 | 2.0371 | 0.010025 | -0.10989 | -0.12205 | -0.012153
5 | YES | 150.707 | 2.1758 | 0.0089689 | -0.15724 | -0.12789 | 0.029352
5 | YES | 124.044 | 2.6831 | 0.012874 | -0.1639 | -0.14926 | 0.01464
5 | YES | 101.474 | 3.998 | 0.011501 | -0.21401 | -0.20466 | 0.0093462
5 | YES | 75.2867 | 5.7981 | 0.037336 | -0.29987 | -0.2805 | 0.019362
5 | YES | 49.8081 | 6.2679 | 0.011322 | -0.21206 | -0.3003 | -0.088236
5 | YES | 25.2931 | 5.7126 | 0.015859 | -0.3694 | -0.2769 | 0.092501
5 | YES | 10.3099 | 5.7041 | 0.026635 | -0.26693 | -0.27654 | -0.0096112
5 | YES | 5.3529 | 5.7156 | 0.02113 | -0.25144 | -0.27703 | -0.025585
6 | NO | 3276.19 | 2.1742 | 0.0052468 | -0.16977 | -0.12782 | 0.041944
6 | NO | 3000.48 | 2.1445 | 0.0049557 | -0.14653 | -0.12657 | 0.019956
6 | NO | 2501.59 | 1.7573 | 0.0045462 | -0.22471 | -0.11026 | 0.11446
6 | NO | 2501.59 | 1.7573 | 0.0045462 | -0.10471 | -0.11026 | -0.0055442
6 | NO | 2000.27 | 1.2232 | 0.0036682 | -0.074834 | -0.087755 | -0.012921
6 | NO | 1501.16 | 0.60927 | 0.0032809 | -0.01373 | -0.061891 | -0.04816
6 | NO | 1250.57 | 0.38709 | 0.0031696 | -0.057913 | -0.05253 | 0.0053831
6 | NO | 1001.93 | 0.22771 | 0.0019075 | -0.011286 | -0.045815 | -0.034529
6 | NO | 801.143 | 0.26251 | 0.0024751 | -0.04849 | -0.047281 | 0.0012084
6 | NO | 600.801 | 0.51929 | 0.0018098 | -0.064705 | -0.0581 | 0.0066055
6 | NO | 400.283 | 1.3473 | 0.0028976 | -0.13473 | -0.092984 | 0.04175
6 | NO | 301.716 | 2.5586 | 0.0042794 | -0.17243 | -0.14402 | 0.028414
6 | NO | 249.814 | 3.273 | 0.0030968 | -0.22396 | -0.17412 | 0.049839
6 | NO | 200.271 | 4.0577 | 0.0056607 | -0.23326 | -0.20718 | 0.026082
6 | NO | 176.597 | 4.3958 | 0.005897 | -0.24818 | -0.22142 | 0.02676
6 | NO | 150.549 | 4.6352 | 0.0091057 | -0.25283 | -0.23151 | 0.021327
6 | NO | 124.879 | 5.0843 | 0.0094488 | -0.26169 | -0.25043 | 0.011263
6 | NO | 100.671 | 5.8131 | 0.0097689 | -0.26988 | -0.28114 | -0.011256
6 | NO | 75.8548 | 6.4835 | 0.015863 | -0.3485 | -0.30938 | 0.039122
6 | NO | 50.9557 | 6.7227 | 0.016748 | -0.3433 | -0.31946 | 0.023842
6 | NO | 25.6344 | 5.9998 | 0.01886 | -0.38818 | -0.289 | 0.099181
6 | NO | 9.8456 | 2.3884 | 0.025795 | -3.5306 | -0.13685 | 3.3937
6 | NO | 5.8409 | 2.106 | 0.0078225 | -3.82 | -0.12495 | 3.695
7 | NO | 3500.19 | 2.5276 | 0.0043321 | -0.22339 | -0.14271 | 0.080677
7 | NO | 2999.58 | 2.2661 | 0.0053773 | -0.19986 | -0.1317 | 0.068161
7 | NO | 2497.23 | 1.7816 | 0.0045839 | -0.12635 | -0.11128 | 0.015068
7 | NO | 2496.9 | 1.7818 | 0.0046856 | -0.12515 | -0.11129 | 0.013856
7 | NO | 2001.58 | 1.175 | 0.0032203 | -0.044012 | -0.085725 | -0.041713
7 | NO | 1501.9 | 0.47704 | 0.0031055 | -0.029963 | -0.056319 | -0.026357
7 | NO | 1251.29 | 0.30083 | 0.0028858 | -0.010174 | -0.048895 | -0.038721
7 | NO | 1000.7 | 0.25587 | 0.0025214 | -0.029133 | -0.047001 | -0.017869
7 | NO | 801.449 | 0.36527 | 0.0024324 | -0.022726 | -0.051611 | -0.028885
7 | NO | 599.478 | 0.57156 | 0.002331 | -0.047439 | -0.060302 | -0.012863
7 | NO | 400.235 | 1.396 | 0.0041091 | -0.11402 | -0.095036 | 0.018979
7 | NO | 300.436 | 2.3084 | 0.0038714 | -0.22357 | -0.13348 | 0.09009
7 | NO | 250.34 | 2.9682 | 0.0091679 | -0.21684 | -0.16127 | 0.055568
7 | NO | 200.034 | 4.1863 | 0.02518 | -0.33474 | -0.21259 | 0.12214
7 | NO | 176.444 | 4.8064 | 0.0094576 | -0.34155 | -0.23872 | 0.10282
7 | NO | 149.804 | 5.4656 | 0.010897 | -0.35639 | -0.26649 | 0.089891
7 | NO | 126.568 | 6.2031 | 0.01604 | -0.40091 | -0.29757 | 0.10334
7 | NO | 101.691 | 6.4238 | 0.0076469 | -0.38922 | -0.30686 | 0.08236
7 | NO | 74.5608 | 6.5239 | 0.024834 | -0.48108 | -0.31108 | 0.17
7 | NO | 49.9787 | 6.593 | 0.018445 | -0.36099 | -0.31399 | 0.046998
7 | NO | 25.418 | 5.6457 | 0.022196 | -0.37828 | -0.27408 | 0.1042
7 | NO | 10.5048 | 5.6245 | 0.016742 | -0.34247 | -0.27319 | 0.069274
7 | NO | 6.1822 | 5.6415 | 0.017912 | -0.33653 | -0.2739 | 0.062632
8 | NO | 4315.72 | 3.0101 | 0.0067191 | -0.21489 | -0.16304 | 0.051851
8 | NO | 3998.96 | 2.9411 | 0.007228 | -0.25786 | -0.16014 | 0.097728
8 | NO | 3501.9 | 2.7317 | 0.0053451 | -0.22328 | -0.15131 | 0.07197
8 | NO | 3000.3 | 2.3371 | 0.0058576 | -0.20088 | -0.13469 | 0.066197
8 | NO | 2500.89 | 1.7898 | 0.004994 | -0.16321 | -0.11163 | 0.05158
8 | NO | 2003.55 | 1.1834 | 0.0039803 | -0.11965 | -0.086078 | 0.033567
8 | NO | 1501.95 | 0.49161 | 0.0036682 | -0.00739 | -0.056933 | -0.049543
8 | NO | 1250.92 | 0.32903 | 0.0029901 | -0.012967 | -0.050084 | -0.037117
8 | NO | 999.715 | 0.35476 | 0.0024273 | -0.023241 | -0.051168 | -0.027927
8 | NO | 802.475 | 0.4331 | 0.0019361 | -0.6209 | -0.054468 | 0.56643
8 | NO | 599.257 | 0.56649 | 0.0024068 | -0.060506 | -0.060088 | 0.00041777
8 | NO | 399.823 | 1.0657 | 0.0028522 | -0.11028 | -0.081121 | 0.029157
8 | NO | 297.593 | 1.7515 | 0.0051562 | -0.14347 | -0.11002 | 0.033456
8 | NO | 247.972 | 2.5384 | 0.0075519 | -0.21861 | -0.14317 | 0.075439
8 | NO | 200.803 | 3.4453 | 0.013674 | -0.29367 | -0.18138 | 0.11229
8 | NO | 174.613 | 4.1073 | 0.0079549 | -0.30468 | -0.20927 | 0.095415
8 | NO | 149.88 | 5.0372 | 0.018226 | -0.36982 | -0.24844 | 0.12138
8 | NO | 124.694 | 6.0571 | 0.013726 | -0.46995 | -0.29141 | 0.17854
8 | NO | 100.56 | 6.4907 | 0.014136 | -0.43729 | -0.30968 | 0.12761
8 | NO | 73.8024 | 6.4701 | 0.01843 | -0.43592 | -0.30881 | 0.1271
8 | NO | 50.6964 | 6.4631 | 0.0187 | -0.44993 | -0.30852 | 0.14141
8 | NO | 23.8606 | 5.8293 | 0.013286 | -0.41173 | -0.28182 | 0.12991
8 | NO | 11.107 | 5.8095 | 0.018044 | -0.40253 | -0.28098 | 0.12155
8 | NO | 5.5023 | 5.8033 | 0.018351 | -0.57167 | -0.28072 | 0.29095
9 | NO | 4024.09 | 2.7851 | 0.0073442 | -0.39191 | -0.15356 | 0.23834
9 | NO | 3498.86 | 2.6327 | 0.0069541 | -0.24833 | -0.14714 | 0.10119
9 | NO | 3498.86 | 2.6327 | 0.0069541 | -0.20433 | -0.14714 | 0.05719
9 | NO | 2998.76 | 2.3213 | 0.0052467 | -0.16272 | -0.13402 | 0.028698
9 | NO | 2501.18 | 1.794 | 0.0040669 | -0.15396 | -0.11181 | 0.042151
9 | NO | 1999.9 | 1.2011 | 0.004018 | -0.077852 | -0.086827 | -0.0089753
9 | NO | 1501.12 | 0.508 | 0.0028358 | -0.036 | -0.057624 | -0.021624
9 | NO | 1252.06 | 0.33658 | 0.0031631 | -0.0034205 | -0.050402 | -0.046981
9 | NO | 1001.45 | 0.31054 | 0.0026972 | -0.010465 | -0.049305 | -0.03884
9 | NO | 801.223 | 0.42635 | 0.0021842 | -0.044647 | -0.054184 | -0.0095368
9 | NO | 601.026 | 0.5665 | 0.002174 | -0.033502 | -0.060089 | -0.026586
9 | NO | 399.824 | 1.1597 | 0.0042208 | -0.13529 | -0.085081 | 0.050209
9 | NO | 300.177 | 1.884 | 0.0040253 | -0.20196 | -0.1156 | 0.086365
9 | NO | 250.475 | 2.4496 | 0.0055907 | -0.21636 | -0.13943 | 0.076929
9 | NO | 198.379 | 3.6022 | 0.0062745 | -0.31876 | -0.18799 | 0.13077
9 | NO | 175.41 | 4.0755 | 0.013949 | -0.31946 | -0.20793 | 0.11153
9 | NO | 150.732 | 4.9887 | 0.043861 | -0.44525 | -0.2464 | 0.19885
9 | NO | 124.125 | 5.7935 | 0.028387 | -0.49652 | -0.28031 | 0.21621
9 | NO | 99.4924 | 6.2541 | 0.018472 | -0.47289 | -0.29972 | 0.17318
9 | NO | 74.7791 | 6.3995 | 0.023285 | -0.44349 | -0.30584 | 0.13765
9 | NO | 49.3064 | 6.4517 | 0.029448 | -0.48427 | -0.30804 | 0.17622
9 | NO | 24.7034 | 5.6485 | 0.011864 | -0.45955 | -0.2742 | 0.18535
9 | NO | 9.7331 | 5.6612 | 0.021845 | -0.44281 | -0.27473 | 0.16808
9 | NO | 5.1962 | 5.6572 | 0.024126 | -0.45378 | -0.27457 | 0.17921
10 | NO | 1998.83 | 1.1924 | 0.0047216 | -0.097585 | -0.086459 | 0.011126
10 | NO | 1518.6 | 0.54876 | 0.0029639 | -0.014236 | -0.059341 | -0.045105
10 | NO | 1249.67 | 0.32417 | 0.0025906 | 0.013174 | -0.049879 | -0.063053
10 | NO | 999.208 | 0.27827 | 0.0030413 | 0.024274 | -0.047945 | -0.072219
10 | NO | 799.271 | 0.35488 | 0.0025987 | -0.033116 | -0.051173 | -0.018057
10 | NO | 599.385 | 0.54376 | 0.0024845 | -0.023237 | -0.059131 | -0.035894
10 | NO | 402.41 | 1.0239 | 0.0073432 | -0.07805 | -0.079362 | -0.0013118
10 | NO | 301.932 | 1.4423 | 0.0027367 | -0.10473 | -0.096986 | 0.007744
10 | NO | 250.688 | 1.6911 | 0.006699 | -0.12986 | -0.10747 | 0.022392
10 | NO | 199.64 | 2.0769 | 0.0036755 | -0.15214 | -0.12372 | 0.028418
10 | NO | 175.483 | 2.3036 | 0.0074563 | -0.15136 | -0.13328 | 0.01808
10 | NO | 150.246 | 2.5619 | 0.0096489 | -0.20505 | -0.14416 | 0.060886
10 | NO | 124.167 | 3.7627 | 0.032554 | -0.40133 | -0.19475 | 0.20658
10 | NO | 99.5251 | 5.319 | 0.016428 | -0.39605 | -0.26032 | 0.13574
10 | NO | 74.3439 | 6.1191 | 0.011027 | -0.46193 | -0.29403 | 0.16791
10 | NO | 49.3314 | 6.4523 | 0.025608 | -0.49372 | -0.30806 | 0.18566
10 | NO | 40.9976 | 6.5152 | 0.035805 | -0.48879 | -0.31072 | 0.17808
10 | NO | 29.6041 | 5.6367 | 0.0091063 | -0.44334 | -0.2737 | 0.16963
10 | NO | 22.1111 | 5.612 | 0.025038 | -0.45695 | -0.27266 | 0.18428
10 | NO | 22.1111 | 5.612 | 0.025038 | -0.45495 | -0.27266 | 0.18228
10 | NO | 10.6242 | 5.5937 | 0.019454 | -0.41628 | -0.27189 | 0.14439
10 | NO | 4.5235 | 5.5706 | 0.032389 | -0.42942 | -0.27092 | 0.1585
10 | NO | 3.0114 | 5.5692 | 0.03293 | -0.45884 | -0.27086 | 0.18798
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.3 | Run at 2009/10/06 09:03:22
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX_BOT | DOX_CTD |  | DOX_CTD - DOX_BOT
Average Size: | 2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX_BOT | DOX_CTD | DOX_CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2009-10-0001.mrg | 2009-10-0001.sam | 201 | 220.267 | 220.267 | 0.035526 | -0.00016785 | 0 | 0.050766 | 0.001127 | 0.050766
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 200.729 | 200.729 | 0.022664 | 0.00010681 | 0 | 0.047012 | 0.0015261 | 0.047012
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 176.236 | 176.236 | 0.021474 | 0.00032043 | 0 | 0.044834 | 0.00091597 | 0.044834
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 151.113 | 151.113 | 0.023265 | -0.00018311 | 0 | 0.044527 | 0.00097484 | 0.044527
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 125.529 | 125.529 | 0.026691 | 0.00027466 | 0.143 | 0.21221 | 0.0015754 | 0.069208
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 100.564 | 100.564 | 0.022357 | 6.1035e-05 | 0.322 | 0.42818 | 0.018345 | 0.10618
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 49.576 | 49.5756 | 0.025434 | -0.0004158 | 2.671 | 2.8843 | 0.0062989 | 0.21328
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 25.596 | 25.5956 | 0.023418 | -0.00042343 | 3.512 | 3.6486 | 0.0090746 | 0.13656
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 10.637 | 10.6371 | 0.023221 | 9.1553e-05 | 5.403 | 5.3874 | 0.048179 | -0.015605
2009-10-0001.mrg | 2009-10-0001.sam | 241 | 5.969 | 5.9689 | 0.021007 | -7.0572e-05 | 6.591 | 6.5613 | 0.022672 | -0.029714
2009-10-0003.mrg | 2009-10-0003.sam | 482 | 200.994 | 200.987 | 0.025742 | -0.0065155 | 0.009 | 0.063886 | 0.0017874 | 0.054886
2009-10-0003.mrg | 2009-10-0003.sam | 482 | 200.981 | 200.987 | 0.025742 | 0.006485 | 0 | 0.063886 | 0.0017874 | 0.063886
2009-10-0003.mrg | 2009-10-0003.sam | 299 | 99.349 | 98.3678 | 0.52729 | -0.98125 | 0.514 | 0.51929 | 0.0053849 | 0.0052876
2009-10-0003.mrg | 2009-10-0003.sam | 299 | 98.127 | 98.3678 | 0.52729 | 0.24075 | 0.55 | 0.51929 | 0.0053849 | -0.030712
2009-10-0003.mrg | 2009-10-0003.sam | 299 | 51.388 | 51.3838 | 0.023359 | -0.0041924 | 2.609 | 2.8127 | 0.075542 | 0.20366
2009-10-0003.mrg | 2009-10-0003.sam | 299 | 51.383 | 51.3838 | 0.023359 | 0.00080872 | 2.633 | 2.8127 | 0.075542 | 0.17966
2009-10-0003.mrg | 2009-10-0003.sam | 299 | 11.168 | 11.1691 | 0.021696 | 0.0010567 | 5.913 | 6.1357 | 0.027563 | 0.22272
2009-10-0003.mrg | 2009-10-0003.sam | 299 | 11.169 | 11.1691 | 0.021696 | 5.722e-05 | 5.675 | 6.1357 | 0.027563 | 0.46072
2009-10-0006.mrg | 2009-10-0006.sam | 241 | 106.372 | 106.372 | 0.69877 | 0.00049591 | 1.551 | 1.5767 | 0.011722 | 0.025656
2009-10-0006.mrg | 2009-10-0006.sam | 373 | 75.968 | 76.1436 | 0.30133 | 0.17562 | 2.193 | 2.128 | 0.0067447 | -0.064981
2009-10-0006.mrg | 2009-10-0006.sam | 373 | 76.241 | 76.1436 | 0.30133 | -0.097374 | 2.193 | 2.128 | 0.0067447 | -0.064981
2009-10-0006.mrg | 2009-10-0006.sam | 241 | 51.405 | 51.4045 | 0.53512 | -0.00048447 | 2.696 | 2.6019 | 0.0058068 | -0.094104
2009-10-0006.mrg | 2009-10-0006.sam | 241 | 26.303 | 26.3033 | 0.54378 | 0.00032806 | 4.621 | 4.2925 | 0.070084 | -0.32851
2009-10-0006.mrg | 2009-10-0006.sam | 241 | 11.177 | 11.1766 | 0.24738 | -0.00041485 | 5.821 | 5.5508 | 0.030052 | -0.2702
2009-10-0006.mrg | 2009-10-0006.sam | 241 | 5.525 | 5.5253 | 0.23008 | 0.000278 | 5.945 | 5.7243 | 0.011576 | -0.22071
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 1300.62 | 1300.62 | 0.34523 | 0.00024414 | 0.55 | 0.49522 | 0.0021859 | -0.054784
2009-10-0012.mrg | 2009-10-0012.sam | 336 | 1250.13 | 1250.05 | 0.14024 | -0.084961 | 0.446 | 0.41302 | 0.0030083 | -0.032976
2009-10-0012.mrg | 2009-10-0012.sam | 336 | 1250.01 | 1250.05 | 0.14024 | 0.034058 | 0.44 | 0.41302 | 0.0030083 | -0.026976
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 1000.54 | 1000.54 | 0.22001 | -0.00018311 | 0.251 | 0.23775 | 0.0018914 | -0.013245
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 801.075 | 801.075 | 0.19628 | 0.00036621 | 0.237 | 0.21223 | 0.0020806 | -0.024768
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 600.918 | 600.918 | 0.28339 | 6.1035e-05 | 0.464 | 0.42334 | 0.0028851 | -0.040664
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 400.248 | 400.248 | 0.12057 | -3.0518e-05 | 1.07 | 0.99714 | 0.001509 | -0.072863
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 301.823 | 301.823 | 0.22578 | 9.1553e-05 | 1.363 | 1.2875 | 0.002412 | -0.075485
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 201.201 | 201.201 | 0.16453 | 0.00032043 | 1.735 | 1.6274 | 0.00573 | -0.1076
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 174.556 | 174.556 | 0.26223 | 0.00042725 | 1.952 | 1.8759 | 0.0068785 | -0.076058
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 150.455 | 150.455 | 0.30495 | -0.00018311 | 2.366 | 2.2244 | 0.011745 | -0.14162
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 125.295 | 125.295 | 0.21595 | 3.8147e-05 | 2.89 | 2.7552 | 0.0045848 | -0.13478
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 100.512 | 100.512 | 0.13887 | -0.00028992 | 3.534 | 3.3938 | 0.0073132 | -0.14018
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 75.746 | 75.7457 | 0.17574 | -0.00032043 | 4.902 | 4.6783 | 0.018784 | -0.22371
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 51.446 | 51.446 | 0.13503 | 2.2888e-05 | 6.373 | 6.1723 | 0.010765 | -0.20073
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 24.725 | 24.7249 | 0.1211 | -8.7738e-05 | 6.599 | 6.2779 | 0.017082 | -0.32107
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 10.395 | 10.3947 | 0.097903 | -0.00033665 | 6.722 | 6.4598 | 0.044508 | -0.26224
2009-10-0012.mrg | 2009-10-0012.sam | 241 | 5.239 | 5.239 | 0.16664 | 4.1485e-05 | 6.86 | 6.5395 | 0.016358 | -0.32054
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 2003.65 | 2003.65 | 1.0063 | 0.00048828 | 1.451 | 1.2821 | 0.003906 | -0.1689
2009-10-0020.mrg | 2009-10-0020.sam | 307 | 1500.45 | 1500.62 | 0.48245 | 0.17639 | 0.681 | 0.61915 | 0.0025496 | -0.06185
2009-10-0020.mrg | 2009-10-0020.sam | 307 | 1500.67 | 1500.62 | 0.48245 | -0.049561 | 0.682 | 0.61915 | 0.0025496 | -0.06285
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 1250.17 | 1250.17 | 0.86145 | -0.00036621 | 0.355 | 0.33626 | 0.0025662 | -0.018739
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 1001.62 | 1001.62 | 0.54564 | -0.00048828 | 0.237 | 0.21613 | 0.0024958 | -0.020871
2009-10-0020.mrg | 2009-10-0020.sam | 112 | 800.414 | 800.329 | 1.1435 | -0.085388 | 1.005 | 0.21353 | 0.0022177 | -0.79147
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 599.855 | 599.855 | 0.23631 | -0.00018311 | 0.428 | 0.37322 | 0.0022839 | -0.054776
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 401.219 | 401.219 | 0.2799 | -0.00039673 | 1.1 | 1.0396 | 0.0021928 | -0.060402
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 299.827 | 299.827 | 1.0458 | 3.0518e-05 | 1.91 | 1.8098 | 0.0070129 | -0.10022
2009-10-0020.mrg | 2009-10-0020.sam | 146 | 249.435 | 249.318 | 1.1201 | -0.11693 | 2.034 | 1.9105 | 0.0063301 | -0.12346
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 200.212 | 200.212 | 0.88315 | 0.00041199 | 2.047 | 1.9582 | 0.0035649 | -0.088759
2009-10-0020.mrg | 2009-10-0020.sam | 85 | 173.907 | 173.88 | 1.1507 | -0.02684 | 2.147 | 2.0371 | 0.010025 | -0.10989
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 150.707 | 150.707 | 0.98167 | -0.00015259 | 2.333 | 2.1758 | 0.0089689 | -0.15724
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 124.044 | 124.044 | 0.4846 | -0.0001297 | 2.847 | 2.6831 | 0.012874 | -0.1639
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 101.474 | 101.474 | 0.27531 | -0.00027466 | 4.212 | 3.998 | 0.011501 | -0.21401
2009-10-0020.mrg | 2009-10-0020.sam | 179 | 75.31 | 75.2867 | 1.2255 | -0.0233 | 6.098 | 5.7981 | 0.037336 | -0.29987
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 49.808 | 49.8081 | 0.46081 | 9.5367e-05 | 6.48 | 6.2679 | 0.011322 | -0.21206
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 25.293 | 25.2931 | 0.16434 | 0.00012207 | 6.082 | 5.7126 | 0.015859 | -0.3694
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 10.31 | 10.3099 | 0.41325 | -5.4359e-05 | 5.971 | 5.7041 | 0.026635 | -0.26693
2009-10-0020.mrg | 2009-10-0020.sam | 241 | 5.353 | 5.3529 | 0.81521 | -6.2466e-05 | 5.967 | 5.7156 | 0.02113 | -0.25144
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 3276.19 | 3276.19 | 0.82445 | 0 | 2.344 | 2.1742 | 0.0052468 | -0.16977
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 3000.48 | 3000.48 | 0.49474 | 0 | 2.291 | 2.1445 | 0.0049557 | -0.14653
2009-10-0029.mrg | 2009-10-0029.sam | 385 | 2501.41 | 2501.59 | 0.59658 | 0.17773 | 1.982 | 1.7573 | 0.0045462 | -0.22471
2009-10-0029.mrg | 2009-10-0029.sam | 385 | 2501.7 | 2501.59 | 0.59658 | -0.10718 | 1.862 | 1.7573 | 0.0045462 | -0.10471
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 2000.27 | 2000.27 | 0.77713 | -0.00036621 | 1.298 | 1.2232 | 0.0036682 | -0.074834
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 1501.16 | 1501.16 | 0.59756 | 0.00024414 | 0.623 | 0.60927 | 0.0032809 | -0.01373
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 1250.57 | 1250.57 | 0.59708 | 0 | 0.445 | 0.38709 | 0.0031696 | -0.057913
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 1001.93 | 1001.93 | 0.46401 | 0.00042725 | 0.239 | 0.22771 | 0.0019075 | -0.011286
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 801.143 | 801.143 | 0.6649 | 0.00048828 | 0.311 | 0.26251 | 0.0024751 | -0.04849
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 600.801 | 600.801 | 0.31219 | 6.1035e-05 | 0.584 | 0.51929 | 0.0018098 | -0.064705
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 400.283 | 400.283 | 0.68747 | -0.00039673 | 1.482 | 1.3473 | 0.0028976 | -0.13473
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 301.716 | 301.716 | 0.23463 | -0.00042725 | 2.731 | 2.5586 | 0.0042794 | -0.17243
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 249.814 | 249.814 | 0.62569 | -3.0518e-05 | 3.497 | 3.273 | 0.0030968 | -0.22396
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 200.271 | 200.271 | 0.18418 | 0.00018311 | 4.291 | 4.0577 | 0.0056607 | -0.23326
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 176.597 | 176.597 | 0.39284 | -0.00036621 | 4.644 | 4.3958 | 0.005897 | -0.24818
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 150.549 | 150.549 | 0.61058 | -0.00041199 | 4.888 | 4.6352 | 0.0091057 | -0.25283
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 124.879 | 124.879 | 0.27473 | 0.00047302 | 5.346 | 5.0843 | 0.0094488 | -0.26169
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 100.671 | 100.671 | 0.29383 | 0.00011444 | 6.083 | 5.8131 | 0.0097689 | -0.26988
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 75.855 | 75.8548 | 0.43504 | -0.00016022 | 6.832 | 6.4835 | 0.015863 | -0.3485
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 50.956 | 50.9557 | 0.4349 | -0.00034332 | 7.066 | 6.7227 | 0.016748 | -0.3433
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 25.634 | 25.6344 | 0.26073 | 0.00043106 | 6.388 | 5.9998 | 0.01886 | -0.38818
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 9.846 | 9.8456 | 0.20419 | -0.00036907 | 5.919 | 2.3884 | 0.025795 | -3.5306
2009-10-0029.mrg | 2009-10-0029.sam | 241 | 5.841 | 5.8409 | 0.29992 | -9.1553e-05 | 5.926 | 2.106 | 0.0078225 | -3.82
2009-10-0037.mrg | 2009-10-0037.sam | 77 | 3500.16 | 3500.19 | 1.3474 | 0.03833 | 2.751 | 2.5276 | 0.0043321 | -0.22339
2009-10-0037.mrg | 2009-10-0037.sam | 233 | 2999.5 | 2999.58 | 1.3096 | 0.07544 | 2.466 | 2.2661 | 0.0053773 | -0.19986
2009-10-0037.mrg | 2009-10-0037.sam | 397 | 2497.47 | 2497.23 | 1.1098 | -0.23438 | 1.908 | 1.7816 | 0.0045839 | -0.12635
2009-10-0037.mrg | 2009-10-0037.sam | 395 | 2496.93 | 2496.9 | 1.0375 | -0.024658 | 1.907 | 1.7818 | 0.0046856 | -0.12515
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 2001.58 | 2001.58 | 1.1745 | -0.00024414 | 1.219 | 1.175 | 0.0032203 | -0.044012
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 1501.9 | 1501.9 | 0.70793 | 0.00012207 | 0.507 | 0.47704 | 0.0031055 | -0.029963
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 1251.29 | 1251.29 | 0.87408 | -0.00012207 | 0.311 | 0.30083 | 0.0028858 | -0.010174
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 1000.71 | 1000.7 | 0.48301 | -0.00036621 | 0.285 | 0.25587 | 0.0025214 | -0.029133
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 801.449 | 801.449 | 0.70801 | -0.00042725 | 0.388 | 0.36527 | 0.0024324 | -0.022726
2009-10-0037.mrg | 2009-10-0037.sam | 205 | 599.093 | 599.478 | 1.3289 | 0.38483 | 0.619 | 0.57156 | 0.002331 | -0.047439
2009-10-0037.mrg | 2009-10-0037.sam | 201 | 399.749 | 400.235 | 0.99769 | 0.48614 | 1.51 | 1.396 | 0.0041091 | -0.11402
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 300.436 | 300.436 | 0.52614 | 9.1553e-05 | 2.532 | 2.3084 | 0.0038714 | -0.22357
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 250.34 | 250.34 | 0.53921 | 0.00036621 | 3.185 | 2.9682 | 0.0091679 | -0.21684
2009-10-0037.mrg | 2009-10-0037.sam | 228 | 200.149 | 200.034 | 0.91388 | -0.11519 | 4.521 | 4.1863 | 0.02518 | -0.33474
2009-10-0037.mrg | 2009-10-0037.sam | 238 | 176.47 | 176.444 | 0.80401 | -0.026154 | 5.148 | 4.8064 | 0.0094576 | -0.34155
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 149.804 | 149.804 | 0.70377 | -6.1035e-05 | 5.822 | 5.4656 | 0.010897 | -0.35639
2009-10-0037.mrg | 2009-10-0037.sam | 116 | 126.487 | 126.568 | 1.2545 | 0.081024 | 6.604 | 6.2031 | 0.01604 | -0.40091
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 101.691 | 101.691 | 0.88376 | 0.00023651 | 6.813 | 6.4238 | 0.0076469 | -0.38922
2009-10-0037.mrg | 2009-10-0037.sam | 109 | 74.735 | 74.5608 | 1.2013 | -0.17425 | 7.005 | 6.5239 | 0.024834 | -0.48108
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 49.979 | 49.9787 | 0.51724 | -0.00030136 | 6.954 | 6.593 | 0.018445 | -0.36099
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 25.418 | 25.418 | 0.55624 | -3.2425e-05 | 6.024 | 5.6457 | 0.022196 | -0.37828
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 10.505 | 10.5048 | 0.49839 | -0.00022411 | 5.967 | 5.6245 | 0.016742 | -0.34247
2009-10-0037.mrg | 2009-10-0037.sam | 241 | 6.182 | 6.1822 | 0.43915 | 0.00023651 | 5.978 | 5.6415 | 0.017912 | -0.33653
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 4315.72 | 4315.72 | 1.0682 | 0.00048828 | 3.225 | 3.0101 | 0.0067191 | -0.21489
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 3998.97 | 3998.96 | 0.9869 | -0.00048828 | 3.199 | 2.9411 | 0.007228 | -0.25786
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 3501.9 | 3501.9 | 0.87917 | 0.00024414 | 2.955 | 2.7317 | 0.0053451 | -0.22328
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 3000.3 | 3000.3 | 0.73549 | -0.00048828 | 2.538 | 2.3371 | 0.0058576 | -0.20088
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 2500.89 | 2500.89 | 0.50414 | -0.00024414 | 1.953 | 1.7898 | 0.004994 | -0.16321
2009-10-0044.mrg | 2009-10-0044.sam | 228 | 2003.43 | 2003.55 | 1.1508 | 0.11585 | 1.303 | 1.1834 | 0.0039803 | -0.11965
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 1501.95 | 1501.95 | 1.2056 | 0.00048828 | 0.499 | 0.49161 | 0.0036682 | -0.00739
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 1250.92 | 1250.92 | 0.77534 | -0.00012207 | 0.342 | 0.32903 | 0.0029901 | -0.012967
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 999.715 | 999.715 | 1.0938 | 0.00042725 | 0.378 | 0.35476 | 0.0024273 | -0.023241
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 802.475 | 802.475 | 0.6178 | 6.1035e-05 | 1.054 | 0.4331 | 0.0019361 | -0.6209
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 599.257 | 599.257 | 0.19004 | -6.1035e-05 | 0.627 | 0.56649 | 0.0024068 | -0.060506
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 399.823 | 399.823 | 0.59737 | -0.00039673 | 1.176 | 1.0657 | 0.0028522 | -0.11028
2009-10-0044.mrg | 2009-10-0044.sam | 144 | 297.414 | 297.593 | 1.1634 | 0.17871 | 1.895 | 1.7515 | 0.0051562 | -0.14347
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 247.972 | 247.972 | 0.95337 | -1.5259e-05 | 2.757 | 2.5384 | 0.0075519 | -0.21861
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 200.803 | 200.803 | 0.8706 | 0.00028992 | 3.739 | 3.4453 | 0.013674 | -0.29367
2009-10-0044.mrg | 2009-10-0044.sam | 221 | 174.522 | 174.613 | 1.2048 | 0.090774 | 4.412 | 4.1073 | 0.0079549 | -0.30468
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 149.88 | 149.88 | 0.71767 | 0.00045776 | 5.407 | 5.0372 | 0.018226 | -0.36982
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 124.694 | 124.694 | 1.0512 | 0.00016785 | 6.527 | 6.0571 | 0.013726 | -0.46995
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 100.56 | 100.56 | 0.79801 | 9.1553e-05 | 6.928 | 6.4907 | 0.014136 | -0.43729
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 73.802 | 73.8024 | 0.87496 | 0.00035095 | 6.906 | 6.4701 | 0.01843 | -0.43592
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 50.696 | 50.6964 | 1.0482 | 0.00035095 | 6.913 | 6.4631 | 0.0187 | -0.44993
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 23.861 | 23.8606 | 0.74928 | -0.0004406 | 6.241 | 5.8293 | 0.013286 | -0.41173
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 11.107 | 11.107 | 0.41243 | 3.2425e-05 | 6.212 | 5.8095 | 0.018044 | -0.40253
2009-10-0044.mrg | 2009-10-0044.sam | 241 | 5.502 | 5.5023 | 0.42543 | 0.00030708 | 6.375 | 5.8033 | 0.018351 | -0.57167
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 4024.08 | 4024.09 | 0.86144 | 0.00048828 | 3.177 | 2.7851 | 0.0073442 | -0.39191
2009-10-0061.mrg | 2009-10-0061.sam | 292 | 3498.44 | 3498.86 | 0.46943 | 0.41992 | 2.881 | 2.6327 | 0.0069541 | -0.24833
2009-10-0061.mrg | 2009-10-0061.sam | 292 | 3498.95 | 3498.86 | 0.46943 | -0.091065 | 2.837 | 2.6327 | 0.0069541 | -0.20433
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 2998.76 | 2998.76 | 0.69856 | -0.00024414 | 2.484 | 2.3213 | 0.0052467 | -0.16272
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 2501.18 | 2501.18 | 0.86306 | -0.00024414 | 1.948 | 1.794 | 0.0040669 | -0.15396
2009-10-0061.mrg | 2009-10-0061.sam | 223 | 1999.73 | 1999.9 | 1.0618 | 0.16809 | 1.279 | 1.2011 | 0.004018 | -0.077852
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 1501.12 | 1501.12 | 0.4335 | 0.00024414 | 0.544 | 0.508 | 0.0028358 | -0.036
2009-10-0061.mrg | 2009-10-0061.sam | 176 | 1252.42 | 1252.06 | 1.0576 | -0.35718 | 0.34 | 0.33658 | 0.0031631 | -0.0034205
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 1001.45 | 1001.45 | 0.98054 | 0.00042725 | 0.321 | 0.31054 | 0.0026972 | -0.010465
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 801.223 | 801.223 | 0.90857 | 0 | 0.471 | 0.42635 | 0.0021842 | -0.044647
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 601.026 | 601.026 | 0.65532 | 0.00024414 | 0.6 | 0.5665 | 0.002174 | -0.033502
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 399.824 | 399.824 | 0.75186 | -3.0518e-05 | 1.295 | 1.1597 | 0.0042208 | -0.13529
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 300.177 | 300.177 | 0.88513 | -0.00021362 | 2.086 | 1.884 | 0.0040253 | -0.20196
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 250.475 | 250.475 | 0.68298 | -0.00033569 | 2.666 | 2.4496 | 0.0055907 | -0.21636
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 198.379 | 198.379 | 0.84765 | 0.00045776 | 3.921 | 3.6022 | 0.0062745 | -0.31876
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 175.41 | 175.41 | 1.0878 | 0.00038147 | 4.395 | 4.0755 | 0.013949 | -0.31946
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 150.732 | 150.732 | 0.68734 | -0.00039673 | 5.434 | 4.9887 | 0.043861 | -0.44525
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 124.125 | 124.125 | 0.85266 | -0.00020599 | 6.29 | 5.7935 | 0.028387 | -0.49652
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 99.492 | 99.4924 | 0.51298 | 0.00036621 | 6.727 | 6.2541 | 0.018472 | -0.47289
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 74.779 | 74.7791 | 0.60745 | 8.3923e-05 | 6.843 | 6.3995 | 0.023285 | -0.44349
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 49.306 | 49.3064 | 0.55475 | 0.00039673 | 6.936 | 6.4517 | 0.029448 | -0.48427
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 24.703 | 24.7034 | 0.44537 | 0.00043297 | 6.108 | 5.6485 | 0.011864 | -0.45955
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 9.733 | 9.7331 | 0.69777 | 0.00011635 | 6.104 | 5.6612 | 0.021845 | -0.44281
2009-10-0061.mrg | 2009-10-0061.sam | 241 | 5.196 | 5.1962 | 0.25438 | 0.00017405 | 6.111 | 5.6572 | 0.024126 | -0.45378
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 1998.83 | 1998.83 | 0.82366 | 0.00024414 | 1.29 | 1.1924 | 0.0047216 | -0.097585
2009-10-0071.mrg | 2009-10-0071.sam | 127 | 1517.89 | 1518.6 | 1.0538 | 0.71167 | 0.563 | 0.54876 | 0.0029639 | -0.014236
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 1249.68 | 1249.67 | 0.45674 | -0.00048828 | 0.311 | 0.32417 | 0.0025906 | 0.013174
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 999.208 | 999.208 | 0.357 | 0.00018311 | 0.254 | 0.27827 | 0.0030413 | 0.024274
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 799.271 | 799.271 | 0.50629 | -0.00012207 | 0.388 | 0.35488 | 0.0025987 | -0.033116
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 599.385 | 599.385 | 0.43077 | 0 | 0.567 | 0.54376 | 0.0024845 | -0.023237
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 402.41 | 402.41 | 0.57904 | -0.00039673 | 1.102 | 1.0239 | 0.0073432 | -0.07805
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 301.932 | 301.932 | 0.53611 | 0.00039673 | 1.547 | 1.4423 | 0.0027367 | -0.10473
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 250.688 | 250.688 | 0.97115 | -7.6294e-05 | 1.821 | 1.6911 | 0.006699 | -0.12986
2009-10-0071.mrg | 2009-10-0071.sam | 207 | 199.994 | 199.64 | 1.3391 | -0.35413 | 2.229 | 2.0769 | 0.0036755 | -0.15214
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 175.483 | 175.483 | 0.89544 | 0.00035095 | 2.455 | 2.3036 | 0.0074563 | -0.15136
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 150.246 | 150.246 | 0.64677 | 1.5259e-05 | 2.767 | 2.5619 | 0.0096489 | -0.20505
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 124.167 | 124.167 | 0.72301 | 0.00040436 | 4.164 | 3.7627 | 0.032554 | -0.40133
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 99.525 | 99.5251 | 0.99821 | 0.00011444 | 5.715 | 5.319 | 0.016428 | -0.39605
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 74.344 | 74.3439 | 0.33272 | -9.1553e-05 | 6.581 | 6.1191 | 0.011027 | -0.46193
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 49.331 | 49.3314 | 0.90908 | 0.00035095 | 6.946 | 6.4523 | 0.025608 | -0.49372
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 40.998 | 40.9976 | 0.65047 | -0.0003624 | 7.004 | 6.5152 | 0.035805 | -0.48879
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 29.604 | 29.6041 | 0.47331 | 0.00010872 | 6.08 | 5.6367 | 0.0091063 | -0.44334
2009-10-0071.mrg | 2009-10-0071.sam | 482 | 22.851 | 22.1111 | 0.82155 | -0.73991 | 6.069 | 5.612 | 0.025038 | -0.45695
2009-10-0071.mrg | 2009-10-0071.sam | 482 | 21.371 | 22.1111 | 0.82155 | 0.74009 | 6.067 | 5.612 | 0.025038 | -0.45495
2009-10-0071.mrg | 2009-10-0071.sam | 241 | 10.624 | 10.6242 | 0.86183 | 0.00022411 | 6.01 | 5.5937 | 0.019454 | -0.41628
2009-10-0071.mrg | 2009-10-0071.sam | 305 | 4.899 | 4.5235 | 0.74701 | -0.37555 | 6 | 5.5706 | 0.032389 | -0.42942
2009-10-0071.mrg | 2009-10-0071.sam | 285 | 2.717 | 3.0114 | 0.73559 | 0.2944 | 6.028 | 5.5692 | 0.03293 | -0.45884
Create Fit Sheet