## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 24 Nov 2021 | ### Corrected Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 17 Dec 2020 | ### Added HPLC Data. S.H. |
| 18 October 2010 | Recalibrated Transmissivity Data – see end of report for details. |

## PROCESSING NOTES

Cruise: 2009-14

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdeF

Party Chief: Chandler P.

Platform: Vector

Date: September 16, 2009 – September 20, 2009

Processed by: Germaine Gatien

Date of Processing: January 7, 2009 – May 20, 2009

Number of original CTD casts:  73	Number of CTD casts processed: 73

Number of bottle casts:		21	Number of bottle casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2356) with a 3X or 10X cable, a Biospherical QSP-400 PAR sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter (#1252).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book lacked an equipment list. Since the fluorometer was not entered into the configuration file, there is no way to be sure which instrument was used. For the two previous cruises using CTD #0550 the fluorometer was said to be #2355, but there is no record of such a fluorometer at IOS, and Scott Rose believes it was #2356. The log was otherwise in good order with notes about problems encountered during the cruise. The rosette sheets include many comments, mostly concerning DO analysis.

Cast #26 contained only upcast data; there was no mention of this in the log so it is unknown what caused the problem. Upcast data were prepared for archiving, but the quality is lower than for other casts, because of the shed wakes caused by the rosette package and the many stops for sampling.

Secondary temperature and salinity were chosen for archiving because the primary temperature data were frequently poor. The comparison with bottles suggests that the secondary salinity is low by about 0.002, but the fit was too noisy to justify recalibration. During cast #61 the secondary pump did not work well, so primary channels were used for that cast only; the primary salinity was recalibrated to bring it into line with the secondary salinity.

For cast #66 acquisition did not begin until the CTD was at 50db. An attempt was made to patch the top 50db from the upcast to the downcast data, but this was unsatisfactory because the upcast data was extraordinarily noisy with many stops for bottles and very noisy data between those stops. So only the downcast data from 50db to the bottom were prepared for the archive.

Sea-Bird have a new algorithm for dissolved oxygen with several parameters that the manufacturer recommends be fine-tuned for each instrument to produce the best data. This requires bottle samples from deep casts. Since these new parameters became available there has been no cruise using this sensor that has sufficient deep sampling to do this, so nominal values were used. The hysteresis correction was not applied since there are no data below 500db. The Tau correction in the derivation of dissolved oxygen concentration is known to exaggerate residual noise in deep water, but that should not be a problem for this cruise, so that correction was applied.

Fluorescence and Transmissivity data are nominal and unedited except that some records were removed in editing temperature and salinity.

For the cruise as a whole, thhe SBE dissolved oxygen data in the CTD files are considered, roughly:

•	±0.6  ml/l from  0– 25db

•	±0.2  ml/l from 25– 75db

•	±0.1 ml/l below 75db

For cast #1 in Saanich Inlet, the near-surface temperature and DO gradients were very high and the SBE Dissolved Oxygen values are likely too high by up to 1mL/L near 5db.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. A few files had errors in the names; these were corrected.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

A number of problems were noted including the observation that the end cap of bottle #10 was closing due to bungee stretching for the last rosette cast.

There were problems with the primary temperature from cast #45 onwards that may have been due to the CTD hitting the deck a few times.

Extracted chlorophyll, dissolved oxygen, nutrients and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

All the CON files produced in acquisition were the same. However, according to the log the fluorometer cable was 3X for cast #1 and 30X thereafter. The log notes that the configuration file was adjusted, but all the configuration files have 30X. For now it will be assumed that it should be 3X for cast #1 only and 30X for all other casts, but after conversion plots will be examined and when CHL data are available checks will be made to see if this assumption was correct. We can recalibrate FL later if necessary. The fluorometer serial number was not recorded in the configuration file and there is no equipment list in the log book. When this equipment was last used the fluorometer was identified as #2355, but that sensor is unknown. Scott Rose confirms that the serial number was 2356.

The two files to be used are as follows:
       2009-14-ctd1.con – for fluorescence gain 3X: casts  #1

2009-14-ctd2.con – for fluorescence gain 30X: casts  #2-73.

-

The calibration constants were checked for all instruments. There were a number of problems:

- The configuration for DO #0119 used the Owens-Millard parameters. It is better to select the Murphy-Larson parameters, so that the Tau correction can be applied. The nominal values were entered for the E, H1 and H3 since there is insufficient deep sampling to fine-tune them. It is not expected that the hysteresis correction will be applied since there was no very deep sampling.

- The transmissivity date and parameters were wrong as they have been for a few recent cruises. For those other cruises checks were made to see if the serial number might be wrong. The only transmissivity calibrations for that date had different values and no calibration could be found with those values. It will be assumed that this really is 1005DR and the appropriate numbers were entered in the con files.

- The pressure offset for CTD #0550 was entered as 0.1033. In other recent cruises a value of +0.5db was found appropriate, though for 2009-38 it was noted that it might be time to increase the offset a little. Examination of a few test conversions for this cruise suggests a value of +0.7db is probably appropriate because there are clear “in-water” salinity values at pressure -0.5db when the offset was entered as 0.1db. That implies at least 0.6db is needed; including an extra 0.1db is a good idea to avoid losing any useful surface data. So the offset was changed to +0.7db.

##### Conversion of Raw Data

Rosette files were converted using a start time of -5s and duration of 10s and using 2009-14-ctd1.con for cast #1 and 2009-14-ctd2.con for all other casts. The Tau correction was chosen, but not the hysteresis correction for DO calculation.

CNV files were converted using the configuration files 2009-14-ctd1.con for cast #1 and 2009-14-ctd2.con for all other casts. Based on rosette log notes, the bottom depth was changed in the CNV and ROS files for cast #68.

A few CNV casts were examined and all expected channels are present.

- There are no obvious spikes in pressure, but it is noted that there was an initial downcast section in cast #1 that will need to be removed before running DELETE.

- The two temperature channels are mostly in reasonable agreement on the downcasts though there are some excursions. The upcast data are much noisier with some casts especially so. Conductivity is similar. While the log notes that the primary temperature “crashed” I can’t see that in this data – just a lot of noise and it is not always clear which sensor is better.

- The fluorescence values look reasonable, so the gains entered in the two con files are probably correct.

- Dissolved oxygen voltage looks as usual with an offset but some detail to help alignment.

- The descent rate looks reasonably steady except near the mouth of Juan de Fuca Strait where it is very noisy.

- PAR, Surface PAR and transmissivity look fine.

- The altimetry looks good near the bottom.

The rosette files were then converted to IOS SHELL files. The files were renamed as *.BOT. All BOT files were plotted. There were a few outliers in salinity in casts 13 (secondary), 48 (secondary) and 63 (primary). CTDEDIT was used to clean those and the output was copied to *.BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Tests were run on 4 casts using a variety of settings for CELLTM. The differences among the various choices were small. Overall the best choices were found to be (α = 0.0245, β=9.5) for the primary and was (α = 0.02, β=7) for the secondary, the same as found during 2009-26 when the same CTD sensors were used.

CELLTM was run applying those settings to all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts using was plotted to check for agreement between the pairs of T and C sensors. The differences are extremely noisy despite steady descent rates, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 4 | 275 | +0.0015 | -0.00015 | -0.0028 | High, steady |
| 63 | 275 | +0.002 | ~0 VN | -0.0026 VN | High, steady |
| 74 | 275 325 400 | +0.0015 N +0.001 VN +0.0015 N | ~0 N -0.0002 -0.0001 | -0.002 -0.0023 N -0.002 | Moderate, steady |
| 73 | 275 325 | +0.0015 +0.0013 | -0.00004 -0.00008 | -0.0017 -0.0017 | Mod., v.steady |

As was the case for cruise 2009-26 with the same sensors, the temperature differences are larger than usual while conductivity differences are very small.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. No problems were found.

The cross-reference check was compared with the log book, and some errors were found in station names; those were corrected in the IOS, CLN and BOT files. The times and positions vary slightly for many casts, probably because the log data were not entered at the same time as the data acquisition began. Where times do agree closely to the log, the positions are usually very close. It will be assumed that the NMEA header positions and times are correct since the differences are not large.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run and shows the average surface pressure to be 2.3db, typical of the Vector. The lowest reading was 0.64db and the salinity was 0.12 so the CTD really was very close to the surface in that case. This suggests that the pressure offset is ok.

The altimeter readings and water depths were exported from the headers to a spreadsheet. Plots were made of altimetry for a random selection of casts to see if the entries were reasonable and they were fine, even when the descent rate was very noisy. The bottom depths (called water depths in CTD headers) were checked against log entries and there were many discrepancies – where differences were >3db checks were made. In some cases there may have been drift during the cast, so differences do not necessarily indicate error, but we expect the log entries and header entries to be the same.

- Sometimes it looked like a simple typo or forgetting to update, since the log entries for maximum depth sampled and bottom depth are consistent with the maximum depth in the CTD data (estimated from pressure) plus altimeter header entry. The headers were fixed to match the log entries for bottom depth for those cases.

- For a few cases the maximum depth in the log is not in close agreement with the CTD maximum depth, nor is the water depth in the file in agreement with the bottom depth in the log book. These cases are all from a single watch, so it is assumed there was an error in the method used to obtain these data. An estimate was made for bottom depth as the sum of the CTD maximum depth plus the altimeter bottom reading.

These changes were made to both the CLN files and the BOT files.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There were a number of cases of more bottles being fired than are listed on the rosette sheets – these were surface samples, so were removed from the addsamp file assuming they were accidental firings. The addsamp.csv file was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

The salinity data were delivered in spreadsheet 2009-14-salinity.xls; there were no duplicates. The spreadsheet was simplified (unneeded columns removed and headers changed to standard format) and saved as 2009-14-sal.csv which was then converted to individual SAL files. The file for event #7 was misnamed as #6 – that was corrected.

DISSOLVED OXGYEN

The dissolved oxygen data was provided in spreadsheet 2009-14oxy.xls with quality flags and comments. There was an analysis of duplicates. The spreadsheet was simplified and saved as 2009-14oxy.csv which was converted into individual ADD files. These data were reviewed to check

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2009-14nuts.xls which included a report on precision. The spreadsheet was simplified and saved as 2009-14-nuts.csv. File 2009-14-nuts.csv was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file 2009-14-chlarc.xls with c flags assigned where the duplicates differed by >10%. The file was simplified, event numbers were added. All data were duplicates so “f” flags were added to all.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

Compare was run and shows a lot of scatter in the differences between bottles and CTD for both sensor pairs. There is little difference in shape between the primary and secondary plots. Below 300db the differences are less scattered but there are only 7 bottles. For the primary the range of differences is +0.005 to -0.004 with an average difference -0.0004. For the secondary the range is +0.003 to -0.006 with an average of -0.0023. There is never more than 1 sample per cast, so it is impossible to separate the effect of pressure dependence from the fact that the shallower casts are in areas with more active mixing. There is a hint of differences being larger towards the end of the cruise, but that is more likely to be a geographic variation as drift in a sensor. Casts #45 and 48 in the middle of Georgia Strait are probably the most reliable for calibration purposes; the primary are either high by 0.002 or low by 0.0025 and the secondary high by 0.001 or low by 0.0046. So even in the best possible conditions, the scatter is fairly high. There were no notes about sampling or analysis problems. Recalibration is not justified except to bring the two channels into rough agreement. (See 2009-14-sal-comp1.xls.)

Dissolved Oxygen –

COMPARE was run using the SBE DO and the Titrated bottle DO data.

Outliers were gradually removed until a fit was found with a small offset:

CTD-BOT = 1.0414 DOX-CTD + 0.0005

A linear fit suited these data well. When used during 2009-26 (April) the results for the same sensor were:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

These are reasonably close suggesting little drift in the sensor.  (See 2009-14-dox-comp1.xls.)

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity.

The following outliers were considered for flags:

- Cast #1 – 5m and 1.7m. These bottles are from the surface of Saanich Inlet where the SBE DO goes down by more than 4mL/L between 5 and 10db. The large differences from bottles are probably due to the sensor not being able to resolve this gradient. The bottles will not be flagged.

- Cast #7 – The bottles at 126 and 177db are out of line in COMPARE. There is no obvious evidence of problems with the CTD, but there is some structure in the DO at the level of these bottles that the sensor might not have resolved well. These two bottles will not be flagged.

- Cast #45 – The surface bottle is out of line, but the CTD data are noisy there. The bottle will not be flagged.

All the bottles flagged “b” by the analyst were checked and none stood out as outliers in COMPARE, so the flags were removed. In most cases they had been flagged because of low surface values, but the CTD data confirm that surface values were low in Haro Strait and the eastern part of Juan de Fuca Strait.

A final run of COMPARE was used to check on the data from Niskin Bottle #10. The only significant outlier had already been flagged.

Fluorescence

COMPARE was run using the CTD CHL and the Titrated Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD Fluorescence. As mentioned earlier it is not completely clear that the gain has been entered correctly.

- Cast #1 shows the ratio of FL to Extracted CHL as ~1.4 at 20db and very close to 1 at 2db and 10db. So the 3X gain used for conversion for that cast does appear to be correct. The configuration used at sea had 30X entered and if that were correct the fluorescence values would be lower by a factor of 10 which does not look right.

- For cast #7 the extracted CHL at the surface is ~15ug/L but the SBE fluorescence is <1. If the cable were really 10X or 3X then the fluorescence would be higher by a factory of 3 or 10, but that would make the other 2 bottles look out of line. With a gain of 30X the fluorometer would go off scale at 5ug/L and there is no sign of that happening on either the downcast or upcast which both had maxima of ~3ug/L. The Nitrate and Phosphate values are very low for the bottle with the high CHLa. The CHLa is an average of duplicates and was not flagged, so there was reasonable agreement. There were odd conditions at station P4 in early September during cruise 2009-10 with many euphausids in the water, but the CHL/FL ratio was close to 1 there. However, this does suggest there could be unusual patchy conditions near the mouth of Juan de Fuca.

For most of the casts in Juan de Fuca Strait the fluorescence is lower than the extracted CHL, while the opposite is seen in the Strait of Georgia, particularly to the north. A similar pattern was noted during 2009-43 in June though the ratios were different. (See 2009-14-fl-chl-comp.xls.)

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always rough estimate as the upcast data are usually very noisy. The usual shift of +24 records (1s) was found to improve the alignment. This is the shift that has been used in most other cruises, and it was applied to these data. (Output: SHFFL)

Conductivity

Tests were run on a few casts and the best choice varied found 0 to -0.2 for the primary conductivity channel. For the secondary the best choice varied from -0.2 to -0.6 with -0.4 best overall.  SHIFT was not run on the primary, and was run on the secondary with a setting of -0.4 records. (Output: *.SHFC1)

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. In other recent uses a value of +70 was used, but during 2009-26, 2009-34 and 2009-38 using the same sensor a setting of +65 looked better. A setting of +65 looked best overall for this cruise too.

SHIFT was run using +65 records for all casts.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for the following casts:

- #1 – There was an initial partial downcast; this data was removed from the SHFO file and it was put through DELETE again.

- #26 – This is an upcast only. There is no note in the log about this, so it is assumed no one noticed what happened. The SHFO file was put through REVERSE and then through DELETE. A warning was entered into the header since the upcast data will not be of the same quality especially since there were many stops for rosette sampling.

- #48 – The warning concerns surface data at the end of the upcast, so this is of no concern.

There were no further warnings when the two amended files were rerun in DELETE.

There was no warning, but cast #66 has no downcast data in the top 50db. There is no indication in the log book about what went wrong, though wind and the wire dragging at the surface are mentioned. The cast was put through DELETE with output DELREV so that upcast data could be edited.

##### DETAILED EDITING

The primary salinity is closer to the bottles, on average, though there is a lot of noise in the comparison. However, the primary salinity is much noisier with odd spikes not seen in the secondary and later in the cruise there were problems with the primary temperature. So, secondary channels were chosen for archiving, except for cast #61 which appeared to have a problem with the secondary pump.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

The following cast required no editing: 64.

Upcast data were used for the following casts: 26. The data quality is lower than usual for that cast.

Examination of the upcast data for cast #66 showed an extraordinarily noisy trace. The low ascent rate and many stops for bottles make it impossible to distinguish good records from those corrupted by shed wakes. An attempt was made to join the top 50db from the upcast to the downcast data from 50db downwards, but the noise was especially bad in the top section because there were 7 bottle stops in the top 50db. It was decided to use only the downcast data.

As usual for the SoG/JdF cruises, the heaviest editing was required for the casts near the mouth of Juan de Fuca Strait. The descent rate was generally quite steady elsewhere except near the surface and bottom of the casts.

##### Initial Recalibration

File 2009-14-recal1.ccf  was prepared to subtract 0.002 from the primary salinity and to apply the following recalibration to the SBE dissolved oxygen:

CTD-BOT = 1.0414 DOX-CTD + 0.0005

This was applied to the SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the corrections were applied correctly and they were. (See 2009-14-dox-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

The recalibration of the primary salinity is only relevant to cast #61 since that channel will be removed from all other casts. This will bring the primary into line with the secondary based on the results of section 7 and COMPARE.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data is compared to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When obvious outliers are excluded, the differences range from -0.2 to +0.2ml/l. When a few more outliers are removed based on residuals, the fit against pressure is:

DOX(corrected) =  DOX +0.0001*Pressure -0.0237

That is very close to the correction done for 2009-26:

DOX(corrected) =  DOX +0.0001*Pressure -0.0315

However, both are rough fits as the choice of outliers influences the result significantly. The CTD values are high by an average of ~0.02mL/L with values too high near the surface and a little low at depth. There are many surface bottles, so those values have more influence than may be justified. There is no obvious dependence on DO values. As was decided for 2009-26 when the same sort of differences and scatter were found, further recalibration is not justified. (See 2009-14-dox-comp3.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 200db and processed separately for A. Peña. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  The primary conductivity sensor was found to be low by 0.004 salinity units when used for 2009-03 (Feb) but the sampling was limited to 2 depths. During 2009-27 (March) and 2009-38 (May) there was no calibration sampling. During 2009-26 in April it was found to be high by 0.0025 but there was a lot of scatter in the comparison. The secondary conductivity sensor was used for the same cruises and was found to be low by <0.001 for 2009-03 and high by 0.0004 for 2009-26. It was also used on an SBE25 in 2007 but there were doubts about the bottle analyses for those cruises.

2. Dissolved Oxygen – There are only 3 other cruises since the last factory calibration from which there is DO calibration sampling and both had few samples. For 2009-38 some samples came from areas of high variability in DO. The fits from those were:

CTD-BOT = 1.0209 DOX-CTD + 0.0668 (2009-27 - March)

CTD-BOT = 1.0388 DOX-CTD + 0.0088 (2009-26 – April)

CTD-BOT = 1.0700 DOX-CTD – 0.0887 (2009-38 - May)

3. Pressure –The sensor was recalibrated in August 2007 and an offset of +0.5db has been used for the 6 cruises processed since then. However, it was noted that it might soon be time to increase the offset, so that was done for this cruise.

Historic ranges (3 standard deviations) – Profile plots were made with historic ranges of T and S superimposed. For most casts north of Lasqueti Island the deepest parts of the casts had temperatures below the range minima. There were 2 casts at the mouth of Juan de Fuca where temperatures went above the range maxima. Salinity was a little high for one cast and low for another. These excursions are presumed to reflect real conditions rather than calibration problems since they go in both directions.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts except #61: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate, Bottle_Number and Flag.

The following channels were removed from cast #61: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate, Bottle_Number and Flag.

A second SBE DO channel (with umol/kg units) was added. REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files are considered, roughly:

•	±0.6  ml/l from  0– 25db

•	±0.2  ml/l from 25– 75db

•	±0.1 ml/l below 75db

A separate note was place in the header of cast #1 to say that the SBE DO is likely high by up to 1mL/L around 5db because of the very high temperature and DO gradients at that level.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from 65% to 150%. The only very high reading was 150% in Saanich Inlet. The local gradient was very high and the CTD reads higher than the bottle, so the SBE: DO value is probably a little too high, but even the bottle shows very high values. The lowest saturations were in Haro Strait, Juan de Fuca Strait and well-mixed casts in narrow passages at the northern end of the survey. Values in the Strait of Georgia were generally close to 100%.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

2. Log mentions change of cable from 3X to 30X. This is said to have been entered in the con file, but the con file for cast #1 indicates 30X and there was no change. So presumably there was a different con file that was changed, not the one used in acquisition.

45. Primary Temperature sensor crashed at 100m during upcast and again for the top 3 bottles. There had been a some hard deck landings. It was flushed and tube was removed at the inlet after the cast.

46. Temperature seemed better except at the surface.

48. Bad temperature difference.

66. Wind, wire dragged at surface. Problems with DO sampling - bubbles in Alk Iod.

68. Bad temperature on upcast. Bottom depth not changed. Bottle 10 – end cap closes at times due to bungee stretch. Flag all bottle 10.

REVISION – October 18, 2010

After these data had been processed and archived it was discovered that the calibration for transmissometer #1005DR had been drifting significantly since about September 2008. The drift was fairly steady, but there was a sudden drop to very low values around August 2009, followed by a further steady drift.  A correction factor was derived by forcing deep offshore data for 2 cruises in 2010 to have transmissivity of 62%/m and then applying a linear fit to time.

For full details on how the correction was derived see the document:

Osd_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissivity 1005DR Corrections.doc

The correction produces more realistic values in deep water. Given that this cruise occurred close to the time of the sudden drop, there is some doubt about this step, so values must be considered to be very rough estimates.

It should also be noted that this transmissometer produced values decreasing slightly from about 500m downwards in 2010, so there may be a pressure related problem with the sensor.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| PAR | 4565 | 29Jan2009 | IOS |  |  |
| Surface PAR | 16504 | 24Feb2009 |  |  |  |
| Fluorometer | 2356 |  |  |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |