## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 24 Nov 2021 | ### Corrected Salinity:Bottle precision that was lost during HPLC addition. S.H. |
| ### 18 Dec 2020 | ### Added HPLC Data. Changed and flagged a few Turner Chl values. S.H. |
| 3 May 2010 | See note at end for explanation of problem that led to many bad DO bottle values. |
| 18 December 2009 | Replaced Oxygen:Dissolved data in CHE files with corrected values. See note at end for details. |
| 18 August 2010 | Corrected transmissivity; see end of report for details. DO sample data restored in file 2009-26-0011.CHE |

## PROCESSING NOTES

Cruise: 2009-26

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdeF

Party Chief: Chandler P.

Platform: Vector

Date: April 11, 2009 – April 15, 2009

Processed by: Germaine Gatien

Date of Processing: June 22, 2009 – July 2, 2009

Number of original CTD casts:  72	Number of CTD casts processed: 72

Number of bottle casts:	21		Number of bottle casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2355) with a 3X or 10X cable (DO and FL both on the primary pump), a Biospherical QSP-400 PAR sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter (#1252). The deck unit was a model 911+ (#0424).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Log Book and rosette log sheets were generally in good order with notes about problems encountered during the cruise, bit a few entries were ambiguous.

The calibration data originally entered for the dissolved oxygen sensor were incorrect. The same values were used for a series of cruises in spring 2009 and come from a pre-repair calibration; the post-repair calibrations should be used. An error was made in the configuration file when the fluorometer cable was changed from 10X to 3X, with 30X being entered. This must have been confusing at sea, and appears to have led to an inappropriate change back to a 10X cable. Later the 3X cable was reinstalled. The errors are fixed in processing so that the concentrations are correct, but the choice of an inappropriate cable can’t be fixed. For cast #39 it is not certain what cable was used. The fluorescence is very low compared to nearby casts with either a 3X or a 10X gain, but 3X is closer to the extracted chlorophyll data, so that was selected for conversion.

Sea-Bird have a new algorithm for dissolved oxygen with several parameters that the manufacturer recommends be fine-tuned for each instrument to produce the best data. This requires bottle samples from deep casts. There has been no cruise (since these new parameters became available) using this sensor that has sufficient deep sampling to do this, so nominal values were used. The hysteresis correction was not applied since there are no data below 500db. The Tau correction in the derivation of dissolved oxygen concentration is known to exaggerate residual noise in deep water, but that should not be a problem for this cruise, so that correction was applied.

One data file had been overwritten by data from another cast. Fortunately the chief scientist had backed up the first file, so it was not lost.

There were many errors in the NMEA dates; sometimes the error was 24 hours, sometimes 48 hours.

The file names were non-standard.

For the first time dissolved oxygen data were delivered by the analyst in spreadsheet format (with quality flag and comment columns). This is a very convenient format for processing rather than having the analysts prepare individual ADD files which often have format errors that are time-consuming to fix.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. A number of problems were noted.

Extracted chlorophyll, dissolved oxygen, nutrients and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

There were two instrument configurations used during the cruise due to the switch in fluorescence cable. The two files are to be used as follows:
       2009-26-ctd1.con – for fluorescence gain 10X: casts  #1-38 & 50-51

2009-26-ctd2.con – for fluorescence gain 3X: casts  #40-49 & 52-74

NOTE: The choice for cast #39 is uncertain, so 2 conversions were done, one with each choice. They will be compared later.

The calibration constants were checked for all instruments. There were a number of problems:

- The configuration for DO #0119 came from a calibration done before the membrane was replaced. There is a later calibration that should be used. It is also better to select the Murphy-Larson parameters, so that the Tau correction can be applied. The nominal values were entered for the E, H1 and H3 since there is insufficient deep sampling to fine-tune them. It is not expected that the hysteresis correction will be applied.

- The transmissivity date and parameters were wrong in 2009-26-ctd1.con, so checks were made to see if the serial number might be wrong. The only transmissivity calibrations for that date had different values and no calibration could be found with those values. It will be assumed that this really is 1005DR since it is attached to the rosette itself, so not likely to have been changed when the CTD was changed.

- The pressure offset for CTD #0550 was changed from 0.1033 to +0.5db which is the value found on other recent cruises.

- The parameters for the Surface PAR were confirmed correct by Scott Rose.

##### Conversion of Raw Data

Data were converted using the configuration files 2009-26-ctd1.con and 2009-26-ctd2.con.

A few casts were examined and all expected channels are present.

- There are some spikes in pressure, temperature and conductivity.

- The two temperature channels are in reasonable agreement on the downcast. The upcast data are much noisier so there are significant differences. Conductivity is similar.

- The fluorescence is frequently off-scale. Cast #39, station 46, looks oddly low compared to nearby casts, though it is higher at around 40db. The conversion using 3X instead of the 10X gain (which was used at sea) still looks out of line, though less so. The notes for this cast are quite unclear about whether the cable was changed before or after, though the con file was definitely changed after the cast. For now, the data with file named 2009-26-0039 has the 10X fluorometer gain and a second conversion (called 2009-26-9039) has a 3X gain for fluorescence so this can be investigated further later. Perhaps something was wrong with the fluorometer or this was just an unusual cast. The temperature, salinity, DO and transmissivity all look a little out of line with the nearby casts as well. (NOTE: the final file named 2009-26-0039 has a 3X gain as this was later determined to be the best choice.)

- Dissolved oxygen voltage looks as usual with an offset but some detail to help alignment.

- PAR, Surface PAR and transmissivity look fine.

- The altimetry looks good near the bottom.

Rosette files were converted using a start time of -5s and duration of 10s and bottom depths were corrected in 2 casts. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers with output *.BOT. All BOT files were plotted. There were a few outliers in salinity in casts 57, 60 and 64. CTDEDIT was used to clean those and the output was copied to *.BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on 4 casts using a variety of settings for CELLTM. The differences among the various choices were small. Overall the best choices were found to be (α = 0.0245, β=9.5) for the primary and was (α = 0.02, β=7) for the secondary.

CELLTM was run applying those settings to all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts using was plotted to check for agreement between the pairs of T and C sensors. The differences are extremely noisy despite steady descent rates, so these are very rough averages and in some cases no reasonable estimate could be made.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 4 | 200 250 | +0.0013 +0.002 XN | -0.0002 -0.0002 | +0.0067 -0.002 VN | Steady |
| 63 | 200 250 | +0.001 VN +0.001 | -0.00015 -0.00010 | +0.009 -0.012 | Steady |
| 74 | 200 250 300 | Too noisy +0.001 XN +0.001 | -0.0004 XN -0.0002 XN -0.00012 | Too noisy -0.0026 -0.0022 | Steady |

The temperature differences are larger than usual while conductivity differences are quite small.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. Missing cast #s included 25, 58 and 62. The first two were expected, but #62 is listed in the log book with no mention of problems. The headers of file #61 match the log details for event #62, and not #61. It is likely that the data from event #61 was overwritten. File #61 was renamed as #62 and the event number changed. The chief scientist found back up copies of the files from event #61.

The cross-reference check was compared with the log book, and a few errors in station names were found and corrected. The dates are wrong for casts 15-60 and 74, sometimes out by one day, sometimes by two days. Similar time problems have been seen before. The CLN files were copied to ATC and then ADD TIME CHANNEL was used to advance the header times by 24 or 48 hours as needed for the casts with errors. The cross-reference check was rerun and no further errors found.

The dates were also fixed in the BOT files.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run and shows the average surface pressure to be 2db, typical of the Vector. A few casts were examined in detail and there are some cases of negative pressures with “in-water” transmissivity and conductivity. In some cases these look like pressure spikes. At the end of the upcast for event #8 there are some negative pressures. Between +0.1db and -0.1db the conductivity is sometimes low enough to suggest the CTD was out of the water, but it is sometimes clearly in the water. This suggests that it was very close to the surface and perhaps being sprayed. The transmissivity goes to zero at pressure ~+0.1db with higher transmissivity values at -0.1db. It looks as if the pressure is within 0.1db.

The altimeter readings and water depths were exported from the headers to a spreadsheet. Plots were made of altimetry for many casts to see if the entries were reasonable and they were for all except cast #1 when the CTD did not get close to the bottom and a spike was misinterpreted. The altimeter header was removed for that cast. The water depths were checked against log entries and the one for cast #11 was changed to match the log. (The discrepancy was noted in the log.)

The cross-reference list, track plots and header check were rerun after all corrections and the addition of cast #61 data. No further errors were found.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.  The addsamp.csv file was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

There is a discrepancy between the rosette log sheet and the Daily Log book for station 69. The log records this as Event #11. It is incorrectly labelled as #10 on the rosette sheet and this led to all samples being mislabelled. All samples were renamed as #11.

SALINITY

The salinity data were delivered in spreadsheet 2009-26-salinity.xls; there were no duplicates. The spreadsheet was simplified (unneeded columns removed and headers changed to standard format) and saved as 2009-26-sal.csv which was then converted to individual SAL files.

DISSOLVED OXGYEN

The dissolved oxygen data was provided in spreadsheet 2009-26final.xls with quality flags and comments. The spreadsheet was simplified and saved as 2009-26-oxy.csv which was converted into individual ADD files. This is a new way of delivering DO data. This is much easier to work with than having the analysts prepare ADD files since those usually contain errors in formatting. There were no duplicates.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2009-26nuts.xls which included a report on precision. The spreadsheet was simplified and saved as 2009-26nuts.csv. File 2009-26-nuts.csv was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file chl2009-26-26May2009.xls which included a comparison of duplicates. The file was edited to include only the averaged values for each sample and to remove extraneous lines and columns and change header names to standard format. The file was sorted on sample number, and saved as 2009-26-chl.csv. The simplified spreadsheet file was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

Compare was run and shows a lot of scatter in the differences between bottles and CTD for both sensor pairs. There is little difference in shape between the primary and secondary plots. Below 250db the differences are less scattered but there are only 5 bottles; both CTD channels are higher than the bottles with the primary high by an average of 0.0025 and the secondary by 0.0004. When excluding all bottles with standard deviation in the CTD salinity data >0.001 the average is the same for the primary and closer to zero for the secondary. There is never more than 1 sample per cast, so it is impossible to separate the effect of pressure dependence from the fact that the shallower casts are in areas with more active mixing.  (See 2009-26-sal-comp1.xls.)

Dissolved Oxygen –

An initial run of COMPARE turned up many severe outliers. In many cases the analyst had not entered the values on the rosette sheet and had made note of a bad end point and/or that the value was clearly no good. For those cases the simplified spreadsheet was edited to change the “d” to an “e” and a pad value was entered. There were a few other cases where a value was written on the rosette sheet, but it is noted that there was a bad end point and again the values are way out of line. Those were also replaced with pad values and an “e” flag in the edited spreadsheet. In each case COMPARE was checked to ensure it really was a severe outlier. Samples edited in this way were:

- Cast #20, sample #74

- Cast #31, sample #120

- Cast #45, samples #165-167

- Cast #52, samples #196 & 198

- Cast #64, sample #256

- Cast #66, sample #266, 271 and 272

After these changes the spreadsheet was reconverted and merged and COMPARE was run again.

A few more outliers were identified and the following actions taken:

- Cast #1, sample 1, outlier, flagged “c”

- Cast #1, sample 8, outlier, flagged “c”

- Cast #1, sample 9, minor outlier, high standard deviation in CTD data, surface, left unflagged

- Cast #23, sample 88, severe outlier, flagged “d”

- Cast #39, sample 126, outlier, already flagged “c” due to titration problem – outlier in DO versus salinity plot - changed to “d”.

- Cast #41, sample 148, outlier, already flagged “c”

- Cast #48, sample 182, outlier, already flagged “c”

- Cast #52, sample 184, severe outlier, already flagged “c”; changed to “d”.

- Cast #52, sample 197, outlier, already flagged “c”, changed to “d”

- Cast #55, sample 214, outlier, flagged “c”

The analyst noted “No DMQ on oxy flasks” for casts 57-64. A fit of differences versus file pair number suggests there might be slightly lower bottle DO for those casts, but given the geographic variability for this cruise, this is far from clear, so no further flags will be assigned. The analysts concerns are already noted in the header comments.

For casts #66 and 68 there is a note in the DO analysis spreadsheet indicating that a Niskin bottle was leaky for one DO sample each. The nutrient data was examined to see if this was likely a problem for all samples and there is no evidence of a large error and it is not the same Niskin bottle so this is unlikely to be a systematic problem.

After identifying major outliers, more points were removed based on residuals to achieve a reasonably tight fit which was:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

(See 2009-26-dox-comp1.xls.)

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity. The only outliers were from bottles that had already been noted above.

After the flags and comments were added to the ADD files, the merge process was repeated and COMPARE was rerun to ensure no errors had been made.

Fluorescence

COMPARE was run using the CTD CHL and the Titrated Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD CHL. For the casts in Juan de Fuca Strait both are fairly low and fairly close. In the southern Strait of Georgia the gain was changed on the CTD fluorometer from 10X to 3X. For some casts there is a good correspondence between the two even when CHL is high, but for others there is not. The downcast fluorescence is highly variable, but where the bottles and fluorometer are very different, even the downcast does not see the high values seen in the CHL data.

(See 2009-26-chl-fluor-comp.xls.)

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always rough estimate as the upcast data are usually noisy, but for this cruise it was extraordinarily difficult to find sections of data useful for the tests. The usual shift of +24 records (1s) was found to improve the alignment, but whether it is the best choice is impossible to say. This is the shift that has been used in most other cruises, so it was applied to these data. (Output: SHFFL)

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. In other recent uses a values of +70 was used, though during 2009-38 which followed this cruise +65 looked better. A setting of +65 looked best overall for this cruise too.

SHIFT was run using +65 records for all casts.

Conductivity

Tests were run on a few casts and the best choice was found to be -0.3 for both conductivity channels. The same values were found for 2009-38 which followed this cruise and used the same equipment.

SHIFT was run using those settings. (Output: *.SHFC0 and SHFC1).

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

At this point a study was made of cast #39 to determine if the correct configuration was used for the fluorescence conversion. The note in the log is clear that the gain in the configuration file was changed after that cast, but it is not clear if the cable was changed before or afterwards.

The first thing to note is that this cast stands out from other nearby casts, with lower temperature and dissolved oxygen values and higher transmissivity near the surface than at #33 and 38. Low fluorescence values are therefore not necessarily a sign of having the wrong gain entered. However, the fluorescence is also very low compared to the extracted chlorophyll from the same cast. A test conversion was done using gain 3X for comparison. The 3X gain brings the fluorescence more into line with the extracted chlorophyll overall.

The following table shows a comparison of extracted chlorophyll and CTD fluorescence for nearby rosette casts (#31 and 41) and #39 with the two possible gains (3X and 10X).

|  | cast 31 |  |  | cast 39 (10X) |  |  | cast 39(3X) |  |  | cast 41 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 20db | 10db | 1db | 20db | 10db | 1db | 20db | 10db | 1db | 20db | 10db | 1db |
| CHL | 2.20 | 2 | 1.8 | 0.7 | 2.3 | 6.6 | 0.7 | 2.3 | 6.6 | 4.3 | 41.7 | 43.7 |
| FL | 1.40 | 1.2 | 0.6 | 0.3 | 0.6 | 0.7 | 0.9 | 2 | 2.1 | 2.7 | 49.4 | 49.3 |

When ratios of CHL to Fluorescence are compared from nearby rosette casts there is a lot of variability but again the 3X seems closer.

|  |  | Ratio of extracted CHL to Fluorescence |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Cast | 31 | 39 (10X) | 39 (3X) | 41 |
| Pressure |  |  |  |  |  |
| 20db |  | 1.57 | 2.33 | 0.78 | 1.59 |
| 10db |  | 1.67 | 3.83 | 1.15 | 0.84 |
| 1db |  | 3.00 | 9.43 | 3.14 | 0.89 |

Looking at downcast files changing the gain to 3X does not increase the fluorescence enough to look like nearby casts #33 and 38 near the surface, though it does bring the deeper values into line with those casts.

The balance of evidence is that the 3X gain is appropriate, so that was selected. A note was placed in the headers to indicate that there is some doubt about the fluorescence gain. From this point onwards the file named 2009-26-0039 has a 3X gain.

##### DETAILED EDITING

There is little difference in noise level between the two T/S channel pairs and the secondary pair produced salinity that is closer to the bottles, so the secondary temperature and salinity channels were selected for editing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

The following casts required no editing: 2, 28, 29, 31, 38,

The following casts required fairly heavy editing: 19, 20, 21, 24, 54, 55, 57,

All other casts required light editing.

##### Initial Recalibration

File 2009-26-recal1.ccf  was prepared to apply the following recalibration to the SBE dissolved oxygen:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

This was applied to the SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the corrections were applied correctly and they were. (See 2009-26-dox-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data is compared to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When obvious outliers are excluded, the differences range from -0.3 to +0.4ml/l. When a few more outliers are removed based on residuals, the fit against pressure is:

DOX(corrected) =  DOX +0.0001*Pressure -0.0315

So the values are slightly low below 300db and high near the surface. The differences are small and because there is a lot of scatter in the differences slight variations in how outliers are selected makes a significant difference to the fit. Further recalibration is not justified. (See 2009-26-dox-comp3.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 200db and processed separately for A. Peña. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  The primary conductivity sensor was found to be low by 0.003 when used for 2009-03 (Feb) but the sampling was limited to 2 depths. During 2009-27 (March) and 2009-38 (May) there was no calibration sampling. The secondary conductivity sensor was used for the same cruises and was found to be low by <0.001 for 2009-03. It was also used on an SBE25 in 2007 but there were doubts about the bottle analyses for those cruises.

2. Dissolved Oxygen – There are only 2 other cruises since the last factory calibration from which there is DO calibration sampling and both had few samples. For 2009-38 some samples came from areas of high variability in DO. The fits from those were:

CTD-BOT = 1.0209 DOX-CTD + 0.0668 (2009-27 - March)

CTD-BOT = 1.0700 DOX-CTD – 0.0887 (2009-38 - May)

3. Pressure –The sensor was recalibrated in August 2007 and an offset of +0.5db has been used for the four cruises processed since then.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were only 2 minor excursions, with salinity slightly below the historic minimum at one spot and above the maximum at another, so these are suggestive of calibration problems.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files are considered:

•	±0.4  ml/l from  0– 200db

•	±0.1 ml/l below 200 db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from 70% to 145%. The lowest saturations were in Juan de Fuca Strait. Most of the casts in the Strait of Georgia had saturations >130%, with slightly lower values closer to Vancouver Island. This cruise occurred near the time of spring bloom and the near-surface nutrients are very low for those casts with very high DO saturation.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

25. Cast abandoned

26. Partial cast, returned to surface and then down again

39. Not clear what cable used, think 3X, but could be 10X.

40. Fluorescence cable changed to 3X, but con file says 30X.

42. Log note “Syringes on” but presume this was between casts as data looks ok.

49. Fluorescence cable changed to 10X

52. Fluorescence cable changed to 3X

55. Bottle #16 did not fill, samples filled with bucket at surface

66. Bottle 13 leaked

72. Log note “Syringes on” but presume this was between casts as data looks ok

# NOTE: December 18, 2009. It was found that the spreadsheet provided by the analyst of the Oxygen:Dissolved data had been reprocessed through IOSShell-Derived Quantities-Dissolved Oxygen using an out-of-date version of the flask file. As a result many of the values were incorrectly recalculated. These values were removed and replaced with the original values from the .oxy file which were deemed to be correct. New DO values were calculated as required where two titers had to be added together. It is imperative that only .oxy NOT reprocessed .doxx values be used until the problem with the flask file in IOSShell is corrected.

# May 2010: There was also a major problem with the 450 nm filter on the colorimeter probe such that starting transmissivities for all samples were reading ~60% regardless of their actual value. This problem was not noticed during the cruise by the analyst but led to many aborted endpoints as the software did not register any changes in transmissivity below this value. As a result there is a higher than average number of missing values. Samples that were successfully titrated do not appear affected by this problem and all standardization parameters are reasonable.

REVISION: August 18, 2010

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly but steadily until July 2009; then a sudden shift occurred, so that maximum values between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study was made of transmissivity that led to a decision to apply post-processing corrections to all cruises between March 2008 and June 2010.

Transmissivity data from this cruise were corrected by multiplying the original values by correction factor 1.23. This was based on assumptions that deep offshore transmissivity from a June 2009 cruise should be about 62%/m and that drift was linear with time between March 2008 and July 2009. The corrections produced reasonable results for all cruises in that period.

For details on how the correction factor was derived see:

OSD_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer 1005DR Corrections.doc

These data should be considered estimates.

It was also noted at this time that the bottle dissolved oxygen data were missing from file 2009-26-0011.che, likely due to an error made in applying the May 2010 correction; these data were restored.

Revisions done by: Germaine Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| PAR | 4565 | 29Jan2009 | IOS |  |  |
| Surface PAR | 16504 | 24Feb2009 |  |  |  |
| Fluorometer | 2355 | ? | IOS |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |