## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 03-Oct-2017 | Corrected MISSION metadata field in header. R.H. |
| 27 May 2010 | Salinity data from 2 bottles were found and added to CHE files. The CTD salinity data are higher than the bottle data by 0.008 and 0.005. No recalibration was applied since these differences are out of line with other cruises using the same sensors. |
| 18 August | Transmissivity corrected; see note at end of report for details. |

## PROCESSING NOTES

Cruise: 2009-27

Agency: OSD

Location: Strait of Georgia

Project: ERI-SOB Moorings

Party Chief: Wright C.

Platform: CCGS Vector

Date: March 31, 2009 – April 2, 2009

Processed by: Germaine Gatien

Date of Processing: 19 May 2009 – 29 May 2009

Number of original CTD casts:  2	Number of CTD casts processed: 2

Number of bottle casts: 2		Number of bottle casts processed: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#0997/#1119), a Seapoint Fluorometer (#2228/#2356) with a 10X cable and an altimeter (#1252). The deck unit was a model 911+ (#0424).

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order with notes about problems encountered. The results of the dissolved oxygen titrations were sent as OXY files with some confusion about sample numbers. Notes in the log helped resolved the ambiguities.

SBE Dissolved Oxygen calibration sampling was all from below 150m. So no estimate can be made for accuracy above that level; the DO is considered	±0.1ml/l below 150db.

There was no salinity calibration sampling. The only other use of the T/C sensors since their latest factory calibration was cruise 2009-03 when the calibration sampling was severely limited. While the 2009-03 results suggest that the salinity is within 0.001, it may be necessary to reconsider recalibration of the data later should further calibration information become available.

There were many errors in the calibration control file used at sea. The Dissolved Oxygen parameters were from before some repairs to the instrument and the transmissivity parameters were very old.

The NMEA dates were wrong as were the dates from the Computer. All dates were changed to match the records in the log.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. The 2nd CTD cast is called Event #3 in the Daily Log Book and #2 on the rosette sheet. It was assumed that the Daily Log is correct.

The file names were non-standard. The events labelled SOGN and SOGS were renamed 2009-27-0001 and 2009-27-0003 respectively.

The titrated dissolved oxygen files were provided in # individual OXY files. The sample numbers were not all clear. There were replicates for 4 bottles.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The configuration files did not vary; one was saved as 2009-27-ctd.con.

The calibration constants were checked for all instruments. There were a number of problems:

- The configuration for DO had the wrong calibrations – they were from a post-cruise calibration before repairs. The new calibrations were entered including the Murphy-Larson parameters, using nominal values for the H1, H2 and H3 and E.

- The transmissivity date was very old and the parameters were wrong, so they were replaced by the latest ones available, 5 March 08. It is possible there is a more recent calibration.

- The offset for CTD #0550 was changed to +0.5db which was used for the last two cruises using this sensor. Checks will be done later to ensure it is appropriate.

- A Surface PAR is included in the configuration file, but is not mentioned in the log book and a few casts tested had no signal in that channel, so it will not be converted. This was also found during 2009-03.

The new SeaBird dissolved oxygen algorithm has a few parameters that are meant to be fine-tuned, but from this cruise the sampling is not deep enough to do the tests, so nominal values will be used. For shallow data this should not make any significant difference.

##### Conversion of Raw Data

Data were converted using configuration file 2009-27-ctd.con.

The casts were examined and all expected channels are present.

- The descent rate was steady for the downcasts.

- Spikes in the primary temperature are unusually large especially during the upcast. Conductivity is similar. This was also noted for 2009-03.

- The fluorescence and transmissivity channels look a little spiky but ok.

- Dissolved oxygen voltage looks ok.

- The altimetry looks fine.

Rosette files were converted using a start time of -5s and duration of 10s. The rosette files were then converted to IOS header files. The event numbers were present so CLEAN was not run, but the IOS files were renamed as *.BOT. Bottom depths were added to both and a station name was added to cast #1.

Temperature and conductivity in the BOT files were plotted. There were no outliers that required editing,

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

A variety of settings were tested by running on both casts. The best results were with α = 0.0245 and β=9.5 for both sensors, though other settings were very similar. CELLTM was run on both casts using that setting.

##### DERIVE

Program DERIVE was run twice:

on both casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on both casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The data were plotted to check for agreement between the pairs of T and C sensors. The differences were quite noisy especially for the second cast and there are no data below 300db so the results cannot be compared very well with 2009-03. But in rough terms the temperature differences seem similar, but the conductivity differences have the opposite signs. There were suspicions about the primary channels during 2009-03. That might have been due to pump problems which might have been resolved before this cruise. The following show the comparison between 1 cast from each of those cruises:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-03 | 1200 | +0.0015 | +0.0003 | +0.0018 | Steady |
| 2009-27 | 295 | +0.001 | -0.0007 | -0.0019 | High, X noisy |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and checked. The speed between the two stations was <0.

The cross-reference check was compared with the log book. The times and positions are correct, but the dates were wrong. These were corrected to match the log, though it is odd to have computer and NMEA both get it wrong in the same way.

The station name was missing from cast #1. That was added.

The bottom depth was missing from both casts, so that was added to the headers based on log entries.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run; the two values are 1.2 and 2.0db. For the first cast the pressure at the end of the cast seems quite low, ~0.3db with salinity ~29.8, though the pumps were not on at that point. This definitely indicates that an offset of at least 0.5db is needed.

The altimeter header entries were checked against plots and look fine.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.  There is a rosette bottle not listed on the rosette sheet (surface, cast #1). It is assumed this was an unintended firing, so it was not processed further.

The file was converted to CST files to be used as a framework for the bottle files. The SAM files were then bin-averaged.

DISSOLVED OXGYEN

Dissolved oxygen files (*.oxy) were provided in 3 files with no flags or comments, but there are remarks from the analyst on the rosette sheet. There were a few problems in relating samples to rosette bottles:

- There were 2 samples numbered 100 and 1001. The DO values show these are from near the surface, so they are assumed to actually be replicates for sample #10 which is indicated to be a TEST on the rosette sheet. Given the two values are quite different from each other and from the CTD DO data from that bottle, it is assumed these are not to be processed, truly being just tests. Those values were not entered into the spreadsheet.

- Other samples had irregular names, mainly because there were replicates; from the notes of the analyst their origins were clear.

Because there are so few samples, the easiest way to create ADD files with flags was to copy the data into a spreadsheet.

The first step after assembling the data was to do a study of the replicates. There were 4 pairs, three of which differed by <0.3% of the average value while one differed by 2.8%. Sp=0.135 where Sp is defined as:

Sp = Square Root (sum of squares of differences / 2*number of pairs)

This is a fairly high value compared to other cruises, but examination of the bottle stop with the 2.8% difference showed that temperature and CTD Dissolved Oxygen varied significantly throughout that stop, so the local variability must have been high. Hence the water in the Niskin may not have been as homogeneous as usual. If that value is not included, then Sp=0.022. The analyst noted a problem with one of the endpoints, but that replicate was not the outlier. (See 2009-27-DO-replicates.xls.)

The replicates were averaged and entered into the main DO spreadsheet with flag “f” to indicate replicates. (See 2007-27-oxy.csv.) The spreadsheet was then converted to ADD files. The analyst noted problems with the first sample, so it was flagged “c” and will be checked in COMPARE.

The ADD files were merged with CST files. The output file was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files.

##### Compare

Dissolved Oxygen

COMPARE was run for the DO sensor. There was one major outlier, the sample that was flagged as “c” above –this was changed to flag “d”. Excluding the outliers, there are only 5 points, with trendline:

CTD-BOT = 1.0209 DOX-CTD + 0.0668

(See 2009-27-dox-comp1.xls.)

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity. The only notable outlier was the one already flagged “d”.

##### 12. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The usual shift of +24 records (1s) was found to have good results, so it was applied to these data. (Output: SHFFL)

Dissolved Oxygen

Tests were run on the 2 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. A setting of +70 looked reasonable.

SHIFT was run using +70 records for all casts.

Conductivity

During 2009-03 the best choice of conductivity shift was found to be no shift for the primary and that proved to be the same for this cruise. For 2009-03 a shift of -0.3s looked best for the secondary, and again tests run on these data showed the same result. So SHIFT will not be applied to primary conductivity. SHIFT was run using -0.3s for the secondary conductivity only. (Output: *.SHFC1).

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The noise level in the two T/S channel pairs look very similar. For 2009-03 the secondary salinity appeared to be closer to the bottles, though there was little sampling available.

The secondary temperature and salinity channels were selected for editing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. Both casts required some editing.

##### Initial Recalibration

Salinity will not be recalibrated because there are no bottles and cruise 2009-03 suggested that the calibration was ok for the secondary channel.

For dissolved oxygen we have little to go by. The results of fitting CTD against 5 bottles showed:

CTD-BOT = 1.0209 DOX-CTD + 0.0668

This compares with the result of 2009-03 which was also based on very little - a repeat cast to compare to another sensor whose calibration was better known. That produced the following fit:

DOX(corrected) = 1.018 * DOX +0.04

The two are not all that different which is surprising and offers at least a little confidence in the fit.

File 2009-27-recal1.ccf was prepared and applied to the SAM file to create SAMCOR1.

COMPARE was rerun to check that the correction was done correctly. (See 2009-27-dox-comp2.xls.)

The recalibration was then applied to the MRGCLN2 and the EDT files to create MRGCOR1 and COR1 files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the levels of the bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When the one extreme outlier was excluded, the differences range from -0.07 to +0.03ml/l, with an average difference of -0.01. There is too much scatter to conclude anything from this except to say that the DO is within 0.01 below 150db. No further recalibration will be applied. (See 2009-27-dox-comp3.xls.)

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no problems were found.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity - Since its last factory calibration the primary conductivity sensor has been used only for 2009-03. There were only a few calibration samples but what was available suggests the CTD was low by 0.004. The secondary sensor has been used on many cruises, but for most of those there were doubts about the salinity analysis. From 2009-03 the few bottles available suggest it is within 0.001 of the bottles.

2. Dissolved Oxygen – The sensor was used for 2009-03 when there were no bottles. The only calibration information was a comparison between 2 casts at one site with different DO sensors mounted. The comparison led to the following correction being applied:

DOX(corrected) = 1.018 * DOX +0.04

3. Pressure –The sensor on CTD #0550 was recalibrated in August 2007 at which time an offset of +0.1db was found. It has been used for 2 other cruises since then and an offset of +0.5db was found appropriate. Some drift in the offset is expected from this type of sensor as it ages.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. All data fell within the local climatology.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

There was no salinity calibration sampling. During 2009-03 the CTD was

within 0.001 but there were few samples. No recalibration was applied to

the salinity.

For the Sea-Bird dissolved oxygen recalibration was based on only 5 bottles. The

results of that comparison are quite close to those of 2009-03. However, for

2009-03 recalibration was based on intercomparison with another sensor.

Confidence is lower than usual in the SBE DO calibration.

Dissolved Oxygen appears to be within 0.01ml/l below 150db. There was no

DO sampling above that level, so no estimate can be made

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were between 90% and 100%.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

REVISION: August 18, 2010

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly but steadily until July 2009; then a sudden shift occurred, so that maximum values between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study was made of transmissivity that led to a decision to apply post-processing corrections to all cruises between March 2008 and June 2010.

Transmissivity data from this cruise were corrected by multiplying the original values by correction factor 1.212. This was based on assumptions that deep offshore transmissivity from a June 2009 cruise should be about 62%/m and that drift was linear with time between March 2008 and July 2009. The corrections produced reasonable results for all cruises in that period.

For details on how the correction factor was derived see:

OSD_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer 1005DR Corrections.doc

These data should be considered estimates.

Revisions done by: Germaine Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| Fluorometer | 2356 | ? | IOS |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |