## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| May 13, 2010 | Note added to Section 10 explaining colorimeter problem leading to bad DO bottle values. |
| May 13, 2010 | Off-scale values in fluorescence data replaced with pad values in CTD and CHE files. |
| Aug 18, 2010 | Transmissivity corrected; see end of report for more details. |

## PROCESSING NOTES

Cruise: 2009-34

Agency: MEAD

Location: Strait of Georgia

Project: Coastal Ecosystem Research Program

Party Chief: Sutherland T.

Platform: CCGS Vector

Date: April 17, 2009 – April 22, 2009

Processed by: Germaine Gatien

Date of Processing: 12 September 2009 – 16 September 2009

Number of original CTD casts:  33	Number of CTD casts processed: 33

Number of bottle casts: 		32	Number of bottle casts processed: 32 (only 20 with samples)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119 probably on the primary pump), a Seapoint Fluorometer with a 3X cable (#2355 on the primary pump), a Biospherical QSP-400 PAR sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter (#1252). The deck unit was a model 911+ (#0424).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Log Book and rosette log sheets were generally in good order.

The calibration data originally entered for the dissolved oxygen sensor were incorrect. The same values were used for a series of cruises in spring 2009 and come from a pre-repair calibration; the post-repair calibrations should be used.

Sea-Bird have a new algorithm for dissolved oxygen with several parameters that the manufacturer recommends be fine-tuned for each instrument to produce the best data. This requires bottle samples from deep casts. There has been no cruise (since these new parameters became available) using this sensor that has sufficient deep sampling to do this, so nominal values were used. The hysteresis correction was not applied since there are no data below 500db. The Tau correction in the derivation of dissolved oxygen concentration is known to exaggerate residual noise in deep water, but that should not be a problem for this cruise, so that correction was applied. It improves the fit in the high gradient near the surface.

As for some other spring cruises using this equipment, there were many errors in the NMEA dates; sometimes the error was 72 hours, sometimes 96 hours.

The file names were non-standard.

For casts #18 and 21 no downcast data were recorded. The upcast data were used so quality is lower than usual, especially since there were stops for bottles.

There were many problems with the shipboard titration of near-surface dissolved oxygen samples, so that the comparison between bottles and the SBE DO sensor was limited to a very narrow range. The results of the DO comparison for 2009-26 (April 11-15) were applied to this cruise; even though similar problems were noted during that cruise, there were enough surface data to do a reasonable fit.

The DO data are considered, roughly:

# ±0.4 ml/l from 0– 100db

# ±0.2 ml/l from 100– 250db

# ±0.05 ml/l below 250 db

No recalibration was applied to the salinity. While some calibration sampling was done, the bottles could not be found. The conductivity sensors have been used on many other cruises, but none of those included a really good comparison with bottles. The best was from 2009-26 when the average difference showed the secondary salinity to be very close to the bottles, but the scatter was high. The calibration of salinity for this cruise should be revisited when a drift report is received from the next factory recalibration of the secondary conductivity sensor.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. The only problems noted in the log are that archiving only occurred for the upcast of events #18 and 21.

Dissolved oxygen data were obtained in individual OXY files with no flags or comments. The analyst did make notes on the rosette sheets; there were many titration problems.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

A single instrument configuration was used during the cruise. The calibration constants were checked for all instruments. There were a number of problems:

- The configuration for DO #0119 came from a calibration done before the membrane was replaced. There is a later calibration that should be used. It is also better to select the Murphy-Larson parameters, so that the Tau correction can be applied. The nominal values were entered for the E, H1 and H3 since there is insufficient deep sampling to fine-tune them. It is not expected that the hysteresis correction will be applied.

- The transmissivity date and parameters were wrong. Checks were made to see if the serial number might be wrong. The only transmissivity calibrations for that date had different values and no calibration could be found with those values. It will be assumed that this really is 1005DR.

- The pressure offset for CTD #0550 was changed from 0.1033 to +0.5db which is the value found on other recent cruises.

- The parameters for the Surface PAR were confirmed correct by Scott Rose.

Because the same problems have been noted and corrected for several other recent cruises, the configuration file for cruise 2009-26 was copied and saved as 2009-34-ctd.con.

##### Conversion of Raw Data

The file names had the wrong format. An extra zero was added to all file names.

Data were converted using the configuration files 2009-34-ctd.con.

A few casts were examined and all expected channels are present.

- The descent rate is generally steady.

- There are some spikes in temperature and conductivity.

- The two temperature channels are in reasonable agreement on the downcast. The upcast data are much noisier so there are significant differences. Conductivity is similar to temperature.

- The fluorescence has many spikes to high values. The 3X cable looks like a good choice, though there are still some off-scale values.

- Dissolved oxygen voltage looks as usual with an offset but some detail to help alignment.

- PAR and Surface PAR look fine.

- The transmissivity is generally lower at the bottom, but only a few spikes to very low values are seen there.

- The altimetry looks usable near the bottom but there are large spikes that may interfere with the header algorithm; all values should be checked.

- Bottle stops look like they are all at least 30s.

Rosette files were converted using a start time of -5s and duration of 10s and bottom depths were corrected in 2 casts. The rosette files were then converted to IOS SHELL files and named *.BOT. Bottles had been fired during all casts except #8. All BOT files were plotted. The data at the bottom frequently look quite noisy. Since the transmissivity does not usually get extremely low this is assumed to not be due to hitting bottom. There were a few outliers in salinity in casts 9, 11 and 14. CTDEDIT was used to clean those and the output was copied to *.BOT. There were other noisy patches that upon examination look believable; those were left unedited.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

This cruise is a poor one for testing CELLTM parameters. Many of the casts were very shallow, so temperature does not vary much, and for all casts except #8 there were stops for bottles, which makes it difficult to compare the upcast with downcast. Tests were run on cast #8. For the primary the best results were either with (α = 0.03, β=9) or (α = 0.0245, β=9.5). For the secondary (α = 0.02, β=7) looked best.

Since exactly the same equipment was used for cruise 2009-26 it is not surprising that these results are in good agreement with those of the earlier cruise when (α = 0.0245, β=9.5) was applied to the primary and (α = 0.02, β=7) to the secondary.

CELLTM was run applying the same settings as for 2009-26 to all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The deep casts were all early in the cruise. Cast #8 was studied to see how it compared with the results of cruise 2009-26. The differences are noisy despite fairly steady descent rates, so these are very rough averages. The 250m values are highlighted to make comparison easier.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-26-0004 | 200 250 | +0.0013 +0.002 XN | -0.0002 -0.0002 | +0.0067 -0.002 VN | Steady |
| 2009-26-0063 | 200 250 | +0.001 VN +0.001 | -0.00015 -0.0001 | +0.009 -0.012 | Steady |
| 2009-26-0074 | 200 250 300 | Too noisy +0.001 XN +0.001 | -0.0004 XN -0.0002 XN -0.00012 | Too noisy -0.0026 -0.0022 | Steady |
| 2009-34-0008 | 200 250 | +0.0016 N +0.0015 | -0.0003 N -0.0001 | -0.0034 N -0.0026 | Moderate Very steady |

There is no evidence of significant drift, though the results from both cruises are very noisy and not really deep enough for a good check.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. There was a missing station name for cast #1; this was added to the IOS and BOT files. The speed check looks odd. A check of the cross-reference list against the log book showed why. Most of the header dates were wrong. This occurred for some other spring cruises. The dates will be fixed later.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run and shows the average surface pressure to be 2.3db, typical of the Vector. The only instances of negative pressures noted were small spikes in pressure near the surface. There is no evidence that the pressure offset needs adjusting.

The altimeter readings and water depths were exported from the headers to a spreadsheet. Plots were made of altimetry for many casts to see if the entries were reasonable and despite some spiky data the algorithm appears to have worked well.

##### BOTTLE FILE PREPARATION

Add Time Channel was run to fix the times in the headers, with output *.ATC.

The ATC files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.  The addsamp.csv file was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the ATC files to produce SAM files which were then bin-averaged.

The only samples for this cruise are dissolved oxygen. There was some salinity sampling, but the bottles could not be found.

DISSOLVED OXGYEN

The dissolved oxygen data were provided in individual OXY files with no quality flags or comments. The data from those files were exported to spreadsheet 2009-34-oxy.csv. The headers were edited, formats fixed and columns were added for flags and comments based on comments on the rosette log sheets. One error was corrected (flask number had been entered where the sample #244 should have been.) There were many surface samples early in the cruise that were entered in the OXY files but had a note that the endpoint was too early. ”d” flags were added, but if the values prove to be severe outliers in COMPARE they should be changed to “e” flags and pad values entered. Eventually the analyst gave up entering values for the surface bottles until cast #33 when a good surface value was achieved.

The spreadsheet was converted into individual ADD files.

There were no duplicates.

Note Added May 13, 2010:

There was a major problem with the 450 nm filter on the colorimeter probe such that starting transmissivities for all samples were reading ~60% regardless of their actual value (as seen later during 2009-08). This problem was not noticed during the cruise by the analyst but led to a few aborted endpoints as the software did not register any changes in transmissivity below this value.  Samples that were successfully titrated do not appear affected by this problem and all standardization parameters are reasonable.

The ADD files were merged with CST files. (Output: MRG3), MRG3 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Dissolved Oxygen –

An initial run of COMPARE turned up many severe outliers from near the surface, which also happened during 2009-26. This is presumably a problem with high DO values. There were cases in which:

- There is a value in the OXY file, but the analyst had not entered the values on the rosette sheet and had made note of a bad end point and/or that the value was clearly no good. For those cases the simplified spreadsheet was edited to enter an “e” flag and a pad value was entered.

- There is a value written on the rosette sheet, but it is noted that there was a bad end point and again the values are way out of line. Those were also replaced with pad values and the “d” flags entered earlier were changed to “e” flags in the edited spreadsheet.

- There is no value on the rosette log sheet for near-surface bottles and no note about a bad endpoint, but the values are extreme outliers, so obviously were affected by the same problem as for the others. Again an “e” flag and pad value were entered.

- There were a few other problems with mislabelling of event number (caused by multiple casts in single OXY files) or two values for a single sample. Those were fixed.

After the spreadsheet was edited to enter these changes the bottle files were recreated and COMPARE was rerun. There were no further extreme outliers, but the scatter is much higher than usual. When outliers were excluded based on residuals the following fit was found:

CTD-BOT = 1.0308 DOX-CTD + 0.0702

which is reasonably close to that found for 2009-26:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

Cast #33 is the only cast with more than 3 bottles that did not have end point problems. When just cast #33 is plotted the scatter makes a fit impossible. But that picture changes if the 4 bottles are excluded that were identified as outliers on the basis of residuals when the whole data set was analyzed. That leaves near-surface data and data below 100m. The fit of those 9 points is:

CTD-BOT = 1.0357 DOX-CTD + 0.0555

which is close to the 2009-26 results. So there is some consistency, but that is all we can say. (It is also noted for cast #33 that profiles of Bottle DO and CTD DO versus Salinity look very similar.)

There were a lot more data available for 2009-26 and a wider range of DO values despite the problems with DO analysis during that cruise. So 2009-26 provides the best calibration information. (See 2009-34-dox-comp1.xls.)

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity. The only outliers were from bottles that had already been noted above. For cast #33 there is an excellent correspondence between the DO versus Salinity plots for Bottle DO and CTD DO.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always a rough estimate as the upcast data are usually noisy, and with so many bottle stops these data were not very useful for such tests. But the usual shift of +24 records (1s) did seem to improve the alignment, so it was applied to these data. (Output: SHFFL)

Conductivity

During 2009-26 and 2009-38 when the same equipment was used (before and after this cruise) the best choice was found to be -0.3 for both conductivity channels. Tests were run on one cast varying that value slightly and -0.3 records looks best for both conductivity channels.

SHIFT was run using those settings. (Output: *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. In other recent uses a values of +70 was used, but during 2009-26 and 2009-38 a shift of +65 looked better. A setting of +65 looked best overall for this cruise too.

SHIFT was run using +65 records for all casts.

##### 14. DELETE

Before running DELETE casts #18 and 21 were put through REVERSE since there are no downcast data available.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

There is little difference in noise level between the two T/S channel pairs and the secondary pair produced salinity that is closer to the bottles during 2009-26, so the secondary temperature and salinity channels were selected for editing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

All casts required light editing, except for the two reversed files which needed a little heavier editing. For many of the shallow casts there were unstable features around 6m. They are deeper than the corruption often seen around 3 or 4m as the CTD first starts down. These could be natural overturns or perhaps some ship effect caused them. Since there was no obvious instrumental cause, the features were left unedited.

##### Initial Recalibration

File 2009-34-recal1.ccf  was prepared to apply the following recalibration (from 2009-26) to the SBE dissolved oxygen:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

This was applied to the SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see what effect the corrections had. As expected the fit is not perfect – the CTD DO is found to be low by an average of 0.2ml/L when outliers are excluded. However, the scatter is very large, so a further recalibration is not justified. (See 2009-34-dox-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps partially correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data are compared to bottle data from the same pressure. However, for this cruise it was decided to skip this step given the problems with the bottles and the fact that for 2009-26 no further recalibration was found appropriate.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

For casts 1-9 and 33, the following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000*	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

*For casts 9-32 the same parameters were used except that the averaging interval was 0.5.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  The primary conductivity sensor was found to be low by 0.004 when used for 2009-03 (Feb) but the sampling was limited to 2 depths. During 2009-26 the salinity was higher than bottles by 0.0025. During 2009-27 (March) and 2009-38 (May) there was no calibration sampling. The secondary conductivity sensor was used for the same cruises and was found to be within 0.001 for 2009-03 and 2009-26. It was also used on an SBE25 in 2007 but there were doubts about the bottle analyses for those cruises.

2. Dissolved Oxygen – There are only 3 cruises that have been processed since the last factory calibration for which there was DO calibration sampling. The fits from those were:

CTD-BOT = 1.0209 DOX-CTD + 0.0668 (2009-27 - March) – only 5 bottles

CTD-BOT = 1.0388 DOX-CTD + 0.0088 (2009-26 - April)

CTD-BOT = 1.0700 DOX-CTD – 0.0887 (2009-38 - May) – poor duplicates/lot of scatter

3. Pressure –The sensor was recalibrated in August 2007 and an offset of +0.5db has been used for the five cruises since then.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were no excursions from the local climatology.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

As noted earlier most of the dates were out by 3 or 4 days. ADD TIME CHANNEL was used to advance the header times by 72 or 96 hours as required, with output files named *.ATC.

(The dates were fixed adjusted earlier for the BOT files, with output *.ATC.)

The header check and cross-reference listing were rerun after the time fix. The results look good. No further errors were found.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files are considered:

- ±0.4  ml/l from      0– 100db

- ±0.2  ml/l from  100– 250db

- ±0.05 ml/l below 250 db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from 100% to 160% with the highest values in the first half of the cruise. Unfortunately there are very few near surface bottles available to check that these values are reasonable. This cruise occurred only a few days after 2009-26 when saturation values in the Strait of Georgia were between 130% and 145%, so the results look reasonable. Surface nutrients were also very low during 2009-26 so these high DO saturations appear to be due to a spring bloom.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure. REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

18. Acquisition for upcast only.

21. Acquisition for upcast only.

REVISION: August 18, 2010

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly but steadily until July 2009; then a sudden shift occurred, so that maximum values between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study was made of transmissivity that led to a decision to apply post-processing corrections to all cruises between March 2008 and June 2010.

Transmissivity data from this cruise were corrected by multiplying the original values by correction factor 1.230. This was based on assumptions that deep offshore transmissivity from a June 2009 cruise should be about 62%/m and that drift was linear with time between March 2008 and July 2009. The corrections produced reasonable results for all cruises in that period.

For details on how the correction factor was derived see:

OSD_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer 1005DR Corrections.doc

These data should be considered estimates.

Revisions done by: Germaine Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| PAR | 4565 | 29Jan2009 | IOS |  |  |
| Surface PAR | 16504 | 24Feb2009 |  |  |  |
| Fluorometer | 2355 | ? | IOS |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |