## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 11March2010 | Dissolved oxygen bottle data were recalculated and reassessed by Nina Nemcek. For details see notes in BOLD TYPE mostly in sections 10 and 11. The changes are not systematic or so large as to justify recalibration of the CTD DO channel. The titrated DO data in the CHE files were updated. |
| 18 August 2010 | Corrected Transmissivity; for details see end of report. |

## PROCESSING NOTES

Cruise: 2009-38

Agency: OSD

Location: Saanich Inlet

Project: Tracer Release Experiment

Party Chief: Cummins P.

Platform: CCGS Vector

Date: May 1, 2009 – May 8, 2009

Processed by: Germaine Gatien

Date of Processing: 18 June 2009 – 17 September 2009

Number of original CTD casts: 	9

Number of CTD casts processed: 	9

Number of bottle casts: 		 	8

Number of bottle casts processed: 	8

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2356) with a 10X cable and an altimeter (#1252). The deck unit was a model 911+ (#0424).

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order with notes about problems encountered. The CTD listed in the log was not the one actually in use, but the sensors listed were correct.

Samples numbers were repeated during two casts. To avoid confusion, sample numbers for cast #12 were changed by adding a leading “9”.

The comparison of bottles and SBE Dissolved Oxygen was very noisy and the fit of differences against DO looked much different than from other spring cruises using the same equipment. While there were many calibration samples, most came from a few depths at a few sites, which may also limit the quality of the comparison. The results of 2009-26 were used to recalibrate the data. The results produce SBE DO too low near the surface, too high at mid-depths and about right at depth. The recalculation of bottle dissolved oxygen data in March 2010 did not lead to further recalibration of the SBE DO channel.

SBE DO values are considered

•	 ±0.25 ml/L from 0-125db

•	 ±0.06 ml/L below 125db.

There was no salinity calibration sampling. When used for two other cruises the salinity was found to be within 0.001 of the bottles, but there were only 5 bottles for one comparison and a lot of scatter for the other. It may be necessary to reconsider recalibration of the data later should further calibration information become available.

There were many errors in the calibration control file used at sea. The Dissolved Oxygen parameters used were from before some repairs to the instrument and the transmissivity parameters were very old.

The data in file 2009-38-0001.ctd is from the upcast because data were not logged on the downcast. The quality of this data will be lower than for other casts since there were many stops for bottles.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and Rosette Log sheets were obtained.

The equipment list in the front of the Log Book identified the CTD as #0506, but the pressure sensor serial number is the one on CTD #0550, so the log entry is assumed to be wrong. The instruments used for this cruise are assumed to be the same as for cruise 2009-27 when CTD #0506 was used.

There was no salinity calibration sampling.

The nutrient data were provided in spreadsheet format.

The titrated dissolved oxygen files were provided in 3 individual OXY files. There were many replicates and from one cast there were 2 samples for each bottle.

NOTE: After the post-cruise corrections were done to the Dissolved Oxygen bottle data, spreadsheet 2009-38oxy.xls was received; this included a duplicate study.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The configuration files did not vary; one was saved as 2009-38-ctd.con.

The calibration constants were checked for all instruments. There were a number of problems:

- The configuration for DO had the wrong calibrations – they were from a post-cruise calibration before repairs. The new calibrations were entered including the Murphy-Larson parameters, using nominal values for the H1, H2 and H3 and E.

- The transmissivity date was very old and the parameters were wrong, so they were replaced by the latest ones available, 5 March 08. Scott Rose is unaware of any more recent calibration.

- The offset for CTD #0550 was changed to +0.5db which was used for the last two cruises using this sensor. Checks will be done later to ensure it is appropriate.

- A Surface PAR is included in the configuration file, but is not mentioned in the log book and a few casts tested had no signal in that channel, so it will not be converted. This was also found during 2009-03 and 2009-27.

- A PAR sensor is included in the configuration file, but is not mentioned in the log book. Test conversion showed no signal in that channel so it will not be converted.

- The fluorometer serial number was not entered in the log or the configuration file, but #2356 was the one used for 2009-27 so is probably the one used since all other sensors are the same.

The new SeaBird dissolved oxygen algorithm has a few parameters that are meant to be fine-tuned, but since there has been no cruise with deep enough sampling to do the tests nominal values will be used. For shallow data this should not make any significant difference.

##### Conversion of Raw Data

Data were converted using configuration file 2009-38-ctd.con.

The casts were examined and all expected channels are present.

- The descent rate was steady for the downcasts.

- When these sensors were used during 2009-03 and 2009-27 it was noted that there were unusually large spikes in the primary temperature and conductivity especially during the upcasts. For this cruise there is a lot of noise in the upcast traces and the primary are slightly noisier though it is not as obvious which channels are better. The DO and fluorescence channels were both on the primary pump, which is not usual. It is possible that the pump was not working as efficiently as usual, and that seems to show up on upcasts more than downcasts. The secondary pair are probably a better choice.

- The fluorescence and transmissivity channels look normal.

- Dissolved oxygen voltage looks ok with the usual vertical offset.

- The altimetry looks good.

Rosette files were converted using a start time of -5s and duration of 10s. The rosette files were then converted to IOS header files. The event numbers were present so CLEAN was not run, but the IOS files were renamed as *.BOT. The bottom depth was fixed for cast #7.

Temperature and conductivity in the BOT files were plotted. There were no outliers that required editing,

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

A variety of settings were tested on 2 casts with few stops. The best results were with α = 0.0245 & β=9.5 for the primary sensors, and α = 0.02 & β=7 for the secondary sensors, though other settings were very similar. CELLTM was run on all casts using those settings.

##### DERIVE

Tests were run on the derivation of dissolved oxygen to see if the use of the TAU correction was useful and it did produce a more detailed profile near the surface. Tests were also run varying the value of E; the changes were slight with no clear indication of which should be chosen, so the nominal value was selected.

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration (with Tau correction).

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The data were plotted to check for agreement between the pairs of T and C sensors. The differences were quite noisy as expected with shallow casts. For comparison results from sample casts from three earlier cruises that used the same sensors are included in the table. The temperature differences are quite close, and the conductivity and salinity are similar to 2009-27 and 2009-26, but have the opposite signs to 2009-03. There were suspicions about the primary channels during 2009-03. That might have been due to pump problems which might have been resolved before the other two cruises.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-03 | 1200 | +0.0015 | +0.0003 | +0.0018 | Steady |
| 2009-27 | 295 | +0.001 | -0.0007 | -0.0019 | High, X noisy |
| 2009-26 | 250 | +0.002 | -0.0002 | -0.002 | Very Noisy |
| Cast 2 | 170 | +0.0015 | -0.0003 | -0.004 | Steady |
| Cast 9 | 170 | +0.001 | -0.0002 | -0.003 | Steady |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and no errors were found.

The cross-reference check was compared with the log book. No errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run and the average value was 2db which is reasonable. From cast #1 there is some data with negative pressure at the end of the upcast; the transmissivity goes to zero when pressure is ~-0.8db. This may indicate that it is time to increase the pressure offset, but the pumps were off, so there is no confirming information from conductivity. Upcast data is a little hard to interpret with the CTD being affected by the rosette package dragging deeper water with it. One other approach is to add the altimeter reading to the maximum pressure and compare that with the bottom depth entry in the log. Two casts were examined in that way and the results showed no error in the pressure. However, it is possible that the bottom depth entry is a rough estimate, During cruise 2009-26 the pressure looked to be within ±0.1db. An eye should be kept on this pressure sensor in future, but it does look like any error is small.

The altimeter and water depth header entries were checked against plots and logs. The only problem found was an error in the water depth for cast #7; that was corrected. There were very slight differences in positions and water depths for a few other casts, but those likely represent a difference in time of entering information in the log and time of actual cast so they were not changed in the headers.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.  A few problems were found:

- Cast #9 – there are 10 bottles listed on the rosette log sheet, but only 5 bottles were fired.

- Cast #10 – there are 10 bottles listed on the rosette log sheet, but only 5 bottles were fired.

- Casts 10/12 - Sample numbers 61-65 were used for both casts #10 and #11, so those for cast #12 were renamed to #961 to 965 to avoid confusion. This affects nutrients only and the nutrient analyst made the same correction.

The file was converted to CST files to be used as a framework for the bottle files. The SAM files were then bin-averaged.

DISSOLVED OXGYEN

All titration data and endpoints were checked and DO values recalculated as necessary by Nina Nemcek in Mar 2010. There was a major problem with the 450 nm filter on the colorimeter probe such that starting transmissivities for all samples were reading ~60% regardless of their actual value (as seen later during 2009-08). This problem was not noticed during the cruise by the analyst but led to a few aborted endpoints as the software did not register any changes in transmissivity below this value.  Samples that were successfully titrated do not appear affected by this problem and all standardization parameters are reasonable.

Dissolved oxygen files (*.oxy) were provided in 3 files with no flags or comments, but there are remarks from the analyst on the rosette sheet. One of the files has 2 casts included in it, and another has Niskin numbers instead of sample numbers. There were many duplicates and often there were samples from two Niskin bottles fired at the same depth. Cast 14 was sampled as an intercomparison between the IOS oxygen titrator and the Scripps Institution of Oceanography titrator, as well as between 2 different sampling methods thus 4 oxygen samples were taken from each Niskin and 6 replicate Niskins were fired at each of 3 depths. Only 2 samples from each bottle were analyzed on the IOS system and only these are presented here for consistency.

The OXY files were exported to a spreadsheet 2009-38-oxy.csv. The duplicates were extracted to sheet 2009-38-oxy-duplicates.xls. The duplicates were averaged and flag “f” added. Flag “e” a PAD value, and a comment were added to samples 41-44 based on notes from the analyst on the rosette log sheet. The event number for samples 61 to 65 were changed to #10. And the sample numbers for cast #14 were corrected to match the rosette log sheet. The file was then converted to individual ADD files.

The pooled standard deviation of duplicates (Sp) was calculated with Sp defined as

Sp = SQRT{sum (d*d)/2k}

where k = no. of pairs and d = difference between pairs

There were 24 duplicate samples taken from individual Niskins including those from Cast 14. Although the duplicates from Cast 14 were sampled using slightly different methods the difference in DO value between pairs was no different than for regular duplicates sampled with a single method so they were included in the analysis. The pooled standard deviation (Sp) for all 24 duplicates was 0.031. Of these, one was determined to be an outlier based on Chauvenet’s criterion and was excluded resulting in an Sp= 0.027.

There were 9 sample pairs collected from duplicate Niskins fired at the same depth excluding the intercomparison cast #14. The Sp for these replicates was 0.044. One outlier was identified based on Chauvenet’s criterion and excluded reducing the Sp to 0.016 for the 8 remaining pairs.

The ADD files were merged with CST files. (In order to run COMPARE the following steps were run, the output file was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files. Later new MRG files were prepared with Nutrients.)

NUTRIENTS

Nutrient data were received in file CumminsNutsJuly17_2009Saanich Inlet.xls. This was simplified, reordered on sample number and saved as 2009-09-nuts.csv. That file was converted to individual NUT files. The data were merged with the dissolved oxygen data and, after cleaning to reduce the headers to File and Comment sections only, with the SAMAVG files to form MRG files. Those were cleaned again to remove the Sea-Bird headers and the comments from the SAMAVG files.

##### Compare

Dissolved Oxygen

COMPARE was run for the DO sensor. There were 4 major outliers that had already been flagged by the analyst; the flags were changed from “c” to “d”. There were 3 other samples flagged “c” because the duplicates differed by >10%. Excluding the outliers, the trendline is:

CTD-BOT = 1.070 DOX-CTD – 0.088

but there is a lot of scatter and only 43 bottles and many are from the same depths and casts. Some samples look out of line and an examination of profiles offers an explanation. The samples from around 75db come from just below the temperature minimum and are in an area of high DO variability. At 110db and 125db there is also a lot of variability. If those samples are excluded the fit is:

CTD-BOT = 1.066 DOX-CTD – 0.020

The average difference below 150db where DO values are very low (<0.06ml/l) is 0.02ml/l. Removing some of the repeat bottles to find the ones most likely to have equilibrated properly reduces the slope slightly, to about 1.06, but the difference is not striking. (See 2009-38-dox-comp1.xls.)

This is a much larger correction than was found during 2009-27 (March 2009):

CTD-BOT = 1.0209 DOX-CTD + 0.0668

but there were only 5 samples available from that cruise. A more useful comparison is 2009-26 (April 2009) when the trendline was:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

This is reasonably close to the results for 2009-27, and also far from the results of 2009-38.

Saanich Inlet has a very high overall DO gradient with values ranging from 0 to 9ml/l in 200m and many small-scale features. At the end of this report there is a plot illustrating how difficult it would be to get bottles and CTD DO to match. There are overturning features at mid-depths seen in both temperature and dissolved oxygen and very high gradients at the surface. For other cruises, the Saanich Inlet results are buried in those of other sites. For 2009-09, for example, the scatter in the Saanich Inlet data is very high and the trendline is notably different from that for all the other casts. However, it should be noted that the trendline is highly dependent on the exclusion of individual points. Cruise 2009-38 is better designed to produce a good calibration result for Saanich Inlet, but the duplicate study suggests the sampling or analysis is not as reliable as usual. That might be because the high variability leads to Niskin bottle contents not being well-mixed. Perhaps a good fit for Saanich Inlet is impossible.

The duplicate Sp value of 0.13 (0.04 with 3 outliers excluded) is much higher than usually found, so it seems wisest to use the results of 2009-26. This will produce much smaller corrections than using the 2009-38 results. This could be revisited later if a large drift is noted when this sensor is next used or returned to the factory for recalibration.

Plots were made of CTD Dissolved Oxygen and Titrated Dissolved Oxygen versus salinity. The only notable outliers had already been flagged “d”.

March 11, 2010: COMPARE was rerun after replacing bottle values as detailed in section 10. The comparison does not suggest that further calibration is justified and the assessment of potential errors in the CTD DO remains the same. Flags were changed for a few samples based on the analyst’s remarks and the results of COMPARE.  (See 2009-10-dox-comp-final.xls.)

##### 12. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The usual shift of +24 records (1s) had reasonably good results, so it was applied to these data.

Dissolved Oxygen

Tests were run on the 2 casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. A setting of +60 looked best in some areas and +70 in others.

SHIFT was run using +65 records for all casts.

Conductivity

Tests were run using a variety of shifts and the best choice for both sensors was -0.3 records. SHIFT was run using -0.3s for both channels.

13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning was for cast #1 which had no downcast data. The file for that cast was put through REVERSE and then put through DELETE.

##### DETAILED EDITING

The noise level in the two T/S channel pairs look very similar. For 2009-03 and 2009-26 the secondary salinity appeared to be closer to the bottles. The secondary temperature and salinity channels were selected for editing. Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. All casts required light editing, except for cast #1 (from upcast) which needed heavier editing.

##### Initial Recalibration

Salinity will not be recalibrated because there are no bottles and cruises 2009-03 and 2009-26 suggest that no recalibration is needed for the secondary channel. Pressure will not be recalibrated.

File 2009-38-recal1.ccf  was prepared to apply the following correction to the dissolved oxygen channel based on the results of cruise 2009-26:

CTD-BOT = 1.0388 DOX-CTD + 0.0088

This was applied to the SAM file to create SAMCOR1.  COMPARE was rerun. The results show that the correction does not provide a perfect fit. The CTD DO values look low at the surface, high at mid-depths and about right near the bottom. The average difference between CTD and bottles is low, but the details are not a good as usual. This is probably due to the complex profiles and temporal variations in Saanich Inlet. Values are within 0.25ml/L. (See 2009-38-dox-comp2.xls.)

The recalibration was applied to the MRGCLN2 and the EDT files to create MRGCOR1 and COR1 files.

##### Final Calibration of DO

A further recalibration is sometimes applied to the downcast SBE Dissolved Oxygen. The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction is sometimes found appropriate to further correct for response time by comparing downcast CTD data to bottle data from the same pressure.  However, since recent improvements to the algorithm, it is usually not necessary, and for this cruise the bottle data is insufficient to resolve such small errors. So no further study was done.

##### Special Fluorometer Processing

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no problems were found.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity - Since its last factory calibration the primary conductivity sensor has been used for 3 other cruises. During 2009-03 it was found to be low by 0.004 and for 2009-26 it was high by 0.0025, though there was a lot of scatter in the latter. There was no calibration sampling for the 3rd cruise. The secondary sensor has been used on many cruises, but for most of those there were doubts about the salinity analysis. From 2009-03 the few bottles available suggest it is within 0.001 of the bottles and it was even closer for 2009-26 though there was a lot of scatter in that comparison.

2. Dissolved Oxygen – The sensor was used for 2009-03 when there were no bottles. The only calibration information was a comparison between 2 casts at one site with different DO sensors mounted. The comparison led to the following correction being applied:

DOX(corrected) = 1.018 * DOX +0.04

During 2009-26 the fit was

DOX(corrected) = 1.0388 * DOX +0.0088

The results of 2009-27 were similar but were based on only 5 bottles.

3. Pressure –The sensor on CTD #0550 was recalibrated in August 2007 at which time an offset of +0.1db was found. It has been used for 2 other cruises since then and an offset of +0.5db was found appropriate. Some drift in the offset is expected from this type of sensor as it ages.

Historic ranges – There is local climatology available only for cast #3; the temperature & salinity data for that cast fell within the historic ranges.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

There was no salinity calibration sampling for this cruise. During use in

the previous month (cruise 2009-26) the CTD salinity was within 0.001, but

there was a lot of scatter in the comparison.

For the SBE Dissolved Oxygen channel recalibration was based on the results

of 2009-26. There was calibration sampling from this cruise but the results

are not trusted due to the complex DO profiles in this area and larger than

usual differences between duplicate calibration samples.

SBE DO values are considered

•	 ±0.25 ml/L from 0-125db

•	 ±0.06 ml/L below 125m.

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found. Plots of all the S5 stations together show great variability in all CTD properties. The T-S plots look mostly stable, but there is a lot of structure.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. For station S5 the near-surface values were highly variable, starting at ~150%, ranging from 125-135% through the next 3 casts, then from 95% to 110% for the last four. Temperature and fluorescence also shows large swings in surface values. There are very few surface DO bottles to investigate these results, but where comparison is possible the SBE DO looks low as mentioned earlier. For station S10 the surface saturation is 83%; this is a well-mixed cast with no DO bottles.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and no problems were found.

A cross-reference list was produced and turned up no errors.

March 11, 2010: CHE files were put through CLEAN to update headers.

Header Check was rerun for CHE files.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars

1. Down cast not logged.

REVISION: August 18, 2010

Transmissometer #1005DR was calibrated in March 2008, and drifted significantly but steadily until July 2009; then a sudden shift occurred, so that maximum values between September 2009 and July 2010 were very low, ~25%/m. In August 2010 a study was made of transmissivity that led to a decision to apply post-processing corrections to all cruises between March 2008 and June 2010.

Transmissivity data from this cruise were corrected by multiplying the original values by correction factor 1.247. This was based on assumptions that deep offshore transmissivity from a June 2009 cruise should be about 62%/m and that drift was linear with time between March 2008 and July 2009. The corrections produced reasonable results for all cruises in that period.

For details on how the correction factor was derived see:

OSD_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer 1005DR Corrections.doc

These data should be considered estimates.

Revisions done by: Germaine Gatien

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| Fluorometer | 2356? | ? | IOS |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |