## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2009-40

Agency: OSD

Location: WCVI/Strait of Georgia

Project: La Perouse / Effingham / Georgia Strait

Party Chief: Juhasz T.

Platform: John P. Tully

Date: May 14, 2009 – May 25, 2009

Processed by: Germaine Gatien

Date of Processing: 27 January 2010 – 30 June 2010

Number of original CTD casts:  61   	Number of CTD casts processed: 60 (pumps off for 1 cast)

Number of bottle casts: 	40                	Number of bottle casts processed: 40 (1 cast has no pumped

CTD channels but bottle file was produced.)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#0997), a Seapoint Fluorometer (#2228) with a 10X cable (DO and FL both on the primary pump), a Biospherical QSP-400 PAR sensor (#4615), a pH sensor (#0692) and an altimeter (#1252). Two deck units were used, a model 11 s/n 424 was replaced with s/n 508 due to remote display problems. The data logging computer was #02588. The salinometer used was a model 8400B Autosal, serial # 69086.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were generally in good order. The file names were non-standard. There were errors in some file names – there was a warning about that in the log.

All CTD casts are reported to be ROS files in the log and there were BL files for all, but in many cases there was no sampling, so the ROS files contain only header information, no data. The log makes it clear that no sampling was done, so such ROS files were deleted.

For cast #1 the log record of how bottles were fired does not agree with what is in the data files.

For cast #20 the pumps were not turned on, so the pumped channels are not usable. The profile data were not processed, but a CHE file was prepared with Pressure, Transmissivity, PAR and the bottle sample data for salinity and dissolved oxygen. There was no note in the log about this.

The salinity was not recalibrated. The comparison between bottles and CTD sensor was noisy, but both salinity channels appear to be reading low by ~0.04. This is one of many recent cruises that show very large differences, and it is thought that the likely problem is poor seals on bottles and/or chipped bottles which would allow evaporation of samples. Analysis of these samples was 8 months after the cruise, so if evaporation were occurring it would be particularly serious in this case. When these sensors are next checked at the factory, this issue should be revisited.

There was a pH sensor mounted but as there is no pH sampling to enable calibration checks, that data will not be archived. Special files were prepared for the use of Cindy Wright with pH data included.

The comparison of bottles and CTD DO was unusual and suggests some problems with the SBE dissolved oxygen sensor. Time-dependent recalibration was applied to the data. The data are estimated, roughly, to be:

•	±1ml/l from 0 - 50db

•	±0.5ml/l from 50 -200db

•	±0.2ml/l below 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There were some problems reported with the acquisition computer crashing.

Extracted chlorophyll, nutrients, dissolved oxygen and salinity were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, pressure and DO sensors were obtained.

The calibration constants were checked for all instruments. There were a number of problems:

- The calibration information for dissolved oxygen sensor #0997 had the old Owens-Millard parameters. These were replaced with the new Sea-Bird parameters. The values for E, H1 and H3 were set equal to those used for 2009-09 when deep sampling enabled fine-tuning of those parameters. The values are expected to remain constant at least until recalibration of the sensor.

- The transmissivity date and parameters used at sea are wrong. These same values have been used for a few other cruises and the source of the information is unknown. The correct parameters were entered.

- The pressure offset for CTD #0443 has been drifting, and for other recent cruises a setting of +6.7db. For three cruises that followed this one an offset +6.8db was used, so that was selected for this mission as well.

After those corrections were made, the file was saved as 2009-40-ctd.con.

##### Conversion of Raw Data

The first file had a different configuration from all the others. It was converted using 2009-40-001.ctd.

Data for all other events were converted using the configuration file 2009-40-ctd.con. The only difference in the configurations is that NMEA positions were not acquired for the first cast.

The file names were non-standard; they were missing a digit in the event number section. A zero was added to all file names and the file names for 4 casts had to have event numbers adjusted, so

- 2009-40-017 became 2009-40-0018

- 2009-40-018 became 2009-40-0019

- 2009-40-019 became 2009-40-0020

- 2009-40-020 became 2009-40-0021.

The file named 2009-40-0116 contains the same data as 2009-40-017.hex, so will not be processed.

The file named 2009-016.hex contains little data; there is a note in the log that during event #18 there was a computer crash and the times show that is where this file came from – it will not be processed further.

The station name for event #59 was fixed as noted in the log book.

Cast #1 was an oxygen sampling test cast. The header had no position in it, but the rosette log does, so those values were put in the header. The time is wrong in the header, so it was changed to match the rosette log, though that time is for the bottom of the cast, rather than the beginning, but that is not a large error.

A channel existed in the configuration file for SPAR, but there appears to be no signal from that sensor, so it will be removed later.

A few casts were examined and all expected channels are present.

- The two temperature channels are generally very close during the downcast. The upcast data are much noisier with spikes in both channels. Conductivity is similar in spikiness. The two C channels are further apart than usual.

- The fluorescence data look reasonable; the dark value is ~0.13.

- Dissolved oxygen voltage has the usual offset between downcast and upcast.

- pH traces are similar to dissolved oxygen.

- PAR look fine.

- The transmissivity has little hysteresis.

- The altimetry looks usable.

- The descent rate was generally high, and varied from very steady to noisy.

Rosette files were converted using a start time of -5s and duration of 10s. The TAU correction was used for all files. The hysteresis correction was applied only to events #52, 53, 58 and 59 because those were deeper than 1000m. The file names were changed in the same way as described for the CNV files, and the station name corrected for cast #59.

The rosette files were then converted to IOS SHELL files.

Temperature and salinity were plotted for all IOS files and a few outliers were noted: secondary salinity for casts #7, 9, 24, 49 and 55.

CTDEDIT was used to clean or remove a few points in salinity for those casts.

Editing details were added to the header comments.

The output files were then copied to *.IOS.

CLEAN was run to add event numbers, with output named *.BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

When this CTD was used during 2009-08 and 2009-09 which followed this cruise, the best choice of CELLTM parameters proved to be (α = 0.03, β=9) for the primary and (0.02, 7) for the secondary. Tests were run to ensure they work well for this data as well and they did.

CELLTM was run on all casts using those values.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on the 4 deepest casts (52, 53, 58, 59) to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The differences are often noisy and there is some pressure-dependence in the salinity differences, but these are small differences overall. Some casts from 2009-08 which immediately followed this cruise are included to show the contrast.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 52 | 1000 1700 | ~0 ~0 | -0.00004 +0.00005 XN | -0.0004 +0.0004 | High, Noisy |
| 53 | 1000 1700 | +0.0001 +0.0002 | -0.00003 +0.00005 | -0.0004 +0.0003 | High, Noisy |
| 58 | 1000 | -0.0003 N | -0.00003 | -0.0001 | High |
| 59 | 1000 | +0.0002 | +0.00002 | +0.0002 | High |

During 2009-08 which followed 2009-40 larger differences were seen in conductivity and salinity and those differences grew through the cruise. That drift continued through 2009-09.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 08-32 | 1000 1500 | -0.0001 ~0 | +0.0003 +0.0003 | +0.0037 +0.0043 | High, f. steady |
| 08-39 | 1000 1400 | -0.0002 -0.0005 VN | +0.0003 +0.0004 VN | +0.0040 +0.0044VN | High, moderate |
| 08-67 | 1000 1500 1950 | -0.0001 +0.0002 +0.0001 | +0.0004 +0.0004 +0.0004 | +0.0048 +0.0053 +0. | High, v. noisy |
| 08-68 | 1000 1500 1950 | -0.0001 -0.0001 ~0 | +0.0004 +0.0004 +0.00045 | +0.0045 +0.0052 +0.0054 | High, noisy |
| 08-89 | 1000 1500 1950 | ~0 ~0 +0.0001 | +0.00045 +0.00045 +0.0005 | +0.0053 +0.0061 +0.0061 | High, noisy ~1000 Very noisy P>1250 |
| 08-94 | 1000 1500 1950 | ~0 VN -0.0001 ~0 | +0.0005 +0.00045 +0.0005 | +0.0055 +0.0057 +0.0062 | High, v. noisy |

So whatever went wrong with the salinity sensors must have occurred after 2009-40-0059. There is no note of trouble in the log.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. A problem was found with the first file – it was missing positions so those were added based on the log records. The time is wrong in that file, and further checks show all computer clock times (Upload Time) are wrong. Since the start time was not recorded in the log book, the time the CTD reached bottom was used – this is a shallow cast, so the error is likely only 5 minutes. The rosette file time and position for the first cast had been corrected earlier.

The cross-reference check was compared with the log book. The only error found was for cast #11. The header time differs from the log book by an hour, but the log time only allows 7 minutes to travel from the previous station, so the header is presumed correct.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run. The average surface pressure was 1.4db. This is low for the Tully, but the surface salinity values are low enough to suggest it is reasonably accurate.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet and many casts were checked looking at plots and the log book records. The algorithm worked well with no corrections required to the downcast files. The water depth was missing from the headers of files #10 and 11; these were added to the CLN and BOT files based on the log book entries. A number of bottom depths don’t agree with the log, so those were investigated based on maximum sampling + altimetry and the log looks right in most cases, but not all. Where there was no log entry, a change was only made if the header entry was clearly wrong (ex. bottom depth < maximum sampling). The header entry was changed to match the log in CLN and BOT for casts: 13, 26, 28, 32, 50, 51, 55, 56, 57, 59, 60 and 66.

##### BOTTLE FILE PREPARATION

The altimeter header entries were checked for a random sample of BOT files and no problems were found. There is sometimes a problem with the altimeter header when only surface bottles are fired, but there were no such bottles casts during this cruise.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets, and it was then converted to form CST files to be used as a framework for the bottle files.

The ADD SAMPLE routine was then run to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

Salinity analysis was done at IOS using a Guildline Autosal Model 8410B, serial # 69086. The analyst did not attach any flags. There were no duplicates. The spreadsheet was simplified by removing unneeded columns, headers were changed to standard format, the information from the bottle label column was separated into Station_Name and Sample_Number columns and an Event_Number column was created and filled in with information from the log. This was saved as 2009-40-sal.csv.

DISSOLVED OXGYEN

Titrated dissolved oxygen data were provided in spreadsheet format and included a precision analysis. There was a problem with the colorimeter which resulted in some loss of data, but the samples that were successfully titrated should not be affected by this.  The spreadsheet was simplified, event #s were added and some formats were changed; it was then converted into individual ADD files.

##### Compare

Salinity

Compare was run. The data are very noisy. When data with CTD salinity standard deviation >0.001 and differences >0.05 were removed, the primary salinity was found to be low by an average of 0.018 and the secondary low by 0.029; standard deviations in the differences were ~0.01 for each. There is a very slight time-dependence in the fits, but it looks very similar for both sensor pairs, so this may be more geographic than temporal. If the shallow casts of Effingham Inlet are removed plus the outliers identified previously, the fit is quite flat with time and with pressure. Removing those bottles from the comparison suggests that the primary and secondary CTD salinity channels are low by 0.016 and 0.023. Given the results of section 7 it is surprising to see such differences between the primary and secondary. However, for the 2 deepest bottles both salinity channels are low by ~0.04 and ~0.02. And if we restrict the comparison to 4 bottles for which the primary and secondary CTD averages differ by <0.001, the primary is low by 0.0383 and the secondary by 0.0379. The scatter in differences between primary and secondary CTD data reflects the noise level in these data. The large difference between bottles and CTD even when restricted to cases where primary and secondary are close shows that there is either a large calibration drift or problems with bottles. Based on other recent cruise results, the latter is more likely.

It is notable that the shallow casts in Effingham show larger differences from the CTD than those in open waters. This does not appear to be the result of particularly noisy CTD data.

The comparison of bottles with CTD turned up the following significant outliers in the fit against pressure:

- Cast 13, Sample #123 – higher than CTD salinity by 0.27 at 150db, CTD data look fine; flagged “d”.

- Cast #20, Sample #146 – pumps were not on for this cast, so bottle may be fine, but CTD data are not reliable. A bottle file will be prepared with pressure from CTD and bottle data, but all pumped channels should be removed (Temp, Cond, FL, DO & Salinity.)

There were other outliers, but some are associated with noisy CTD salinity data, or are very shallow. And there is so much noise in the fit that it is not entirely clear which are the outliers and which not. (See 2009-40-sal-comp1.xls.)

Dissolved Oxygen –

COMPARE was run for Dissolved Oxygen. The first impression of the fit of differences against SBE DO is that there appear to be two groups. Plots were made of casts with some anoxic areas in Effingham since that seemed an area likely to be different, and while it shows a lot of scatter, it is not notably out of line with the bulk of the points. Saanich Inlet is also a likely source of variation but there were only 3 independent values from that cast. There were 6 bottles at each of 3 depths and they were collected in an unusual way with CTD moved up and down repeatedly and not allowed to wait 30s before firing. The surface bottles are from an area of very high DO gradient and not surprisingly are outliers. One bottle from each of the other Saanich Inlet levels were selected. Given the problems with how they were collected, it seems unwise to put too much weight on these samples by including all of them.

When points were gradually removed until the odd group stand out, most of casts 1-14 were included.

The results are baffling. If the problem were due to the sensor doing badly in hypoxic conditions and then gradually recovering we would expect to see similar problems in the anoxic parts of Effingham Inlet and we don’t. When this sensor was used for 2009-51 (in September 2009) it was remarked that the fit was odd, but all bottles were from a narrow DO range. The odd fit of 2009-51 had the opposite slope to the one seen here, and speculation focused on the chemicals used for the analysis. The odd casts here also have narrow range, but that would seem more likely to lead to a smaller correction than a larger one, since there is less sensitivity to response time, alignment and the distance between bottles and CTD. The problematic casts are exclusively in the Strait of Georgia and Saanich Inlet. When the ship moved to the west coast of Vancouver Island the first rosette cast had no DO sampling. The next cast (#20) is useless because the pumps were not on. The rest of the cruise has smaller corrections.

There is no note in the log about the pumps being off, but it is possible that there was a problem that had been affecting the cruise up to that point that was then corrected. Looking in detail in the Strait of Georgia shows a gradually reducing slope, with the casts in the southern Strait having the largest slope, a little lower at the northern end and then as the ship moved to the central part lower again. Later in the cruise there are some variations from cast to cast, but they do not seem to be a matter of temporal drift. So a single fit for Events #22-69 looks reasonable. How many fits to use for the earlier group is not obvious because single casts don’t have enough data to be reliable, so as a compromise fits were done in 3 groups. Because there are few low DO samples, setting the offset is problematic. Only cast #1 has low values in Saanich Inlet, and as noted before, it is unwise to put too much stress on these values. The offset was fixed at +0.02 for the Strait fits since that was found for the offshore group and looks reasonable for the others. That offset works well for values >0.1mL/L, but gives DO that is a little high below that. That might be due to a limitation of the instrument or problems with bottom sampling.

So the fits found are:

Casts 1-5:

CTD_BOT = 1.1812 DOX_CTD + 0.02

Casts 6-9

CTD_BOT = 1.1375 DOX_CTD + 0.02

Casts 10-18

CTD_BOT = 1.1162 DOX_CTD + 0.02

Casts 19-66

CTD_BOT = 1.0772 DOX_CTD + 0.0211

These fits are rough, and there is no obvious explanation for the results. It appears that the calibration settled down after cast #20, but there is sufficient noise that there remains some doubt about that. It is possible that the problem is with the sampling or analysis, but this does not seem likely as the same analyst did all the work and there are no log notes indicating problems. Conditions were calm in the Strait. The differences from bottles are larger than for 2009-08 (May) and 2009-09 (June) when the same sensor was used. During 2009-51 in July when this sensor was used again, there were even more unusual results with the slope in the opposite direction to the usual which suggested a possible problem with chemicals.

If the problem is with the chemicals used for this cruise, then it would be best to apply the fit used for 2009-09, but there is no reason to believe that was the case. It is possible that there was an issue with the plumbing that gradually cleared. This would lead to gradually reducing slopes, and it was noted that alignment varied from cast to cast. Those tests are time-consuming and difficult to interpret when there are stops for bottles, so the variable recalibration is probably the most efficient and most reliable approach to fixing the problem. (See 2009-40-dox-comp1.xls.)

Another confusing issue is that the log notes do not agree with what the CTD recorded for cast #1. According to the log, the CTD was raised to the sampling depth, stopped 20s, then moved up 3m, down 6m, up 3m then wait 30s. This procedure was repeated 5 times. What is not stated is whether the bottles were fired at the end of each 30s wait, but I presumed that was the point. But in the BL files, the firings are within a few seconds of each other, and examination of the full file indicates there was a wait of no more than 11s before the first firing and no more than 24s before the last of them. Only 2 stops were for more than 20s. See file “study of stops – event #1.xls” for plots of what happened during the stops.

The following outliers were examined to see if flags are appropriate; plots of titrated DO and CTD DO versus CTD salinity were examined to help decide:

- Cast #1 – The surface samples were extreme outliers, but the CTD data were noisy and the local gradient was extremely high, so the distance between CTD and Niskin bottle is likely significant. No flags were assigned.

- Cast #20 – Most samples are extreme outliers, but since the CTD pumps were off, this is not surprising. No flags were assigned.

- Cast #24 – Sample #162 was flagged “b” by the analyst and is a mild outlier, so the flag was changed to “d”.

- Cast #29 – Sample #194 was flagged “b” by the analyst and is a mild outlier, so the flag was changed to “d”.

- Cast #30 – The sample from 28db is a mild outlier, but the CTD data are very noisy. No flag assigned.

- Cast #31 – The sample from 30db is a mild outlier, but the CTD data are noisy. No flag assigned.

- Cast #32 – The sample from 30db is a mild outlier, but the CTD data are noisy. No flag assigned.

- Cast #36 – Sample #244 had “bf” flag. Neither the average nor either of the duplicate values would show up as an outlier, though the higher of the duplicates is closer to the trendline.

- Cast #37 – Sample #250 from 2db is an extreme outlier. There is no evidence that the CTD fired at the wrong depth – salinity values are right for the surface. The bottle value is unrealistically low (0.73mL/L). Flag “d” assigned.

Notes were added to the headers for all samples previously flagged by the analyst to indicate whether COMPARE confirmed a problem. Because there were doubts about the data for this cruise, the “b” flags were left in place even when they looked ok in COMPARE. Whether the problem is CTD or bottles, it seems wisest to draw attention to potential problems.

Plots were made of dissolved oxygen (SBE and bottles) versus salinity and the only odd points had already been flagged or problems with the SBE data noted.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. This is a rough estimate as the upcast data is affected by the rosette weak, but the usual shift of 1s looks appropriate.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Conductivity

Tests were run on the conductivity sensors for 3 casts to find the best settings for aligning conductivity with temperature judged by the reduction in salinity noise.

SHIFT was run applying settings of +0.2s and -0.7s for the primary and secondary channels, respectively.

Dissolved Oxygen

Tests were run on a few casts to check what SHIFT value best aligns the Dissolved Oxygen channel to the temperature. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. The best values varied from one cast to another between +60 and +80 records. SHIFT was run using +70 records for all casts.

pH

Test were run on a few casts to determine that a shift of +50 records was appropriate based on matching the offset between upcast and downcast of the temperature and pH traces.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary temperature and salinity channels were selected for editing based on observations on later cruises that there was significant temporal drift in the secondary. The noise level is similar in both salinity channels.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. All casts required some editing.

Special notes:
Cast #18 – unstable features around 7-9db are considered real by Diane Masson, so were left unedited.

Cast #19 – descent rate was very low in the top 6db with several major reversals of direction in the CTD. This appears to be heavily corrupted by shed wakes to the point that it is impossible to confirm if any of the data are useful. All data were removed to 6.4db.

Cast #20 – pumps were off so this cast will not be processed further. (A bottle file will be prepared without pumped channels.)

After editing T-S plots were made to check if more editing was required of T-S data and while there are unstable near-surface features, these are not unexpected, especially close to shore.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:

- The primary conductivity sensor had repairs done in January 2009 followed by recalibration. It was used during 2009-08 and 2009-09, when it was found to produce salinity low by about 0.023 and 0.020.

- The secondary T and C sensors were used before this cruise for 2008-61, 2009-03 and afterwards on 2009-08 and 2009-09. The comparison for 2008-61 was very noisy and the result was not trusted. For 2009-03 the salinity was found to be high by 0.006. For 2009-08 it was low by 0.019 and there was a suggestion of time-dependence in the fit. During 2009-09 the salinity was found to be low by 0.0007 with some time dependence. The temporal drift is very odd as the differences are getting smaller, which is not the usual pattern.

2. Dissolved Oxygen – The sensor was used for 2009-03 when there were serious problems with deep DO sampling.  It was also used for 2009-08 and 2009-09. During the latter cruise there were sufficient bottles for a reasonable fit and recalibration was applied; calibration parameters were fine-tuned at that time. The fits for 2009-08 and 2009-09 are reasonably close. For 2009-51 there was a strange fit, but the range of DO and pressure sampled was unusual as well.

3. Pressure – The sensor is an older one prone to drift. An offset of +6.8db has been used most recently.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed, but there were problems with the climatology data – some were not available and for others only the 1 standard deviation data were available, so this step was not useful.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

While the primary salinity was found to be significantly lower than bottles, there are concerns about the bottle values based on recent cruises. Whether the problem is in sampling or analysis is not known, but the likeliest explanation is that the bottle liners or chips in the bottle lips may be allowing some evaporation to occur. It is notable that these salinity samples sat for 8 months before being analyzed. So if evaporation is a problem this is a cruise that should have been affected more than most. No recalibration will be applied at this point. When these sensors are next recalibrated, this can be revisited.

For the dissolved oxygen channel recalibration fits were done in groups to allow for the temporal drift. The MRG and SAM files were recalibrated using file 2009-40-recal1.ccf to apply the following DO corrections:

Casts 1-5:

CTD_BOT = 1.1812 DOX_CTD + 0.02

Casts 6-9

CTD_BOT = 1.1375 DOX_CTD + 0.02

Casts 10-18

CTD_BOT = 1.1162 DOX_CTD + 0.02

Casts 19-66

CTD_BOT = 1.0772 DOX_CTD + 0.0211

COMPARE was rerun using the recalibrated dissolved oxygen data and showed the recalibration was applied properly. (See 2009-40-dox-comp2.xls.)

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may partly correct for response time errors, but a further correction may be appropriate depending on how downcast CTD data compare to upcast bottle data from the same pressure. Sometimes there is a systematic difference, though in recent use this is often not the case. Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When outliers are excluded based on residuals, the average difference is +0.1ml/l with slight dependence on pressure and DO, but no significant temporal dependence. That is encouraging. A few attempts to apply a second recalibration to these data produced good results in some areas, but made others worse; some fits favour deep versus shallow casts or high values versus low. There is no obvious best choice, the offset is small and the fits noisy and based on a rough subset of the downcast data. So no further recalibration is justified.

##### Special Fluorometer Processing

Peak fluorescence values were examined to see if there were any instances of off-scale values. For casts #1 and #27 there are some data with values >14.5 but after surface records were removed in editing, no values that high are left.

There was no extracted chlorophyll data collected for this cruise, so no special files were prepared for A. Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots and profiles were examined on screen. Large spikes were noted in transmissivity data for cast #59 around 1224db, so those were replaced with pad values using a text editor. A comment was put in the header about the editing.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, pH, PAR:Reference, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

The PAR channel was removed from casts #52-60 only.

Separate files were prepared for Cindy Wright with pH channel for casts 1-51 and 61-66.

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity, Fluorescence, Dissolved Oxygen and PAR are nominal and unedited (unless

otherwise noted) except that some records were removed in editing temperature and salinity.

The comparison of SBE DO and bottles was unusual, and the results are not

trusted as much as usual. A time-dependent recalibration scheme was used for

the SBE DO.

The data are estimated, roughly, to be:

•	±1.0 mL/L from 0 - 50db

•	±0.5 mL/L from 50 -200db

•	±0.2 mL/L below 200db

The Standards Check routine was run and no problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The values ranged from ~95% to 135%. We expect spring values to be fairly high, but the question arises as to whether they are unrealistically high due to the method of recalibration. If the assumption was wrong that the odd results of COMPARE are due to sensor problems rather than analysis problems, then the recalibration scheme used would lead to values that are much too high in the southern Strait of Georgia, a little high in the middle of the Strait and just slightly too high for the off-shore and Effingham Inlet. But this does not appear to be the case, because the saturation was ~100% for 5 of the casts in the Strait of Georgia and some of the highest values were in Effingham and the offshore region. The value in Saanich Inlet was just under 130% which is not unusual there; the fluorescence was very high and transmissivity very low indicating intense biological activity.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

REMOVE was run to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter, PAR:Reference, pH and Flag.

The PAR channel was removed from casts #52-60.

Files with extension CHEPH were prepared for Cindy Wright; these included the pH channel for casts 1-51 and 61-66.

All pumped channels were removed from cast #20.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

The bottle data were exported to a spreadsheet to check that all data are present. All samples indicated on the rosette sheets were found, but there were a few salinity samples that were not shown on the rosette sheets. A few of these samples were checked to ensure they had not been mislabelled; they look reasonable as identified.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### 24. Producing final files

The sensor history was updated.

Particulars:

1. Test cast for Dissolved Oxygen sampling protocol comparison.

9. Deck unit was changed before this cast. At about this time a problem with bottle #1 was noted.

18. Data computer crashed at start.

18-21 – Errors in original file names

20. Pumps not turned on. Not processed.

30. A CTD with Optode was mounted in the rosette for testing that was to be done later.

36. Duplicate O2 test using different methods.

43. CTD stopped at 90m of upcast, returned to bottom to fire bottle #1

44. Particle study cast

47. Bottle #7 not tripped.

49. Computer crash – started cast again.

59. Labelled as LE09 – should be LE08.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 06May08 | Factory “ |  |  |
| Conductivity | 2128 | 30Jan09 | “ |  |  |
| Secondary Temp. | 2449 | 6May08 | “ |  |  |
| Secondary Cond. | 2424 | 7May08 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 0997 | 01Mar2008 | Factory |  |  |
| PAR | 4615 | 15Dec2000 | IOS |  |  |
| Fluorometer | 2228 | ? | IOS |  |  |
| pH | 0692 | 6Feb09 |  |  |  |
| Pressure Sensor | 63507 | 25/Oct/2004 | Factory |  |  |
| Altimeter | 1252 | ? | ? |  |  |