## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 24 Nov 2021 | ### Corrected Salinity:Bottle precision that was lost during HPLC addition. S.H. |
| ### 18 Dec 2020 | ### Added HPLC Data. Changed and flagged a Turner Chl value. S.H. |
| 03-Oct-2017 | Corrected MISSION metadata field in header. R.H. |

## PROCESSING NOTES

Cruise: 2009-43

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdeF

Party Chief: Chandler P.

Platform: Vector

Date: June 17, 2009 – June 24, 2009

Processed by: Germaine Gatien

Date of Processing: August 17, 2009 – April 26, 2009

Number of original CTD casts:  74	Number of CTD casts processed: 74 (1 very shallow)

Number of bottle casts:	22		Number of bottle casts processed: 22

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0941) was mounted in a rosette and attached were a Wetlabs   CSTAR transmissometer (#983DR), an SBE 43 DO sensor (#1117 or #1438), a Seapoint Fluorometer (#2336) with a 10X cable, a Biospherical QSP-400 PAR sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter. The salinometer used was a model 8400B Autosal, serial # 69086.

# SUMMARY OF QUALITY AND CONCERNS

Some of the file names were non-standard, having a space before the period.

The Daily Log Book contained no information about what equipment was used. There are log entries about some problems encountered, but it was not noted that there were many cases of jellyfish getting caught on the rosette. Recording these incidents can aid data processing; the salinity data from this cruise were frequently bad and biological fouling of the TC ducts or pump is a possible explanation.

The rosette log sheets were incomplete. There were notes about wrong sample numbers for casts #46 and #61; those numbers were corrected by some analysts but not by others. The combination of these two factors makes it difficult to be sure that analysis results are entered with the right CTD data. In addition the SBE Dissolved Oxygen sensor malfunctioned for many casts and there were few salinity samples, so the usual forensic tool bag is almost empty. Fortunately, there does not seem to be any question that the dissolved oxygen, extracted chlorophyll and nutrients are properly associated for cast #61, since if they followed the wrong labels the profiles would be very strange, but there remains some doubt about the salinity. Unfortunately, cast #46 had the additional problem that 16 bottles fired, but only 14 were listed on the rosette log sheet. It is believed that the correct association has been made, based on a low extracted chlorophyll value that looks like the 30db level with which it is associated, rather than the usual 20db that would have happened if the rosette had not been fired twice around 125m.

The process of confirming calibration parameters has become very inefficient. There is no central place where one can access this information. To confirm the parameters for this cruise required contacting 3 different people, not an easy task where sea-going personnel are involved.

The membrane on the dissolved oxygen sensor that was used for most of the cruise failed leading to values that are much too low; this appears to have happened before the cruise or almost immediately on deployment. The problem was not noticed until just before cast #53 when a different sensor was installed. CTD dissolved oxygen was removed from casts #1 to #52. In future a quick check of DO values early in the cruise is recommended.

The temperature and conductivity sensors were new, and the differences between the salinity channels were often <0.001, so there does not appear to have been any significant calibration drift. However, both conductivity channels were affected by intermittent problems that appear to be pump or duct related and lead to differences between salinity channels as high as 0.02. Generally, it is clear which conductivity sensor is providing more reliable data, but for event #6 (station 17) this is not the case. Decisions were made individually on which sensor pair to archive for each cast. There may be patches of salinity data of lower accuracy than usual where local variability masks problems or where both sensor pairs are affected. Comparison to bottle salinity is not helpful in resolving these problems, since there are few samples and never more than one per cast.

The comparison of bottles with the CTD salinity show differences ~0.005 for both CTD sensors (excluding cases where the two CTD channels did not agree well). This seems unlikely given the sensors were both recently recalibrated. This may indicate problems with sampling or analysis. If the seals on bottles were not excellent, then the 8-month wait between sampling and analysis may be relevant. No recalibration was applied, but this should be revisited after further use or factory checks.

Because of the problems with sample numbers checks were made of the salinity bottles that were outliers in the comparison to ensure that the large differences were not just due to mislabelling, but that does not appear to have been the case. It was helpful that the analyst did not perform the analysis in the order the samples were collected. The outliers appear randomly in the analysis order, so Autosal drift does not explain them. The scatter in the differences between bottles and CTD salinity increases somewhat late in the cruise when conditions were rougher, but this does not explain all large outliers. Outliers were flagged where there was no clear evidence that the CTD was producing noisy data, or the two channels were not in reasonable agreement.

One bottle was fired for cast #11 but no sample number was assigned. A CHE file was created in case the data are needed.

For events #53-74 the dissolved oxygen is considered:

•	±0.3   ml/l from 0– 25db

•	±0.08 ml/l from 25–150db

•	±0.03 ml/l below 150 db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

Some of the file names had a space before the period.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There was no list of equipment used, and no record of which pump the external sensors were mounted on. A number of problems were noted in the log, but the fact that the DO sensor was changed between casts 1 and 2 is not mentioned, so it is unknown why this was done or if this was just a correction of the configuration file. (Based on later observations, it looks like only the configuration file changed, not the sensor.) Later in the cruise the DO sensor was definitely changed, and note is made of that. There are many notes on the rosette sheets that samples were labelled wrong and new numbers are written beside the originally assigned numbers. The Daily Log Book has the corrected numbers. These rosette notes are confusing since it is not clear if the labels themselves were corrected, or just the rosette sheet. A quick check of analysis results shows that the sample numbers were corrected for some either by the analyst or those collecting the sample. For other samples the correction had not been made.

Dissolved oxygen bottle data were obtained in separate cast files with quality flags and comments, but this was later replaced with a spreadsheet file.

Extracted chlorophyll, nutrients and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The conductivity, DO and pressure sensors were either new or freshly recalibrated, so there was no history available.

There were two instrument configurations used during the cruise due to the switch in dissolved oxygen sensors. The two files are to be used as follows:
       2009-43-ctd1.con – for DO sensor #1438: casts #1 and 53-74.

2009-43-ctd2.con – for DO sensor 1117: casts #2-52.

The calibration constants were checked and the only error was found in one of the transmissometer parameters. That was fixed.

The parameters for the Surface PAR were confirmed correct by Melanie Quenneville. This information used to be available on OSDSHARE instrument calibration site, but Melanie has been unable to access the site to update it. The transmissometer information is also out of date on the site and there are some errors in what is there. The process of confirming calibrations has become very inefficient.

##### Conversion of Raw Data

Data were converted using the configuration files 2009-43-ctd1.con and 2009-43-ctd2.con.

The files names were fixed by removing a space before the “.CNV”.

A few casts were examined and all expected channels are present.

- The two temperature channels are mostly in reasonable agreement on the downcasts but there are some odd jogs. The upcast data are much noisier so there are significant differences. The primary temperature looks especially odd during upcasts. Conductivity is similar. There are small spikes in both C and T channels.

- The fluorescence was sometimes off-scale. The dark value is ~0.125

- Dissolved oxygen voltage looks very low for some casts.

- PAR, Surface PAR and transmissivity look fine.

- The altimetry looks good near the bottom.

Rosette files were converted using a start time of -5s and duration of 10s and bottom depths were corrected in 2 casts. The file names were fixed where a space was entered before the period. The rosette files were then converted to IOS SHELL files. CLEAN was run to add event numbers with output *.BOT. All BOT files were plotted. There were outliers in salinity in casts 10, 16 and 18. CTDEDIT was used to clean small spikes. Most data from the secondary channel for the bottle stop at 150db of cast #16 were removed. It looked like flow was impeded and cleared suddenly after a spike. The output files from the editor were copied to *.BOT.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Tests were run on 4 casts using a variety of settings for CELLTM. The differences among the various choices were small. Overall the best choices were found to be (α = 0.02, β=9) for both the primary and secondary conductivity channels.

CELLTM was run applying those settings to all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts using was plotted to check for agreement between the pairs of T and C sensors. The differences were mostly small, but the variability is odd and there are many cases of sudden excursions to larger differences.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2 | 250 325 | +0.0008 +0.0008 | -0.0003 -0.0003 | +0.003 -0.0026 | High, Steady |
| 6 | 200 250 | +0.0001 +0.0003 | -0.0012 -0.0010 | -0.013 -0.011 | High, Steady |
| 14 | 250 325 | -0.0008 ~0 XN | -0.0002 -0.05 | -0.0018 -0.5 | High, Steady |
| 25 | 250 325 400 | ~0 XN +0.0004 +0.0003 | +0.0003 +0.00007 +0.00005 | +0.0004 +0.0004 +0.0004 | High, Stead |
| 48 | 250 | Too noisy to judge | Too noisy | Too noisy | High, moderate |
| 64 | 170 | ~0 Noisy | ~0 Noisy | ~0 Noisy | Moderate, noisy |
| 74 | 200 | ~0 X Noisy | -0.0001 | -0.0007 | Extremely noisy |

These differences were from the downcast sections of the casts. Frequently the upcasts looked even more surprising. A close examination of the 2 channels showed sometimes problems in the primary, sometimes in the secondary. Here are a few examples:

- During cast #2 the primary salinity is bad from 140 to 200db of the downcast. While the temperature is noisy in that section, it appears to be the primary conductivity that is responsible for the jog to high salinity values. Then during the upcast, the opposite occurs with the secondary conductivity leading to very low salinity above 200db.

- For cast #6 the whole profile has high salinity and conductivity differences, with no obvious evidence as to which channel is at fault. The fluorescence is also full of spikes

- During cast #14 there is a similar problem with large differences below 250db of the downcast and above 250 in the upcast. Both sensors have problems, but the secondary looks worse. The only salinity bottle is from 349db so not helpful in this analysis, but it does show the secondary salinity as closer to the CTD at that level, but more out of line with the average values whereas the primary fits the general average.

- For cast #17, there is a shift in differences just before the bottom was reached, and it is the secondary conductivity that is a problem. There are blips in other channels but no obvious offset.

- Cast #20 has problems in the secondary conductivity below 245db and they get worse on the upcast.

- Cast #22 has clear problems in secondary conductivity but both may well be poor in places.

- Cast #30 has problems in upcast in both conductivity channels. This affects bottle stops too.

- Cast #45 looks ok at 275db, salinity difference ~0.001 and other casts at about that time have similar values.

- For casts after #52 there is no hint of this sort of problem, though the data are quite noisy due both to real conditions and ship motion. So a jog in values might not be seen. It is possible to see sections in which the two salinity channels are within 0.001 of each other.

Someone on the cruise reported that there were many jellyfish, though this is not noted in the log – could fouling of both pumps be a source of such noise? Might the TC ducts or pumps have been cleared at the same time as the dissolved oxygen sensor was replaced?

It will be necessary to look at each cast individually to decide which sensor pair is better for the downcasts. Not all are rosette casts, and even when salinity bottles are available the comparison is not likely to help since the problems are intermittent and there is at most 1 bottle sample per rosette cast.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. No problems were found.

The cross-reference check was compared with the log book, and two errors in station names were found and corrected – these were both mentioned in the log book. No other errors were found.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program was run and shows the average surface pressure to be 1.3db, which is a little low for the Vector. This is a new sensor, so we don’t expect an offset.  There is no evidence of “out-of-water” data near the surface, so it does look like it was just deployed higher than usual.

The altimeter readings and water depths were exported from the headers to a spreadsheet. A few errors were found in the bottom depths and corrected based on log records and altimetry. None of these affect rosette casts. Plots were made of altimetry for a selection of casts to see if the entries were reasonable, and despite frequent noisy patches near the bottom the algorithm worked very well.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. A few changes had to be made.

- The bottle fired for cast #11 has no sample # so it was removed from the list.

- Many sample numbers were changed on the rosette sheets with a note that the samples were labelled wrong; some analysts corrected the numbers, while others used the label information. This was difficult to sort out in the absence of complete sampling notes on the sheets.

One of the casts with sample number problems was #46. This is unfortunate because there were 16 bottles fired for cast #46 but the rosette sheet shows only 14. That sheet shows the target pressures but does not indicate where sampling was intended. There were 2 bottles fired at 125db and at 5db, whereas there is only 1 for each of those on the rosette sheet. It is unlikely that the drawers of samples knew this and chose Niskin bottles appropriately, but it is possible. The only salinity sample suggests that the 2nd sample was taken from Niskin #2. The first Extracted Chlorophyll sample has quite a low value, so was probably from Niskin #11 which was meant to be fired at 20db (the usual deepest CHL sample for this cruise) but was actually closed at 30db. This suggests that the first 14 bottles were sampled. There was one set of nutrient values (from the first bottle fired at 125db) that were all flagged “c” without comment. They are lower than all the other values in the profile including the other bottle at approximately the same level. It could have been mislabelled, but it would be just as easy to think that others in the profile are out of line. This was at station 59, a very complex area. So “c” flags will be left as is. The SBE Dissolved Oxygen data were bad for that cast, so they don’t help. Based on the evidence it seems likely that the data are entered for the right levels.

The addsamp.csv file was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

The salinity data were delivered in spreadsheet 2009-43.xls; there were no duplicates. There was no flag channel, but there was a channel for comments. No comments were entered so it was assumed that no flags were required. The analysis was done in February 2010, about 8 months after the samples were collected. The spreadsheet was simplified (unneeded columns removed and headers changed to standard format) and saved as 2009-43-sals.csv which was then converted to individual SAL files.

DISSOLVED OXGYEN

The dissolved oxygen data were originally provided in individual ADD files. There were many problems with those files. Later, the data were all reassessed due to problems found in the software. The new data were provided in spreadsheet 2009-43oxy.xls. This included an analysis of duplicates. The file was edited so the formats were standard and it was saved as 2009-43-oxy.csv.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2009-43nuts.xls which included a report on precision. The spreadsheet was simplified, ordered on sample number and saved as 2009-43nuts.csv. File 2009-43-nuts.csv was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file 2009-43chlarc.xls. There was no comparison of duplicates but “c” flags were attached where variability was >10% with variability assessed as the % (std dev/mean*100). The file was edited to add “f” to the flags where duplicates were averaged and a “c” flag was removed from the one case where there was no duplicate sample, since this is not usually cause for flagging and the absence of an “f” flag does give the message. The comment will be left to make it clear that a duplicate was not rejected for some reason since there are duplicates indicated on the rosette sheet. The file was edited to remove extraneous lines and columns and change header names to standard format and to add the event number. Sample numbers were changed by the analyst for cast #61. On the rosette sheet the samples had been identified as 219-228 but on the samples they were 218-227. The nutrient analyst changed the numbers to match the original plan because otherwise there would be a duplicate sample #. So the numbers in the CHL file were also changed to match what the nutrient analyst did. The dissolved oxygen and salinity samples also had those sample numbers. The file was sorted on sample number, and saved as 2009-43-Chl.csv. The simplified spreadsheet file was converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

COMPARE was run and shows a lot of scatter in the differences between bottles and CTD for both sensor pairs. There is little difference in shape between the primary and secondary plots. Based on the discussions of section 7, we must allow for sporadic problems with the CTD data. It appears that when operating well the two salinity channels are within 0.001 of each other, so a test was done by excluding any cases where the differences were larger than that. This included casts 14, 18, 27 and 57. Only the cast #57 sample stood out as an outlier, and for that case the differences were only 0.003 which is not significant. The other three only differed by 0.005 which is also not particularly noteworthy. There was something more notable about the cast #57 sample, and that was extremely high standard deviation in the CTD channels so this is a sample we would normally reject from the comparisons anyway.

When points are excluded based on the standard deviation in the CTD salinity a few outliers are removed (including cast #57), but there is still a lot of scatter and pressure-dependence, with the same pattern for both salinity channels. Removing more points based on differences produces a flat fit with the primary salinity low by about 0.0047 and the secondary by 0.005. This suggests that the noise is due to problems with bottles rather than then the CTD. As described earlier there are serious concerns about how the CTD sensors performed, but the problems don’t appear to explain the outliers in COMPARE because the differences are similar for both salinity channels for most of the outliers and the checks described in section 3 turned up shift and spike problems for only 1 bottle cast, and that one was edited to remove the questionable data. So the sporadic CTD problem is not expected to affect this comparison.

Two other explanations that deserve consideration are errors in sample numbers, but investigation turned up no evidence that this happened. The lack of tick marks on the rosette sheets means not all sample labels could be confirmed, and in a few cases it is possible that the wrong bottle label was assigned, but the bottle values are not significantly closer to the CTD data from any other bottle. Details for the outliers follow:

- Sample #55, Event #16 – Bottle much saltier than any CTD salinity in the profile, CTD lower than bottle by 0.060. Sampling ticked off on rosette sheet matches sample # and CTD data not noisy, long enough stop.  No obvious explanation for this outlier. There were problems with the CTD for this cast, but not at this level. Flag “d”.

- Sample #143, Event #34 – Bottle saltier than any CTD salinity in the profile, CTD lower than bottle by 0.018. Sample # agrees with rosette entry. Bottom two bottles have similar salinity. CTD data not noisy. No obvious explanation for this outlier. Flag “d”.

- Sample #171, Event #43 – CTD lower than bottle by 0.018. Rosette sheet confirms the sample is from the correct depth. CTD data not noisy. No obvious explanation for this outlier. Flag “d”.

- Sample #198, Event #46 – CTD slightly noisy, CTD lower than bottle by 0.021. No flag.

- Sample #210, Event #57 – CTD data quite noisy, CTD higher than bottle by 0.031. Rosette sheet confirms the sample is from the correct depth. Probably CTD is the problem here, bottle ok. Worth a note in the header. No flag.

- Sample #230, Event #64 – Bottle much saltier than any CTD salinity in the profile, CTD lower than bottle by 0.034. Flag “d”.

- Sample #243, Event #67 – Bottle saltier than any CTD salinity in the profile, CTD lower than bottle by 0.015. There is no rosette entry for this sample so can’t confirm the sample is from the right depth. There are two bottles below the one indicated by the sample number – the usual sampling pattern was to sample Bottle #2 whereas this sample # is from Bottle #3. If it were really from #2 the difference would be 0.011, so not much better.  Flag “d”.

- Sample #269, Event #73 – Bottle saltier than any CTD salinity in the profile, CTD lower than bottle by 0.014. Just inside the cut-off for a flag.

Fortunately, the samples were analyzed out of order and there is no grouping of bad samples, so this does not look like it was due to drift in the salinometer. The only case for which the bottle values are lower than the CTD is one for which the CTD data are poor. So it appears that the bottle values are too high. This could be due to the salinometer performing poorly, but there is no obvious evidence that was so. A more likely explanation is that the bottles were not sufficiently air-tight for 8 months storage; evaporation would explain the results.

The differences appear to be getting smaller with time, but the noise in the comparison gets worse. This could be related to the fact that the intermittent CTD problems were not seen in the second half of the cruise, and that the sea state was much worse leading to a lot of motion in the CTD and difficulty in gathering samples. There is no justification for recalibration of the salinity channels. One question that remains is which channels to archive for the bottle files; the problems seen in the downcast data were often not present in the upcast or replaced with problems in the other T/C sensor pair. As mentioned earlier the examination in section 4 turned up a problem with only one bottle from cast #16 and the bad data were removed, so using the primary for all channels is ok for the bottle files. (See 2009-43-sal-comp1.xls.)

Dissolved Oxygen –

COMPARE was run using all the data, but because there were 2 different sensors used, the results were analyzed in two groups: Group1: CTD #1438: casts 1 and 53-74

Group2: CTD #1117: casts 2-52

All data from the first cast stood out as outliers in the first group. When cast #1 is excluded the results from DO sensor 1438 look as expected. When a few outliers are excluded the fit is:

CTD-BOT = 1.0551 DOX-CTD + 0.0204

For that group there were only minor outliers that don’t justify quality flags, particularly given that they were all very close to the surface.

.

The data from the second group look most unusual with a very large offset. Clearly the DO sensor was reading much too low. When cast #1 data are plotted with Group 2, the first cast stands well apart from the others. Cast #1 does seem more like Group 2 than Group1, so it is possible that sensor #1117 was used for that cast. To test the idea that the wrong DO sensor is entered in the CON file, the data were reconverted using the other DO sensor’s values. The results were only slightly different and still much lower than the bottle values. It is possible that the same sensor was used for cast #1 as for casts 2-52, but if so, whatever was wrong got much worse between casts #1 and #2. The only trendline that come close to fitting the Group 2 data is a polynomial.

(See 2009-43-dox-comp1.xls.)

Plots were made of the titrated Dissolved Oxygen and SBE DO against pressure and salinity. Both suggest that the CTD data has the right shape, just the wrong range and/or offset. The voltages are extremely low so this is not a matter of the conversion being done wrong. Is there some setting that could be applied in acquisition so that the voltage range is reduced? Mellissa Hennekes reports that there were a lot of jellyfish found on the rosette during this cruise, so reduced flow to the sensor is a possibility, though this would not seem to be consistent with the distribution having the right shape. This seems more likely to be an electrical problem.

Fluorescence

COMPARE was run using the CTD CHL and the Titrated Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD CHL. When all data are included the Extracted Chlorophyll is about 30% higher than the CTD fluorescence, but when cases where fluorescence is >2ug/l are excluded, plus a few outliers, the Extracted CHL are about 80% of the CTD fluorescence. When Extracted CHL and CTD Fluorescence are plotted against file pair number the picture is clearer. In the earlier part of the cruise in the Strait of Georgia, Extracted CHL was low and CTD Fluorescence slightly higher. In Juan de Fuca Strait the CHL was higher and the CTD fluorescence was mostly lower than the extracted CHL. (See 2009-43-chl-fluor-comp.xls.)

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always rough estimate but the usual shift of +24 records (1s) looks appropriate. A shift of 24 records was applied to all casts. (Output: SHFFL)

Conductivity

Tests were run on a few casts and the best choice was found to be -0.3 for the primary conductivity channel and -0.5s for the secondary conductivity.

SHIFT was run using those settings. (Output: *.SHFC0 and SHFC1).

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment.

For the Group 2 casts the best setting is +55 records though the only casts deeper than 200db were marked by very noisy descent rates, and thus noisy data.

For the Group 1 casts, the membrane failed so there is impossible to align or recalibrate the data, and the data should not be archived. Cast #1 does not look useful, so it will not be aligned either. While it is still unclear which sensor was used it is likely that it was sensor #1117. In any case the data is bad using either configuration file, so this data should not be archived.

SHIFT was run using +55 records for casts #53-74.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning is for cast #54 because the maximum pressure was <10db. This is not a problem as that was the plan for this cast.

##### DETAILED EDITING

The choice of which sensor pair to archive is a difficult one because there are problems with both. Plots were made of every cast looking at the difference between the conductivity channels. Where the differences were >0.001 or shifted notably, careful inspection generally led to a conclusion of which channels were a better choice. Details for all the casts examined:

- #2 – secondary looks better, but there is a 0.004 shift downwards around 199db. The problems in the primary are greater – with large shifts. Transmissivity had a big spike during the downcast.

- #5 – both spiky but no obvious shifts. Secondary bit spikier, so choose primary.

- #6 – differences very high 0.01 and higher. The only obvious shift is in the secondary around 255dbof the upcast. The downcast is more like the upcast for the primary than for the secondary. The primary was selected for archiving, but a note will be put in the header to indicate that the salinity may be off by as much as 0.02.

- #14 – both have offsets, but the primary is less affected than the secondary. The primary was selected, but some T/S records will be removed, others cleaned. The only calibration sample came from the bottom where the two sensors were in closer agreement.

- #15 – the primary salinity has large offset in bottom 80db, so secondary was selected.

- #17 – Secondary poor at bottom, and probably upcast. Use primary.

- #20 – the secondary has an offset below 240db until just above the bottom. Primary was selected.

- #22 – primary has an offset from 200-240; secondary was selected.

- #25 – the secondary is bad just at the bottom and for the upcast; primary was selected.

- #27 – the secondary is bad at the bottom, both are very noisy in upcast; primary was selected.

- #30 - both have bad patches in the upcast, but the secondary is worse. Use primary.

- #31 – both have noisy salinity sections, but temperature is also noisy.

- #48 – the primary T and S had a large spike around 45db with no obvious cause– secondary looked better, so the secondary channels were selected.

Generally the primary channels were the better choice, but the secondary are better for casts #2, 15 and 22. The primary will be selected for all the others.

Even in the channels selected for editing there are cases of sudden shifts in salinity that are not explained by sudden changes in temperature or shed wake corruption. It looks as though the flow to the conductivity sensor is impeded and then suddenly clears, but the beginning of such an event is not easy to identify, so no attempt was made to edit the data around the shifts. The accuracy of the salinity will be affected by these shifts. It is likely that there are other shifts masked by local variability.

Since there are reports that there were many jellyfish found on the rosette, the transmissivity was examined for a few casts. During cast #2 the salinity differences increased from 0.005 to 0.01 at the same time as the transmissivity spiked to 0 values. Later, a second transmissivity spike is associated with another shift to even higher salinity differences, ~0.02. The salinity differences then went down to 0.004. Cast #30 had many transmissivity spikes and problems with salinity, though the spikes are not as clearly lined up with the changes in salinity differences as they were for cast #2.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

Note made of shifts: 2 (1 shift),

The following casts required no editing: 44, 62 and 54; the last one had data only in the top 3db.

The following casts required fairly heavy editing: 49-50 and 68-74.

All other casts required light editing.

##### Initial Recalibration

NOTE: This step and those that follow were run originally in August 2009, but after reanalysis of the dissolved oxygen bottle data, they were repeated for casts #53-74 using the new information. For the earlier casts the SBE:DO data is not usable, so those were not included in the reprocessing.

File 2009-43-recal1.ccf was prepared to apply the following recalibration to the SBE dissolved oxygen for files 53-74 only:

CTD-BOT = 1.0551 DOX-CTD + 0.0204

This was applied to the temporary SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the corrections were applied correctly and they were. (See 2009-43-dox-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

##### Final Calibration of DO

The test described here was run on the data before the DO correction was made. There is no reason to expect the outcome of this test to affect the results, so it was not repeated with the new data.

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data are compared to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When obvious outliers are excluded, the differences range from -0.5 to +0.4ml/l.with an average of =0.04ml/l. The fits were flat against pressure and DO values. The differences are small and because there is not a lot of data and quite a lot of scatter in the differences, further recalibration is not justified. (See 2009-43-dox-comp3.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 150db and processed separately for A. Peña. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

NOTE: After the FCTD and BOF files were prepared a reanalysis of the SBE Dissolved Oxygen was done and it was decided to recalibrate that channel for casts #53-74. This was a small change, ~0.02ml/l, so this step was not repeated since that channel is not critical to the work Dr. Pena will do with these files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity:  Both conductivity sensors were new so there is no history of use.

2. Dissolved Oxygen – Both sensors were calibrated at the factory early in 2009 and no other data has been processed that used either of them. Sensor #1438 was a new one calibrated in February 2009. Sensor #1117 was recalibrated in February 2009.

3. Pressure –The sensor was calibrated in March 2009 and no other data has been processed that used it.  No offsets are expected in such a new sensor.

Historic ranges – Profile plots were made with historic ranges of T and S superimposed. There were many excursions. In the northern part of the Strait of Georgia below 100db temperatures were generally lower than the historic maxima, while in the south-east the salinity was sometimes high. Juan de Fuca Strait data west of Victoria were mostly within the climatology except at station 101 where the deep salinity was lower than the minimum. Since these excursions go in both directions, there is no suggestion of calibration problems.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

REMOVE was run as follows -

- The following channels were removed from casts #2, 15, 22 and 48: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- The following channels were removed from casts #3-14, 16-21, 23-47, 49-52: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- The following channels were removed from casts #53-74: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel (with umol/kg units) was added for casts 53-74.

HEADER EDIT was used to fix formats and channel names and to add the following comments as follows:

-for casts #1-52:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

While the two salinity channels were generally in very good agreement, there

are cases where the differences between them were significant, up to 0.02.

The best salinity pair to archive was usually obvious, but there may be

patches of salinity data of lower accuracy than usual.

Oxygen:Dissolved:SBE data were removed because the sensor malfunctioned.

-for casts #53-73:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

The dissolved oxygen data in the CTD files are considered:

•	±0.3     ml/l from 0– 25db

•	±0.08   ml/l from 25–150db

•	±0.03 ml/l below 150 db

While the two salinity channels were generally in very good agreement, there are cases where the differences between them were significant, up to 0.02.The best salinity pair to archive was usually obvious, but there may be patches of salinity data of lower accuracy than usual.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made. Fluorescence has some odd features, with zeros (lower than the general dark value), spikiness near the surface and for a few casts the values went off-scale. Cast #2 has spikes in transmissivity. Temperature, salinity and PAR look ok.

T-S plots look fine.

The track plot looks fine.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from 70% to 115%. Most of the casts had saturations between 85% and 105%, with a few slightly lower values in Haro Strait and slightly higher in Juan de Fuca Strait. Where possible surface bottles were used to check the outliers and no evidence was found of poor Oxygen:Dissolved:SBE performance.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

As explained in section 11, the primary channels look best for all bottle files.

Remove was run as follows -

- The following channels were removed from cast #1: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- The following channels were removed from casts #2-52: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Oxygen:Dissolved:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

- The following channels were removed from casts #53-74: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units (casts 53-73 only).

At this point it was discovered that the CHL data had somehow been missed.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

##### Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

3. Rosette touched bottom lightly.

18. Problem with bottle tripping – had to go back down.

24. Header mislabelled change to Stn 25.

30. Mud on bottom.

33. Pump problem – restarted deck unit

36. Wrong stn # in header should be stn 46.

52. Unsure of DO readings so replaced with sensor #1438 after this cast.

53. DO readings better.

54. Very shallow cast.

61. HPLC and CHL samples not put in fridge and became warm

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 941 | Yes | Yes |

| Calibration Information CTD |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 5048 | 21Feb09 | Factory “ |  |  |
| Conductivity | 3579 | 18Mar09 | “ |  |  |
| Secondary Temp. | 5073 | 21Feb09 | “ |  |  |
| Secondary Cond. | 3581 | 17Mar09 | “ |  |  |
| Transmissometer | 983DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1438 | 3Feb09 | Factory |  |  |
| SBE 43 DO sensor | 1117 | 12Feb2008 | Factory |  |  |
| PAR | 4565 | 29Jan2009 | IOS |  |  |
| Surface PAR | 16504 | 24Feb2009 |  |  |  |
| Fluorometer | 2336 | ? | IOS |  |  |
| Pressure Sensor | 0941 | 16/Mar/2009 | Factory |  |  |
| Altimeter | ? | ? | ? |  |  |