## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2009-56

Agency: PBS, Marine Ecosystems and Aquaculture Division, Nanaimo, BC

Location: West Coast Vancouver Island

Project: Sardine Survey

Party Chief: McFarlane S.

Platform: W.E. Ricker

Date: July 22, 2009 – August 4, 2009

Processed by: Germaine Gatien

Date of Processing: March 6, 2013 – March 27, 2013

Number of original HEX casts: 42	Number of CTD casts processed: 41 (Pumps off for 1 cast)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with a Seapoint Fluorometer (#2229) with a 10X cable, a transmissometer (#723DR) and an SBE43 Dissolved Oxygen sensor (#0047).

# SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard. They have been changed to standard format. The last 4 digits contain the event number as recorded on the log sheets.

The CTD probably touched bottom during cast #34 and the secondary temperature and salinity data looked poor after that, so primary channels were selected for archiving except for cast #14 during which the primary channels were poor below 200m. Based on a post-cruise factory calibration in 2010 and the comparison with the other sensors, it appears likely that the primary salinity calibration is good to ±0.002 but this is based on assumption of linear drift over 2.5 years and many cruises. The secondary salinity at the time of cast #14 is likely good to ±0.002 as well; the secondary conductivity and temperature sensors had only been used on 2 previous cruises. No recalibration was applied to salinity.

The pumps were never turned on for cast #3 so it was not processed.

Pumps came on late for cast #5 so upcast data were selected for archiving; the quality of upcast data is somewhat lower than for downcast data.

There was some bottle sampling, but no data or samples could be found.

The dissolved oxygen data were recalibrated using the results of 2008-01 which is the only cruise that did DO sampling between the 2007 and the return to the factory in 2011. In 2011 no post-cruise calibration was provided because of damage noted. Profiles of DO were examined carefully to see if that damage occurred when the CTD hit bottom during cast #34. The DO signal was bad on the way up for cast #34, but thereafter the signals look reasonable with no obvious evidence of damage.

# PROCESSING SUMMARY

##### 1. Seasave

This step was completed at sea; the raw data files have extension HEX.

The file names were non-standard. Correcting the names was left until after conversion.

##### 2. Preliminary Steps

Copies of the log sheets were obtained. The only problem noted was the failure of bottle firing for one cast. That cast was repeated.

A spreadsheet with the log details was also provided.

Sampling was done but no sample data were available at the time of processing. Either the samples or the data are presumed to have been lost.

The cruise summary sheet was completed. The histories of the dissolved oxygen, conductivity and pressure sensors were obtained.

The configuration file used at sea had an error in the slope of the pressure sensor and the oxygen parameters were the old style (Owens-Millard). The file was corrected and saved as 2009-56-ctd.con. Nominal values were entered for the hysteresis parameters E, H1 and H3 since there has so far been no opportunity to fine-tune these parameters for this sensor. Given there was no very deep sampling this is of no importance.

##### 3. Conversion of Raw Data

All data were converted using the con file given above.

File names were then changed to standard format based on the log information.

A check to see if a rosette was used produced just one record at the surface of each cast, not from the level at which the Niskin was closed. No processing will be applied to those files.

A few casts were examined and all expected channels are present and look reasonable.

Downcast temperature and conductivity look good.

Both primary and secondary temperature and conductivity look noisy during upcasts, with the primary T and C having some very odd excursions.

The fluorescence looks ok with upcast and downcast being offset vertically by ~2m more than temperature, so a shift of 1s or 24 records should fix that.

The dark value of the CTD fluorescence is reasonable (<0.16ug/L).

##### 4. WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure, temperature and conductivity channels only.    Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

5. ALIGN DO

Tests were done on a few casts applying a variety of settings to make the offset between the upcast and downcast DO channel close to that for the temperature. The noisy upcast data make this hard to judge, but an advance of 4.5s appears to have the best effect overall.

ALIGNCTD was used to advance the DO Voltage by 4.5s relative to the pressure.

##### 6. CELLTM

Tests were run on 3 casts using a variety of settings for CELLTM to determine the best choice of parameters. The best choice varied for different depths and casts. For the primary there was a lot of noise in the upcasts making it hard to pick, but overall (0.02, 9) seems best. For the secondary the choice of (0.03, 7) looked best. All settings were a big improvement over not applying any correction.

CELLTM was run using (0.02, 9) for the primary and (0.03, 7) for the secondary conductivity for all casts.

##### 6. DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration; the tau correction was applied.

on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### 7. Test Plots and Channel Check

Tests were run on 3 casts to at least 350m to find the differences between the pairs of T, C and S channels. When a sudden change was noted 2 more casts were added to try to track down when the problem arose. The following values are rough estimates from downcast data with 1 cast from 2009-01 included since the same equipment was used then, though all casts for that cruise were < or =250db. Note that there is one comparison from deeper in the water column for cast #20 – most are from 390db:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-01-0259 | 240 | +0.0004 | +0.0002 | -0.0002 | High, fairly steady |
| 2009-56-0009 | 390 | +0.0004 | -0.00006 | -0.001 | very high, very noisy |
| 2009-56-0020 | 390 | +0.0007VN | -0.00007 | -0.0011 | very high, very noisy |
| 2009-56-0020 | 790 | +0.0004 | +0.0001 | +0.0008 | very high, very noisy |
| 2009-56-0034 | 390 | +0.0009 | +0.00004 | -0.0003 | very high, very noisy |
| 2009-56-0039 | 390 | +0.0008 | -0.00004 | -0.0047 XN | very high, very noisy |
| 2009-56-0040 | 390 | +0.0006VN | -0.00003 | -0.0049 N | very high, very noisy |

The temperature differences are very noisy, and there is pressure variation in the conductivity and salinity, though that is not unusual at these depths. The change between casts #34 and 39 is odd. There were spikes in all channels at the bottom of cast #34 and the secondary salinity was bad for the upcast and dissolved oxygen had virtually no signal. Most channels spiked at the bottom of the cast. Transmissivity reached values <10%/m, not quite zero, but it seems likely that the CTD may have touched bottom, perhaps in such a way that only the secondary sensors were seriously affected. After that cast the secondary salinity looks quite noisy at times and temperature and salinity differences were significantly higher. Conductivity differences are small and show no obvious trend, so perhaps only the temperature sensor was affected.

##### 8. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### 9. Checking Headers

The header check was run and it was noted that station names are missing from a few files and look wrong for a few others. The CLN files were fixed to match the log sheets. Bottom depths were included in the log sheets but not in the headers so those were added manually.

Some odd salinity values in cast #3 turned out to be due to the pumps being turned off for the whole cast. This file was not processed further.

The track plots (using event #s and station names) were produced and added to the end of this report.

The average surface pressure is ~0.3db which is lower than usual for the Ricker, but the associated salinity values are extremely low. The surface pressure shows no obvious change after cast #34. During the upcast of cast #11 there are records with pumps on near the surface. The values at 0.3db are clearly in water but at 0.1db they could be in or out. For cast #41 the downcast values show the CTD was clearly in the water for pressures very close to 0db.

##### 10. SHIFT

Fluorescence

In the past it has been found appropriate to advance the Fluorescence by +24 records (1s) relative to the other channels. A few casts were examined to compare the offset between the upcast and downcast fluorescence with that of the temperature traces to ensure this setting is appropriate and +24s looks like a reasonable choice.

A shift of +24 records (1s) was applied. (Output: SHFFL)

Conductivity

Tests were run on the primary and secondary conductivity channels using a variety of settings. There were variations from feature to feature but the best results overall were with a choice of -0.5 records for the primary and +0.5 records for the secondary. (The primary setting is the same as the one used for 2009-01 but the secondary was -1.1 records for the earlier cruise.)

All casts were put through SHIFT using -0.5s for the primary and +0.5s for the secondary conductivity.

Dissolved Oxygen

Tests were run on a few casts to compare the offset between downcast and upcast oxygen compared to that in the temperature traces and the alignment step taken earlier worked well so no further shift is needed.

##### 11. REVERSE

##### All files were put through REVERSE in case they are needed later. REVERSE files may be needed in case sample data become available later and are not from the bottom of the casts.

##### 12. DELETE

The full files were put through DELETE using the following parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0              		Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: All warnings pertain to the upcast surface data so are of no concern.

All DEL files were copied to *.EDT.

DELETE was run on the reversed files; the output files, *.REVDEL, will be saved in case the bottle data are later found and they can be used for any cases of bad downcast data.

##### 13. COMPARISON WITH BOTTLES

Bottle data were not available.

##### 14. Other Comparisons

Previous experience with these sensors –

- The primary conductivity sensor was recalibrated in May 2008 and was used many times over the following 2.5 years with few bottle calibrations and none that were really trusted. The secondary sensors were used less often. The post-cruise calibration at the factory in Dec. 2010 indicated that the errors in the primary and secondary sensors were equivalent to salinity being low by 0.003 and 0.005 respectively. The report on the secondary temperature did not show any damage, but more drift than usual. If we assume linear drift, then in July 2009 we would expect errors on the order of 0.001 and 0.002 for primary and secondary salinity respectively. But if something went wrong with the secondary, as the results of section 7 suggest, then the drift in the secondary was probably not linear.

- The pressure sensor was recalibrated in May 2006. The pressure offset was adjusted in spring 2010 by adding 0.8db to the factory offset.

- The dissolved oxygen sensor was repaired and recalibrated in April 2011. The only cruise with bottle sampling was 2008-01 in January 2008. Those results were used for the following 3 uses. Cruise 2009-56 was the last known use before the sensor went back to the factory. It had been damaged and no post-cruise drift estimate was given.

Historic ranges – Profile plots were made of T and S with local climatology superimposed where available. The only data that fell outside the historic ranges were salinity in the bottom 10db of one cast. The salinity range looked unbelievably narrow at that level, so there are likely few data points as deep as this in the climatology. There is no hint of severe calibration problems.

##### 15. DETAILED EDITING

The primary sensors appear to be the best choice for archiving. The results of section 7 suggest something went wrong during cast #34 that probably involved the secondary temperature. The secondary salinity was quite noisy thereafter. To make sense of the section 7 results we could presume that the primary drift was negligible at the time of this cruise and that most of the drift in the secondary occurred during this cruise. A difference between primary and secondary salinity of 0.001 early in the cruise and 0.005 at the end is consistent with that theory. The primary sensors were used many more times after cruise 2009-56 than the secondary so it is likely that most of the primary drift occurred after this cruise.

All casts were required some editing. Salinity was noisy near the surface for many casts which may suggest that there were plumbing problems with effects worst in high-gradient zones, but the noise is small-scale, so the associated errors are likely not large. Special problems were investigated for the following casts:

- For cast #5 the pumps were off from the surface to 11.2db of the downcast, so the upcast was chosen for editing and archiving. The quality of upcast data are considered a little lower than downcast data, but that seems less serious than the loss of a quarter of the data.

- For cast #6 both primary and salinity data and downcast and upcast look odd between 12db and 17db. The temperature seems ok and the data are marginally stable in T-S space. The pumps were on. Given that the upcast looks similar would indicate this is not an “overshoot” caused by poor alignment and the temperature gradients do not seem high enough to suggest that the CELLTM correction could cause this feature, so the data will not be removed. Downcast data were selected for editing and archiving.

- Cast #14 had bad primary salinity from 200db to the bottom. The secondary was ok, so that so secondary channels were chosen for editing and archiving.

On-screen plots of descent rate and pump status were used to guide editing.

Note was made of the editing details in the headers of the relevant files.

##### 16. Initial Recalibration

The dissolved oxygen data were recalibrated using the results of 2008-01 since that was the only cruise with bottle sampling between factory visits.

The pressure does not appear to need adjustment.

There was insufficient information on which to base recalibration of salinity and it is likely both channels were within 0.002.

File 2009-56-recal1.ccf  was used to apply the following correction to the DO channel:

Dissolved Oxygen (Corrected) = 1.0633 * Dissolved Oxygen - 0.0194

##### 17. Fluorescence Processing and special files for Angelica Peña

The edited files were put through a median filter, size 11, applied to the fluorescence channel only.

##### 18.  BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used. Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. No further editing was considered necessary.

Transmissivity and Fluorescence profiles were plotted on-screen and no problems were noted.

##### 19. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts except #14: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag.

The following channels were removed from cast #14: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate and Flag.

A second CTD dissolved oxygen channel was derived with mass units (umol/kg).

REORDER was run to put the two DO channels together.

Dissolved Oxygen saturation was calculated and surface values were plotted. The values varied greatly with no obvious pattern except most of the lower values (70-90%) were in areas where tidal mixing was a possible factor. The highest values 130-170% were close in time and space. The cruise track was complex so any analysis is difficult. 

An attempt was made to see if there was any obvious damage to the oxygen sensor after cast #34. First, a profile of up and down casts together shows similar offsets before and after cast #34. But for the upcast of cast #34 the DO signal was bad on the upcast with values near 0. For casts thereafter the DO values and saturation are relatively low. Plots of DO saturation versus transmissivity show that earlier, when the DO saturation was >130%, transmissivity was usually low, often <15%/m and always <30%/m. After cast #34, DO saturation is relatively low, but transmissivity is also relatively high, so it may just be that conditions had changed. There is no evidence that the dissolved oxygen sensor was performing badly. In the absence of bottle data that is all we can determine.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

Dissolved Oxygen data were recalibrated based on the results of cruise 2008-01

because there was no calibration sampling during this cruise. Since it is likely

that some drift had occurred, these data must be considered less reliable than usual.

Salinity was not recalibrated because there was no calibration sampling available

The history of the sensors suggests that calibration drift was likely small in

temperature and conductivity.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

##### 20. Producing final files

A cross-reference listing was produced for CTD files.

HEADER CHECK was rerun and a cruise track plotted and no errors found.

The sensor history was updated for the CTD sensors.

Particulars:

3. Pumps off – not processed

4. Bottle mechanism failed

5. Repeat of #4 for bottle sample. Pumps off until 11db down, so upcast processed for archive.

14. Primary sensor data poor near bottom. Used secondary for archive.

34. CTD appears to have hit bottom

# CRUISE SUMMARY

| Cruise ID#: 2009-56 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Dates: Start: 22 July 2009 End: 4 August 2009 |  |  |  |  |  |
| Location: WCVI |  |  |  |  |  |
| Vessel: W.E.Ricker Party Chief: McFarlane S. |  |  |  |  |  |
| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506         Cruise ID#:		2009-56

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4752 | 06/03/07 | Factory |  |  |
| Conductivity | 2173 | 07/05/08 | “ |  |  |
| Secondary Temp. | 2968 | 22/08/07 | “ |  |  |
| Secondary Cond. | 2399 | 13/06/08 | “ |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Transmissometer | 723DR | 24/01/08 | IOS |  |  |
| Pressure Sensor | 69698 | 26/05/2006 | Factory |  |  |