## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2009-59

Agency: OSD

Location: WCVI/Strait of Georgia

Project: La Perouse, Effingham, Moorings

Party Chief: Juhasz T.

Platform: John P. Tully

Date: September 15, 2009 – September 25, 2009

Processed by: Germaine Gatien

Date of Processing: 18 May 2010 – 19 August 2010

Number of original CTD casts: 44	Number of CTD casts processed: 44

Number of bottle casts: 	21		Number of bottle casts processed: 20 (no sample #s for 1)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was used during this cruise. It was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#983DR), an SBE 43 DO sensor (#1176 on the primary pump), a Seapoint Fluorometer (#2228) with a 10X cable (on the primary pump), a Biospherical QSP-400 PAR sensor (#4615) and an altimeter (#1252). All casts were done using the LARS mid-ship station.

# SUMMARY OF QUALITY AND CONCERNS

The CTD and rosette logs were in good. There were extensive notes on CTD problems and trouble with Niskin #7 (also noted during 2009-11 which preceded this cruise.)

The choice of T/S channels to archive was difficult. There was no rosette salinity sampling (because bottles were not loaded.) There were many cases of very bad data in the primary channels in the top 20db. However, the secondary data are often very noisy in the top 30db. While the primary data are smoother there, it is not clear that they are better. During 2009-10 when the same equipment was used, there were problems that were attributed to the primary pump. On balance, it looked best to choose the secondary channels for the archive.

The dissolved oxygen gradients were often very high near the surface and variability was high throughout the profile for many of the casts. The precision of the SBE dissolved oxygen channel is estimated on the basis of a comparison to bottle data to be, roughly:

•	±2.5 ml/l  from   0-20 db

•	±0.5 ml/l  from  20-50 db

•	±0.3 ml/l  from 50-250 db

•	±0.1 ml/l  below   250 db

A long-standing error in the calibrations used for the pressure sensor was discovered after these data were converted, so the proper parameters were used. There is some question of what offset to apply. For this cruise a net offset of -0.2db was applied based on the results of 2008-04. When further experience has been gained, this choice can be revisited.

The transmissometer has not been recalibrated since 2006 and the values appear to be too low, but there is insufficient evidence to enable correction.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. There were a number of CTD computer problems. Rosette bottle #7 often failed to close. The chief scientist confirms that the deployments were done with the LARS based on some comments in the log.

Titrated dissolved oxygen data were obtained in spreadsheet format.

No salinity samples were gathered because no bottles were loaded.

The cruise summary sheet was completed.

The histories of the pressure, oxygen and conductivity sensors were obtained.

The calibration constants were checked for all instruments. A few changes were made:

- The transmissometer offset is wrong – it should be -1.1772, not -1.2737.

- The oxygen calibration data was updated with the new values for parameters E, H1 and H3 found during 2009-10.

- The pressure calibration was changed to the proper parameters for 08 March 2000. Shortly before processing this cruise it was discovered that the wrong parameters have been used for many years. There will be a need to adjust the offset since there has almost certainly been considerable drift through the time. For 2008-04 an offset of -0.2db was used (thus adding +0.9155db to the factory offset of -1.1155db); the results suggested that surface pressures are a little too low and with another year’s drift it looks reasonable to add a further 0.2db. So, for this cruise, an offset of 0 was applied, and the results will be studied later.

##### Conversion of Raw Data

Data were converted using configuration file 2009-59-ctd.con.

A few casts were examined and all expected channels are present.

- The two temperature channels are generally very close during the downcast but there are a few odd sections near the surface. The upcast data are much noisier with significant differences. Sometimes one channel is noisiest, sometimes the other.

- Conductivity channels are similar in spikiness and neither is obviously better.

- The fluorescence looks ok with a dark value ~0.1 to ~0.17, which is higher than during 2009-11 using the same sensor. The offset between Fl and Temperature taking into account the descent rate, suggests a 1 or 2s offset for FL.

- Dissolved oxygen voltage has the usual offset.

- PAR data look fine except that it looks like the sensor was probably removed for the deep cast, #24, and mounted again for cast #30. There is no note in the log that it was removed, but there is no noticeable signal. There are other casts with little signal due to it being night, but they have a slight variation. The channel won’t be removed since even if mounted there would have been no significant signal because casts #24-29 were all at night. So it will not be misleading and just possibly the sensor was present.

- The transmissivity has no noticeable hysteresis; there are isolated spikes at depth, but mostly small ones.

- The altimetry looks noisy near the bottom for some casts, so checks will have to be made to ensure the header entries are valid.

- The descent rate was kept fairly high. It was very steady for many casts, but very noisy for the offshore casts where there were many complete reversals of the CTD.

Rosette files were converted using a start time of -5s and duration of 10s. The TAU and hysteresis corrections were selected.

The rosette files were then converted to IOS SHELL files.

CLEAN was run on all the IOS files to ensure the event number is in every header, output files were named *.BOT.

Temperature and salinity were plotted for all BOT files. There were many outliers in shallow water, but none looks like editing is appropriate.

Header Check was run to see if fluorescence went off-scale and no evidence of that was found.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.

Parameters used were: 	Pass 1    Std Dev = 2 	Pass 2    Std Dev = 5 	Points per block = 50

##### CELLTM

Finding casts suitable for testing the CELLTM settings was difficult because those with a steady descent rate and no bottle sampling are mostly shallow. Tests were run on the best available and the best choice overall proved to be (α = 0.03, β=9) for both the primary and secondary.

The same choice was found appropriate for 2009-10 and 2009-11 when the same equipment was used.

CELLTM was run on all casts using that setting.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. There were only 2 deep casts and those were consecutive and had very noisy descent rates. One shallow cast was also examined to see if there is any hint of temporal changes. Data from 2009-11 are included in the table to check for temporal changes.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-11-37 | 1850 | +0.001 XN | -0.0002 VN | -0.0037 XN | V High, V Noisy |
| 2009-11-75 | 1850 | +0.0009 XN | -0.00022 VN | -0.0039 XN | V High, V Noisy |
| 2009-11-81 | 1850 | +0.0009XN | -0.00024 XN | -0.004 XXN | High, X Noisy |
| 2009-59-14 | 200 | +0.001 | -0.0004 | -0.005 | High, Steady |
| 2009-59-23 | 900 | +0.0012 N | -0.0004 | -0.005 N | Mod, X Noisy |
| 2009-59-24 | 1000 1500 | +0.0007 N +0.0007 N | -0.0004 -0.00035 | -0.005 N -0.005 N | Mod, X Noisy “ |

The temperature differences are smaller than seen in the two earlier cruises, while the conductivity and salinity differences are higher. This suggests that the trend of increasing conductivity and salinity differences seen during 2009-10 and 2009-11 is continuing. While there is no hint of time-dependent changes within the 2009-59 data, that is not surprising because of the limited information. There is no evidence of pressure dependence in the differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run and no problems were found. Fluorescence values were only near the maximum for two casts and only at the surface as the pumps were turned on. That data will be removed either by DELETE or in editing, so there is on need for a separate step to remove it.

The cross-reference check was compared with the log book, and the only problems noted were one missing station name and format errors in the naming of a few others. Those were fixed.

After those corrections, cruise track plots were prepared and added to the end of this report. No further problems were found.

The surface values program was run. The average surface pressure was 2.4db which is low for the Tully. For many casts the near-surface salinity looks very low. For example, for cast #41 at 1.75db the salinity is very low but the pumps had not been turned on. There are other casts with similar surface pressures that have salinity values >28 before the pumps came on and the cases of low surface salinity are all very close to shore.  Cast #17 has a salinity of ~30 at 1.75db after some time soaking, but pumps off. So the pressures are probably not too low.

A few casts were checked to see if the maximum depth measured (converted from pressure) plus the altimeter readings at that depth is close to the bottom depth recorded in the log. In each case the results were very close with the pressure usually a little too high rather than too low. This is a very rough check with unknown errors in altimetry and the bottom depth measurements, but offers some reassurance that the pressure calibration is close to correct.

This pressure sensor has been used a lot, but a recent discovery of errors in the calibrations used for most of the cruises means the history is not very helpful. Cruise 2010-03 suggests that using a 0db offset leads to values that are too high, Given the uncertainty, it seems wisest to use the setting applied to 2008-04 (a net offset of -0.2db). This choice can be revisited when we have more experience using the sensor with the correct calibrations. We will expect the offset to increase with time, but maybe very slowly.

The altimeter readings from the headers of the CLN files were exported to a spreadsheet and all casts were checked since the signal was very noisy.

Corrections were made to the header entries based on the plots and a note was made that the estimate was based on plots for the following casts: 2, 5 and 6.

The header entry was removed from the headers of the following casts because the signal was too noisy to make a reasonable estimate: 15.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Two bottles were fired during cast #13, but this was only for test purposes with no sample numbers, so this file will not be processed further.

After those corrections were made to the addsamp.csv file, it was converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged. The SAMAVG files were checked to ensure that the altimetry header makes sense and no problems were found.

SALINITY

There was no salinity sampling.

DISSOLVED OXGYEN

Nina Nemcek revisited all the DO analyses of 2009 because of a software problem that led to values being a little low. New data were delivered in spreadsheet 2009-59oxy.xlsx in July 2010. This included a duplicate analysis, averaged values for duplicates and appropriate flags and comments.

The spreadsheet was edited to change headers to standard format and to add the prefix “DO:” to the comments. It was then converted to individual ADD files.

The ADD files were merged with CST files and the output files were put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output: MRG).

##### Compare

Dissolved Oxygen –

COMPARE was run for Dissolved Oxygen. There were many outliers. First, the most extreme outliers were removed, then gradually more were excluded based on residuals. The resulting fit using all casts was:

DOX_BOT = 1.0689 * DOX_CTD – 0.0113

Using a polynomial fit was no better with the R2 value ~0.85 for both.

When divided into three sections the fits were:

DOX_BOT = 1.0537 * DOX_CTD + 0.0457 (For casts #2-16 – Broughton excluding 1 outlier)

DOX_BOT = 1.0502 * DOX_CTD + 0.0588 (For casts #2-16 – Broughton excluding samples with ambiguous endpoints and 1 outlier)

DOX_BOT = 1.0773 * DOX_CTD - 0.0328 (For casts #2-16 – Broughton excluding 3 outliers)

DOX_BOT = 1.0740 * DOX_CTD +0.0127 (For casts #19-23 – Off-shore excluding 2 extreme outliers)

DOX_BOT = 1.0803 * DOX_CTD - 0.0187 (For casts #19-23 – Off-shore with 6 excluded)

DOX_BOT = 1.0780 * DOX_CTD - 0.0146 (Offshore – deep cast only #23)

DOX_BOT = 1.0576 * DOX_CTD – 0.0154 (For casts #29-40 - Effingham Inlet excluding ~31 pts)

DOX_BOT = 1.0702 * DOX_CTD + 0.0028 (For casts #29-40 - Effingham Inlet excluding 11 pts)

The question arises as to whether these groups should be recalibrated using separate fits or just 1.

- The range of DO samples is very limited for the Broughton samples, with most between 2.7 and 4mL/L. Even within that range most values are tightly packed at about 3.2 and 3.8mL/L. The only values outside the 2.7-4mL/L range are 5 cases above 4mL/L; for the 3 highest of those, the analyst noted ambiguous endpoints. Several of the casts are very well mixed. The choice of outliers has a huge impact on the fit and the offsets seem unrealistic which is no surprise with no DO readings <2.7mL/L. So recalibration based on only these samples appears unwise.

- There are samples from only 3 offshore-casts. There is a lot of scatter but no matter how outliers are identified the slope is higher, though only slightly so when all but extreme outliers are included. The offshore casts are different in character from the other groups, having fewer small-scale features and more motion during stops. There was 1 deep rosette cast. Recalibrating this group separately might be justified, but examination of the offshore CTD casts without bottle sampling show higher DO values near the surface than were captured at these 3 stations, so they may not do well for all offshore casts.

- There is a lot of scatter in the Effingham casts much of which is explained by the very sharp gradients in the top 20m where there is a narrow subsurface maximum. There are other unusual features making these very challenging profiles for the SBE sensor. A few casts were examined in detail and stops for bottles were all at least 30s long. In Effingham Inlet the temperature and oxygen often display a lot of variability even after the wait, even at the bottom.

Another fit was done by looking at plots to exclude bottles where DO was still changing significantly at the end of the bottle stop. It is noted that even when the standard deviation in the CTD data is low, there was sometimes a smooth drift in values so that the match between what was in the bottle and what the CTD measured could be significantly different. Particularly in Effingham Inlet there is evidence of changes at the bottom suggesting a bottom current. When a few other outliers were excluded the fit was:

DOX_BOT = 1.0705 * DOX_CTD – 0.013

This is remarkably close to the initial fit.

It appears best to use the fit with all casts for the whole cruise since the early casts have the least noise but lack a broad range in the bottle samples, while the Effingham casts are very noisy with challenging conditions for the sensor and noisy temperature measurements. The offshore has steadier gradients and deeper sampling but noisier sea-state and the 3 casts sampled may not be representative of the whole offshore area. It is hoped that combining the groups will minimize the errors from each. Broughton has high gradients in the top 25m, then well-mixed below. Effingham has a mix of features, with sharp gradients common at the surface, and steady gradients below that for some or little or complex variation for others. The Effingham data dominates the fit by virtue of the most samples.

Outliers were examined to decide whether flags were needed.

Four major outliers were identified, but for none do flags appear appropriate:

Cast #29 – sample #98 – high gradient, there is no sample #98 on rosette log – Value moved to #99 – and #92 should have been assigned to this cast, not cast #28 – there was no sampling for #28 and no sample #s assigned. All data for #29 are misaligned. Fixed CHE file. None of these data had been included in the comparison.

Cast #31 – sample #104 – DO was drifting at the bottom, so a mismatch between CTD and bottle likely

Cast #36 – sample #148 - high gradient, so problem likely mismatch or poor CTD response

Cast #40 – sample #172 – CTD data noisy, high gradient in DO

The analyst assigned some “b” flags to indicate that there were minor problems in the analysis. The values are probably fine. Normally COMPARE might be used to determine if they are ok. Unfortunately, the comparison was so noisy that it is hard to make a judgment. In only 2 cases were the “b” flagged points clear outliers in COMPARE. For one of those, the CTD data did not look reliable and for the other, it was found that the sample had been mislabelled.

The analyst has asked that the “b” flags be left in place with a note of explanation.

All MRG files were put through CLEAN to remove Sea-Bird headers and comments from the secondary files.

Plots of DO versus salinity turned up only 1 major outlier which was from the sample that is thought to have been mislabelled. Once that was fixed there were no major outliers.

##### 13. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles for a few casts to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the sum of the descent and ascent rates to find the shift (in seconds) to remove that offset. The usual shift of 1s looks reasonable though the noisy upcast temperature make the judgment a very rough one.

SHIFT was run on all casts to advance the fluorescence channel by +24 records. (Output: SHFFL)

Dissolved Oxygen

Tests were run on a few casts for each sensor to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. A value of +90 seemed best as was found during 2009-10 and 2009-11 and looks reasonable for this cruise, though in many cases there is too little DO variability to judge. SHIFT was run using +90 records for all casts.

Conductivity

Tests were run on each conductivity channel for a few casts using a variety of shift parameters and a setting of +0.5 records looks best for the primary and +1.1 records for the secondary conductivity..

SHIFT was run on the primary conductivity using +0.5 records.

SHIFT was run on the secondary conductivity using +1.1 records.

##### 14. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

Note that cast #32 was rerun with the low drop rate featured turned off because it was discovered in editing that too many good data points had been rejected. A note was made about this in the header.

##### DETAILED EDITING

There is no salinity calibration sampling for this cruise and no obvious difference in spikiness between the two T/C sensor pairs. There is a lot of fine-scale noise in the secondary salinity. For 2009-11 and 2009-46 which bracketed this cruise, the primary sensors were selected. An initial attempt to edit the primary channels turned up a serious problem in the top 20db of many casts, so the secondary channels will be edited instead.

The following casts have suspicious data in either primary conductivity or fluorescence channels in the top 20db: 1, 5 (secondary a little odd too), 6, 7, 16, 30-32, 34-37, 40.

Those casts are all close to shore. The problem looks like a plumbing blockage that clears by 20db. That could be a failure of the primary pump to turn on due to low salinity but the pump status shows them as on and sometimes the salinity is too high rather than too low. Perhaps it is related to surface debris though why that should be more of a problem for the primary and rarely for the secondary is not obvious. Similar problems were noted for 2009-46 which suggests a plumbing problem as the issue rather than debris. Since the fluorometer and DO sensor were on the primary pump, we might expect some problems there too, but there is no obvious evidence that fluorescence or dissolved oxygen were affected. They are both noisy enough near the surface that we might not see the effect

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used. Editing was used to remove spikes where they are systematic in direction and/or likely to affect the metre-averaged results. Records were removed that were corrupted by shed wakes or near the surface before the pumps were turned on. Editing was also applied to salinity data in some very spiky sections in the top 30db; the cause is not known but is likely to involve plumbing/pump problems.

The following cast required no editing: 7.

All other casts required some editing. Notes were entered in the headers on editing applied.

The altimeter header entry was adjusted for casts #12 and 32 because ~1m of bottom data were removed.

For cast #32 DELETE was rerun with the “delete low drop rate” turned off so that the 20m of data at the bottom would not be lost; the descent rate was very low but steady there.

##### Initial Recalibration

Based on the results of 2008-04 it was decided to subtract 0.2db from the pressure. The same offset was used for 2009-03. If further cruises show this to have been a poor decision corrections can be applied later. (This is a net offset of -0.2db since 0db was used in conversion – the last factory offset was -1.1db so the sensor has drifted by about +0.9db.)

Based on the sensor history no calibration will be applied to the salinity.

File 2009-59-recal1.ccf was prepared to subtract 0.2db from the pressure and to apply the following correction to the Oxygen:Dissolved:SBE channel:

DOX (Corrected) = 1.0689 * DOX (CTD) – 0.0113

The SAM and MRGCLN2 files were recalibrated first. COMPARE was run using the recalibrated dissolved oxygen data and showed the recalibration was applied properly and it was. (See 2009-59-dox-comp2.xls.)

The EDT files were then recalibrated.

##### Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps may also partly correct for response time errors, but a further correction is sometimes applied by comparing downcast CTD data to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. As found in the original comparison there is a lot of noise, but when outliers are excluded the comparison looks good near the surface. At low DO values the SBE DO looks very slightly high. The correction looks too small (~0.02) to apply given the uncertainties in how outliers are excluded and knowing that at the bottom bottle stops there was more variability than usual. No further correction will be applied. (See 2009-59-dox-comp3.xls.)

##### Special Fluorometer Processing

An examination of the fluorescence channel shows a dark value of ~0.17mg/m3 . The data were checked for off-scale values and none were found.

There were no extracted chlorophyll data from this cruise, so no special files were prepared for Dr. Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. A few casts were examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and the only unstable features noted are very small-scale, so no further editing appeared to be necessary.

Profile were then examined on-screen. As expected from these complex regions, there are some odd features, but they all look reasonable and will not be removed.

##### Other Comparisons

Previous experience with these sensors –

1. Salinity -The conductivity sensors were used for 3 other cruises in 2009 that have been processed. Only 2009-10 had a good comparison with salinity bottles. The primary salinity was found to be high by ~0.001 and the secondary high by 0.004.

2. Dissolved Oxygen – The sensor was used for 2009-53 which has not been processed yet. It was also used for 2009-10 when temporal variations were noted, and 2009-11 when the slope was roughly 1.075 but the fit was odd.

3. Pressure – The sensor is an older one whose calibration is drifting. There was a long-standing error in the pressure configuration used on most of the cruises that used this sensor. While the error was small and mostly fixed by applying offsets, it means they can’t easily be compared with cruises with the correct parameters. For 2008-04 an offset of -0.2db was applied. (The factory offset was -1.1 so there is a net drift of +0.9db).

Historic ranges – Profile plots were made with historic ranges of T and S superimposed where available. There were frequent excursions from the climatology near the surface which is to be expected. For the LC and LG lines there were small excursions in the top 100m with high temperature and low salinity near shore and low temperatures and high salinity well offshore. Deeper excursions were found along the LG line near the shelf break, with temperatures high from 400 to 800db at LG07 and from 700-900db at LG06. Salinity was a little high from 200-500db at LG07. During several other cruises between June and September 2009 it was observed that temperatures were a little high from Brooks Peninsula south, just beyond the shelf break. The only other excursions were at 2 stations close to shore that are probably not represented well in the climatology. There is no evidence of equipment malfunction.

Quality Checks – There were no repeat casts. Plots were made of casts along lines. There is a lot of variability but none that seems unlikely.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, PAR, Status:Pump, Descent_Rate and Flag and Bottle:Position and Bottle_Number. (The latter two channels were converted inadvertently.)

A second SBE DO channel (with umol/kg units) was added.

REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names, to add the location of the winch (Mid-Ship) and to add the following comments:

Transmissivity and fluorescence are nominal and unedited except that

some records were removed in editing temperature and salinity.

The transmissometer has not been recalibrated since 2006 and the values appear

to be too low, but there is insufficient evidence to enable correction.

The dissolved oxygen gradients were often very high near the surface and

variability was high throughout the profile for many of the casts.

The precision of the SBE dissolved oxygen channel is estimated on the basis of

a comparison to bottle data to be, roughly:

•	±2.5 ml/l  from   0-20 db

•	±0.5 ml/l  from  20-50 db

•	±0.3 ml/l  from 50-250 db

•	±0.1 ml/l  below   250 db

There were no salinity samples available, so no recalibration was applied, but

based on the history of the sensor it is believed to be within ±0.002.

For details on the processing see processing report: 2009-59-proc.doc

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied greatly. For the LG2 to LG7 and EFF12 – EFF16 they were between 90 and 110%, but in the inlets and close to shore values were from 35% to 140%. For a few cases bottle values were checked to ensure the CTD data were not out of line. For cast #2 the DO and saturation are very low in the top 2m, but below that they are both higher and in good agreement with the bottle from 3m. Whether the top 2m are correct or not is impossible to confirm, but all variables are notably different in the top 2db and the salinity is very low there. So these values are either correct or there was a total failure of the CTD. The other cases checked have SBE:DO downcast values at ~3m very close to the bottle values from the upcast.

##### Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Status:Pump, Descent_Rate and Flag.

A second SBE DO channel was added with different units and REORDER to get the 2 SBE DO channels together.

HEADER EDIT was run to fix formats and units, to add the winch location and to add a comment about quality flags and analysis methods. A note was added to the analysis method section to mention that “b” flags were used for poor endpoints because values are probably ok, but the comparison to CTD data was to noisy to establish that. And a note was entered about the transmissivity values which are low, but there is insufficient information to correct them.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### 24. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars – including notes from log, rosette sheets:

2. Niskin #7 – spigot pushed in

3. and frequently thereafter - Many computer problems – software.

10. Aborted but rerun with same cast #

12. Bottle #7 did not fire. Bottle 3 failed integrity

15. Cancelled due to strong currents.

28. Rosette cast but only sample was for Cindy Wright.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4054 | 24Dec08 | Factory “ |  |  |
| Conductivity | 3321 | 16Jan09 | “ |  |  |
| Secondary Temp. | 4700 | 24Dec08 | “ |  |  |
| Secondary Cond. | 1766 | 16Jan09 | “ |  |  |
| Transmissometer | 983DR | 27June06 | IOS |  |  |
| SBE 43 DO sensor | 1176 | 10Mar2009 | Factory |  |  |
| PAR | 4615 | 15Dec2000 | IOS |  |  |
| Fluorometer | 2228 |  |  |  |  |
| Pressure Sensor | 77511 | 08/Mar/2000 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |