## File List
Number of Files= |  | 20
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 11 | 22 | 2009-59-0002.mrgcor1 & 2009-59-0002.samcor1
2 | 2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 24 | 39 | 2009-59-0006.mrgcor1 & 2009-59-0006.samcor1
3 | 2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 41 | 46 | 2009-59-0009.mrgcor1 & 2009-59-0009.samcor1
4 | 2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 48 | 57 | 2009-59-0012.mrgcor1 & 2009-59-0012.samcor1
5 | 2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 59 | 70 | 2009-59-0016.mrgcor1 & 2009-59-0016.samcor1
6 | 2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 72 | 78 | 2009-59-0019.mrgcor1 & 2009-59-0019.samcor1
7 | 2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 80 | 88 | 2009-59-0021.mrgcor1 & 2009-59-0021.samcor1
8 | 2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 90 | 105 | 2009-59-0023.mrgcor1 & 2009-59-0023.samcor1
9 | 2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 107 | 112 | 2009-59-0029.mrgcor1 & 2009-59-0029.samcor1
10 | 2009-59-0030.mrgcor1 | 2009-59-0030.samcor1 | 114 | 116 | 2009-59-0030.mrgcor1 & 2009-59-0030.samcor1
11 | 2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 118 | 123 | 2009-59-0031.mrgcor1 & 2009-59-0031.samcor1
12 | 2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 125 | 131 | 2009-59-0032.mrgcor1 & 2009-59-0032.samcor1
13 | 2009-59-0033.mrgcor1 | 2009-59-0033.samcor1 | 133 | 137 | 2009-59-0033.mrgcor1 & 2009-59-0033.samcor1
14 | 2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 139 | 149 | 2009-59-0034.mrgcor1 & 2009-59-0034.samcor1
15 | 2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 151 | 159 | 2009-59-0035.mrgcor1 & 2009-59-0035.samcor1
16 | 2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 161 | 167 | 2009-59-0036.mrgcor1 & 2009-59-0036.samcor1
17 | 2009-59-0037.mrgcor1 | 2009-59-0037.samcor1 | 169 | 173 | 2009-59-0037.mrgcor1 & 2009-59-0037.samcor1
18 | 2009-59-0038.mrgcor1 | 2009-59-0038.samcor1 | 175 | 179 | 2009-59-0038.mrgcor1 & 2009-59-0038.samcor1
19 | 2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 181 | 187 | 2009-59-0039.mrgcor1 & 2009-59-0039.samcor1
20 | 2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 189 | 194 | 2009-59-0040.mrgcor1 & 2009-59-0040.samcor1

## (DOX_CTD - DOX_BOT)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=DOX_CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00015429 |  | X | CALC Y
m= | 0.00072124 | 0.0048383 |  | Std Dev of Resid |  | 0.089299 |  | 0.058157 | -0.0014698
b= | -0.0015117 | 0.018177 |  | Resid Sum of Squares |  | 1.1483 |  | 7.779 | 0.0040988
# of Values Extracted = |  |  | 165
# of Values Used in Fit = |  |  | 146
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | DOX_CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 255.965 | 3.378 | 0.003888 | -0.10598 | 0.00092464 | 0.10691 |  | 3.378 | -0.10598
1 | YES | 200.325 | 3.4178 | 0.0074498 | -0.085182 | 0.00095334 | 0.086135 |  | 3.4178 | -0.085182
1 | YES | 175.822 | 3.3591 | 0.004829 | 0.042116 | 0.00091101 | -0.041205 |  | 3.3591 | 0.042116
1 | YES | 150.895 | 3.1969 | 0.0043566 | -0.051107 | 0.000794 | 0.051901 |  | 3.1969 | -0.051107
1 | YES | 125.147 | 3.079 | 0.0040082 | -0.0060332 | 0.00070895 | 0.0067421 |  | 3.079 | -0.0060332
1 | YES | 100.395 | 2.9716 | 0.0040331 | -0.0034461 | 0.00063148 | 0.0040776 |  | 2.9716 | -0.0034461
1 | YES | 75.4347 | 2.9627 | 0.0045818 | 0.016661 | 0.00062507 | -0.016036 |  | 2.9627 | 0.016661
1 | YES | 50.4708 | 3.1498 | 0.0093615 | 0.02381 | 0.00076005 | -0.02305 |  | 3.1498 | 0.02381
1 | YES | 30.0569 | 3.5679 | 0.006075 | 0.039859 | 0.0010616 | -0.038798 |  | 3.5679 | 0.039859
1 | YES | 20.4931 | 4.623 | 0.029302 | 0.12302 | 0.0018226 | -0.12119 |  | 4.623 | 0.12302
1 | YES | 10.6532 | 5.6938 | 0.011727 | 0.12477 | 0.0025948 | -0.12217 |  | 5.6938 | 0.12477
1 | YES | 2.9716 | 5.9592 | 0.039647 | -0.02681 | 0.0027863 | 0.029596 |  | 5.9592 | -0.02681
2 | YES | 533.704 | 3.2719 | 0.0045081 | -0.053099 | 0.0008481 | 0.053947 |  | 3.2719 | -0.053099
2 | YES | 500.868 | 3.2582 | 0.0047777 | -0.0028017 | 0.00083822 | 0.0036399 |  | 3.2582 | -0.0028017
2 | YES | 400.379 | 3.254 | 0.0088426 | 0.015975 | 0.00083518 | -0.01514 |  | 3.254 | 0.015975
2 | YES | 300.955 | 3.2373 | 0.0056921 | 0.0012727 | 0.00082312 | -0.00044956 |  | 3.2373 | 0.0012727
2 | YES | 250.688 | 3.2223 | 0.0018718 | 0.016322 | 0.00081234 | -0.01551 |  | 3.2223 | 0.016322
2 | YES | 200.797 | 3.2094 | 0.0060047 | 0.017389 | 0.00080302 | -0.016586 |  | 3.2094 | 0.017389
2 | YES | 176.138 | 3.1964 | 0.0073403 | 0.012355 | 0.00079362 | -0.011562 |  | 3.1964 | 0.012355
2 | YES | 150.165 | 3.1802 | 0.0048576 | 0.0042312 | 0.00078199 | -0.0034492 |  | 3.1802 | 0.0042312
2 | YES | 125.23 | 3.1651 | 0.0061023 | -0.003942 | 0.00077104 | 0.0047131 |  | 3.1651 | -0.003942
2 | YES | 100.474 | 3.0821 | 0.0072873 | 0.022058 | 0.00071118 | -0.021347 |  | 3.0821 | 0.022058
2 | YES | 76.2912 | 2.9707 | 0.0081251 | 0.013727 | 0.00063089 | -0.013096 |  | 2.9707 | 0.013727
2 | YES | 50.5887 | 3.3496 | 0.0078726 | 0.032587 | 0.00090413 | -0.031683 |  | 3.3496 | 0.032587
2 | YES | 30.8936 | 3.5286 | 0.0035083 | -0.011397 | 0.0010332 | 0.01243 |  | 3.5286 | -0.011397
2 | YES | 20.6527 | 3.7514 | 0.0077098 | 0.057397 | 0.0011939 | -0.056203 |  | 3.7514 | 0.057397
2 | YES | 10.8736 | 4.6585 | 0.010207 | 0.17545 | 0.0018481 | -0.1736 |  | 4.6585 | 0.17545
2 | NO | 2.9215 | 6.8374 | 0.021916 | 0.37336 | 0.0034196 | -0.36994 |  | 4.0301 | -0.089925
3 | YES | 95.7032 | 4.0301 | 0.0034476 | -0.089925 | 0.0013949 | 0.09132 |  | 4.01 | -0.0049586
3 | YES | 75.4049 | 4.01 | 0.0068937 | -0.0049586 | 0.0013805 | 0.0063391 |  | 3.9606 | -0.0044215
3 | YES | 50.735 | 3.9606 | 0.0077047 | -0.0044215 | 0.0013448 | 0.0057663 |  | 3.9541 | 0.01114
3 | YES | 30.3056 | 3.9541 | 0.0050105 | 0.01114 | 0.0013402 | -0.0098001 |  | 3.9691 | 0.021058
3 | YES | 20.2846 | 3.9691 | 0.007309 | 0.021058 | 0.0013509 | -0.019707 |  | 3.9833 | 0.018281
3 | YES | 10.402 | 3.9833 | 0.0076488 | 0.018281 | 0.0013612 | -0.01692 |  | 4.1348 | -0.0071573
4 | YES | 207.526 | 4.1348 | 0.010839 | -0.0071573 | 0.0014705 | 0.0086278 |  | 4.1348 | 0.001843
4 | YES | 176.417 | 4.1348 | 0.0075266 | 0.001843 | 0.0014705 | -0.00037249 |  | 4.139 | -0.0050249
4 | YES | 151.188 | 4.139 | 0.0068538 | -0.0050249 | 0.0014735 | 0.0064984 |  | 4.136 | -0.014033
4 | YES | 125.741 | 4.136 | 0.0032273 | -0.014033 | 0.0014713 | 0.015505 |  | 4.1236 | -0.024388
4 | YES | 101.456 | 4.1236 | 0.0072714 | -0.024388 | 0.0014624 | 0.025851 |  | 4.1157 | -0.004281
4 | YES | 75.74 | 4.1157 | 0.010306 | -0.004281 | 0.0014567 | 0.0057377 |  | 4.0316 | -0.017364
4 | YES | 30.5111 | 4.0316 | 0.005503 | -0.017364 | 0.0013961 | 0.01876 |  | 4.0346 | 0.0076199
4 | YES | 21.0084 | 4.0346 | 0.011241 | 0.0076199 | 0.0013982 | -0.0062217 |  | 4.1242 | 0.016166
4 | YES | 10.6474 | 4.1242 | 0.0082405 | 0.016166 | 0.0014628 | -0.014703 |  | 4.0854 | -0.036628
4 | YES | 2.7629 | 4.0854 | 0.0068278 | -0.036628 | 0.0014348 | 0.038063 |  | 4.084 | -0.033009
5 | YES | 254.963 | 4.084 | 0.0092317 | -0.033009 | 0.0014338 | 0.034442 |  | 4.0999 | -0.02205
5 | YES | 205.14 | 4.0999 | 0.0061818 | -0.02205 | 0.0014453 | 0.023495 |  | 4.1125 | -0.058471
5 | YES | 178.131 | 4.1125 | 0.011698 | -0.058471 | 0.0014544 | 0.059926 |  | 4.111 | -0.058975
5 | YES | 155.898 | 4.111 | 0.0049539 | -0.058975 | 0.0014533 | 0.060429 |  | 4.1115 | -0.047471
5 | YES | 128.731 | 4.1115 | 0.01044 | -0.047471 | 0.0014537 | 0.048925 |  | 4.111 | -0.024008
5 | YES | 101.971 | 4.111 | 0.0073524 | -0.024008 | 0.0014533 | 0.025462 |  | 4.131 | -0.014041
5 | YES | 75.4901 | 4.131 | 0.0061297 | -0.014041 | 0.0014677 | 0.015509 |  | 4.116 | -0.028
5 | YES | 51.4701 | 4.116 | 0.0092286 | -0.028 | 0.0014569 | 0.029457 |  | 4.12 | -0.067049
5 | YES | 30.5239 | 4.12 | 0.0099288 | -0.067049 | 0.0014597 | 0.068509 |  | 4.1287 | -0.020306
5 | YES | 20.3946 | 4.1287 | 0.007001 | -0.020306 | 0.0014661 | 0.021772 |  | 4.1337 | 0.0067024
5 | YES | 10.3998 | 4.1337 | 0.0041264 | 0.0067024 | 0.0014697 | -0.0052328 |  | 4.1441 | -0.019859
5 | YES | 2.8575 | 4.1441 | 0.011314 | -0.019859 | 0.0014772 | 0.021336 |  | 1.462 | -0.16003
6 | YES | 95.4426 | 1.462 | 0.002523 | -0.16003 | -0.00045729 | 0.15958 |  | 1.8638 | -0.13621
6 | YES | 75.742 | 1.8638 | 0.010315 | -0.13621 | -0.00016748 | 0.13604 |  | 3.1588 | -0.22321
6 | YES | 50.1009 | 3.1588 | 0.0091896 | -0.22321 | 0.00076652 | 0.22397 |  | 4.4047 | -0.12329
6 | YES | 31.0488 | 4.4047 | 0.0049991 | -0.12329 | 0.0016651 | 0.12495 |  | 5.6163 | -0.036661
6 | YES | 20.1762 | 5.6163 | 0.009266 | -0.036661 | 0.002539 | 0.0392 |  | 6.0945 | -0.049471
6 | YES | 10.3967 | 6.0945 | 0.011578 | -0.049471 | 0.0028839 | 0.052355 |  | 6.2992 | -0.059793
6 | YES | 3.1758 | 6.2992 | 0.009842 | -0.059793 | 0.0030315 | 0.062825 |  | 1.4575 | 0.044488
7 | YES | 140.439 | 1.4575 | 0.0027053 | 0.044488 | -0.00046052 | -0.044948 |  | 1.4844 | -0.026562
7 | YES | 126.026 | 1.4844 | 0.0030629 | -0.026562 | -0.00044108 | 0.026121 |  | 2.7444 | -0.062595
7 | YES | 100.851 | 2.7444 | 0.012684 | -0.062595 | 0.00046766 | 0.063063 |  | 3.7151 | -0.019868
7 | YES | 75.6848 | 3.7151 | 0.0032144 | -0.019868 | 0.0011678 | 0.021035 |  | 3.2712 | -0.01081
7 | YES | 50.6874 | 3.2712 | 0.0039398 | -0.01081 | 0.00084759 | 0.011657 |  | 4.128 | 0.10203
7 | YES | 29.8926 | 4.128 | 0.03187 | 0.10203 | 0.0014656 | -0.10057 |  | 5.3017 | 0.078735
7 | YES | 20.2401 | 5.3017 | 0.018275 | 0.078735 | 0.0023121 | -0.076423 |  | 5.6083 | -0.12668
7 | YES | 10.9548 | 5.6083 | 0.01656 | -0.12668 | 0.0025332 | 0.12921 |  | 5.8747 | -0.10126
7 | YES | 3.426 | 5.8747 | 0.011276 | -0.10126 | 0.0027254 | 0.10398 |  | 0.22783 | 0.018827
8 | YES | 961.794 | 0.22783 | 0.0026455 | 0.018827 | -0.0013474 | -0.020174 |  | 0.24736 | 0.013364
8 | YES | 751.086 | 0.24736 | 0.0019365 | 0.013364 | -0.0013333 | -0.014697 |  | 0.67177 | -0.026231
8 | YES | 502.25 | 0.67177 | 0.0020484 | -0.026231 | -0.0010272 | 0.025204 |  | 0.95184 | -0.010157
8 | YES | 400.356 | 0.95184 | 0.0033541 | -0.010157 | -0.00082521 | 0.0093318 |  | 1.4734 | -0.023545
8 | YES | 298.823 | 1.4734 | 0.0029326 | -0.023545 | -0.00044901 | 0.023096 |  | 1.8324 | -0.01562
8 | YES | 250.097 | 1.8324 | 0.0036612 | -0.01562 | -0.00019014 | 0.01543 |  | 2.9126 | -0.066389
8 | YES | 175.039 | 2.9126 | 0.021708 | -0.066389 | 0.00058897 | 0.066978 |  | 3.5235 | -0.055537
8 | YES | 149.859 | 3.5235 | 0.0034906 | -0.055537 | 0.0010295 | 0.056567 |  | 3.8073 | 0.018331
8 | YES | 124.937 | 3.8073 | 0.013329 | 0.018331 | 0.0012343 | -0.017096 |  | 3.3094 | -0.0076447
8 | YES | 100.298 | 3.3094 | 0.011562 | -0.0076447 | 0.00087512 | 0.0085198 |  | 3.9631 | 0.051116
8 | YES | 74.7543 | 3.9631 | 0.010648 | 0.051116 | 0.0013466 | -0.049769 |  | 4.9783 | -0.089669
8 | NO | 50.6341 | 4.2094 | 0.0061852 | 0.77939 | 0.0015243 | -0.77786 |  | 6.0468 | -0.06024
8 | NO | 30.739 | 4.2215 | 0.02516 | 0.46352 | 0.001533 | -0.46199 |  | 6.1441 | -0.043934
8 | YES | 20.9145 | 4.9783 | 0.02119 | -0.089669 | 0.0020788 | 0.091748 |  | 5.2519 | -0.16314
8 | YES | 10.4181 | 6.0468 | 0.0073599 | -0.06024 | 0.0028494 | 0.063089 |  | 5.4546 | 0.081578
8 | YES | 4.4139 | 6.1441 | 0.034137 | -0.043934 | 0.0029196 | 0.046854 |  | 5.7552 | 0.096207
9 | NO | 91.1854 | 1.1333 | 0.004596 | -0.36567 | -0.00069432 | 0.36497 |  | 4.5961 | 0.1821
9 | NO | 75.5848 | 1.5443 | 0.0033744 | -1.1137 | -0.0003979 | 1.1133 |  | 4.2814 | -0.033595
9 | NO | 50.9383 | 2.6878 | 0.0029701 | -1.2302 | 0.00042683 | 1.2306 |  | 4.4348 | -0.2562
9 | NO | 30.483 | 3.9731 | 0.058771 | -0.44093 | 0.0013538 | 0.44229 |  | 0.062297 | 0.062297
9 | NO | 20.7521 | 4.4183 | 0.0070355 | -0.43669 | 0.0016749 | 0.43837 |  | 0.058157 | 0.008157
9 | NO | 10.8585 | 4.7674 | 0.017284 | -1.7626 | 0.0019267 | 1.7645 |  | 0.90846 | -0.13054
10 | YES | 28.9931 | 5.2519 | 0.034625 | -0.16314 | 0.0022761 | 0.16542 |  | 3.8201 | -0.089901
10 | YES | 20.2194 | 5.4546 | 0.030849 | 0.081578 | 0.0024223 | -0.079156 |  | 4.116 | -0.10104
10 | YES | 10.1846 | 5.7552 | 0.019724 | 0.096207 | 0.0026392 | -0.093568 |  | 4.1589 | 0.0019503
11 | NO | 66.0995 | 0.13668 | 0.0025272 | -2.7783 | -0.0014131 | 2.7769 |  | 2.4787 | -0.26434
11 | NO | 50.6574 | 0.85786 | 0.0077967 | 0.63386 | -0.000893 | -0.63475 |  | 3.7075 | 0.22945
11 | YES | 30.3693 | 4.5961 | 0.068019 | 0.1821 | 0.0018032 | -0.1803 |  | 3.9769 | -0.0040495
11 | YES | 20.2553 | 4.2814 | 0.018565 | -0.033595 | 0.0015762 | 0.035171 |  | 4.0781 | 0.043091
11 | YES | 10.4717 | 4.4348 | 0.089573 | -0.2562 | 0.0016868 | 0.25788 |  | 0.75659 | -0.024413
11 | NO | 3.3607 | 7.5419 | 0.044105 | 0.49486 | 0.0039278 | -0.49093 |  | 1.0535 | 0.016479
12 | YES | 117.394 | 0.062297 | 0.0021666 | 0.062297 | -0.0014668 | -0.063764 |  | 0.96864 | -0.056355
12 | YES | 75.5174 | 0.058157 | 0.00082672 | 0.008157 | -0.0014698 | -0.0096268 |  | 0.87331 | -0.014694
12 | YES | 50.0859 | 0.90846 | 0.027715 | -0.13054 | -0.00085651 | 0.12969 |  | 0.80339 | 0.019388
12 | YES | 30.2482 | 3.8201 | 0.013842 | -0.089901 | 0.0012435 | 0.091144 |  | 1.0309 | 0.038893
12 | YES | 20.1725 | 4.116 | 0.016943 | -0.10104 | 0.0014569 | 0.1025 |  | 1.9747 | 0.0047025
12 | YES | 10.5469 | 4.1589 | 0.0039091 | 0.0019503 | 0.0014879 | -0.00046238 |  | 3.6182 | 0.019157
12 | NO | 2.762 | 6.6568 | 0.11503 | 0.68778 | 0.0032894 | -0.68449 |  | 3.9773 | -0.059661
13 | YES | 50.096 | 2.4787 | 0.022388 | -0.26434 | 0.00027599 | 0.26461 |  | 4.02 | -0.0020246
13 | YES | 30.4983 | 3.7075 | 0.031619 | 0.22945 | 0.0011622 | -0.22829 |  | 0.80195 | 0.01695
13 | YES | 20.5725 | 3.9769 | 0.026364 | -0.0040495 | 0.0013566 | 0.0054062 |  | 0.77711 | 0.024107
13 | YES | 10.4191 | 4.0781 | 0.0054894 | 0.043091 | 0.0014296 | -0.041662 |  | 0.84628 | 0.054281
13 | NO | 3.1766 | 7.3193 | 0.049272 | -0.48175 | 0.0037672 | 0.48552 |  | 1.0436 | 0.051653
14 | YES | 198.302 | 0.75659 | 0.011808 | -0.024413 | -0.00096604 | 0.023447 |  | 2.3662 | 0.11222
14 | YES | 174.92 | 1.0535 | 0.0023773 | 0.016479 | -0.00075191 | -0.017231 |  | 3.5273 | 0.10933
14 | YES | 150.377 | 0.96864 | 0.0020811 | -0.056355 | -0.00081309 | 0.055542 |  | 3.9708 | 0.03676
14 | YES | 125.358 | 0.87331 | 0.0025879 | -0.014694 | -0.00088186 | 0.013812 |  | 4.2248 | 0.19877
14 | YES | 100.146 | 0.80339 | 0.002055 | 0.019388 | -0.00093228 | -0.020321 |  | 6.9339 | 0.10494
14 | YES | 75.4503 | 1.0309 | 0.0019992 | 0.038893 | -0.0007682 | -0.039661 |  | 0.82157 | -0.05843
14 | YES | 50.4665 | 1.9747 | 0.024709 | 0.0047025 | -8.7489e-05 | -0.0047899 |  | 1.0072 | 0.063198
14 | YES | 30.7058 | 3.6182 | 0.0075421 | 0.019157 | 0.0010978 | -0.018059 |  | 2.4668 | 0.14477
14 | YES | 20.944 | 3.9773 | 0.0062896 | -0.059661 | 0.0013569 | 0.061018 |  | 3.7887 | 0.29165
14 | YES | 10.395 | 4.02 | 0.008954 | -0.0020246 | 0.0013876 | 0.0034123 |  | 3.9818 | 0.0078266
14 | NO | 2.7024 | 6.7825 | 0.19227 | 1.3705 | 0.0033801 | -1.3671 |  | 4.0376 | -0.026356
15 | YES | 151.069 | 0.80195 | 0.0017217 | 0.01695 | -0.00093332 | -0.017884 |  | 0.86898 | -0.013017
15 | YES | 126.032 | 0.77711 | 0.0023903 | 0.024107 | -0.00095124 | -0.025059 |  | 2.4354 | 0.21838
15 | YES | 102.007 | 0.84628 | 0.0023102 | 0.054281 | -0.00090135 | -0.055182 |  | 3.9437 | -0.016322
15 | YES | 77.4279 | 1.0436 | 0.0026762 | 0.051653 | -0.000759 | -0.052412 |  | 4.2025 | 0.04948
15 | YES | 51.8949 | 2.3662 | 0.002663 | 0.11222 | 0.00019489 | -0.11203 |  | 3.8888 | 0.11482
15 | YES | 32.0492 | 3.5273 | 0.024443 | 0.10933 | 0.0010323 | -0.1083 |  | 3.9936 | 0.027645
15 | YES | 22.3515 | 3.9708 | 0.0031676 | 0.03676 | 0.0013521 | -0.035408 |  | 4.1482 | 0.030215
15 | YES | 9.8759 | 4.2248 | 0.0071107 | 0.19877 | 0.0015353 | -0.19723 |  | 7.779 | 0.088016
15 | YES | 2.4622 | 6.9339 | 0.087571 | 0.10494 | 0.0034893 | -0.10145 |  | 0.37323 | -0.048769
16 | YES | 117.41 | 0.82157 | 0.003575 | -0.05843 | -0.00091917 | 0.057511 |  | 1.3937 | 0.024669
16 | YES | 76.7026 | 1.0072 | 0.0017777 | 0.063198 | -0.00078529 | -0.063984 |  | 3.1478 | 0.13583
16 | YES | 51.3712 | 2.4668 | 0.021281 | 0.14477 | 0.00026741 | -0.1445 |  | 3.8647 | 0.026719
16 | YES | 31.1189 | 3.7887 | 0.0057223 | 0.29165 | 0.0012208 | -0.29043 |  | 3.9952 | 0.020182
16 | YES | 21.1007 | 3.9818 | 0.0080526 | 0.0078266 | 0.0013601 | -0.0064664 |  | 4.2022 | 0.066215
16 | YES | 11.0205 | 4.0376 | 0.014521 | -0.026356 | 0.0014004 | 0.027756 |  | 6.6938 | -0.31423
16 | NO | 2.6684 | 7.8774 | 0.023852 | 1.6284 | 0.0041697 | -1.6242 |  | 1.1648 | -0.18821
17 | YES | 82.9455 | 0.86898 | 0.0025722 | -0.013017 | -0.00088497 | 0.012132 |  | 3.1995 | 0.14245
17 | YES | 50.5728 | 2.4354 | 0.010424 | 0.21838 | 0.00024477 | -0.21814 |  | 3.8623 | 0.085339
17 | YES | 20.6341 | 3.9437 | 0.0087533 | -0.016322 | 0.0013326 | 0.017655 |  | 4.0688 | 0.07181
17 | YES | 10.7493 | 4.2025 | 0.01858 | 0.04948 | 0.0015193 | -0.04796 |  | 4.4483 | 0.22226
17 | NO | 2.9496 | 7.247 | 0.025047 | 0.76096 | 0.0037151 | -0.75724
18 | NO | 57.0129 | 1.3522 | 0.0057507 | -0.56078 | -0.00053645 | 0.56025
18 | YES | 30.6883 | 3.8888 | 0.0051088 | 0.11482 | 0.001293 | -0.11352
18 | YES | 20.7091 | 3.9936 | 0.0050742 | 0.027645 | 0.0013686 | -0.026276
18 | YES | 10.5712 | 4.1482 | 0.0095936 | 0.030215 | 0.0014801 | -0.028735
18 | YES | 3.0257 | 7.779 | 0.042462 | 0.088016 | 0.0040988 | -0.083918
19 | YES | 95.8626 | 0.37323 | 0.0038789 | -0.048769 | -0.0012425 | 0.047526
19 | YES | 76.039 | 1.3937 | 0.011351 | 0.024669 | -0.00050655 | -0.025176
19 | YES | 50.8867 | 3.1478 | 0.043395 | 0.13583 | 0.00075862 | -0.13507
19 | YES | 30.7941 | 3.8647 | 0.014472 | 0.026719 | 0.0012757 | -0.025443
19 | YES | 20.716 | 3.9952 | 0.0064794 | 0.020182 | 0.0013698 | -0.018812
19 | YES | 10.6236 | 4.2022 | 0.017533 | 0.066215 | 0.0015191 | -0.064696
19 | YES | 3.0009 | 6.6938 | 0.09238 | -0.31423 | 0.0033161 | 0.31755
20 | YES | 80.1208 | 1.1648 | 0.0046156 | -0.18821 | -0.00067163 | 0.18754
20 | YES | 50.3501 | 3.1995 | 0.0062061 | 0.14245 | 0.00079585 | -0.14165
20 | YES | 30.3512 | 3.8623 | 0.0044376 | 0.085339 | 0.0012739 | -0.084065
20 | YES | 20.254 | 4.0688 | 0.0031578 | 0.07181 | 0.0014229 | -0.070387
20 | YES | 10.0848 | 4.4483 | 0.0049679 | 0.22226 | 0.0016965 | -0.22057
20 | NO | 3.366 | 8.7559 | 0.044073 | 3.0049 | 0.0048033 | -3.0001
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX_CTD - DOX_BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX_CTD - DOX_BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX_CTD - DOX_BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.0062399 |  | X | CALC Y
m= | -5.2256e-05 | 5.4954e-05 |  | Std Dev of Resid |  | 0.089027 |  | 2.4622 | 0.0059406
b= | 0.0060693 | 0.0091167 |  | Resid Sum of Squares |  | 1.1413 |  | 961.794 | -0.04419
# of Values Extracted = |  |  | 165
# of Values Used in Fit = |  |  | 146
File | Use | Pressure | DOX_CTD | DOX_CTD | Diff (DOX_CTD - DOX_BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 255.965 | 3.378 | 0.003888 | -0.10598 | -0.0073063 | 0.098677 |  | 255.965 | -0.10598
1 | YES | 200.325 | 3.4178 | 0.0074498 | -0.085182 | -0.0043988 | 0.080783 |  | 200.325 | -0.085182
1 | YES | 175.822 | 3.3591 | 0.004829 | 0.042116 | -0.0031184 | -0.045234 |  | 175.822 | 0.042116
1 | YES | 150.895 | 3.1969 | 0.0043566 | -0.051107 | -0.0018158 | 0.049292 |  | 150.895 | -0.051107
1 | YES | 125.147 | 3.079 | 0.0040082 | -0.0060332 | -0.00047034 | 0.0055628 |  | 125.147 | -0.0060332
1 | YES | 100.395 | 2.9716 | 0.0040331 | -0.0034461 | 0.00082309 | 0.0042692 |  | 100.395 | -0.0034461
1 | YES | 75.4347 | 2.9627 | 0.0045818 | 0.016661 | 0.0021274 | -0.014534 |  | 75.4347 | 0.016661
1 | YES | 50.4708 | 3.1498 | 0.0093615 | 0.02381 | 0.0034319 | -0.020378 |  | 50.4708 | 0.02381
1 | YES | 30.0569 | 3.5679 | 0.006075 | 0.039859 | 0.0044986 | -0.035361 |  | 30.0569 | 0.039859
1 | YES | 20.4931 | 4.623 | 0.029302 | 0.12302 | 0.0049984 | -0.11802 |  | 20.4931 | 0.12302
1 | YES | 10.6532 | 5.6938 | 0.011727 | 0.12477 | 0.0055126 | -0.11926 |  | 10.6532 | 0.12477
1 | YES | 2.9716 | 5.9592 | 0.039647 | -0.02681 | 0.005914 | 0.032724 |  | 2.9716 | -0.02681
2 | YES | 533.704 | 3.2719 | 0.0045081 | -0.053099 | -0.02182 | 0.03128 |  | 533.704 | -0.053099
2 | YES | 500.868 | 3.2582 | 0.0047777 | -0.0028017 | -0.020104 | -0.017302 |  | 500.868 | -0.0028017
2 | YES | 400.379 | 3.254 | 0.0088426 | 0.015975 | -0.014853 | -0.030828 |  | 400.379 | 0.015975
2 | YES | 300.955 | 3.2373 | 0.0056921 | 0.0012727 | -0.0096573 | -0.01093 |  | 300.955 | 0.0012727
2 | YES | 250.688 | 3.2223 | 0.0018718 | 0.016322 | -0.0070305 | -0.023353 |  | 250.688 | 0.016322
2 | YES | 200.797 | 3.2094 | 0.0060047 | 0.017389 | -0.0044235 | -0.021812 |  | 200.797 | 0.017389
2 | YES | 176.138 | 3.1964 | 0.0073403 | 0.012355 | -0.0031349 | -0.01549 |  | 176.138 | 0.012355
2 | YES | 150.165 | 3.1802 | 0.0048576 | 0.0042312 | -0.0017777 | -0.0060089 |  | 150.165 | 0.0042312
2 | YES | 125.23 | 3.1651 | 0.0061023 | -0.003942 | -0.00047468 | 0.0034673 |  | 125.23 | -0.003942
2 | YES | 100.474 | 3.0821 | 0.0072873 | 0.022058 | 0.00081896 | -0.021239 |  | 100.474 | 0.022058
2 | YES | 76.2912 | 2.9707 | 0.0081251 | 0.013727 | 0.0020826 | -0.011645 |  | 76.2912 | 0.013727
2 | YES | 50.5887 | 3.3496 | 0.0078726 | 0.032587 | 0.0034257 | -0.029161 |  | 50.5887 | 0.032587
2 | YES | 30.8936 | 3.5286 | 0.0035083 | -0.011397 | 0.0044549 | 0.015852 |  | 30.8936 | -0.011397
2 | YES | 20.6527 | 3.7514 | 0.0077098 | 0.057397 | 0.0049901 | -0.052407 |  | 20.6527 | 0.057397
2 | YES | 10.8736 | 4.6585 | 0.010207 | 0.17545 | 0.0055011 | -0.16995 |  | 10.8736 | 0.17545
2 | NO | 2.9215 | 6.8374 | 0.021916 | 0.37336 | 0.0059166 | -0.36745 |  | 95.7032 | -0.089925
3 | YES | 95.7032 | 4.0301 | 0.0034476 | -0.089925 | 0.0010683 | 0.090994 |  | 75.4049 | -0.0049586
3 | YES | 75.4049 | 4.01 | 0.0068937 | -0.0049586 | 0.002129 | 0.0070876 |  | 50.735 | -0.0044215
3 | YES | 50.735 | 3.9606 | 0.0077047 | -0.0044215 | 0.0034181 | 0.0078396 |  | 30.3056 | 0.01114
3 | YES | 30.3056 | 3.9541 | 0.0050105 | 0.01114 | 0.0044856 | -0.0066547 |  | 20.2846 | 0.021058
3 | YES | 20.2846 | 3.9691 | 0.007309 | 0.021058 | 0.0050093 | -0.016049 |  | 10.402 | 0.018281
3 | YES | 10.402 | 3.9833 | 0.0076488 | 0.018281 | 0.0055257 | -0.012755 |  | 207.526 | -0.0071573
4 | YES | 207.526 | 4.1348 | 0.010839 | -0.0071573 | -0.0047751 | 0.0023822 |  | 176.417 | 0.001843
4 | YES | 176.417 | 4.1348 | 0.0075266 | 0.001843 | -0.0031495 | -0.0049925 |  | 151.188 | -0.0050249
4 | YES | 151.188 | 4.139 | 0.0068538 | -0.0050249 | -0.0018311 | 0.0031938 |  | 125.741 | -0.014033
4 | YES | 125.741 | 4.136 | 0.0032273 | -0.014033 | -0.00050138 | 0.013532 |  | 101.456 | -0.024388
4 | YES | 101.456 | 4.1236 | 0.0072714 | -0.024388 | 0.00076765 | 0.025156 |  | 75.74 | -0.004281
4 | YES | 75.74 | 4.1157 | 0.010306 | -0.004281 | 0.0021114 | 0.0063925 |  | 30.5111 | -0.017364
4 | YES | 30.5111 | 4.0316 | 0.005503 | -0.017364 | 0.0044749 | 0.021838 |  | 21.0084 | 0.0076199
4 | YES | 21.0084 | 4.0346 | 0.011241 | 0.0076199 | 0.0049715 | -0.0026484 |  | 10.6474 | 0.016166
4 | YES | 10.6474 | 4.1242 | 0.0082405 | 0.016166 | 0.0055129 | -0.010653 |  | 2.7629 | -0.036628
4 | YES | 2.7629 | 4.0854 | 0.0068278 | -0.036628 | 0.0059249 | 0.042553 |  | 254.963 | -0.033009
5 | YES | 254.963 | 4.084 | 0.0092317 | -0.033009 | -0.0072539 | 0.025755 |  | 205.14 | -0.02205
5 | YES | 205.14 | 4.0999 | 0.0061818 | -0.02205 | -0.0046504 | 0.017399 |  | 178.131 | -0.058471
5 | YES | 178.131 | 4.1125 | 0.011698 | -0.058471 | -0.003239 | 0.055232 |  | 155.898 | -0.058975
5 | YES | 155.898 | 4.111 | 0.0049539 | -0.058975 | -0.0020772 | 0.056898 |  | 128.731 | -0.047471
5 | YES | 128.731 | 4.1115 | 0.01044 | -0.047471 | -0.00065762 | 0.046813 |  | 101.971 | -0.024008
5 | YES | 101.971 | 4.111 | 0.0073524 | -0.024008 | 0.00074073 | 0.024749 |  | 75.4901 | -0.014041
5 | YES | 75.4901 | 4.131 | 0.0061297 | -0.014041 | 0.0021245 | 0.016166 |  | 51.4701 | -0.028
5 | YES | 51.4701 | 4.116 | 0.0092286 | -0.028 | 0.0033797 | 0.03138 |  | 30.5239 | -0.067049
5 | YES | 30.5239 | 4.12 | 0.0099288 | -0.067049 | 0.0044742 | 0.071524 |  | 20.3946 | -0.020306
5 | YES | 20.3946 | 4.1287 | 0.007001 | -0.020306 | 0.0050035 | 0.02531 |  | 10.3998 | 0.0067024
5 | YES | 10.3998 | 4.1337 | 0.0041264 | 0.0067024 | 0.0055258 | -0.0011766 |  | 2.8575 | -0.019859
5 | YES | 2.8575 | 4.1441 | 0.011314 | -0.019859 | 0.00592 | 0.025779 |  | 95.4426 | -0.16003
6 | YES | 95.4426 | 1.462 | 0.002523 | -0.16003 | 0.0010819 | 0.16111 |  | 75.742 | -0.13621
6 | YES | 75.742 | 1.8638 | 0.010315 | -0.13621 | 0.0021113 | 0.13832 |  | 50.1009 | -0.22321
6 | YES | 50.1009 | 3.1588 | 0.0091896 | -0.22321 | 0.0034512 | 0.22666 |  | 31.0488 | -0.12329
6 | YES | 31.0488 | 4.4047 | 0.0049991 | -0.12329 | 0.0044468 | 0.12774 |  | 20.1762 | -0.036661
6 | YES | 20.1762 | 5.6163 | 0.009266 | -0.036661 | 0.005015 | 0.041676 |  | 10.3967 | -0.049471
6 | YES | 10.3967 | 6.0945 | 0.011578 | -0.049471 | 0.005526 | 0.054997 |  | 3.1758 | -0.059793
6 | YES | 3.1758 | 6.2992 | 0.009842 | -0.059793 | 0.0059033 | 0.065697 |  | 140.439 | 0.044488
7 | YES | 140.439 | 1.4575 | 0.0027053 | 0.044488 | -0.0012694 | -0.045757 |  | 126.026 | -0.026562
7 | YES | 126.026 | 1.4844 | 0.0030629 | -0.026562 | -0.00051627 | 0.026046 |  | 100.851 | -0.062595
7 | YES | 100.851 | 2.7444 | 0.012684 | -0.062595 | 0.00079926 | 0.063394 |  | 75.6848 | -0.019868
7 | YES | 75.6848 | 3.7151 | 0.0032144 | -0.019868 | 0.0021143 | 0.021982 |  | 50.6874 | -0.01081
7 | YES | 50.6874 | 3.2712 | 0.0039398 | -0.01081 | 0.0034206 | 0.01423 |  | 29.8926 | 0.10203
7 | YES | 29.8926 | 4.128 | 0.03187 | 0.10203 | 0.0045072 | -0.097526 |  | 20.2401 | 0.078735
7 | YES | 20.2401 | 5.3017 | 0.018275 | 0.078735 | 0.0050116 | -0.073724 |  | 10.9548 | -0.12668
7 | YES | 10.9548 | 5.6083 | 0.01656 | -0.12668 | 0.0054968 | 0.13217 |  | 3.426 | -0.10126
7 | YES | 3.426 | 5.8747 | 0.011276 | -0.10126 | 0.0058903 | 0.10715 |  | 961.794 | 0.018827
8 | YES | 961.794 | 0.22783 | 0.0026455 | 0.018827 | -0.04419 | -0.063017 |  | 751.086 | 0.013364
8 | YES | 751.086 | 0.24736 | 0.0019365 | 0.013364 | -0.033179 | -0.046543 |  | 502.25 | -0.026231
8 | YES | 502.25 | 0.67177 | 0.0020484 | -0.026231 | -0.020176 | 0.0060554 |  | 400.356 | -0.010157
8 | YES | 400.356 | 0.95184 | 0.0033541 | -0.010157 | -0.014852 | -0.0046945 |  | 298.823 | -0.023545
8 | YES | 298.823 | 1.4734 | 0.0029326 | -0.023545 | -0.0095459 | 0.014 |  | 250.097 | -0.01562
8 | YES | 250.097 | 1.8324 | 0.0036612 | -0.01562 | -0.0069997 | 0.0086202 |  | 175.039 | -0.066389
8 | YES | 175.039 | 2.9126 | 0.021708 | -0.066389 | -0.0030775 | 0.063311 |  | 149.859 | -0.055537
8 | YES | 149.859 | 3.5235 | 0.0034906 | -0.055537 | -0.0017617 | 0.053776 |  | 124.937 | 0.018331
8 | YES | 124.937 | 3.8073 | 0.013329 | 0.018331 | -0.00045937 | -0.01879 |  | 100.298 | -0.0076447
8 | YES | 100.298 | 3.3094 | 0.011562 | -0.0076447 | 0.00082816 | 0.0084728 |  | 74.7543 | 0.051116
8 | YES | 74.7543 | 3.9631 | 0.010648 | 0.051116 | 0.002163 | -0.048953 |  | 20.9145 | -0.089669
8 | NO | 50.6341 | 4.2094 | 0.0061852 | 0.77939 | 0.0034234 | -0.77596 |  | 10.4181 | -0.06024
8 | NO | 30.739 | 4.2215 | 0.02516 | 0.46352 | 0.004463 | -0.45906 |  | 4.4139 | -0.043934
8 | YES | 20.9145 | 4.9783 | 0.02119 | -0.089669 | 0.0049764 | 0.094646 |  | 28.9931 | -0.16314
8 | YES | 10.4181 | 6.0468 | 0.0073599 | -0.06024 | 0.0055249 | 0.065765 |  | 20.2194 | 0.081578
8 | YES | 4.4139 | 6.1441 | 0.034137 | -0.043934 | 0.0058386 | 0.049773 |  | 10.1846 | 0.096207
9 | NO | 91.1854 | 1.1333 | 0.004596 | -0.36567 | 0.0013043 | 0.36697 |  | 30.3693 | 0.1821
9 | NO | 75.5848 | 1.5443 | 0.0033744 | -1.1137 | 0.0021196 | 1.1158 |  | 20.2553 | -0.033595
9 | NO | 50.9383 | 2.6878 | 0.0029701 | -1.2302 | 0.0034075 | 1.2336 |  | 10.4717 | -0.2562
9 | NO | 30.483 | 3.9731 | 0.058771 | -0.44093 | 0.0044764 | 0.44541 |  | 117.394 | 0.062297
9 | NO | 20.7521 | 4.4183 | 0.0070355 | -0.43669 | 0.0049849 | 0.44168 |  | 75.5174 | 0.008157
9 | NO | 10.8585 | 4.7674 | 0.017284 | -1.7626 | 0.0055019 | 1.7681 |  | 50.0859 | -0.13054
10 | YES | 28.9931 | 5.2519 | 0.034625 | -0.16314 | 0.0045542 | 0.16769 |  | 30.2482 | -0.089901
10 | YES | 20.2194 | 5.4546 | 0.030849 | 0.081578 | 0.0050127 | -0.076566 |  | 20.1725 | -0.10104
10 | YES | 10.1846 | 5.7552 | 0.019724 | 0.096207 | 0.0055371 | -0.09067 |  | 10.5469 | 0.0019503
11 | NO | 66.0995 | 0.13668 | 0.0025272 | -2.7783 | 0.0026152 | 2.7809 |  | 50.096 | -0.26434
11 | NO | 50.6574 | 0.85786 | 0.0077967 | 0.63386 | 0.0034222 | -0.63044 |  | 30.4983 | 0.22945
11 | YES | 30.3693 | 4.5961 | 0.068019 | 0.1821 | 0.0044823 | -0.17762 |  | 20.5725 | -0.0040495
11 | YES | 20.2553 | 4.2814 | 0.018565 | -0.033595 | 0.0050108 | 0.038606 |  | 10.4191 | 0.043091
11 | YES | 10.4717 | 4.4348 | 0.089573 | -0.2562 | 0.0055221 | 0.26172 |  | 198.302 | -0.024413
11 | NO | 3.3607 | 7.5419 | 0.044105 | 0.49486 | 0.0058937 | -0.48897 |  | 174.92 | 0.016479
12 | YES | 117.394 | 0.062297 | 0.0021666 | 0.062297 | -6.5203e-05 | -0.062363 |  | 150.377 | -0.056355
12 | YES | 75.5174 | 0.058157 | 0.00082672 | 0.008157 | 0.0021231 | -0.0060339 |  | 125.358 | -0.014694
12 | YES | 50.0859 | 0.90846 | 0.027715 | -0.13054 | 0.003452 | 0.134 |  | 100.146 | 0.019388
12 | YES | 30.2482 | 3.8201 | 0.013842 | -0.089901 | 0.0044886 | 0.09439 |  | 75.4503 | 0.038893
12 | YES | 20.1725 | 4.116 | 0.016943 | -0.10104 | 0.0050152 | 0.10606 |  | 50.4665 | 0.0047025
12 | YES | 10.5469 | 4.1589 | 0.0039091 | 0.0019503 | 0.0055181 | 0.0035679 |  | 30.7058 | 0.019157
12 | NO | 2.762 | 6.6568 | 0.11503 | 0.68778 | 0.005925 | -0.68185 |  | 20.944 | -0.059661
13 | YES | 50.096 | 2.4787 | 0.022388 | -0.26434 | 0.0034515 | 0.26779 |  | 10.395 | -0.0020246
13 | YES | 30.4983 | 3.7075 | 0.031619 | 0.22945 | 0.0044756 | -0.22498 |  | 151.069 | 0.01695
13 | YES | 20.5725 | 3.9769 | 0.026364 | -0.0040495 | 0.0049943 | 0.0090438 |  | 126.032 | 0.024107
13 | YES | 10.4191 | 4.0781 | 0.0054894 | 0.043091 | 0.0055248 | -0.037566 |  | 102.007 | 0.054281
13 | NO | 3.1766 | 7.3193 | 0.049272 | -0.48175 | 0.0059033 | 0.48766 |  | 77.4279 | 0.051653
14 | YES | 198.302 | 0.75659 | 0.011808 | -0.024413 | -0.0042931 | 0.02012 |  | 51.8949 | 0.11222
14 | YES | 174.92 | 1.0535 | 0.0023773 | 0.016479 | -0.0030713 | -0.019551 |  | 32.0492 | 0.10933
14 | YES | 150.377 | 0.96864 | 0.0020811 | -0.056355 | -0.0017887 | 0.054567 |  | 22.3515 | 0.03676
14 | YES | 125.358 | 0.87331 | 0.0025879 | -0.014694 | -0.00048137 | 0.014213 |  | 9.8759 | 0.19877
14 | YES | 100.146 | 0.80339 | 0.002055 | 0.019388 | 0.0008361 | -0.018552 |  | 2.4622 | 0.10494
14 | YES | 75.4503 | 1.0309 | 0.0019992 | 0.038893 | 0.0021266 | -0.036766 |  | 117.41 | -0.05843
14 | YES | 50.4665 | 1.9747 | 0.024709 | 0.0047025 | 0.0034321 | -0.0012703 |  | 76.7026 | 0.063198
14 | YES | 30.7058 | 3.6182 | 0.0075421 | 0.019157 | 0.0044647 | -0.014692 |  | 51.3712 | 0.14477
14 | YES | 20.944 | 3.9773 | 0.0062896 | -0.059661 | 0.0049748 | 0.064636 |  | 31.1189 | 0.29165
14 | YES | 10.395 | 4.02 | 0.008954 | -0.0020246 | 0.0055261 | 0.0075507 |  | 21.1007 | 0.0078266
14 | NO | 2.7024 | 6.7825 | 0.19227 | 1.3705 | 0.0059281 | -1.3645 |  | 11.0205 | -0.026356
15 | YES | 151.069 | 0.80195 | 0.0017217 | 0.01695 | -0.0018249 | -0.018775 |  | 82.9455 | -0.013017
15 | YES | 126.032 | 0.77711 | 0.0023903 | 0.024107 | -0.00051659 | -0.024624 |  | 50.5728 | 0.21838
15 | YES | 102.007 | 0.84628 | 0.0023102 | 0.054281 | 0.00073885 | -0.053542 |  | 20.6341 | -0.016322
15 | YES | 77.4279 | 1.0436 | 0.0026762 | 0.051653 | 0.0020232 | -0.04963 |  | 10.7493 | 0.04948
15 | YES | 51.8949 | 2.3662 | 0.002663 | 0.11222 | 0.0033575 | -0.10887 |  | 30.6883 | 0.11482
15 | YES | 32.0492 | 3.5273 | 0.024443 | 0.10933 | 0.0043945 | -0.10494 |  | 20.7091 | 0.027645
15 | YES | 22.3515 | 3.9708 | 0.0031676 | 0.03676 | 0.0049013 | -0.031859 |  | 10.5712 | 0.030215
15 | YES | 9.8759 | 4.2248 | 0.0071107 | 0.19877 | 0.0055532 | -0.19322 |  | 3.0257 | 0.088016
15 | YES | 2.4622 | 6.9339 | 0.087571 | 0.10494 | 0.0059406 | -0.099001 |  | 95.8626 | -0.048769
16 | YES | 117.41 | 0.82157 | 0.003575 | -0.05843 | -6.6039e-05 | 0.058364 |  | 76.039 | 0.024669
16 | YES | 76.7026 | 1.0072 | 0.0017777 | 0.063198 | 0.0020611 | -0.061137 |  | 50.8867 | 0.13583
16 | YES | 51.3712 | 2.4668 | 0.021281 | 0.14477 | 0.0033849 | -0.14138 |  | 30.7941 | 0.026719
16 | YES | 31.1189 | 3.7887 | 0.0057223 | 0.29165 | 0.0044431 | -0.28721 |  | 20.716 | 0.020182
16 | YES | 21.1007 | 3.9818 | 0.0080526 | 0.0078266 | 0.0049667 | -0.0028599 |  | 10.6236 | 0.066215
16 | YES | 11.0205 | 4.0376 | 0.014521 | -0.026356 | 0.0054934 | 0.031849 |  | 3.0009 | -0.31423
16 | NO | 2.6684 | 7.8774 | 0.023852 | 1.6284 | 0.0059298 | -1.6224 |  | 80.1208 | -0.18821
17 | YES | 82.9455 | 0.86898 | 0.0025722 | -0.013017 | 0.0017349 | 0.014752 |  | 50.3501 | 0.14245
17 | YES | 50.5728 | 2.4354 | 0.010424 | 0.21838 | 0.0034266 | -0.21495 |  | 30.3512 | 0.085339
17 | YES | 20.6341 | 3.9437 | 0.0087533 | -0.016322 | 0.004991 | 0.021313 |  | 20.254 | 0.07181
17 | YES | 10.7493 | 4.2025 | 0.01858 | 0.04948 | 0.0055076 | -0.043972 |  | 10.0848 | 0.22226
17 | NO | 2.9496 | 7.247 | 0.025047 | 0.76096 | 0.0059151 | -0.75504
18 | NO | 57.0129 | 1.3522 | 0.0057507 | -0.56078 | 0.00309 | 0.56388
18 | YES | 30.6883 | 3.8888 | 0.0051088 | 0.11482 | 0.0044656 | -0.11035
18 | YES | 20.7091 | 3.9936 | 0.0050742 | 0.027645 | 0.0049871 | -0.022657
18 | YES | 10.5712 | 4.1482 | 0.0095936 | 0.030215 | 0.0055169 | -0.024698
18 | YES | 3.0257 | 7.779 | 0.042462 | 0.088016 | 0.0059112 | -0.082105
19 | YES | 95.8626 | 0.37323 | 0.0038789 | -0.048769 | 0.0010599 | 0.049829
19 | YES | 76.039 | 1.3937 | 0.011351 | 0.024669 | 0.0020958 | -0.022574
19 | YES | 50.8867 | 3.1478 | 0.043395 | 0.13583 | 0.0034102 | -0.13242
19 | YES | 30.7941 | 3.8647 | 0.014472 | 0.026719 | 0.0044601 | -0.022259
19 | YES | 20.716 | 3.9952 | 0.0064794 | 0.020182 | 0.0049868 | -0.015195
19 | YES | 10.6236 | 4.2022 | 0.017533 | 0.066215 | 0.0055141 | -0.0607
19 | YES | 3.0009 | 6.6938 | 0.09238 | -0.31423 | 0.0059125 | 0.32014
20 | YES | 80.1208 | 1.1648 | 0.0046156 | -0.18821 | 0.0018825 | 0.1901
20 | YES | 50.3501 | 3.1995 | 0.0062061 | 0.14245 | 0.0034382 | -0.13901
20 | YES | 30.3512 | 3.8623 | 0.0044376 | 0.085339 | 0.0044833 | -0.080856
20 | YES | 20.254 | 4.0688 | 0.0031578 | 0.07181 | 0.0050109 | -0.066799
20 | YES | 10.0848 | 4.4483 | 0.0049679 | 0.22226 | 0.0055423 | -0.21672
20 | NO | 3.366 | 8.7559 | 0.044073 | 3.0049 | 0.0058934 | -2.999
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. DOX_CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2 |  |  |  | COMPARE Version 3.3 | Run at 2010/08/09 09:00:53
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX_BOT | DOX_CTD |  | DOX_CTD - DOX_BOT
Average Size: | 1.5
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX_BOT | DOX_CTD | DOX_CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 255.96 | 255.965 | 0.029216 | 0.0052033 | 3.484 | 3.378 | 0.003888 | -0.10598
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 200.32 | 200.325 | 0.021724 | 0.0048676 | 3.503 | 3.4178 | 0.0074498 | -0.085182
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 175.82 | 175.822 | 0.019977 | 0.0022278 | 3.317 | 3.3591 | 0.004829 | 0.042116
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 150.9 | 150.895 | 0.028723 | -0.0046234 | 3.248 | 3.1969 | 0.0043566 | -0.051107
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 125.15 | 125.147 | 0.028269 | -0.0033875 | 3.085 | 3.079 | 0.0040082 | -0.0060332
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 100.39 | 100.395 | 0.024736 | 0.004715 | 2.975 | 2.9716 | 0.0040331 | -0.0034461
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 75.43 | 75.4347 | 0.021987 | 0.0047073 | 2.946 | 2.9627 | 0.0045818 | 0.016661
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 50.47 | 50.4708 | 0.030865 | 0.00082398 | 3.126 | 3.1498 | 0.0093615 | 0.02381
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 30.06 | 30.0569 | 0.029468 | -0.0031395 | 3.528 | 3.5679 | 0.006075 | 0.039859
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 20.49 | 20.4931 | 0.022092 | 0.0030575 | 4.5 | 4.623 | 0.029302 | 0.12302
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 10.65 | 10.6532 | 0.033272 | 0.0032234 | 5.569 | 5.6938 | 0.011727 | 0.12477
2009-59-0002.mrgcor1 | 2009-59-0002.samcor1 | 121 | 2.97 | 2.9716 | 0.026405 | 0.001653 | 5.986 | 5.9592 | 0.039647 | -0.02681
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 533.71 | 533.704 | 0.023169 | -0.0061646 | 3.325 | 3.2719 | 0.0045081 | -0.053099
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 500.87 | 500.868 | 0.021576 | -0.0021362 | 3.261 | 3.2582 | 0.0047777 | -0.0028017
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 400.38 | 400.379 | 0.024572 | -0.0011597 | 3.238 | 3.254 | 0.0088426 | 0.015975
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 300.96 | 300.955 | 0.019324 | -0.0048523 | 3.236 | 3.2373 | 0.0056921 | 0.0012727
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 250.69 | 250.688 | 0.023549 | -0.0017395 | 3.206 | 3.2223 | 0.0018718 | 0.016322
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 200.8 | 200.797 | 0.03079 | -0.0026398 | 3.192 | 3.2094 | 0.0060047 | 0.017389
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 176.14 | 176.138 | 0.025819 | -0.0020599 | 3.184 | 3.1964 | 0.0073403 | 0.012355
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 150.17 | 150.165 | 0.025952 | -0.0049591 | 3.176 | 3.1802 | 0.0048576 | 0.0042312
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 125.23 | 125.23 | 0.021462 | 0.00041199 | 3.169 | 3.1651 | 0.0061023 | -0.003942
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 100.48 | 100.474 | 0.024233 | -0.0061188 | 3.06 | 3.0821 | 0.0072873 | 0.022058
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 76.29 | 76.2912 | 0.02494 | 0.001152 | 2.957 | 2.9707 | 0.0081251 | 0.013727
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 50.59 | 50.5887 | 0.027293 | -0.0013237 | 3.317 | 3.3496 | 0.0078726 | 0.032587
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 30.89 | 30.8936 | 0.025658 | 0.0036373 | 3.54 | 3.5286 | 0.0035083 | -0.011397
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 20.65 | 20.6527 | 0.021291 | 0.0027275 | 3.694 | 3.7514 | 0.0077098 | 0.057397
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 10.87 | 10.8736 | 0.021403 | 0.0035543 | 4.483 | 4.6585 | 0.010207 | 0.17545
2009-59-0006.mrgcor1 | 2009-59-0006.samcor1 | 121 | 2.92 | 2.9215 | 0.033977 | 0.0014875 | 6.464 | 6.8374 | 0.021916 | 0.37336
2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 121 | 95.71 | 95.7032 | 0.026431 | -0.0067749 | 4.12 | 4.0301 | 0.0034476 | -0.089925
2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 121 | 75.41 | 75.4049 | 0.025235 | -0.0051269 | 4.015 | 4.01 | 0.0068937 | -0.0049586
2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 121 | 50.74 | 50.735 | 0.021375 | -0.0049591 | 3.965 | 3.9606 | 0.0077047 | -0.0044215
2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 121 | 30.3 | 30.3056 | 0.024217 | 0.005621 | 3.943 | 3.9541 | 0.0050105 | 0.01114
2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 121 | 20.28 | 20.2846 | 0.01775 | 0.0046272 | 3.948 | 3.9691 | 0.007309 | 0.021058
2009-59-0009.mrgcor1 | 2009-59-0009.samcor1 | 121 | 10.4 | 10.402 | 0.023686 | 0.0019846 | 3.965 | 3.9833 | 0.0076488 | 0.018281
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 207.53 | 207.526 | 0.028489 | -0.0044708 | 4.142 | 4.1348 | 0.010839 | -0.0071573
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 176.42 | 176.417 | 0.030701 | -0.0029755 | 4.133 | 4.1348 | 0.0075266 | 0.001843
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 151.19 | 151.188 | 0.020053 | -0.0024872 | 4.144 | 4.139 | 0.0068538 | -0.0050249
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 125.74 | 125.741 | 0.027148 | 0.00074768 | 4.15 | 4.136 | 0.0032273 | -0.014033
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 101.46 | 101.456 | 0.029435 | -0.0042953 | 4.148 | 4.1236 | 0.0072714 | -0.024388
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 75.74 | 75.74 | 0.023628 | 0 | 4.12 | 4.1157 | 0.010306 | -0.004281
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 30.51 | 30.5111 | 0.033387 | 0.0010738 | 4.049 | 4.0316 | 0.005503 | -0.017364
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 21.01 | 21.0084 | 0.019106 | -0.0015697 | 4.027 | 4.0346 | 0.011241 | 0.0076199
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 10.64 | 10.6474 | 0.024274 | 0.0074377 | 4.108 | 4.1242 | 0.0082405 | 0.016166
2009-59-0012.mrgcor1 | 2009-59-0012.samcor1 | 121 | 2.76 | 2.7629 | 0.025995 | 0.0028925 | 4.122 | 4.0854 | 0.0068278 | -0.036628
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 254.96 | 254.963 | 0.094167 | 0.0033875 | 4.117 | 4.084 | 0.0092317 | -0.033009
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 205.14 | 205.14 | 0.074511 | -0.00033569 | 4.122 | 4.0999 | 0.0061818 | -0.02205
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 178.13 | 178.131 | 0.09684 | 0.0010681 | 4.171 | 4.1125 | 0.011698 | -0.058471
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 155.9 | 155.898 | 0.074959 | -0.0018921 | 4.17 | 4.111 | 0.0049539 | -0.058975
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 128.73 | 128.731 | 0.063878 | 0.0011597 | 4.159 | 4.1115 | 0.01044 | -0.047471
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 101.97 | 101.971 | 0.033467 | 0.0009079 | 4.135 | 4.111 | 0.0073524 | -0.024008
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 75.49 | 75.4901 | 0.027883 | 8.3923e-05 | 4.145 | 4.131 | 0.0061297 | -0.014041
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 51.47 | 51.4701 | 0.026252 | 8.0109e-05 | 4.144 | 4.116 | 0.0092286 | -0.028
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 30.52 | 30.5239 | 0.021617 | 0.0038853 | 4.187 | 4.12 | 0.0099288 | -0.067049
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 20.4 | 20.3946 | 0.027356 | -0.0053711 | 4.149 | 4.1287 | 0.007001 | -0.020306
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 10.4 | 10.3998 | 0.020856 | -0.00016499 | 4.127 | 4.1337 | 0.0041264 | 0.0067024
2009-59-0016.mrgcor1 | 2009-59-0016.samcor1 | 121 | 2.86 | 2.8575 | 0.024162 | -0.0024793 | 4.164 | 4.1441 | 0.011314 | -0.019859
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 95.44 | 95.4426 | 0.32996 | 0.0026398 | 1.622 | 1.462 | 0.002523 | -0.16003
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 75.74 | 75.742 | 0.21257 | 0.0019836 | 2 | 1.8638 | 0.010315 | -0.13621
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 50.1 | 50.1009 | 0.49114 | 0.00091171 | 3.382 | 3.1588 | 0.0091896 | -0.22321
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 31.05 | 31.0488 | 0.13716 | -0.0011558 | 4.528 | 4.4047 | 0.0049991 | -0.12329
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 20.17 | 20.1762 | 0.097213 | 0.0061989 | 5.653 | 5.6163 | 0.009266 | -0.036661
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 10.4 | 10.3967 | 0.25409 | -0.0033054 | 6.144 | 6.0945 | 0.011578 | -0.049471
2009-59-0019.mrgcor1 | 2009-59-0019.samcor1 | 121 | 3.17 | 3.1758 | 0.10521 | 0.005785 | 6.359 | 6.2992 | 0.009842 | -0.059793
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 140.44 | 140.439 | 0.098003 | -0.0013275 | 1.413 | 1.4575 | 0.0027053 | 0.044488
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 126.03 | 126.026 | 0.4028 | -0.0044632 | 1.511 | 1.4844 | 0.0030629 | -0.026562
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 100.85 | 100.851 | 0.66544 | 0.00099182 | 2.807 | 2.7444 | 0.012684 | -0.062595
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 75.69 | 75.6848 | 0.53761 | -0.0052109 | 3.735 | 3.7151 | 0.0032144 | -0.019868
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 50.69 | 50.6874 | 0.25269 | -0.0025597 | 3.282 | 3.2712 | 0.0039398 | -0.01081
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 29.89 | 29.8926 | 0.16169 | 0.0025616 | 4.026 | 4.128 | 0.03187 | 0.10203
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 20.24 | 20.2401 | 0.5035 | 8.2016e-05 | 5.223 | 5.3017 | 0.018275 | 0.078735
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 10.95 | 10.9548 | 0.22517 | 0.0047932 | 5.735 | 5.6083 | 0.01656 | -0.12668
2009-59-0021.mrgcor1 | 2009-59-0021.samcor1 | 121 | 3.43 | 3.426 | 0.14875 | -0.003967 | 5.976 | 5.8747 | 0.011276 | -0.10126
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 104 | 961.51 | 961.794 | 0.77993 | 0.2843 | 0.209 | 0.22783 | 0.0026455 | 0.018827
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 751.09 | 751.086 | 0.75895 | -0.0042725 | 0.234 | 0.24736 | 0.0019365 | 0.013364
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 502.25 | 502.25 | 0.14286 | -0.00024414 | 0.698 | 0.67177 | 0.0020484 | -0.026231
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 400.36 | 400.356 | 0.69444 | -0.0042725 | 0.962 | 0.95184 | 0.0033541 | -0.010157
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 298.82 | 298.823 | 0.41508 | 0.0026245 | 1.497 | 1.4734 | 0.0029326 | -0.023545
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 250.1 | 250.097 | 0.33207 | -0.0034027 | 1.848 | 1.8324 | 0.0036612 | -0.01562
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 175.04 | 175.039 | 0.54212 | -0.0011444 | 2.979 | 2.9126 | 0.021708 | -0.066389
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 149.86 | 149.859 | 0.55771 | -0.00099182 | 3.579 | 3.5235 | 0.0034906 | -0.055537
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 124.94 | 124.937 | 0.1921 | -0.0026474 | 3.789 | 3.8073 | 0.013329 | 0.018331
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 100.3 | 100.298 | 0.86384 | -0.0016556 | 3.317 | 3.3094 | 0.011562 | -0.0076447
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 74.75 | 74.7543 | 0.55665 | 0.0042953 | 3.912 | 3.9631 | 0.010648 | 0.051116
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 50.63 | 50.6341 | 0.3442 | 0.0041313 | 3.43 | 4.2094 | 0.0061852 | 0.77939
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 30.74 | 30.739 | 0.2328 | -0.00099182 | 3.758 | 4.2215 | 0.02516 | 0.46352
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 20.91 | 20.9145 | 0.36215 | 0.0045452 | 5.068 | 4.9783 | 0.02119 | -0.089669
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 10.42 | 10.4181 | 0.08333 | -0.0019007 | 6.107 | 6.0468 | 0.0073599 | -0.06024
2009-59-0023.mrgcor1 | 2009-59-0023.samcor1 | 121 | 4.41 | 4.4139 | 0.22806 | 0.0038843 | 6.188 | 6.1441 | 0.034137 | -0.043934
2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 121 | 91.19 | 91.1854 | 0.07077 | -0.004631 | 1.499 | 1.1333 | 0.004596 | -0.36567
2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 121 | 75.58 | 75.5848 | 0.033371 | 0.0047913 | 2.658 | 1.5443 | 0.0033744 | -1.1137
2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 121 | 50.94 | 50.9383 | 0.043005 | -0.0017357 | 3.918 | 2.6878 | 0.0029701 | -1.2302
2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 121 | 30.48 | 30.483 | 0.090468 | 0.0029755 | 4.414 | 3.9731 | 0.058771 | -0.44093
2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 121 | 20.75 | 20.7521 | 0.077031 | 0.0021496 | 4.855 | 4.4183 | 0.0070355 | -0.43669
2009-59-0029.mrgcor1 | 2009-59-0029.samcor1 | 121 | 10.86 | 10.8585 | 0.083163 | -0.0014877 | 6.53 | 4.7674 | 0.017284 | -1.7626
2009-59-0030.mrgcor1 | 2009-59-0030.samcor1 | 121 | 28.99 | 28.9931 | 0.02772 | 0.0031414 | 5.415 | 5.2519 | 0.034625 | -0.16314
2009-59-0030.mrgcor1 | 2009-59-0030.samcor1 | 121 | 20.22 | 20.2194 | 0.020748 | -0.00057793 | 5.373 | 5.4546 | 0.030849 | 0.081578
2009-59-0030.mrgcor1 | 2009-59-0030.samcor1 | 121 | 10.18 | 10.1846 | 0.024598 | 0.0046282 | 5.659 | 5.7552 | 0.019724 | 0.096207
2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 121 | 66.1 | 66.0995 | 0.023053 | -0.00049591 | 2.915 | 0.13668 | 0.0025272 | -2.7783
2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 121 | 50.66 | 50.6574 | 0.021281 | -0.0026436 | 0.224 | 0.85786 | 0.0077967 | 0.63386
2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 121 | 30.37 | 30.3693 | 0.031695 | -0.00066185 | 4.414 | 4.5961 | 0.068019 | 0.1821
2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 121 | 20.25 | 20.2553 | 0.027872 | 0.0052891 | 4.315 | 4.2814 | 0.018565 | -0.033595
2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 121 | 10.47 | 10.4717 | 0.024164 | 0.0016527 | 4.691 | 4.4348 | 0.089573 | -0.2562
2009-59-0031.mrgcor1 | 2009-59-0031.samcor1 | 121 | 3.36 | 3.3607 | 0.027543 | 0.00074387 | 7.047 | 7.5419 | 0.044105 | 0.49486
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 117.39 | 117.394 | 0.025167 | 0.0036392 | 0 | 0.062297 | 0.0021666 | 0.062297
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 75.52 | 75.5174 | 0.028709 | -0.0025559 | 0.05 | 0.058157 | 0.00082672 | 0.008157
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 50.09 | 50.0859 | 0.034968 | -0.0041313 | 1.039 | 0.90846 | 0.027715 | -0.13054
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 30.25 | 30.2482 | 0.025496 | -0.0018177 | 3.91 | 3.8201 | 0.013842 | -0.089901
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 20.17 | 20.1725 | 0.031578 | 0.0024796 | 4.217 | 4.116 | 0.016943 | -0.10104
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 10.55 | 10.5469 | 0.021175 | -0.0031404 | 4.157 | 4.1589 | 0.0039091 | 0.0019503
2009-59-0032.mrgcor1 | 2009-59-0032.samcor1 | 121 | 2.76 | 2.762 | 0.020026 | 0.0019834 | 5.969 | 6.6568 | 0.11503 | 0.68778
2009-59-0033.mrgcor1 | 2009-59-0033.samcor1 | 121 | 50.1 | 50.096 | 0.024138 | -0.0040474 | 2.743 | 2.4787 | 0.022388 | -0.26434
2009-59-0033.mrgcor1 | 2009-59-0033.samcor1 | 121 | 30.5 | 30.4983 | 0.023759 | -0.0017357 | 3.478 | 3.7075 | 0.031619 | 0.22945
2009-59-0033.mrgcor1 | 2009-59-0033.samcor1 | 121 | 20.57 | 20.5725 | 0.022704 | 0.0024796 | 3.981 | 3.9769 | 0.026364 | -0.0040495
2009-59-0033.mrgcor1 | 2009-59-0033.samcor1 | 121 | 10.42 | 10.4191 | 0.027809 | -0.00090885 | 4.035 | 4.0781 | 0.0054894 | 0.043091
2009-59-0033.mrgcor1 | 2009-59-0033.samcor1 | 121 | 3.18 | 3.1766 | 0.025934 | -0.0033884 | 7.801 | 7.3193 | 0.049272 | -0.48175
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 198.3 | 198.302 | 0.027811 | 0.0023956 | 0.781 | 0.75659 | 0.011808 | -0.024413
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 174.92 | 174.92 | 0.020674 | -7.6294e-05 | 1.037 | 1.0535 | 0.0023773 | 0.016479
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 150.38 | 150.377 | 0.020072 | -0.003479 | 1.025 | 0.96864 | 0.0020811 | -0.056355
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 125.36 | 125.358 | 0.023464 | -0.0018997 | 0.888 | 0.87331 | 0.0025879 | -0.014694
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 100.15 | 100.146 | 0.026576 | -0.0042191 | 0.784 | 0.80339 | 0.002055 | 0.019388
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 75.45 | 75.4503 | 0.026233 | 0.00033569 | 0.992 | 1.0309 | 0.0019992 | 0.038893
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 50.47 | 50.4665 | 0.02312 | -0.0034714 | 1.97 | 1.9747 | 0.024709 | 0.0047025
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 30.71 | 30.7058 | 0.02362 | -0.0042133 | 3.599 | 3.6182 | 0.0075421 | 0.019157
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 20.95 | 20.944 | 0.026223 | -0.0059509 | 4.037 | 3.9773 | 0.0062896 | -0.059661
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 10.4 | 10.395 | 0.030663 | -0.0050411 | 4.022 | 4.02 | 0.008954 | -0.0020246
2009-59-0034.mrgcor1 | 2009-59-0034.samcor1 | 121 | 2.7 | 2.7024 | 0.031886 | 0.0023966 | 5.412 | 6.7825 | 0.19227 | 1.3705
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 151.07 | 151.069 | 0.025883 | -0.0013275 | 0.785 | 0.80195 | 0.0017217 | 0.01695
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 126.03 | 126.032 | 0.032472 | 0.0017395 | 0.753 | 0.77711 | 0.0023903 | 0.024107
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 102.01 | 102.007 | 0.026193 | -0.0029755 | 0.792 | 0.84628 | 0.0023102 | 0.054281
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 77.43 | 77.4279 | 0.026391 | -0.0020676 | 0.992 | 1.0436 | 0.0026762 | 0.051653
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 51.9 | 51.8949 | 0.025952 | -0.0051269 | 2.254 | 2.3662 | 0.002663 | 0.11222
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 32.05 | 32.0492 | 0.026096 | -0.00082398 | 3.418 | 3.5273 | 0.024443 | 0.10933
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 22.35 | 22.3515 | 0.02319 | 0.0014877 | 3.934 | 3.9708 | 0.0031676 | 0.03676
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 9.87 | 9.8759 | 0.025317 | 0.0059509 | 4.026 | 4.2248 | 0.0071107 | 0.19877
2009-59-0035.mrgcor1 | 2009-59-0035.samcor1 | 121 | 2.46 | 2.4622 | 0.038568 | 0.0022314 | 6.829 | 6.9339 | 0.087571 | 0.10494
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 117.41 | 117.41 | 0.023201 | 0.00016022 | 0.88 | 0.82157 | 0.003575 | -0.05843
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 76.7 | 76.7026 | 0.030352 | 0.0026474 | 0.944 | 1.0072 | 0.0017777 | 0.063198
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 51.37 | 51.3712 | 0.028124 | 0.0012398 | 2.322 | 2.4668 | 0.021281 | 0.14477
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 31.12 | 31.1189 | 0.02323 | -0.0010757 | 3.497 | 3.7887 | 0.0057223 | 0.29165
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 21.1 | 21.1007 | 0.035255 | 0.00066185 | 3.974 | 3.9818 | 0.0080526 | 0.0078266
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 11.02 | 11.0205 | 0.026453 | 0.00049496 | 4.064 | 4.0376 | 0.014521 | -0.026356
2009-59-0036.mrgcor1 | 2009-59-0036.samcor1 | 121 | 2.67 | 2.6684 | 0.021174 | -0.0015702 | 6.249 | 7.8774 | 0.023852 | 1.6284
2009-59-0037.mrgcor1 | 2009-59-0037.samcor1 | 121 | 82.94 | 82.9455 | 0.031572 | 0.005455 | 0.882 | 0.86898 | 0.0025722 | -0.013017
2009-59-0037.mrgcor1 | 2009-59-0037.samcor1 | 121 | 50.57 | 50.5728 | 0.031284 | 0.0028114 | 2.217 | 2.4354 | 0.010424 | 0.21838
2009-59-0037.mrgcor1 | 2009-59-0037.samcor1 | 121 | 20.63 | 20.6341 | 0.029145 | 0.0041332 | 3.96 | 3.9437 | 0.0087533 | -0.016322
2009-59-0037.mrgcor1 | 2009-59-0037.samcor1 | 121 | 10.75 | 10.7493 | 0.025355 | -0.0006609 | 4.153 | 4.2025 | 0.01858 | 0.04948
2009-59-0037.mrgcor1 | 2009-59-0037.samcor1 | 121 | 2.95 | 2.9496 | 0.022413 | -0.00041318 | 6.486 | 7.247 | 0.025047 | 0.76096
2009-59-0038.mrgcor1 | 2009-59-0038.samcor1 | 121 | 57.01 | 57.0129 | 0.02593 | 0.0028954 | 1.913 | 1.3522 | 0.0057507 | -0.56078
2009-59-0038.mrgcor1 | 2009-59-0038.samcor1 | 121 | 30.69 | 30.6883 | 0.023782 | -0.0016537 | 3.774 | 3.8888 | 0.0051088 | 0.11482
2009-59-0038.mrgcor1 | 2009-59-0038.samcor1 | 121 | 20.71 | 20.7091 | 0.018885 | -0.0009079 | 3.966 | 3.9936 | 0.0050742 | 0.027645
2009-59-0038.mrgcor1 | 2009-59-0038.samcor1 | 121 | 10.57 | 10.5712 | 0.023294 | 0.0012398 | 4.118 | 4.1482 | 0.0095936 | 0.030215
2009-59-0038.mrgcor1 | 2009-59-0038.samcor1 | 121 | 3.02 | 3.0257 | 0.02148 | 0.0057025 | 7.691 | 7.779 | 0.042462 | 0.088016
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 95.86 | 95.8626 | 0.025123 | 0.0026398 | 0.422 | 0.37323 | 0.0038789 | -0.048769
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 76.04 | 76.039 | 0.025929 | -0.00099182 | 1.369 | 1.3937 | 0.011351 | 0.024669
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 50.89 | 50.8867 | 0.02444 | -0.0033035 | 3.012 | 3.1478 | 0.043395 | 0.13583
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 30.79 | 30.7941 | 0.026822 | 0.0041313 | 3.838 | 3.8647 | 0.014472 | 0.026719
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 20.72 | 20.716 | 0.023579 | -0.0040493 | 3.975 | 3.9952 | 0.0064794 | 0.020182
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 10.62 | 10.6236 | 0.028515 | 0.0035543 | 4.136 | 4.2022 | 0.017533 | 0.066215
2009-59-0039.mrgcor1 | 2009-59-0039.samcor1 | 121 | 3 | 3.0009 | 0.024495 | 0.00090909 | 7.008 | 6.6938 | 0.09238 | -0.31423
2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 121 | 80.12 | 80.1208 | 0.026193 | 0.00082398 | 1.353 | 1.1648 | 0.0046156 | -0.18821
2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 121 | 50.35 | 50.3501 | 0.024409 | 8.3923e-05 | 3.057 | 3.1995 | 0.0062061 | 0.14245
2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 121 | 30.35 | 30.3512 | 0.02504 | 0.0011578 | 3.777 | 3.8623 | 0.0044376 | 0.085339
2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 121 | 20.25 | 20.254 | 0.023473 | 0.0040493 | 3.997 | 4.0688 | 0.0031578 | 0.07181
2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 121 | 10.08 | 10.0848 | 0.024769 | 0.0047932 | 4.226 | 4.4483 | 0.0049679 | 0.22226
2009-59-0040.mrgcor1 | 2009-59-0040.samcor1 | 121 | 3.37 | 3.366 | 0.022821 | -0.0039668 | 5.751 | 8.7559 | 0.044073 | 3.0049
Create Fit Sheet