## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2009-60

Agency: OSD

Location: Strait of Georgia

Project: SoG Moorings

Party Chief: Johannessen S.

Platform: Vector

Date: October 31, 2009 – November 1, 2009

Processed by: Germaine Gatien

Date of Processing: June 29, 2010 – November 5, 2010

Number of original CTD casts:  2	Number of CTD casts processed: 2

Number of bottle casts:	2		Number of bottle casts processed: 2

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (serial number not recorded but probably #2356) with a 3X cable, a pH sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter (#1252).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book lacked a list of personnel. The fluorometer serial number was not recorded, just the gain.

The gain for the fluorometer was given as 3X in both the configuration file and the log book, but in neither case was the sensor serial number recorded, so perhaps no one checked this. The fluorescence values found when 3X is used are very high for late October. For the previous cruise using the same sensor the 3X sensor was used for the 1st cast only, then the 30X cable was used. The fluorescence data were corrected by dividing all values by 10. A note was put in the header to indicate there is some doubt about the fluorescence gain.

No recalibration of salinity was applied because the evidence is weak and there are on-going concerns about the salinity sampling results from many 2009 cruises. There could be a salinometer problem, but more likely there are sampling issues such as worn-out liners, chipped bottles or bottles with salt crystals in them. Recalibration should be revisited when there is a post-cruise calibration report from the factory.

Fluorescence data are nominal and unedited except that some records were removed in editing temperature and salinity.

Transmissivity data are also nominal; however, a problem has been discovered with the transmissometer used for this cruise. From Sept 2009 it was producing extremely low values. A correction was made by multiplying all values by 2.087. This is still a rough estimate, but brings values into line with expectations. For details see file “Transmissometer 1005DR Corrections.Doc” in folder: Osd_Data_Archive:\Cruise_Data\DOCUMENTS

SBE dissolved oxygen is considered ±0.08mL/L in the top 130m and ±0.02 below that.

pH:SBE data will not be archived because the calibration data are not available yet. Both may be

added later.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. A few files had errors in the names; these were corrected.

There were only 2 casts, so calibration information and alignment parameters will be based on cruises 2009-14 and 2009-64 which preceded and followed this cruise by approximately a month respectively.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained. The log indicates that the CTD event #s were 1 and 4, not 1 and 2 as the data files show. File 2009-60-0002.* will be changed to 2009-60-0004.* after conversion.

Dissolved oxygen and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

The following errors were found in the file:

- The configuration for DO #0119 used the Owens-Millard parameters. It is better to select the Murphy-Larson parameters, so that the Tau correction can be applied. The nominal values were entered for the E, H1 and H3 since there is insufficient deep sampling to fine-tune them. It is not expected that the hysteresis correction will be applied since there was no very deep sampling.

- The transmissivity date and parameters were wrong as they have been for a few recent cruises. For those other cruises checks were made to see if the serial number might be wrong. The only transmissivity calibrations for that date had different values and no calibration could be found with those values. It will be assumed that this really is 1005DR and the appropriate numbers were entered in the con files.

- The pressure offset for CTD #0550 was entered as 0.1033. In other recent cruises a value of +0.5db was found appropriate, though for 2009-38 it was noted that it might be time to increase the offset a little and for 2009-14 +0.7db was used. Examination of a few test conversions for this cruise suggests a value of +0.7db is probably appropriate here as well. So the offset was changed to +0.7db.

- The fluoremeter serial number is missing. It is assumed to be #2356 because that was used for 2009-14 and 2009-64 when all other equipment was the same.

##### Conversion of Raw Data

Rosette files were converted using a start time of -5s and duration of 10s and using 2009-60-ctd.con. The Tau correction was chosen, but not the hysteresis correction for DO calculation.

Plots were made of the two casts and there is no evidence of processing problems.

The ROS files were converted to IOS HEADER format. CLEAN was then run to add event numbers to the headers and those files were named *.BOT. The BOT files were plotted to check for outliers and none were found.

CNV files were converted using the configuration files 2009-60-ctd.con for all casts.

The CNV casts were examined and all expected channels are present.

- The two temperature channels are mostly in reasonable agreement on the downcasts though there are some excursions. The upcast data are much noisier, with spikes in both so neither looks better. The conductivity is similar to the temperature.

- The fluorescence values look reasonable in shape and while it is occasionally high, there is no evidence of data being off-scale. One concern is that the values are high for the time of year, so is the gain wrong? The log and con file both indicate that the gain was 3X, but for all but the first cast of 2009-14 and all of 2009-64 the gain was 30X and those two cruises bracket this one. Why would anyone have thought of using the 3X gain in late October? Using 30X would lead to top values of fluorescence of about 3ug/L instead of the 29.8 with the 3X. The lower values are more believable, though cruise 2009-46 in early October found a few subsurface values of up to ~15ug/L in this area, but they were associated with very low transmissivity as well, which is not the case here. This will be investigated further later and can be recalibrated if it is thought necessary.

- Dissolved oxygen voltage looks as usual with an offset but some detail to help alignment.

- The descent rate looks fairly low but steady.

- Transmissivity looks ok in shape, but the values are probably too low. It is hard to judge in this area, but other recent cruises do suggest there as a problem with this transmissometer.

- The altimetry looks good near the bottom.

- Surface PAR is listed in the configuration file, but an initial conversion showed there was just a constant value, so conversion was rerun without it since one cast was in daylight hours, so the sensor was apparently not in use. There was no PAR for this cruise.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Based on tests done for 2009-14 and 2009-64, CELLTM was run on these data using parameters (α = 0.0245, β=9.5) for the primary and (α = 0.02, β=7) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

The differences between primary and secondary channels were larger for 2009-14 than for 2009-64. One cast from this cruise was examined with one from each of the other 2 cruises.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2009-14-0073 | 275 325 | +0.0015 +0.0013 | -0.00004 -0.00008 | -0.0017 -0.0017 | Mod., V. steady |
| 2009-60-0001 | 270 | +0.0013 | +0.00008 | -0.0004 | Low, steady |
| 2009-64-0059 | 275 320 | +0.0007 +0.0008 | +0.00008 +0.00008 | +0.0003 +0.0004 | High, steady |

The temperature difference looks like the earlier cruise, the conductivity like the later cruise, and the salinity difference is intermediate.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. No problems were found.

The cross-reference check was compared with the log book, and no errors were found.

The cruise track was plotted and added to the end of this report.

The surface values program shows the average surface pressure to be 1.9db which is typical for the Vector. There are a few small negative values at the end of cast #1 when the pumps were off. The data are not sufficient to determine if there is an error in pressure, but if there is, it is slight, so no further recalibration will be applied.

The altimeter readings were examined on a plot. There was a lot of noise in the altimetry at the bottom for event #4, but the algorithm worked well and the header entries appear to be correct.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. The addsamp.csv file was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

The salinity data were delivered in spreadsheet 2009-60-salinity.xls; there were only 2 values, both were the average of duplicates that differed by 0.006 and 0.003. The spreadsheet was simplified (unneeded columns removed and headers changed to standard format) and saved as 2009-60-salinity.csv which was then converted to individual SAL files. Both samples had “fc” flags assigned.

DISSOLVED OXGYEN

The dissolved oxygen data was provided in spreadsheet 2009-60oxy.xls with quality flags and comments. There was an analysis of duplicates. The spreadsheet was simplified and saved as 2009-60oxy.csv which was converted into individual ADD files. The analyst recommended reassessing the duplicates after a look at the CTD data.

The SAL and ADD files were merged with CST files in 2 steps. (Output: MRG1 and MRG3) MRG3 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

Plots were made of salinity versus titrated dissolved oxygen from bottles and SBE:Oxygen:Dissolved. No outliers were found.

##### Compare

Salinity

There are too few samples to make COMPARE worth running, and there are duplicates for both bottles sampled, so differences were found between each CTD salinity channel and each sample from the same cast.

| Event | CTD channel | Sample a | Sample b |
| --- | --- | --- | --- |
| 1 | Primary Sal | -0.0091 | -0.0031 |
| 1 | Secondary Sal | -0.0089 | -0.0029 |
| 4 | Primary Sal | -0.0004 | -0.0031 |
| 4 | Secondary Sal | -0.0006 | -0.0033 |

We expect the differences to be similar for the two CTD channels because of the results of section 7. Using sample b for each of the Niskins does result in both CTD channels being low by about 0.003. Choosing sample “a” for both or “a” for one and “b” for the other produce significant differences between primary and secondary. Based on this result, the average of duplicates was replaced by the 2nd sample in each case; the “f” flag was removed but the “c” was left since the evidence is not strong. The differences do not justify recalibration as there are so few data available. While the differences are intermediate between the results for 2009-14 and 2009-64, there are sufficient doubts about salinity samples from all these cruises to await further, stronger evidence before recalibration of salinity.  (For details see 2009-60-salinity-comparison.xls.)

Dissolved Oxygen –

COMPARE was run using the SBE DO and the Titrated bottle DO data. There are only 2 casts, but there were 14 bottles for each. When a few outliers are removed, the trendline is:

CTD-BOT = 1.0213 DOX-CTD + 0.015

This is similar to the result from 2009-64 from the same region:

CTD-BOT = 1.0203 DOX-CTD + 0.0696

(See 2009-60-dox-comp1.xls.)

The two most significant outliers were

Sample #1 – While the CTD DO is a little noisy, this can only account for a small fraction of the difference. Flagged “c”.

Sample #15 - The analyst noted problems with a computer crash and 2 titers added. There is some noise in the CTD DO data, but not nearly enough to account for the difference. Looks like sample problem. Flagged “d”.

The analyst also noted that the following 2 replicate pairs did not compare well:
Sample #1 – One value compared much better with the CTD than the other, so it was selected, the “f” flag was removed and a comment was entered.

Sample #16 - While value 3.073 is slightly closer to the fit, neither is a clear outlier, so kept average value and “f” flag.

##### 12. Shift

##### There are only 2 casts and both had many bottles fired, so there is little evidence on which to base the shift settings, so those from 2009-64 were applied to fluorescence, conductivity and dissolved oxygen.

Fluorescence

The usual shift of +24 records (1s) was found to improve the alignment. The results are satisfactory.

Conductivity

SHIFT was not run on the primary conductivity; it was run on the secondary with a setting of -0.6 records and the results are satisfactory.

Dissolved Oxygen

SHIFT was run with a setting of +65 since it has been used for other cruises with this sensor and the results were satisfactory.

pH

At this stage no decision had been made on whether to archive this channel or not, so it was examined to see if it needed alignment. There is no mention in the log that it was pumped. If it wasn’t we would not expect to see a delay, but there does appear to be some. A shift of +100 records does make the DO trace offset (between downcast and upcast) look closer to that of the temperature trace and the match between features in the two traces are closer. But this is not a simple comparison given there were so many stops. Nonetheless, it looks useful to advance the pH channel by 100 records.

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings..

##### 14. DETAILED EDITING

COMPARE indicates that the two salinity channels are very close. The primary channels were selected for 2009-64 and a quick look using CTDEDIT suggests that those channels have less noise for this cruise too. So the primary channels were selected for editing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

Both casts were edited lightly. Notes were entered in the header with details of editing.

##### 15. Initial Recalibration

While the salinity appears to be low by ~0.003, there are only 2 bottles fired during this cruise and there are doubts about the quality of the bottle comparisons from other cruises using the same equipment in late 2009. So no recalibration will be applied to salinity, but when any of these sensors are recalibrated, this decision can be revisited. A decision on how to handle the fluorescence was difficult. Dr. A. Peña examined the data and concluded that the fluorescence was probably too high, though the evidence was not strong. That channel will be recalibrated.

File 2009-60-recal1.ccf  was prepared to divide the SBE Fluorescence by 10, multiply transmissivity by 2.087 and to apply the following recalibration to the SBE dissolved oxygen:

CTD-BOT = 1.0213 DOX-CTD + 0.015

This was applied to the SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the DO corrections were applied correctly and they were. (See 2009-60-dox-comp2.xls and 2009-60-fl-chl-comp2.xls.)

The EDT files were then recalibrated to create COR1 files.

##### 16. Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data is compared to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. With so few values and considerable scatter, the fit is rather dependent on how outliers are identified, but it is clear that near-surface values are systematically a little high, while near the bottom they are too low, so it is worth attempting to recalibrate. (See 2009-60-dox-comp3.xls)

The following correction was applied

SBE DO (corrected) = SBE DO (after 1st recal) + 0.0002*Pressure -0.0527

COMPARE was rerun and the results were satisfactory with an average difference of -0.004mL/L between bottles and SBE DO, and with the exception of obvious outliers the bottles were within ±0.08mL/L of the SBE DO in the top 130m and ±0.02 below that. (See 2009-60-dox-comp4.xls.)

##### 17. Special Fluorometer Processing

No special files were prepared for Angelica Pena since there are no CHL data available.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### 18. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

##### 19. Other Comparisons

Previous experience with these sensors –

1. Salinity:  The primary conductivity sensor was found to be low by 0.004 salinity units when used for 2009-03 (Feb) but the sampling was limited to 2 depths. During 2009-27 (March) and 2009-38 (May) there was no calibration sampling. During 2009-26 in April it was found to be high by 0.0025 but there was a lot of scatter in the comparison. During 2009-14 there was a lot of scatter but the CTD appeared to be very close to the bottles. During 2009-64 it was low by 0.0044. The secondary conductivity sensor was used for the same cruises and was found to be low by <0.001 for 2009-03, high by 0.0004 for 2009-26,  low by 0.002 for 2009-14 and by 0.0044 for 2009-64. It was also used on an SBE25 in 2007 but there were doubts about the bottle analyses for those cruises.

2. Dissolved Oxygen – There are only 4 other cruises since the last factory calibration from which there is DO calibration sampling and both had few samples. For 2009-38 some samples came from areas of high variability in DO. The fits from those were:

CTD-BOT = 1.0209 DOX-CTD + 0.0668 (2009-27 - March) Offshore

CTD-BOT = 1.0388 DOX-CTD + 0.0088 (2009-26 – April)  Strait of Georgia

CTD-BOT = 1.0700 DOX-CTD – 0.0887 (2009-38 - May) Saanich Inlet

CTD-BOT = 1.0414 DOX-CTD + 0.0005 (2009-14 - September)  Strait of Georgia

CTD-BOT = 1.0203 DOX-CTD + 0.0696 (2009-64 - September) Strait of Georgia

3. Pressure –The sensor was recalibrated in August 2007 and an offset of +0.5db was used for 6 cruises processed since then. However, it was noted that it might soon be time to increase the offset, so +0.7db was used for 2009-14 and 2009-64.

Historic ranges (3 standard deviations) –All data fell within the local climatology.

##### 20. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Altimeter, Descent_Rate, Status:Pump, pH:SBE and Flag.

A second SBE DO channel (with umol/kg units) was added. REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

SBE Fluorescence was originally converted using a 3X gain, but that led to

high values for the time of year, ~25 mg/m^3. There is no evidence that values

were particularly high in the Strait of Georgia at that time, but we do not

have extracted chlorophyll samples, so there remains some doubt.

The fluorescence was corrected by dividing the values by 10.

Transmissivity data from sensor #1005DR have been found to be very low on

all deployments Since July 2009. A correction factor of 2.087 was applied based

on deep offshore measurements from other cruises. For details see

Osd_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissometer #1105DR Correction.Doc

SBE dissolved oxygen is considered ±0.08mL/L in the top 130m and ±0.02 below that.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values were 90% and 95% which is typical of late Autumn in the Strait of Georgia.

##### 22. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### 23. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| pH | 0692 | 06feb09 |  |  |  |
| Surface PAR | 16504 | 27Feb2009 |  |  |  |
| Fluorometer | ??2356 |  |  |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |