## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 24 Nov 2021 | ### Correcting Salinity:Bottle precision lost during HPLC addition. S.H. |
| ### 17 Dec 2020 | ### Added HPLC Data. S.H. |
| 18 October 2010 | Recalibrated Transmissivity Data – see end of report for details. |

## PROCESSING NOTES

Cruise: 2009-64

Agency: OSD

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG/JdeF

Party Chief: Chandler P.

Platform: Vector

Date: November 30, 2009 – December 5, 2009

Processed by: Germaine Gatien

Date of Processing: February 16, 2010 – June 24, 2010

Number of original CTD casts:  77	Number of CTD casts processed: 77

Number of bottle casts:		22	Number of bottle casts processed: 21(1 had no sample #s)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted in a rosette and attached were a Wetlabs CSTAR transmissometer (#1005DR), an SBE 43 DO sensor (#1119), a Seapoint Fluorometer (#2356) with a 30X cable, a Biospherical QSP-400 PAR sensor (#4565), a QSR-2240 Reference PAR sensor (#16504) and an altimeter (#1252).

# SUMMARY OF QUALITY AND CONCERNS

The Daily Science Log Book lacked an equipment list. There was only a note of the gain used for the fluorometer. The log was otherwise in good order with notes about problems encountered during the cruise. The rosette sheets include many comments, mostly concerning DO analysis, but problems also noted with sampling techniques, spigot leaks and other bottle problems.

Event #56 occurred after event #57; labels had already been prepared before the decision was made to change the order of work.

While both T/C sensor pairs appeared very close during stops for bottles, the secondary temperature was quite noisy when the CTD was in motion, so the primary temperature and salinity were chosen for archiving. Both salinity channels were lower than the bottle salinity by an average of 0.0044. No recalibration of salinity was applied because the evidence is weak and there are on-going concerns about the salinity sampling results on many 2009 cruises; there could be a salinometer problem, but more likely there are sampling issues such as worn-out liners or chipped bottles being used. Recalibration should be revisited when there is a post-cruise calibration report from the factory.

Sea-Bird have a new algorithm for dissolved oxygen with several parameters that the manufacturer recommends be fine-tuned for each instrument to produce the best data. This requires bottle samples from deep casts. Since these new parameters became available there has been no cruise using this sensor that has sufficient deep sampling to do this, so nominal values were used. The hysteresis correction was not applied since there are no data below 500db. The Tau correction in the derivation of dissolved oxygen concentration was applied.

The configuration file used for conversion of these data contained an error in the fluorometer gain; that was corrected at the calibration stage. There was also an error in the original derivation of SBE oxygen concentration. Those data were derived again and merged with the other data so the final files are correct.

Fluorescence and Transmissivity data are nominal and unedited except that some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files are considered, roughly:

•	±0.3  ml/l from  0– 150db

•	±0.1  ml/l from 150– 200db

•	±0.03 ml/l below 200db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension HEX. A few files had errors in the names; these were corrected.

##### Preliminary Steps

The Log Book and rosette log sheets were obtained.

A number of problems were noted including DO analysis issues, poor sampling techniques, spigots leaking and other bottle problems.

Extracted chlorophyll, dissolved oxygen, nutrients and salinity data were obtained in spreadsheet format.

The cruise summary sheet was completed.

The histories of the conductivity, DO and pressure sensors were obtained.

All the CON files produced in acquisition were the same. One was saved as 2009-64-ctd.con. Somehow, an error was introduced in the SBE Fluorescence gain. This was fixed later by recalibration.

The calibration constants were checked for all instruments. There were a number of problems:

- The configuration for DO #0119 used the Owens-Millard parameters. It is better to select the Murphy-Larson parameters, so that the Tau correction can be applied. The nominal values were entered for the E, H1 and H3 since there is insufficient deep sampling to fine-tune them. It is not expected that the hysteresis correction will be applied since there was no very deep sampling.

- The transmissivity date and parameters were wrong as they have been for a few recent cruises. For those other cruises checks were made to see if the serial number might be wrong. The only transmissivity calibrations for that date had different values and no calibration could be found with those values. It will be assumed that this really is 1005DR and the appropriate numbers were entered in the con files.

- The pressure offset for CTD #0550 was entered as 0.1033. In other recent cruises a value of +0.5db was found appropriate, though for 2009-38 it was noted that it might be time to increase the offset a little and for 2009-14 +0.7db was used. Examination of a few test conversions for this cruise suggests a value of +0.7db is probably appropriate here as well. So the offset was changed to +0.7db.

##### Conversion of Raw Data

Rosette files were converted using a start time of -5s and duration of 10s and using 2009-64-ctd.con. The Tau correction was chosen, but not the hysteresis correction for DO calculation. After conversion the file names were changed to standard format. There were many ROS files without any data. Those were deleted. There was a ROS file for cast #76 but this bottle was only used to obtain bulk water and no sample number was assigned, so the file will not be processed further. Header Check was run and there are no off-scale fluorescence values and no evidence of processing problems.

The ROS files were converted to IOS HEADER format. CLEAN was then run to add event numbers to the headers and those files were named *.BOT. The BOT files were plotted to check for outliers and none were found.

CNV files were converted using the configuration files 2009-64-ctd.con for all casts. After conversion the non-standard file names were fixed.

There is a note about a pressure test being run during cast #50, but an initial look at the data shows no indication that the CTD was brought through the surface, so it is not clear if the pressure should read 0 at the end of the cast or not. The minimum pressure is 0.5db and while the conductivity is fairly low, there is no indication of “in-air” values.

A few CNV casts were examined and all expected channels are present.

- There are spikes in pressure for cast #23 near the end of the cast.

- The two temperature channels are mostly in reasonable agreement on the downcasts though there are some excursions. The upcast data are much noisier with some casts especially so. The difference between upcast and downcast seems greater as though the pumps were not working well on the upcast or the rosette package was swinging a lot. It is not clear that one T/C pair were performing better than the other.

- The fluorescence values look reasonable in shape, though sometimes the offset between downcast and upcast is larger than expected. The values seem much too high for December – it was later realized that this was due to an error in the gain entered in the configuration file used for conversion. This will be corrected later in the processing..

- Dissolved oxygen voltage looks as usual with an offset but some detail to help alignment.

- The descent rate looks reasonably steady except near the mouth of Juan de Fuca Strait where it is very noisy.

- PAR, Surface PAR and transmissivity look fine.

- The altimetry looks good near the bottom even for casts with noisy descent rate and weight on.

##### WILDEDIT

Program WILDEDIT was run to remove spikes from the pressure, conductivity and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

The spikes in cast #23 were not removed by WILDEDIT. Since these spikes do not affect the downcast, no further runs on WILDEDIT were done.

Tests were run at this point to see if aligning the oxygen voltage might work better at this stage rather than at the usual stage after conversion to IOS format files. The setting that aligned the voltage best produced oxygen concentration that was not well aligned, so this will be run later as usual.

##### CELLTM

Tests were run on 3 casts using a variety of settings for CELLTM. The differences among the various choices were small. Overall the best choices were found to be (α = 0.0245, β=9.5) for the primary and was (α = 0.02, β=7) for the secondary, the same as found during 2009-14 and 2009-26 when the same CTD sensors were used.

CELLTM was run applying those settings to all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and dissolved oxygen concentration.

on a few casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts using was plotted to check for agreement between the pairs of T and C sensors. The differences are extremely noisy despite steady descent rates, so these are very rough averages.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 2 | 275 320 | +0.0010 +0.0009 | +0.00008 +0.00007 | ~0 VN ~0 VN | Mod., steady |
| 54 | 275 320 | +0.0007 +0.0007 | +0.0001 +0.00015 | +0.0004 +0.0004 | High, steady |
| 59 | 275 320 | +0.0007 +0.0008 | +0.00008 +0.00008 | +0.0003 +0.0004 | High, steady |
| 73 | 275 | +0.0009 | -0.0001 N | ~0 N | High, steady |
| 77 | 275 320 | +0.0009 +0.0008 XN | +0.00003 ~0 N | -0.0005 -0.0007 XN | High, steady |

The differences are smaller than for 2009-14, though the temperature differences are still a little larger than usual. Looking at a plot of the two temperature channels with the temperature differences suggests that the differences are due to slight misalignment with differences only significant in gradients. Where there are small gradients, the differences are very small. The sensors look ok.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

##### Checking Headers

The header check was run. No problems were found. The fluorescence did not go off-scale.

The cross-reference check was compared with the log book, and just one error was found. For cast #44 the longitude in the log differs by 1 minute from the header entry. The log entry had been corrected and that correction appears to be an error. No change was made to the header.

The cruise track was plotted and added to the end of this report. No problems were found.

The surface values program shows the average surface pressure to be 1.2db which is lower than usual for the Vector. However, conditions were unusually calm for most casts which would have allowed deployment closer to the surface. There were a number of readings around 0.5db with salinity usually at least 25, but the pumps were not turned on. To get surface readings this low we would expect that the rosette top was out of the water which is unusual, so it might suggest that the pressure reading is too low. But it was very calm and this may have reflected a desire to start the casts very close to the surface.

During cast #50 a test was done to check the pressure offset, but it is not clear if this was to check the CTD or deck readout of pressure. The rosette was brought through the surface with the pumps off. The minimum pressure for that cast was about 0.5db which might suggest that the CTD pressure is reading too high. It is not certain that the CTD itself came right through the surface. Given that the CTD recorded in-water values to the end, it would seem that it was not out of the water. These were not just salinity and conductivity, which are not reliable with the pumps off, but transmissivity also looks like it did not record out of water. Perhaps acquisition stopped before it went through the surface.

As another check on pressure 10 casts were picked randomly and examined in detail to see if the CTD pressure at the bottom looks right. The bottom pressure, bottom altimeter reading and bottom depth from the log book were recorded. By converting the CTD pressure to depth, adding the altimeter reading and comparing that to the bottom depth in the log book, we can get an estimated error. We do not expect these to compare well since there can be drift between the log entry and the CTD reaching the bottom, and the accuracy of the sounder and altimeter are not known, but if there is a systematic difference it would raise a question as to the CTD pressure.

| Cast | Max Pressure | Max Depth | Altimeter reading at bottom | CTD Max Depth + Altimeter | Bottom Depth From Log | CTD depth estimate – Log depth |
| --- | --- | --- | --- | --- | --- | --- |
| 2 | 328.3 | 325.3 | 4.8 | 330.1 | 328 | -1 |
| 10 | 218.1 | 216.1 | 7.6 | 223.7 | 225 | -1 |
| 19 | 114.3 | 113.3 | 2.7 | 116 | 117 | 0 |
| 33 | 183.1 | 181.5 | 4.6 | 186.1 | 186 | -3 |
| 42 | 242.2 | 240 | 3.0 | 243 | 246 | -1 |
| 50 | 419.5 | 415.5 | 5.7 | 421.2 | 422 | -1 |
| 58 | 351.0 | 347.7 | 3.9 | 351.6 | 353 | 2 |
| 60 | 165.0 | 163.5 | 6.1 | 169.6 | 168 | 2 |
| 70 | 145.0 | 143.7 | 4.9 | 148.6 | 147 | 0 |
| 77 | 350.6 | 347.3 | 5.5 | 352.8 | 353 | 0 |

In most cases the bottom depth estimate from the CTD is a little low compared to the log, but it is higher in 4 cases. There is no hint of the differences being pressure-dependent which suggests there is not a problem with the slope in the pressure calibration. It is dangerous to conclude that the CTD pressure is too low, but there is certainly no sign that it is too high as might be suggested by the pressure test at cast #50.

There is no strong evidence that the pressure is either too high or too low, so no further recalibration will be applied.

The altimeter readings and water depths were exported from the headers to a spreadsheet. Plots were made of altimetry for a random selection of casts and then all casts where the descent rate was noisiest. The header entries all looked fine.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. There were some bottles fired that had no sample numbers assigned. Cast #12 had no sampling so will not be processed further. The addsamp.csv file was then converted to CST files to be used as a framework for the bottle files. It was also used to add sample numbers to the BOT files. The SAM files were then bin-averaged.

SALINITY

The salinity data were delivered in spreadsheet 2009-64-salinity.xls; there were no duplicates. The spreadsheet was simplified (unneeded columns removed and headers changed to standard format) and saved as 2009-64-salinity.csv which was then converted to individual SAL files. There were two errors in the file where sample number and station name were reversed for events #3 and 7.

DISSOLVED OXGYEN

The dissolved oxygen data was provided in spreadsheet 2009-64oxygen.xls with quality flags and comments. There was an analysis of duplicates and two were rejected based on Chauvenet’s criterion. The spreadsheet was simplified and saved as 2009-64oxy.csv which was converted into individual ADD files.

NUTRIENTS

The nutrient data were obtained in spreadsheet QF2009-64nuts.xls which included a report on precision. The spreadsheet was simplified and saved as 2009-64-nuts.csv. File 2009-64-nuts.csv was then converted to individual NUT files.

EXTRACTED CHLOROPHYLL

Extracted chlorophyll data were obtained in file 2009-64chlarc.xls with c flags assigned by the analyst where the duplicates differed by >10%. The file was saved as 2009-64-chlarc.csv, header names were edited, “CHL:” was entered before the comments and event numbers were added. All data were the average of duplicates so “f” flags were added to all. The file was then converted to individual CHL files.

The SAL, CHL, ADD and NUT files were merged with CST files in four steps. (Output: MRG1, MRG2, MRG3 and MRG4), MRG4 was put through CLEAN to reduce the headers to File and Comment sections only. That file was then merged with SAMAVG files (Output:MRG).

##### Compare

Salinity

COMPARE was run and shows a lot of scatter in the differences between bottles and CTD for both sensor pairs but this is expected given the fact that this cruise did not sample below 360db. The two CTD salinity channels are remarkably close, both lower than the bottles by an average of 0.0044 when only bottles below 200db are included. The primary salinity has slightly higher dependence on pressure and on time than the secondary. No outliers were severe enough to justify assignment of flags. The 3 points that differed from the mean by >0.02 were associated with CTD salinity with fairly large standard deviations during the bottle stops and all others that differed from the mean by >0.01 were in the higher gradient zones between 0 and 150m. (See 2009-64-sal-comp1.xls.)

Dissolved Oxygen –

COMPARE was run using the SBE DO and the Titrated bottle DO data. The trendline is unusually flat with a large offset. The range of DO is limited, with the only value <1.7ml/L being a major outlier. And for many casts there was little variation in the profile, so response time is probably better than usual. There is no reason to expect a large offset, so fits that produce one are suspicious. On the other hand there is almost no data other than from cast #1 with low values, so forcing a fit based on reasonable results at low values, is not a good idea. Attempts to break the cruise down into sections and study the results produced slight differences, but all had offsets that seem unlikely. The fits that look good for DO>2.5 do not do a good job below 2.5. Attempts were made to fit polynomials with the offset fixed to see if a fit that suits most of the data could be produced. In the end the following linear fit looks as good as any of the others overall:

CTD-BOT = 1.0203 DOX-CTD + 0.0696

This is remarkably similar to the one from 2009-26, but that cruise had only 5 bottles. From the most recent cruise using this sensor the result was:

CTD-BOT = 1.0414 DOX-CTD + 0.0005

There is no evidence of drift through the cruise, though it would be hard to detect if there were, due to scatter and the different ranges of casts. (See 2009-64-dox-comp1.xls.)

The following samples were studied, and the analyst consulted,  to see if flags needed to be added or removed:

- Cast #1 – Sample #1 from 100m seems too high compared to the sensor, but there is no indication of sampling problems. There was a large shed wake as the CTD stopped that would have brought water with higher DO values through the rosette. The CTD had settled down well before the bottle was closed, but in such a protected area it is possible that the Niskin bottle did not flush well. It was flagged “c”.

- Cast #13 – The sample from 184db was a major outlier, but had already been flagged “d” by the analyst.

- Cast #55 – Sample #203 is an outlier. It looks like it might have been drawn from the wrong bottle. The CTD data look very similar at 200 and 177m both during upcast and downcast. The bottle at 177m looks more like the bottle and CTD at 150m. Sample #177 will be flagged "c".

- All the “b” flags assigned by the analyst will be removed since none of those samples stand out in COMPARE. – an explanation was added to the headers.

- Of the bottles flagged “d” by the analyst, 1 was a clear outlier, one slightly out of line. Those had a comment added and flag left unchanged. For the 3 samples with the comment “VENT OPEN” the flag was changed to “b” since they were not significant outliers; comments were added.

The ADD files were corrected and MERGE run again.

NOTE: The analyst’s original spreadsheet was renamed as “2009-64oxygen-original.xls” and a version with the changes mentioned above was saved as 2009-64oxygen.xls.

Fluorescence

COMPARE was run using the CTD CHL and the Titrated Chlorophyll from bottles. Plots were prepared of titrated CHLa versus CTD Fluorescence. As noted earlier, the SBE fluorescence looks too high. It is higher than the extracted chlorophyll by an average factor of 20. A check of the configuration file shows that the gain was set to 3X in the file used for conversion – those used at sea have a 30X gain which is a logical choice for December. When the SBE Fluorescence values are divided by 10 the comparison looks more believable, with fluorescence higher than extracted chlorophyll by an average of 2X and a range of 0.4X to 4X. A plot was made of the ratio of the corrected SBE Fluorescence to Extracted CHL against event number and pressure. From this we can conclude:

- The ratio is lowest in Juan de Fuca Strait and highest in Saanich Inlet and at Station 27 in the middle of the Strait of Georgia.

- Excluding a few outliers, the ratio gradually rises from the mouth of Juan de Fuca Strait to the northern end of the Strait of Georgia with a lot of variability at the northern end.

- The same trend is found at the 3 sampling depths, but there is higher variability at the surface.

(See 2009-64-fl-chl-comp.xls.)

Data from the MRG files were exported to a spreadsheet to check that everything looks ok. No problems were detected.

##### 12. Shift

Fluorescence

The usual method to find what shift is needed for the fluorescence is to examine upcast and downcast profiles to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. This is always rough estimate as the upcast data are usually very noisy. The usual shift of +24 records (1s) was found to improve the alignment. This is the shift that has been used in most other cruises, and it was applied to these data. (Output: SHFFL)

Conductivity

Tests were run on a few casts and the best choice was 0 for the primary conductivity channel. For the secondary the best choice varied from -0.4 to -0.6 with -0.6 best overall.  SHIFT was not run on the primary conductivity; it was run on the secondary with a setting of -0.6 records. (Output: *.SHFC1)

Dissolved Oxygen

Tests were run on a few casts to determine the best SHIFT value to apply to the Dissolved Oxygen channel. This was judged by how the vertical offset between downcast and upcast traces compares with that of the temperature. Because there is an offset in values between upcast and downcast due to the time response, alignment will not produce traces that overlie each other exactly. Distinctive features aid this judgment. In recent uses of this sensor shift values of +65 or +70 were used. For this cruise values from +50 to +65 looked best for different features, but the temperature data were quite noisy making the comparison difficult. SHIFT was run with a setting of +65 since it has been used for other cruises with this sensor and the results were satisfactory.

##### 13. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0               	Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range:  10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning concerned spikes at the end of the upcast for cast #23, so does not affect the DEL files.

##### 14. DETAILED EDITING

COMPARE indicates that the two salinity channels differ from bottles by the same average amount. The primary shows very slightly more dependence on pressure and time, but the difference is probably not significant given the scatter in the comparisons. In motion the channels are quite different with the secondary temperature generally looking noisier. While the secondary channels have been chosen when these sensors have been used over the past year, the performance is different here, possibly due to pump problems with the secondary. So the primary channels were selected for editing.

Graphical editing was done using program CTDEDIT. On-screen plots of descent rate and pump status were also used.

Cast #6 was difficult to edit. The descent rate was extremely noisy with many complete reversals which can result in corruption by shed wakes, but there also appear to be interleaving features that are real. Editing was applied only where it was clearly corrupted.

Cast #46 – The Header Altimeter entry was edited to reflect the fact that the bottom 3m of data were removed in editing.

As usual for the SoG/JdF cruises, the heaviest editing was required for the casts near the mouth of Juan de Fuca Strait. The descent rate was generally quite steady elsewhere except near the surface and bottom of the casts. There is a lot of variability in many parts of the Strait of Georgia, so many unstable features were left unedited since there is no evidence they are due to shed wakes or instrumental noise. T-S plots were examined and no further editing was found appropriate.

##### 15. Corrections to the Dissolved Oxygen data

At this stage it was realized that two different mistakes had been made in the derivation of the dissolved oxygen in the full files – the bottle files were correct. These errors affected only the DO data. The wrong configuration file had been used in the DERIVE stage (it was ok in the original conversion) and the files used for DELETE were missing the SHIFT in DO.

In order to avoid re-editing the T/S data, a separate set of files were produced. First DERIVE was rerun correctly, then those files were converted to IOS HEADERS. They were then put through CONVERT and CLEAN. Choices were made in CLEAN to remove headers that will otherwise get repeated in the merging process that follows. The same SHIFT setting is expected to be appropriate, but a few tests were done to be sure and +65 still looks appropriate. SHIFT was run on the DO channel using +65. Next the SHFO files were put through DELETE with output *.DEL2.

The DEL2 files were then merged with the EDT files selecting the DO from the DEL2 and everything else from EDT. The output files were named EDTMRG. Because it was easy to do so, and reduces file sizes, a few channels were left out of the merge that should not have been converted in the first place, namely Bottle_Position and Bottle_Number. Altimeter, Status:Pump and Descent_Rate were also removed since they are no longer needed.

##### 16. Initial Recalibration

While the salinity appears to be low by ~0.004, there are doubts about the quality of the bottle comparison. There was a lot of scatter, and other recent comparisons have raised some doubts about samples and/or salinometer. (The best explanation so far for what has been observed is that old liners or chips in bottle lids may be allowing some evaporation of sample.) The salinity data will not be recalibrated. When the T/C sensors are next recalibrated this decision should be reconsidered.

File 2009-64-recal1.ccf  was prepared to divide the SBE Fluorescence by 10 and to apply the following recalibration to the SBE dissolved oxygen:

CTD-BOT = 1.0203 DOX-CTD + 0.0696

This was applied to the SAM and MRGCLN2 files to create SAMCOR1 and MRGCOR1 files. COMPARE was rerun to see that the corrections were applied correctly and they were. (See 2009-64-dox-comp2.xls and 2009-64-fl-chl-comp2.xls.)

Using the corrected fluorescence data, it is clear that the SBE values are similar to the extracted CHL in Juan de Fuca Strait, but fluorescence is higher by a factor of about 2.5 in the Strait of Georgia and even higher in Saanich Inlet.

The EDTMRG files were then recalibrated to create COR1 files.

##### 17. Final Calibration of DO

The first recalibration of dissolved oxygen corrects for calibration drift. Shift corrects for transit time errors. Those 2 steps correct for response time errors, but a further correction is sometimes found appropriate. To check for this downcast CTD data is compared to bottle data from the same pressure.

Downcast files were bin-averaged to 0.5m bins for the casts with DO bottle samples. Those files were then thinned to the usual levels for bottles and compared to the bottle values in the MRG files. COMPARE was used to study the differences between the downcast CTD DO data and the upcast bottles. When obvious outliers are excluded the differences range from -0.24 to +0.35ml/l. The fits against DO are quite flat, but there is some obvious pressure dependence. The following correction looks reasonable:

DOX(corrected) =  DOX +0.00003*Pressure -0.0132

This is a small correction but will be applied since it rduces the pressure dependence. The above correction was applied first to the thinned files and COMPARE was rerun. The results show it was effective. (See 2009-64-dox-comp3.xls and 2009-64-dox-comp4.xls.)

The correction was then applied to the COR1 files with out put COR2.

##### 18. Special Fluorometer Processing

The COR1 files were clipped to 200db and processed separately for A. Peña. The clipped files were bin-averaged (0.25db bins), put through REMOVE and HEADEDIT and named as *.FCTD1 and saved. A second set, *.FCTD2, were created by filtering before bin-averaging. The SAMCOR1 files were put through REMOVE and named *.BOF and saved. A readme.doc file was prepared with some notes on the preparation of those files.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### 19. BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure	Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. There were some unstable features, but these are small and in areas of active mixing. No further editing was applied.

##### 20.Other Comparisons

Previous experience with these sensors –

1. Salinity:  The primary conductivity sensor was found to be low by 0.004 salinity units when used for 2009-03 (Feb) but the sampling was limited to 2 depths. During 2009-27 (March) and 2009-38 (May) there was no calibration sampling. During 2009-26 in April it was found to be high by 0.0025 but there was a lot of scatter in the comparison. During 2009-14 there was a lot of scatter but the CTD appeared to be very close to the bottles. The secondary conductivity sensor was used for the same cruises and was found to be low by <0.001 for 2009-03, high by 0.0004 for 2009-26 and low by 0.002 for 2009-14. It was also used on an SBE25 in 2007 but there were doubts about the bottle analyses for those cruises.

2. Dissolved Oxygen – There are only 4 other cruises since the last factory calibration from which there is DO calibration sampling and both had few samples. For 2009-38 some samples came from areas of high variability in DO. The fits from those were:

CTD-BOT = 1.0414 DOX-CTD + 0.0005 (2009-14 - September)

CTD-BOT = 1.0700 DOX-CTD – 0.0887 (2009-38 - May)

CTD-BOT = 1.0388 DOX-CTD + 0.0088 (2009-26 – April)

CTD-BOT = 1.0209 DOX-CTD + 0.0668 (2009-27 - March)

3. Pressure –The sensor was recalibrated in August 2007 and an offset of +0.5db was used for 6 cruises processed since then. However, it was noted that it might soon be time to increase the offset, and that was done for the most recent use of this sensor during cruise 2009-14.

Historic ranges (3 standard deviations) – The 3 standard deviation plots were not available. Using the 1 standard deviation plots was not useful as too much data falls outside the range limits.

##### 21. FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE and Flag.

A second SBE DO channel (with umol/kg units) was added. REORDER was run to get the two DO channels together.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

Transmissivity and fluorescence data are nominal and unedited except that

some records were removed in editing temperature and salinity.

The SBE dissolved oxygen data in the CTD files are considered, roughly:

•	±0.3  ml/l from  0– 150db

•	±0.1  ml/l from 150– 200db

•	±0.03 ml/l below 200db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. 
The cross-reference list was produced and no problems were found.

The final files were named CTD.

Profile plots were made and no problems were found.

The track plot looks ok.

As a final check of dissolved oxygen data, % saturation was calculated and plotted. The near-surface values varied from 65% to 95%. The lowest saturations were in Haro Strait, Juan de Fuca Strait and well-mixed casts in narrow passages at the northern end of the survey. Values in the Strait of Georgia were generally between 90 and 100%.

##### 22. Final Bottle Files

The MRGCOR1 files were put through SORT to order on increasing pressure.

The following channels were removed from all casts: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

A second SBE DO channel was added with different units.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods.

Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved.

A cross-reference list was produced and turned up no errors.

##### 23. Producing final files

A cross-reference listing was produced for CTD and CHE files.

The sensor history was updated.

Particulars:

3. Problem with connector – sea cable

6. Rosette weight on

12. Chains off

50. Before cast replaced O-rings and cleaned spigots on bottles 1, 2, 3, 8 and 11.

50. Pressure test at end of upcast.

53. Before cast replaced bottle 10 O-ring

56/57. Casts done out of order – labels had already been prepared.

REVISION – October 18, 2010

After these data had been processed and archived it was discovered that the calibration for transmissometer #1005DR had been drifting significantly since about September 2008. The drift was fairly steady, but there was a sudden drop to very low values around August 2009, followed by a further steady drift.  A correction factor was derived by forcing deep offshore data for 2 cruises after August 2009 to have transmissivity of 62%/m and then applying a linear fit to time.

For full details on how the correction was derived see the document:

Osd_Data_Archive:\Cruise_Data\DOCUMENTS\Transmissivity 1005DR Corrections.doc

The correction produces more realistic values in deep water but it is just an estimate. It should also be noted that transmissivity was decreasing slightly from about 500m downwards, so there may be a pressure related problem with the sensor.

# Institute of Ocean Sciences

# CRUISE SUMMARY

CTD

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

| Calibration Information CTD #0550 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 4883 | 22Dec07 | Factory “ |  |  |
| Conductivity | 1763 | 11Oct07 | “ |  |  |
| Secondary Temp. | 2095 | 16Oct07 | “ |  |  |
| Secondary Cond. | 2754 | 25Apr07 | “ |  |  |
| Transmissometer | 1005DR | 5Mar08 | IOS |  |  |
| SBE 43 DO sensor | 1119 | 12Feb2008 | Factory |  |  |
| PAR | 4565 | 29Jan2009 | IOS |  |  |
| Surface PAR | 16504 | 27Feb2009 |  |  |  |
| Fluorometer | 2356 |  |  |  |  |
| Pressure Sensor | 75636 | 20/Aug/2007 | Factory |  |  |
| Altimeter | 1252 |  |  |  |  |